## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 211b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.79791 | 0.78708 | 0.80874 | 0.79176 | 0.80405
PhotonAssay (recommended gross mass* 360±15 g)
Au, Gold (ppm) | 0.79993 | 0.79365 | 0.80621 | 0.79813 | 0.80173
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 0.74967 | 0.73822 | 0.76111 | 0.74334 | 0.75599
Cyanide Leach
Au, Gold (ppm) | 0.77639 | 0.76014 | 0.79264 | 0.77407 | 0.7787
4-Acid Digestion
Ag, Silver (ppm) | 0.33357 | 0.30534 | 0.3618 | 0.30798 | 0.35916
Al, Aluminium (wt.%) | 6.9361 | 6.7169 | 7.1553 | 6.827 | 7.0452
As, Arsenic (ppm) | 33.7495 | 31.954 | 35.5451 | 32.47 | 35.0291
Ba, Barium (ppm) | 151.4677 | 146.1111 | 156.8244 | 147.9811 | 154.9543
Be, Beryllium (ppm) | 0.34181 | 0.30161 | 0.382 | 0.32783 | 0.35579
Bi, Bismuth (ppm) | 0.040847 | 0.027544 | 0.05415 | IND | IND
Ca, Calcium (wt.%) | 6.9344 | 6.7151 | 7.1538 | 6.7993 | 7.0695
Cd, Cadmium (ppm) | 0.45264 | 0.40949 | 0.49578 | 0.42118 | 0.48409
Ce, Cerium (ppm) | 11.7322 | 11.1286 | 12.3358 | 11.3705 | 12.0938
Co, Cobalt (ppm) | 44.3712 | 42.6688 | 46.0735 | 43.4246 | 45.3177
Cr, Chromium (ppm) | 121.3588 | 112.2483 | 130.4692 | 116.6876 | 126.03
Cs, Caesium (ppm) | 0.68548 | 0.64781 | 0.72316 | 0.65705 | 0.71392
Cu, Copper (ppm) | 177.0683 | 170.8627 | 183.2738 | 172.6792 | 181.4573
Dy, Dysprosium (ppm) | 3.9042 | 3.6508 | 4.1576 | 3.7894 | 4.0189
Er, Erbium (ppm) | 2.339 | 2.1756 | 2.5023 | 2.2478 | 2.4302
Eu, Europium (ppm) | 0.97963 | 0.86545 | 1.0938 | 0.92003 | 1.0392
Fe, Iron (wt.%) | 7.9712 | 7.7925 | 8.15 | 7.8542 | 8.0883
Ga, Gallium (ppm) | 15.9611 | 15.3249 | 16.5974 | 15.7105 | 16.2117
Gd, Gadolinium (ppm) | 3.2839 | 3.06 | 3.5077 | 3.1013 | 3.4664
Hf, Hafnium (ppm) | 1.5476 | 1.4455 | 1.6496 | 1.4657 | 1.6294
Ho, Holmium (ppm) | 0.78957 | 0.73623 | 0.84291 | 0.74499 | 0.83416
In, Indium (ppm) | 0.073368 | 0.066441 | 0.080295 | 0.068019 | 0.078717
K, Potassium (wt.%) | 0.36168 | 0.34653 | 0.37683 | 0.35255 | 0.3708
La, Lanthanum (ppm) | 4.7405 | 4.4984 | 4.9826 | 4.6282 | 4.8528
Li, Lithium (ppm) | 11.5743 | 10.9245 | 12.2241 | 11.1737 | 11.9749
Lu, Lutetium (ppm) | 0.34045 | 0.31121 | 0.36968 | 0.31836 | 0.36254
Mg, Magnesium (wt.%) | 3.8772 | 3.7808 | 3.9736 | 3.8154 | 3.939
Mn, Manganese (wt.%) | 0.13547 | 0.13056 | 0.14038 | 0.13275 | 0.13819
Mo, Molybdenum (ppm) | 1.5214 | 1.4019 | 1.6408 | 1.4309 | 1.6118
Na, Sodium (wt.%) | 1.9039 | 1.8504 | 1.9575 | 1.8742 | 1.9337
Nb, Niobium (ppm) | 3.4307 | 3.229 | 3.6323 | 3.295 | 3.5664
Nd, Neodymium (ppm) | 8.1287 | 7.5875 | 8.67 | 7.9248 | 8.3326
Ni, Nickel (ppm) | 94.2933 | 90.8741 | 97.7124 | 91.9012 | 96.6853
P, Phosphorus (wt.%) | 0.041496 | 0.040052 | 0.042941 | 0.040218 | 0.042774
Pb, Lead (ppm) | 14.8071 | 13.6251 | 15.989 | 14.0563 | 15.5579
Pr, Praseodymium (ppm) | 1.666 | 1.5822 | 1.7498 | 1.6038 | 1.7283
Rb, Rubidium (ppm) | 10.1312 | 9.5389 | 10.7234 | 9.8347 | 10.4276
Re, Rhenium (ppm) | 0.0022867 | 0.0013485 | 0.0032248 | IND | IND
S, Sulphur (wt.%) | 0.26704 | 0.25393 | 0.28015 | 0.25815 | 0.27593
Sb, Antimony (ppm) | 1.3757 | 1.273 | 1.4785 | 1.3241 | 1.4273
Sc, Scandium (ppm) | 41.1073 | 39.4171 | 42.7976 | 40.0987 | 42.116
Sm, Samarium (ppm) | 2.5163 | 2.2991 | 2.7335 | 2.3852 | 2.6474
Sn, Tin (ppm) | 0.9161 | 0.80739 | 1.0248 | IND | IND
Sr, Strontium (ppm) | 213.3548 | 206.8029 | 219.9067 | 208.6957 | 218.0139
Ta, Tantalum (ppm) | 0.2247 | 0.2057 | 0.2437 | 0.21017 | 0.23923
Tb, Terbium (ppm) | 0.57997 | 0.50753 | 0.65241 | 0.53876 | 0.62118
Te, Tellurium (ppm) | 0.080686 | 0.056633 | 0.10474 | IND | IND
Th, Thorium (ppm) | 0.73983 | 0.69276 | 0.7869 | 0.71991 | 0.75975
Ti, Titanium (wt.%) | 0.63328 | 0.61021 | 0.65635 | 0.62211 | 0.64445
Tl, Thallium (ppm) | 0.15263 | 0.1418 | 0.16345 | IND | IND
Tm, Thulium (ppm) | 0.32838 | 0.29854 | 0.35822 | 0.31226 | 0.3445
U, Uranium (ppm) | 0.20453 | 0.19028 | 0.21879 | 0.18959 | 0.21947
V, Vanadium (ppm) | 281.8985 | 272.6275 | 291.1696 | 276.2257 | 287.5714
W, Tungsten (ppm) | 16.8373 | 16.1856 | 17.4891 | 16.3161 | 17.3586
Y, Yttrium (ppm) | 20.3193 | 19.4155 | 21.2231 | 19.8942 | 20.7445
Yb, Ytterbium (ppm) | 2.1796 | 2.004 | 2.3552 | 2.0839 | 2.2753
Zn, Zinc (ppm) | 126.4197 | 122.0553 | 130.7841 | 123.5418 | 129.2977
Zr, Zirconium (ppm) | 47.1279 | 44.2251 | 50.0307 | 44.1146 | 50.1412
Aqua Regia Digestion
Ag, Silver (ppm) | 0.33853 | 0.3095 | 0.36755 | 0.31922 | 0.35784
Al, Aluminium (wt.%) | 3.2816 | 3.1429 | 3.4203 | 3.2244 | 3.3388
As, Arsenic (ppm) | 32.6363 | 31.2129 | 34.0596 | 31.662 | 33.6105
B, Boron (ppm) | 20.0159 | 18.8069 | 21.2248 | 16.9156 | 23.1161
Ba, Barium (ppm) | 29.3352 | 27.7568 | 30.9135 | 28.2237 | 30.4467
Be, Beryllium (ppm) | 0.19024 | 0.16452 | 0.21595 | IND | IND
Bi, Bismuth (ppm) | 0.032967 | 0.024189 | 0.041745 | IND | IND
Ca, Calcium (wt.%) | 2.3592 | 2.1181 | 2.6002 | 2.2854 | 2.4329
Cd, Cadmium (ppm) | 0.42454 | 0.38877 | 0.4603 | 0.39648 | 0.45259
Ce, Cerium (ppm) | 8.6565 | 8.0637 | 9.2492 | 8.34 | 8.9729
Co, Cobalt (ppm) | 31.0953 | 29.9515 | 32.2391 | 30.1754 | 32.0153
Cr, Chromium (ppm) | 19.2842 | 18.094 | 20.4744 | 18.4521 | 20.1163
Cs, Caesium (ppm) | 0.48688 | 0.46071 | 0.51306 | 0.46253 | 0.51124
Cu, Copper (ppm) | 176.0751 | 170.9045 | 181.2458 | 173.3453 | 178.8049
Eu, Europium (ppm) | 0.49515 | 0.34639 | 0.6439 | 0.45151 | 0.53878
Fe, Iron (wt.%) | 5.6266 | 5.5221 | 5.7311 | 5.5436 | 5.7095
Ga, Gallium (ppm) | 9.9621 | 9.2943 | 10.6298 | 9.6286 | 10.2956
Ge, Germanium (ppm) | 0.13883 | 0.10591 | 0.17176 | IND | IND
Hf, Hafnium (ppm) | 0.43885 | 0.39078 | 0.48691 | 0.41148 | 0.46622
Hg, Mercury (ppm) | 0.051685 | 0.033134 | 0.070236 | IND | IND
Ho, Holmium (ppm) | 0.47861 | 0.36144 | 0.59578 | 0.43664 | 0.52057
K, Potassium (wt.%) | 0.10349 | 0.095345 | 0.11163 | 0.099788 | 0.10719
La, Lanthanum (ppm) | 3.5394 | 3.2857 | 3.7931 | 3.4089 | 3.6698
Li, Lithium (ppm) | 10.0328 | 9.5658 | 10.4998 | 9.79 | 10.2756
Mg, Magnesium (wt.%) | 1.7075 | 1.6558 | 1.7593 | 1.6801 | 1.735
Mn, Manganese (wt.%) | 0.072608 | 0.070402 | 0.074814 | 0.071591 | 0.073625
Mo, Molybdenum (ppm) | 1.4846 | 1.3972 | 1.572 | 1.394 | 1.5751
Na, Sodium (wt.%) | 0.13728 | 0.12876 | 0.14579 | 0.13287 | 0.14169
Nb, Niobium (ppm) | 0.11658 | 0.080375 | 0.15279 | 0.10244 | 0.13073
Ni, Nickel (ppm) | 64.968 | 62.4198 | 67.5162 | 63.6707 | 66.2653
P, Phosphorus (wt.%) | 0.040239 | 0.038742 | 0.041736 | 0.039163 | 0.041315
Pb, Lead (ppm) | 14.3247 | 13.5978 | 15.0515 | 13.7362 | 14.9132
Pr, Praseodymium (ppm) | 1.3244 | 0.96168 | 1.6871 | 1.2062 | 1.4426
Rb, Rubidium (ppm) | 4.3604 | 4.1272 | 4.5935 | 4.1881 | 4.5326
Re, Rhenium (ppm) | 0.0018267 | 0.00052709 | 0.0031262 | IND | IND
S, Sulphur (wt.%) | 0.2657 | 0.25437 | 0.27704 | 0.2558 | 0.27561
Sb, Antimony (ppm) | 0.88319 | 0.78386 | 0.98251 | 0.85027 | 0.9161
Sc, Scandium (ppm) | 5.1708 | 4.6935 | 5.648 | 4.9996 | 5.3419
Se, Selenium (ppm) | 0.58707 | 0.44407 | 0.73008 | IND | IND
Sn, Tin (ppm) | 0.53273 | 0.44854 | 0.61693 | 0.49113 | 0.57433
Sr, Strontium (ppm) | 35.281 | 30.9085 | 39.6535 | 34.045 | 36.517
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.323 | 0.26144 | 0.38456 | 0.29985 | 0.34615
Te, Tellurium (ppm) | 0.060855 | 0.04379 | 0.07792 | IND | IND
Th, Thorium (ppm) | 0.58064 | 0.50962 | 0.65166 | 0.52182 | 0.63945
Ti, Titanium (wt.%) | 0.37124 | 0.34006 | 0.40242 | 0.35737 | 0.38511
Tl, Thallium (ppm) | 0.077424 | 0.069473 | 0.085375 | 0.071119 | 0.083729
U, Uranium (ppm) | 0.14338 | 0.12875 | 0.15801 | 0.12801 | 0.15874
V, Vanadium (ppm) | 136.2931 | 126.4812 | 146.105 | 131.8285 | 140.7577
W, Tungsten (ppm) | 10.9694 | 9.8732 | 12.0656 | 10.4447 | 11.4941
Y, Yttrium (ppm) | 11.2365 | 10.624 | 11.849 | 10.795 | 11.6779
Zn, Zinc (ppm) | 112.8153 | 108.5673 | 117.0632 | 110.5607 | 115.0698
Zr, Zirconium (ppm) | 14.6974 | 13.4754 | 15.9194 | 14.0378 | 15.357
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 211b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 44.6667 | Hg | ppm | 0.12167 | Se | ppm | 0.8363
Ge | ppm | 0.19376 | Pt | ppb | 19.1667
Aqua Regia Digestion
Dy | ppm | 2.1043 | Lu | ppm | 0.16465 | Sm | ppm | 1.4509
Er | ppm | 1.2991 | Nd | ppm | 5.6705 | Tm | ppm | 0.17576
Gd | ppm | 1.9022 | Pd | ppb | 16.7332 | Yb | ppm | 1.3132
In | ppm | 0.0285 | Pt | ppb | 18.7059
Borate Fusion XRF
Al2O3 | wt.% | 13.655 | Fe2O3 | wt.% | 11.7021 | S | wt.% | 0.252
As | ppm | 30 | K2O | wt.% | 0.446 | SiO2 | wt.% | 50.16
BaO | ppm | 240.0475 | MgO | wt.% | 6.61 | Sn | ppm | 45
CaO | wt.% | 10.14 | MnO | wt.% | 0.186 | Sr | ppm | 225
Cl | ppm | 1510 | Na2O | wt.% | 2.56 | TiO2 | wt.% | 1.1085
Co | ppm | 60 | Ni | ppm | 155 | V2O5 | ppm | 526.634
Cr2O3 | ppm | 241.1475 | P2O5 | wt.% | 0.093956 | Zn | ppm | 140
Cu | ppm | 200 | Pb | ppm | 85 | Zr | ppm | 100
Thermogravimetry
LOI1000 | wt.% | 2.955
Infrared Combustion
C | wt.% | 0.115 | S | wt.% | 0.295
Laser Ablation ICP-MS
Ag | ppm | 0.35 | Hf | ppm | 1.845 | Sm | ppm | 2.615
As | ppm | 33.2 | Ho | ppm | 0.845 | Sn | ppm | 0.9
Ba | ppm | 150 | In | ppm | 0.05 | Sr | ppm | 207.5
Be | ppm | 0.6 | La | ppm | 4.77 | Ta | ppm | 0.21
Bi | ppm | < 0.2 | Lu | ppm | 0.33 | Tb | ppm | 0.585
Cd | ppm | 0.5 | Mn | wt.% | 0.1445 | Te | ppm | < 0.2
Ce | ppm | 11.55 | Mo | ppm | 1.8 | Th | ppm | 0.73
Co | ppm | 45.7 | Nb | ppm | 3.43 | Ti | wt.% | 0.6485
Cr | ppm | 142.5 | Nd | ppm | 8.325 | Tl | ppm | < 0.2
Cs | ppm | 0.735 | Ni | ppm | 98 | Tm | ppm | 0.365
Cu | ppm | 179 | Pb | ppm | 14.5 | U | ppm | 0.22
Dy | ppm | 3.88 | Pr | ppm | 1.71 | V | ppm | 299
Er | ppm | 2.485 | Rb | ppm | 9.75 | W | ppm | 17.25
Eu | ppm | 0.94 | Re | ppm | < 0.01 | Y | ppm | 20.75
Ga | ppm | 15.55 | Sb | ppm | 1.45 | Yb | ppm | 2.33
Gd | ppm | 3.21 | Sc | ppm | 41.2 | Zn | ppm | 125
Ge | ppm | 1.375 | Se | ppm | < 5 | Zr | ppm | 63.75
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 211b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.79791 | 0.024855 | 0.7482 | 0.84762 | 0.72334 | 0.87247 | 3.115% | 6.23% | 9.345% | 0.75801 | 0.8378
PhotonAssay
Au, ppm | 0.79993 | 0.031634 | 0.73666 | 0.8632 | 0.70503 | 0.89483 | 3.955% | 7.909% | 11.86% | 0.75993 | 0.83993
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.74967 | 0.027725 | 0.69422 | 0.80512 | 0.66649 | 0.83284 | 3.698% | 7.397% | 11.09% | 0.71219 | 0.78715
Cyanide Leach
Au, ppm | 0.77639 | 0.035236 | 0.70592 | 0.84686 | 0.67068 | 0.8821 | 4.538% | 9.077% | 13.62% | 0.73757 | 0.81521
4-Acid Digestion
Ag, ppm | 0.33357 | 0.021801 | 0.28997 | 0.37717 | 0.26817 | 0.39897 | 6.536% | 13.07% | 19.61% | 0.31689 | 0.35025
Al, wt.% | 6.9361 | 0.25207 | 6.4319 | 7.4402 | 6.1799 | 7.6923 | 3.634% | 7.268% | 10.9% | 6.5893 | 7.2829
As, ppm | 33.7495 | 1.3203 | 31.1089 | 36.3902 | 29.7885 | 37.7105 | 3.912% | 7.824% | 11.74% | 32.062 | 35.437
Ba, ppm | 151.4677 | 3.3902 | 144.6873 | 158.2482 | 141.2971 | 161.6384 | 2.238% | 4.476% | 6.715% | 143.8943 | 159.0411
Be, ppm | 0.34181 | 0.029733 | 0.28234 | 0.40127 | 0.25261 | 0.43101 | 8.699% | 17.4% | 26.1% | 0.32472 | 0.3589
Bi, ppm | 0.040847 | 0.0099207 | 0.021006 | 0.060688 | 0.011085 | 0.070609 | 24.29% | 48.57% | 72.86% | 0.038805 | 0.042889
Ca, wt.% | 6.9344 | 0.25905 | 6.4163 | 7.4525 | 6.1573 | 7.7116 | 3.736% | 7.472% | 11.21% | 6.5877 | 7.2811
Cd, ppm | 0.45264 | 0.036171 | 0.38029 | 0.52498 | 0.34412 | 0.56115 | 7.991% | 15.98% | 23.97% | 0.43 | 0.47527
Ce, ppm | 11.7322 | 0.47251 | 10.7872 | 12.6772 | 10.3147 | 13.1497 | 4.027% | 8.055% | 12.08% | 11.1456 | 12.3188
Co, ppm | 44.3712 | 1.3375 | 41.6961 | 47.0463 | 40.3585 | 48.3838 | 3.014% | 6.029% | 9.043% | 42.1526 | 46.5897
Cr, ppm | 121.3588 | 10.9694 | 99.42 | 143.2975 | 88.4506 | 154.2669 | 9.039% | 18.08% | 27.12% | 115.2908 | 127.4267
Cs, ppm | 0.68548 | 0.028208 | 0.62907 | 0.7419 | 0.60086 | 0.77011 | 4.115% | 8.23% | 12.35% | 0.65121 | 0.71976
Cu, ppm | 177.0683 | 5.3498 | 166.3687 | 187.7679 | 161.0189 | 193.1177 | 3.021% | 6.043% | 9.064% | 168.2149 | 185.9217
Dy, ppm | 3.9042 | 0.16766 | 3.5689 | 4.2395 | 3.4012 | 4.4072 | 4.294% | 8.588% | 12.88% | 3.709 | 4.0994
Er, ppm | 2.339 | 0.10256 | 2.1339 | 2.5441 | 2.0313 | 2.6466 | 4.385% | 8.769% | 13.15% | 2.222 | 2.4559
Eu, ppm | 0.97963 | 0.10152 | 0.77659 | 1.1827 | 0.67506 | 1.2842 | 10.36% | 20.73% | 31.09% | 0.93065 | 1.0286
Fe, wt.% | 7.9712 | 0.23019 | 7.5109 | 8.4316 | 7.2807 | 8.6618 | 2.888% | 5.775% | 8.663% | 7.5727 | 8.3698
Ga, ppm | 15.9611 | 0.64526 | 14.6706 | 17.2517 | 14.0253 | 17.8969 | 4.043% | 8.085% | 12.13% | 15.1631 | 16.7592
Gd, ppm | 3.2839 | 0.14391 | 2.996 | 3.5717 | 2.8521 | 3.7156 | 4.382% | 8.765% | 13.15% | 3.1197 | 3.4481
Hf, ppm | 1.5476 | 0.068 | 1.4116 | 1.6836 | 1.3436 | 1.7516 | 4.394% | 8.788% | 13.18% | 1.4702 | 1.6249
Ho, ppm | 0.78957 | 0.028722 | 0.73213 | 0.84702 | 0.70341 | 0.87574 | 3.638% | 7.275% | 10.91% | 0.7501 | 0.82905
In, ppm | 0.073368 | 0.0045838 | 0.0642 | 0.082536 | 0.059617 | 0.087119 | 6.248% | 12.5% | 18.74% | 0.0697 | 0.077036
K, wt.% | 0.36168 | 0.017039 | 0.3276 | 0.39576 | 0.31056 | 0.4128 | 4.711% | 9.422% | 14.13% | 0.34359 | 0.37976
La, ppm | 4.7405 | 0.26494 | 4.2106 | 5.2704 | 3.9457 | 5.5353 | 5.589% | 11.18% | 16.77% | 4.5035 | 4.9775
Li, ppm | 11.5743 | 1.0639 | 9.4464 | 13.7022 | 8.3825 | 14.7661 | 9.192% | 18.38% | 27.58% | 10.9956 | 12.153
Lu, ppm | 0.34045 | 0.016105 | 0.30824 | 0.37266 | 0.29213 | 0.38876 | 4.73% | 9.461% | 14.19% | 0.32342 | 0.35747
Mg, wt.% | 3.8772 | 0.10413 | 3.669 | 4.0855 | 3.5648 | 4.1896 | 2.686% | 5.372% | 8.057% | 3.6834 | 4.0711
Mn, wt.% | 0.13547 | 0.0063254 | 0.12282 | 0.14812 | 0.1165 | 0.15445 | 4.669% | 9.338% | 14.01% | 0.1287 | 0.14224
Mo, ppm | 1.5214 | 0.097997 | 1.3254 | 1.7174 | 1.2274 | 1.8154 | 6.441% | 12.88% | 19.32% | 1.4453 | 1.5974
Na, wt.% | 1.9039 | 0.050566 | 1.8028 | 2.0051 | 1.7522 | 2.0556 | 2.656% | 5.312% | 7.968% | 1.8087 | 1.9991
Nb, ppm | 3.4307 | 0.1981 | 3.0345 | 3.8269 | 2.8364 | 4.025 | 5.774% | 11.55% | 17.32% | 3.2591 | 3.6022
Nd, ppm | 8.1287 | 0.31552 | 7.4977 | 8.7598 | 7.1822 | 9.0753 | 3.881% | 7.763% | 11.64% | 7.7223 | 8.5352
Ni, ppm | 94.2933 | 3.9038 | 86.4856 | 102.1009 | 82.5818 | 106.0047 | 4.14% | 8.28% | 12.42% | 89.5786 | 99.0079
P, wt.% | 0.041496 | 0.0015159 | 0.038464 | 0.044528 | 0.036948 | 0.046044 | 3.653% | 7.306% | 10.96% | 0.039421 | 0.043571
Pb, ppm | 14.8071 | 1.0153 | 12.7765 | 16.8377 | 11.7612 | 17.853 | 6.857% | 13.71% | 20.57% | 14.0667 | 15.5474
Pr, ppm | 1.666 | 0.047703 | 1.5706 | 1.7614 | 1.5229 | 1.8091 | 2.863% | 5.727% | 8.59% | 1.5827 | 1.7493
Rb, ppm | 10.1312 | 0.67498 | 8.7812 | 11.4811 | 8.1062 | 12.1561 | 6.662% | 13.32% | 19.99% | 9.6246 | 10.6377
Re, ppm | 0.0022867 | 0.00054385 | 0.001199 | 0.0033744 | 0.00065512 | 0.0039182 | 23.78% | 47.57% | 71.35% | 0.0021723 | 0.002401
S, wt.% | 0.26704 | 0.012389 | 0.24226 | 0.29182 | 0.22987 | 0.30421 | 4.639% | 9.279% | 13.92% | 0.25369 | 0.28039
Sb, ppm | 1.3757 | 0.082074 | 1.2116 | 1.5399 | 1.1295 | 1.6219 | 5.966% | 11.93% | 17.9% | 1.3069 | 1.4445
Sc, ppm | 41.1073 | 1.8521 | 37.4032 | 44.8115 | 35.5512 | 46.6635 | 4.505% | 9.011% | 13.52% | 39.052 | 43.1627
Sm, ppm | 2.5163 | 0.096236 | 2.3238 | 2.7088 | 2.2276 | 2.805 | 3.825% | 7.649% | 11.47% | 2.3905 | 2.6421
Sn, ppm | 0.9161 | 0.087738 | 0.74062 | 1.0916 | 0.65289 | 1.1793 | 9.577% | 19.15% | 28.73% | 0.8703 | 0.96191
Sr, ppm | 213.3548 | 6.932 | 199.4908 | 227.2188 | 192.5588 | 234.1508 | 3.249% | 6.498% | 9.747% | 202.6871 | 224.0225
Ta, ppm | 0.2247 | 0.012808 | 0.19908 | 0.25032 | 0.18628 | 0.26312 | 5.7% | 11.4% | 17.1% | 0.21346 | 0.23593
Tb, ppm | 0.57997 | 0.056139 | 0.46769 | 0.69225 | 0.41156 | 0.74839 | 9.68% | 19.36% | 29.04% | 0.55097 | 0.60897
Te, ppm | 0.080686 | 0.020634 | 0.039419 | 0.12195 | 0.018785 | 0.14259 | 25.57% | 51.15% | 76.72% | 0.076652 | 0.08472
Th, ppm | 0.73983 | 0.03231 | 0.67521 | 0.80445 | 0.6429 | 0.83676 | 4.367% | 8.735% | 13.1% | 0.70284 | 0.77682
Ti, wt.% | 0.63328 | 0.033882 | 0.56552 | 0.70104 | 0.53163 | 0.73493 | 5.35% | 10.7% | 16.05% | 0.60162 | 0.66494
Tl, ppm | 0.15263 | 0.0052311 | 0.14216 | 0.16309 | 0.13693 | 0.16832 | 3.427% | 6.855% | 10.28% | 0.14499 | 0.16026
Tm, ppm | 0.32838 | 0.034724 | 0.25893 | 0.39782 | 0.22421 | 0.43255 | 10.57% | 21.15% | 31.72% | 0.31196 | 0.3448
U, ppm | 0.20453 | 0.010728 | 0.18308 | 0.22599 | 0.17235 | 0.23672 | 5.245% | 10.49% | 15.74% | 0.19431 | 0.21476
V, ppm | 281.8985 | 9.0501 | 263.7984 | 299.9987 | 254.7483 | 309.0487 | 3.21% | 6.421% | 9.631% | 267.8036 | 295.9935
W, ppm | 16.8373 | 0.78072 | 15.2759 | 18.3988 | 14.4952 | 19.1795 | 4.637% | 9.274% | 13.91% | 15.9955 | 17.6792
Y, ppm | 20.3193 | 0.82923 | 18.6609 | 21.9778 | 17.8316 | 22.807 | 4.081% | 8.162% | 12.24% | 19.3033 | 21.3353
Yb, ppm | 2.1796 | 0.084038 | 2.0115 | 2.3477 | 1.9275 | 2.4317 | 3.856% | 7.711% | 11.57% | 2.0706 | 2.2886
Zn, ppm | 126.4197 | 4.5704 | 117.279 | 135.5604 | 112.7087 | 140.1308 | 3.615% | 7.23% | 10.85% | 120.0987 | 132.7407
Zr, ppm | 47.1279 | 2.8554 | 41.4171 | 52.8387 | 38.5617 | 55.6941 | 6.059% | 12.12% | 18.18% | 44.7715 | 49.4843
Aqua Regia Digestion
Ag, ppm | 0.33853 | 0.035396 | 0.26774 | 0.40932 | 0.23234 | 0.44472 | 10.46% | 20.91% | 31.37% | 0.3216 | 0.35546
Al, wt.% | 3.2816 | 0.16927 | 2.9431 | 3.6201 | 2.7738 | 3.7894 | 5.158% | 10.32% | 15.47% | 3.1175 | 3.4457
As, ppm | 32.6363 | 1.6523 | 29.3317 | 35.9408 | 27.6795 | 37.593 | 5.063% | 10.13% | 15.19% | 31.0044 | 34.2681
B, ppm | 20.0159 | 1.0607 | 17.8944 | 22.1373 | 16.8337 | 23.198 | 5.299% | 10.6% | 15.9% | 19.0151 | 21.0167
Ba, ppm | 29.3352 | 2.3986 | 24.538 | 34.1323 | 22.1394 | 36.5309 | 8.176% | 16.35% | 24.53% | 27.8684 | 30.8019
Be, ppm | 0.19024 | 0.023632 | 0.14297 | 0.2375 | 0.11934 | 0.26113 | 12.42% | 24.84% | 37.27% | 0.18073 | 0.19975
Bi, ppm | 0.032967 | 0.0045334 | 0.0239 | 0.042033 | 0.019367 | 0.046567 | 13.75% | 27.5% | 41.25% | 0.031318 | 0.034615
Ca, wt.% | 2.3592 | 0.43964 | 1.4799 | 3.2384 | 1.0403 | 3.6781 | 18.64% | 37.27% | 55.91% | 2.2412 | 2.4771
Cd, ppm | 0.42454 | 0.032518 | 0.3595 | 0.48957 | 0.32698 | 0.52209 | 7.66% | 15.32% | 22.98% | 0.40331 | 0.44576
Ce, ppm | 8.6565 | 0.74034 | 7.1758 | 10.1371 | 6.4354 | 10.8775 | 8.552% | 17.1% | 25.66% | 8.2236 | 9.0893
Co, ppm | 31.0953 | 1.0266 | 29.0421 | 33.1485 | 28.0155 | 34.1752 | 3.301% | 6.603% | 9.904% | 29.5406 | 32.6501
Cr, ppm | 19.2842 | 1.5935 | 16.0973 | 22.4711 | 14.5038 | 24.0646 | 8.263% | 16.53% | 24.79% | 18.32 | 20.2484
Cs, ppm | 0.48688 | 0.018325 | 0.45023 | 0.52353 | 0.43191 | 0.54186 | 3.764% | 7.527% | 11.29% | 0.46254 | 0.51123
Cu, ppm | 176.0751 | 5.9778 | 164.1195 | 188.0308 | 158.1417 | 194.0086 | 3.395% | 6.79% | 10.19% | 167.2714 | 184.8789
Eu, ppm | 0.49515 | 0.12506 | 0.24502 | 0.74527 | 0.11995 | 0.87034 | 25.26% | 50.52% | 75.77% | 0.47039 | 0.5199
Fe, wt.% | 5.6266 | 0.16086 | 5.3049 | 5.9483 | 5.144 | 6.1091 | 2.859% | 5.718% | 8.577% | 5.3452 | 5.9079
Ga, ppm | 9.9621 | 0.8726 | 8.2169 | 11.7073 | 7.3443 | 12.5799 | 8.759% | 17.52% | 26.28% | 9.464 | 10.4602
Ge, ppm | 0.13883 | 0.027311 | 0.084212 | 0.19346 | 0.056901 | 0.22077 | 19.67% | 39.34% | 59.02% | 0.13189 | 0.14578
Hf, ppm | 0.43885 | 0.036758 | 0.36533 | 0.51237 | 0.32857 | 0.54912 | 8.376% | 16.75% | 25.13% | 0.41691 | 0.46079
Hg, ppm | 0.051685 | 0.018157 | 0.015372 | 0.087999 | 0 | 0.10616 | 35.13% | 70.26% | 105.4% | 0.049101 | 0.054269
Ho, ppm | 0.47861 | 0.080962 | 0.31668 | 0.64053 | 0.23572 | 0.7215 | 16.92% | 33.83% | 50.75% | 0.45468 | 0.50254
K, wt.% | 0.10349 | 0.0075149 | 0.088457 | 0.11852 | 0.080942 | 0.12603 | 7.262% | 14.52% | 21.79% | 0.098312 | 0.10866
La, ppm | 3.5394 | 0.34667 | 2.846 | 4.2327 | 2.4994 | 4.5794 | 9.795% | 19.59% | 29.38% | 3.3624 | 3.7163
Li, ppm | 10.0328 | 0.56449 | 8.9038 | 11.1618 | 8.3393 | 11.7263 | 5.626% | 11.25% | 16.88% | 9.5312 | 10.5344
Mg, wt.% | 1.7075 | 0.060682 | 1.5862 | 1.8289 | 1.5255 | 1.8896 | 3.554% | 7.108% | 10.66% | 1.6222 | 1.7929
Mn, wt.% | 0.072608 | 0.0029551 | 0.066698 | 0.078518 | 0.063743 | 0.081473 | 4.07% | 8.14% | 12.21% | 0.068978 | 0.076238
Mo, ppm | 1.4846 | 0.084354 | 1.3159 | 1.6533 | 1.2315 | 1.7376 | 5.682% | 11.36% | 17.05% | 1.4103 | 1.5588
Na, wt.% | 0.13728 | 0.0081594 | 0.12096 | 0.1536 | 0.1128 | 0.16176 | 5.944% | 11.89% | 17.83% | 0.13041 | 0.14414
Nb, ppm | 0.11658 | 0.025768 | 0.065048 | 0.16812 | 0.03928 | 0.19389 | 22.1% | 44.21% | 66.31% | 0.11076 | 0.12241
Ni, ppm | 64.968 | 2.7002 | 59.5676 | 70.3684 | 56.8674 | 73.0686 | 4.156% | 8.312% | 12.47% | 61.7196 | 68.2164
P, wt.% | 0.040239 | 0.0017463 | 0.036746 | 0.043731 | 0.035 | 0.045478 | 4.34% | 8.68% | 13.02% | 0.038227 | 0.042251
Pb, ppm | 14.3247 | 0.84102 | 12.6426 | 16.0067 | 11.8016 | 16.8477 | 5.871% | 11.74% | 17.61% | 13.6084 | 15.0409
Pr, ppm | 1.3244 | 0.25757 | 0.80928 | 1.8395 | 0.55171 | 2.0971 | 19.45% | 38.9% | 58.34% | 1.2582 | 1.3906
Rb, ppm | 4.3604 | 0.16449 | 4.0314 | 4.6893 | 3.8669 | 4.8538 | 3.772% | 7.545% | 11.32% | 4.1423 | 4.5784
Re, ppm | 0.0018267 | 0.0004682 | 0.00089027 | 0.0027631 | 0.00042206 | 0.0032313 | 25.63% | 51.26% | 76.89% | 0.0017353 | 0.001918
S, wt.% | 0.2657 | 0.016088 | 0.23353 | 0.29788 | 0.21744 | 0.31397 | 6.055% | 12.11% | 18.16% | 0.25242 | 0.27899
Sb, ppm | 0.88319 | 0.097528 | 0.68813 | 1.0782 | 0.5906 | 1.1758 | 11.04% | 22.09% | 33.13% | 0.83903 | 0.92735
Sc, ppm | 5.1708 | 0.71031 | 3.7501 | 6.5914 | 3.0398 | 7.3017 | 13.74% | 27.47% | 41.21% | 4.9122 | 5.4293
Se, ppm | 0.58707 | 0.13595 | 0.31517 | 0.85898 | 0.17922 | 0.99493 | 23.16% | 46.31% | 69.47% | 0.55772 | 0.61643
Sn, ppm | 0.53273 | 0.096769 | 0.3392 | 0.72627 | 0.24243 | 0.82304 | 18.16% | 36.33% | 54.49% | 0.5061 | 0.55937
Sr, ppm | 35.281 | 6.3284 | 22.6242 | 47.9378 | 16.2958 | 54.2662 | 17.94% | 35.87% | 53.81% | 33.5169 | 37.045
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.323 | 0.0455 | 0.232 | 0.414 | 0.1865 | 0.4595 | 14.09% | 28.17% | 42.26% | 0.30685 | 0.33915
Te, ppm | 0.060855 | 0.0085743 | 0.043706 | 0.078004 | 0.035132 | 0.086578 | 14.09% | 28.18% | 42.27% | 0.057812 | 0.063898
Th, ppm | 0.58064 | 0.075268 | 0.4301 | 0.73117 | 0.35483 | 0.80644 | 12.96% | 25.93% | 38.89% | 0.55161 | 0.60967
Ti, wt.% | 0.37124 | 0.047956 | 0.27533 | 0.46715 | 0.22737 | 0.51511 | 12.92% | 25.84% | 38.75% | 0.35268 | 0.3898
Tl, ppm | 0.077424 | 0.0061768 | 0.06507 | 0.089778 | 0.058893 | 0.095955 | 7.978% | 15.96% | 23.93% | 0.073553 | 0.081295
U, ppm | 0.14338 | 0.01447 | 0.11444 | 0.17232 | 0.099968 | 0.18679 | 10.09% | 20.18% | 30.28% | 0.13621 | 0.15055
V, ppm | 136.2931 | 12.4274 | 111.4383 | 161.1479 | 99.0109 | 173.5753 | 9.118% | 18.24% | 27.35% | 129.4784 | 143.1078
W, ppm | 10.9694 | 1.5907 | 7.788 | 14.1508 | 6.1974 | 15.7414 | 14.5% | 29% | 43.5% | 10.4209 | 11.5179
Y, ppm | 11.2365 | 0.55171 | 10.133 | 12.3399 | 9.5813 | 12.8916 | 4.91% | 9.82% | 14.73% | 10.6746 | 11.7983
Zn, ppm | 112.8153 | 5.1719 | 102.4714 | 123.1591 | 97.2995 | 128.331 | 4.584% | 9.169% | 13.75% | 107.1745 | 118.456
Zr, ppm | 14.6974 | 1.3325 | 12.0325 | 17.3623 | 10.7 | 18.6948 | 9.066% | 18.13% | 27.2% | 13.9625 | 15.4323
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 211b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.15)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 211b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST (SÖGÜT), Sögüt, Bilecik Province, Turkey
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Limited, Manso Nkwanta, Ashanti Region, Ghana
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Laboratoire LABOMINE SARL, Agadir, Souss-Massa, Morocco
Labwest Minerals Analysis, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 211b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 13 | 1 | 267819 | 0.083775 | 0.83206 | 0.81265
2 | 1 | ANSLu | 1 | 9 | 3 | 267820 | 0.08328 | 0.84075 | 0.81312
2 | 1 | ANSLu | 1 | 3 | 6 | 267821 | 0.088182 | 0.78758 | 0.81027
2 | 1 | ANSLu | 1 | 19 | 10 | 267822 | 0.08601 | 0.84966 | 0.81359
2 | 1 | ANSLu | 1 | 14 | 12 | 267823 | 0.08806 | 0.81469 | 0.81172
2 | 1 | ANSLu | 1 | 10 | 7 | 267824 | 0.085693 | 0.73947 | 0.8077
2 | 1 | ANSLu | 1 | 8 | 3 | 267825 | 0.085908 | 0.77035 | 0.80935
2 | 1 | ANSLu | 1 | 16 | 4 | 267826 | 0.084414 | 0.78768 | 0.81028
2 | 1 | ANSLu | 1 | 1 | 10 | 267827 | 0.085811 | 0.86071 | 0.81418
2 | 1 | ANSLu | 1 | 5 | 9 | 267828 | 0.087926 | 0.81016 | 0.81148
2 | 1 | ANSLu | 1 | 7 | 11 | 267829 | 0.083746 | 0.74776 | 0.80814
2 | 1 | ANSLu | 1 | 11 | 2 | 267830 | 0.086994 | 0.8473 | 0.81347
2 | 1 | ANSLu | 1 | 6 | 9 | 267831 | 0.086927 | 0.76202 | 0.80891
2 | 1 | ANSLu | 1 | 20 | 4 | 267832 | 0.087125 | 0.84678 | 0.81344
2 | 1 | ANSLu | 1 | 4 | 1 | 267833 | 0.085818 | 0.85177 | 0.81371
2 | 1 | ANSLu | 1 | 17 | 12 | 267834 | 0.084276 | 0.81427 | 0.8117
2 | 1 | ANSLu | 1 | 18 | 5 | 267835 | 0.086997 | 0.8535 | 0.8138
2 | 1 | ANSLu | 1 | 2 | 8 | 267836 | 0.085823 | 0.79796 | 0.81083
2 | 1 | ANSLu | 1 | 12 | 7 | 267837 | 0.083368 | 0.80745 | 0.81134
2 | 1 | ANSLu | 1 | 15 | 6 | 267838 | 0.085856 | 0.8092 | 0.81143
Mean | [6 empty] | 0.085799 | 0.81156 | 0.81156
Median | [7 empty] | 0.81222 | 0.81159
SD | [7 empty] | 0.036868 | 0.0019717
RSD | [7 empty] | 4.543% | 0.243%
RSD Upscale Factor | [7 empty] | 0.053479
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 211b (Certified Value 0.798 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.28 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 30g | 50g | 50g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 40g | 50g | 30g | 40g | 3
| 1 | 1 | 0.86071 | 0.76 | 0.81 | 0.79671 | 0.752 | 0.81504 | 0.798 | 0.809 | 0.8 | 0.83 | 0.82 | 0.82 | 0.79 | 0.85 | 0.82417 | 0.797 | 0.81 | 0.766 | 0.774 | 0.818 | 0.81 | 0.797 | 0.759 | 0.75 | 1
| 1 | 2 | 0.79796 | 0.77 | 0.785 | 0.80209 | 0.742 | 0.8112 | 0.806 | 0.812 | 0.79 | 0.81 | 0.82 | 0.84 | 0.81 | 0.84 | 0.83295 | 0.809 | 0.81 | 0.758 | 0.763 | 0.792 | 0.78 | 0.796 | 0.769 | 0.79 | #N/A
| 1 | 3 | 0.78758 | 0.69 | 0.804 | 0.79891 | 0.753 | 0.81792 | 0.789 | 0.812 | 0.79 | 0.83 | 0.82 | 0.84 | 0.8 | 0.83 | 0.80955 | 0.802 | 0.81 | 0.794 | 0.772 | 0.781 | 0.79 | 0.803 | 0.758 | 0.79 | 16
| 1 | 4 | 0.85177 | 0.76 | 0.809 | 0.80257 | 0.746 | 0.81024 | 0.784 | 0.813 | 0.79 | 0.83 | 0.81 | 0.82 | 0.79 | 0.83 | 0.83216 | 0.834 | 0.81 | 0.794 | 0.777 | 0.785 | 0.79 | 0.797 | 0.758 | 0.78 | 0.79791
| 1 | 5 | 0.81016 | 0.75 | 0.784 | 0.78759 | 0.759 | 0.85056 | 0.806 | 0.804 | 0.8 | 0.83 | 0.83 | 0.82 | 0.79 | 0.83 | 0.8106 | 0.805 | 0.81 | 0.758 | 0.819 | 0.818 | 0.78 | 0.805 | 0.766 | 0.79 | 7
| 1 | 6 | 0.76202 | 0.73 | 0.808 | 0.78895 | 0.742 | 0.83136 | 0.796 | 0.816 | 0.8 | 0.82 | 0.82 | 0.82 | 0.79 | 0.82 | 0.81178 | 0.82 | 0.81 | 0.797 | 0.767 | 0.792 | 0.79 | 0.816 | 0.758 | 0.77
|  | 7 | 0.74776
|  | 8 | 0.77035
|  | 9 | 0.84075
|  | 10 | 0.73947
|  | 11 | 0.8473
|  | 12 | 0.80745
|  | 13 | 0.83206
|  | 14 | 0.81469
|  | 15 | 0.8092
|  | 16 | 0.78768
|  | 17 | 0.81427
|  | 18 | 0.8535
|  | 19 | 0.84966
|  | 20 | 0.84678
| Mean |  | 0.81156 | 0.74333 | 0.8 | 0.79614 | 0.749 | 0.82272 | 0.7965 | 0.811 | 0.795 | 0.825 | 0.82 | 0.82667 | 0.795 | 0.83333 | 0.8202 | 0.81117 | 0.81 | 0.77783 | 0.77867 | 0.79767 | 0.79 | 0.80233 | 0.76133 | 0.77833
| Median |  | 0.81222 | 0.755 | 0.806 | 0.79781 | 0.749 | 0.81648 | 0.797 | 0.812 | 0.795 | 0.83 | 0.82 | 0.82 | 0.79 | 0.83 | 0.81798 | 0.807 | 0.81 | 0.78 | 0.773 | 0.792 | 0.79 | 0.8 | 0.7585 | 0.785
| Std Dev. |  | 0.036868 | 0.029439 | 0.012182 | 0.0064731 | 0.0068118 | 0.015622 | 0.0088938 | 0.0040988 | 0.0054772 | 0.0083666 | 0.0063246 | 0.010328 | 0.0083666 | 0.010328 | 0.010934 | 0.013615 | 1.2162e-16 | 0.019062 | 0.020383 | 0.016305 | 0.010954 | 0.0076333 | 0.0048854 | 0.016021
| Rel.Std.Dev. |  | 4.543% | 3.96% | 1.523% | 0.8131% | 0.9094% | 1.899% | 1.117% | 0.5054% | 0.689% | 1.014% | 0.7713% | 1.249% | 1.052% | 1.239% | 1.333% | 1.678% | 1.501e-14% | 2.451% | 2.618% | 2.044% | 1.387% | 0.9514% | 0.6417% | 2.058%
| PDM3 |  | 1.71% | -6.84% | 0.2621% | -0.2221% | -6.13% | 3.11% | -0.1765% | 1.641% | -0.3645% | 3.395% | 2.769% | 3.604% | -0.3645% | 4.44% | 2.794% | 1.662% | 1.515% | -2.516% | -2.412% | -0.03029% | -0.9911% | 0.5546% | -4.584% | -2.453%
| Z-Score (Absolute) |  | NA | 1.93 | 0.08 | 0.05 | 1.73 | 0.89 | 0.04 | 0.47 | 0.09 | 0.97 | 0.79 | 1.03 | 0.09 | 1.27 | 0.8 | 0.48 | 0.44 | 0.7 | 0.67 | 0 | 0.27 | 0.17 | 1.29 | 0.69
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 211b (Certified Value 0.8 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | 6.13 | 6.14 | 6.15 | 6.16 | 6.17 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 360±15 g | 3
| 1 | 1 | 0.86071 | 0.83 | 0.8 | 0.874 | 0.75 | 0.77 | 0.781 | 0.809 | 0.82 | 0.791 | 0.8 | 0.777 | 0.768 | 0.795 | 0.741 | 0.805 | 0.78 | 0.803 | 1
| 1 | 2 | 0.79796 | 0.84 | 0.77 | 0.814 | 0.76 | 0.79 | 0.804 | 0.78 | 0.788 | 0.8 | 0.77 | 0.81 | 0.807 | 0.808 | 0.822 | 0.805 | 0.77 | 0.814 | #N/A
| 1 | 3 | 0.78758 | [16 empty] | 0.921 | 16
| 1 | 4 | 0.85177 | [16 empty] | 0.824 | 0.79993
| 1 | 5 | 0.81016 | 0.83 | 0.85 | 0.9 | 0.75 | 0.78 | 0.822 | 0.812 | 0.89 | 0.774 | 0.79 | 0.8 | 0.733 | 0.802 | 0.836 | 0.805 | 0.82 | 0.763 | 9
| 1 | 6 | 0.76202 | 0.71 | 0.83 | 0.823 | 0.69 | 0.81 | 0.696 | 0.784 | 0.783 | 0.768 | 0.8 | 0.781 | 0.835 | 0.784 | 0.857 | 0.805 | 0.8 | 0.856
| 1 | 7 | 0.74776 | [16 empty] | 0.78
| 1 | 8 | 0.77035 | [16 empty] | 0.833
| 1 | 9 | 0.84075 | 0.75 | 0.82 | 0.802 | 0.81 | 0.8 | 0.847 | 0.79 | 0.793 | 0.781 | 0.77 | 0.849 | 0.783 | 0.817 | 0.843 | 0.796 | 0.79 | 0.841
| 1 | 10 | 0.73947 | 0.75 | 0.8 | 0.826 | 0.82 | 0.79 | 0.805 | 0.84 | 0.799 | 0.79 | 0.83 | 0.818 | 0.78 | 0.812 | 0.767 | 0.796 | 0.86 | 0.924
| 1 | 11 | 0.8473 | [16 empty] | 0.836
| 1 | 12 | 0.80745 | [16 empty] | 0.833
|  | 13 | 0.83206
|  | 14 | 0.81469
|  | 15 | 0.8092
|  | 16 | 0.78768
|  | 17 | 0.81427
|  | 18 | 0.8535
|  | 19 | 0.84966
|  | 20 | 0.84678
| Mean |  | 0.81156 | 0.785 | 0.81167 | 0.83983 | 0.76333 | 0.79 | 0.7925 | 0.8025 | 0.81217 | 0.784 | 0.79333 | 0.80583 | 0.78433 | 0.803 | 0.811 | 0.802 | 0.80333 | 0.83567
| Median |  | 0.81222 | 0.79 | 0.81 | 0.8245 | 0.755 | 0.79 | 0.8045 | 0.7995 | 0.796 | 0.7855 | 0.795 | 0.805 | 0.7815 | 0.805 | 0.829 | 0.805 | 0.795 | 0.833
| Std Dev. |  | 0.036868 | 0.055045 | 0.027869 | 0.038369 | 0.047188 | 0.014142 | 0.052103 | 0.022554 | 0.040236 | 0.011883 | 0.022509 | 0.026498 | 0.034651 | 0.012066 | 0.046308 | 0.0046476 | 0.03266 | 0.048235
| Rel.Std.Dev. |  | 4.543% | 7.012% | 3.434% | 4.569% | 6.182% | 1.79% | 6.574% | 2.811% | 4.954% | 1.516% | 2.837% | 3.288% | 4.418% | 1.503% | 5.71% | 0.5795% | 4.066% | 5.772%
| PDM3 |  | 1.453% | -1.867% | 1.467% | 4.988% | -4.575% | -1.241% | -0.929% | 0.3211% | 1.53% | -1.992% | -0.8248% | 0.7378% | -1.95% | 0.3836% | 1.384% | 0.2586% | 0.4253% | 4.467%
| Z-Score (Absolute) |  | NA | 1.22 | 0.64 | 2.6 | 2.73 | 0.87 | 0.7 | 0 | 0.67 | 1.29 | 0.64 | 0.23 | 1.27 | 0.03 | 0.59 | 0.03 | 0.06 | 2.31
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 211b (Certified Value 0.75 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*MS | AR*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 50g | 30g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 10g | 30g | 15g | 10g | 15g | 25g | 30g | 10g | 40g | 25g | 3
| 1 | 1 | 0.86071 | 0.73 | 0.78 | 0.84169 | 0.77 | 0.766 | 0.78 | 0.789 | 0.75 | 0.75 | 0.78 | < 0.001 | 0.73717 | 0.739 | 0.74 | 0.7405 | 0.738 | 0.7149 | 0.828 | 0.74 | 0.76 | 0.754 | 0.71 | 1
| 1 | 2 | 0.79796 | 0.71 | 0.79 | 0.82712 | 0.75 | 0.769 | 0.77 | 0.789 | 0.74 | 0.748 | 0.756 | 0.732 | 0.75213 | 0.706 | 0.72 | 0.7019 | 0.762 | 0.7416 | 0.807 | 0.73 | 0.77 | 0.742 | 0.78 | #N/A
| 1 | 3 | 0.78758 | 0.69 | 0.8 | 0.82149 | 0.8 | 0.762 | 0.77 | 0.794 | 0.74 | 0.749 | 0.765 | 0.721 | 0.71098 | 0.726 | 0.73 | 0.7542 | 0.764 | 0.7215 | 0.815 | 0.8 | 0.72 | 0.756 | 0.74 | 16
| 1 | 4 | 0.85177 | 0.69 | 0.77 | 0.76991 | 0.77 | 0.757 | 0.78 | 0.791 | 0.73 | 0.746 | 0.734 | 0.703 | 0.74209 | 0.739 | 0.73 | 0.7585 | 0.753 | 0.7326 | 0.811 | 0.78 | 0.78 | 0.77 | 0.73 | 0.74967
| 1 | 5 | 0.81016 | 0.67 | 0.79 | 0.83968 | 0.8 | 0.771 | 0.78 | 0.784 | 0.73 | 0.731 | 0.76 | 0.704 | 0.7138 | 0.783 | 0.73 | 0.7381 | 0.761 | 0.7569 | 0.812 | 0.75 | 0.75 | 0.753 | 0.78 | 11
| 1 | 6 | 0.76202 | 0.73 | 0.8 | 0.76966 | 0.78 | 0.796 | 0.79 | 0.767 | 0.75 | 0.752 | 0.756 | 0.709 | 0.72385 | 0.759 | 0.74 | 0.7321 | 0.738 | 0.7556 | 0.818 | 0.7 | 0.7 | 0.754 | 0.75
|  | 7 | 0.74776
|  | 8 | 0.77035
|  | 9 | 0.84075
|  | 10 | 0.73947
|  | 11 | 0.8473
|  | 12 | 0.80745
|  | 13 | 0.83206
|  | 14 | 0.81469
|  | 15 | 0.8092
|  | 16 | 0.78768
|  | 17 | 0.81427
|  | 18 | 0.8535
|  | 19 | 0.84966
|  | 20 | 0.84678
| Mean |  | 0.81156 | 0.70333 | 0.78833 | 0.81159 | 0.77833 | 0.77017 | 0.77833 | 0.78567 | 0.74 | 0.746 | 0.7585 | 0.7138 | 0.73 | 0.742 | 0.73167 | 0.73755 | 0.75267 | 0.73718 | 0.81517 | 0.75 | 0.74667 | 0.75483 | 0.74833
| Median |  | 0.81222 | 0.7 | 0.79 | 0.8243 | 0.775 | 0.7675 | 0.78 | 0.789 | 0.74 | 0.7485 | 0.758 | 0.709 | 0.73051 | 0.739 | 0.73 | 0.7393 | 0.757 | 0.7371 | 0.8135 | 0.745 | 0.755 | 0.754 | 0.745
| Std Dev. |  | 0.036868 | 0.024221 | 0.01169 | 0.033256 | 0.019408 | 0.013615 | 0.0075277 | 0.0097091 | 0.0089443 | 0.0076158 | 0.014963 | 0.012438 | 0.016427 | 0.026608 | 0.0075277 | 0.020128 | 0.011961 | 0.017384 | 0.0073052 | 0.035777 | 0.030768 | 0.0089536 | 0.027869
| Rel.Std.Dev. |  | 4.543% | 3.444% | 1.483% | 4.098% | 2.494% | 1.768% | 0.9672% | 1.236% | 1.209% | 1.021% | 1.973% | 1.742% | 2.25% | 3.586% | 1.029% | 2.729% | 1.589% | 2.358% | 0.8962% | 4.77% | 4.121% | 1.186% | 3.724%
| PDM3 |  | 8.255% | -6.181% | 5.158% | 8.26% | 3.824% | 2.734% | 3.824% | 4.802% | -1.29% | -0.4894% | 1.178% | -4.785% | -2.623% | -1.023% | -2.401% | -1.617% | 0.3999% | -1.665% | 8.737% | 0.04421% | -0.4004% | 0.6889% | -0.1781%
| Z-Score (Absolute) |  | NA | 2.1 | 1.79 | 2.86 | 1.33 | 0.96 | 1.33 | 1.67 | 0.42 | 0.14 | 0.43 | 7.06 | 0.88 | 0.33 | 0.8 | 0.53 | 0.16 | 0.55 | 3.02 | 0.04 | 0.11 | 0.26 | 0.04
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 211b (Certified Value 0.776 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | 1.23 | 1.27 | 1.29 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 30g | 200g | 50g | 200g | 50g | 3
| 1 | 1 | 0.86071 | 0.74 | 0.75 | 0.69892 | 0.873 | 0.82369 | 0.8 | 0.81 | 0.78 | 0.83 | 0.73 | 0.57 | 0.82 | 0.7957 | 0.74 | 0.7 | 0.8 | 0.784 | 0.737 | 0.75 | 0.78 | 1
| 1 | 2 | 0.79796 | 0.78 | 0.78 | 0.71331 | 0.884 | 0.77125 | 0.76 | 0.832 | 0.77 | 0.82 | 0.77 | 0.32 | 0.81 | 0.81302 | 0.77 | 0.7 | 0.79 | 0.819 | 0.724 | 0.79 | 0.78 | #N/A
| 1 | 3 | 0.78758 | 0.79 | 0.82 | 0.70835 | 0.882 | 0.75082 | 0.76 | 0.813 | 0.79 | 0.84 | 0.74 | 0.73 | 0.81 | 0.79866 | 0.74 | 0.8 | 0.77 | 0.8 | 0.743 | 0.76 | 0.77 | 16
| 1 | 4 | 0.85177 | 0.74 | 0.77 | 0.72628 | 0.88 | 0.75523 | 0.76 | 0.853 | 0.78 | 0.82 | 0.77 | 0.7 | 0.82 | 0.80659 | 0.75 | 0.7 | 0.78 | 0.772 | 0.745 | 0.76 | 0.75 | 0.77639
| 1 | 5 | 0.81016 | 0.72 | 0.77 | 0.70457 | 0.869 | 0.82034 | 0.76 | 0.816 | 0.78 | 0.83 | 0.72 | 0.74 | 0.79 | 0.80356 | 0.77 | 0.8 | 0.78 | 0.793 | 0.736 | 0.77 | 0.77 | 13
| 1 | 6 | 0.76202 | 0.73 | 0.83 | 0.70795 | 0.889 | 0.7239 | 0.8 | 0.838 | 0.78 | 0.83 | 0.76 | 0.66 | 0.8 | 0.81452 | 0.76 | 0.8 | 0.77 | 0.827 | 0.743 | 0.76 | 0.76
|  | 7 | 0.74776
|  | 8 | 0.77035
|  | 9 | 0.84075
|  | 10 | 0.73947
|  | 11 | 0.8473
|  | 12 | 0.80745
|  | 13 | 0.83206
|  | 14 | 0.81469
|  | 15 | 0.8092
|  | 16 | 0.78768
|  | 17 | 0.81427
|  | 18 | 0.8535
|  | 19 | 0.84966
|  | 20 | 0.84678
| Mean |  | 0.81156 | 0.75 | 0.78667 | 0.7099 | 0.8795 | 0.7742 | 0.77333 | 0.827 | 0.78 | 0.82833 | 0.74833 | 0.62 | 0.80833 | 0.80534 | 0.755 | 0.75 | 0.78167 | 0.79917 | 0.738 | 0.765 | 0.76833
| Median |  | 0.81222 | 0.74 | 0.775 | 0.70815 | 0.881 | 0.76324 | 0.76 | 0.824 | 0.78 | 0.83 | 0.75 | 0.68 | 0.81 | 0.80507 | 0.755 | 0.75 | 0.78 | 0.7965 | 0.74 | 0.76 | 0.77
| Std Dev. |  | 0.036868 | 0.028284 | 0.031411 | 0.0093282 | 0.0073417 | 0.040059 | 0.020656 | 0.0169 | 0.0063246 | 0.0075277 | 0.02137 | 0.15937 | 0.01169 | 0.0075584 | 0.013784 | 0.054772 | 0.01169 | 0.020856 | 0.007746 | 0.013784 | 0.01169
| Rel.Std.Dev. |  | 4.543% | 3.771% | 3.993% | 1.314% | 0.8348% | 5.174% | 2.671% | 2.043% | 0.8108% | 0.9088% | 2.856% | 25.71% | 1.446% | 0.9385% | 1.826% | 7.303% | 1.496% | 2.61% | 1.05% | 1.802% | 1.522%
| PDM3 |  | 4.53% | -3.399% | 1.324% | -8.564% | 13.28% | -0.2814% | -0.3935% | 6.519% | 0.4651% | 6.691% | -3.614% | -20.14% | 4.115% | 3.729% | -2.755% | -3.399% | 0.6798% | 2.934% | -4.945% | -1.467% | -1.038%
| Z-Score (Absolute) |  | NA | 0.65 | 0.34 | 1.74 | 2.84 | 0 | 0.02 | 1.43 | 0.16 | 1.46 | 0.7 | 4.17 | 0.92 | 0.84 | 0.52 | NA | 0.2 | 0.67 | 0.98 | 0.25 | 0.16
Results from laboratory 1.16 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 211b (Certified Value 0.334 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.35 | 0.31536 | 0.4 | < 0.5 | 0.33 | 0.35 | 0.34 | 0.31 | 0.34 | 0.3729 | < 0.3 | 0.3 | 0.309 | 0.35 | 0.294 | 0.32 | < 0.5 | 0.4 | 0.652 |  | 1
| 1 | 2 | 0.35 | 0.33114 | 0.4 | < 0.5 | 0.32 | 0.36 | 0.33 | 0.32 | 0.33 | 0.3817 | < 0.3 | 0.3 | 0.294 | 0.34 | 0.329 | 0.34 | < 0.5 | 0.4 | 0.619 |  | 22
| 1 | 3 | 0.34 | 0.29278 | 0.3 | < 0.5 | 0.35 | 0.32 | 0.33 | 0.31 | 0.32 | 0.37839 | < 0.3 | 0.3 | 0.262 | 0.35 | 0.312 | 0.33 | < 0.5 | 0.4 | 0.727 |  | 16
| 1 | 4 | 0.35 | 0.30083 | < 0.3 | < 0.5 | 0.32 | 0.36 | 0.31 | 0.31 | 0.35 | 0.35229 | < 0.3 | 0.4 | 0.319 | 0.32 | 0.305 | 0.35 | < 0.5 | 0.4 | 3.08 |  | 0.33357
| 1 | 5 | 0.35 | 0.24328 | < 0.3 | < 0.5 | 0.34 | 0.38 | 0.35 | 0.31 | 0.33 | 0.37024 | < 0.3 | 0.3 | 0.313 | 0.35 | 0.304 | 0.35 | < 0.5 | 0.4 | 0.772 |  | 15
| 1 | 6 | 0.36 | 0.3062 | 0.4 | < 0.5 | 0.34 | 0.35 | 0.34 | 0.33 | 0.34 | 0.35812 | < 0.3 | 0.3 | 0.309 | 0.35 | 0.321 | 0.36 | < 0.5 | 0.4 | 0.65
| Mean |  | 0.35 | 0.29826 | 0.375 |  | 0.33333 | 0.35333 | 0.33333 | 0.315 | 0.335 | 0.36894 |  | 0.31667 | 0.301 | 0.34333 | 0.31083 | 0.34167 |  | 0.4 | 1.0833
| Median |  | 0.35 | 0.30352 | 0.4 |  | 0.335 | 0.355 | 0.335 | 0.31 | 0.335 | 0.37157 |  | 0.3 | 0.309 | 0.35 | 0.3085 | 0.345 |  | 0.4 | 0.6895
| Std Dev. |  | 0.0063246 | 0.029984 | 0.05 |  | 0.012111 | 0.019664 | 0.013663 | 0.0083666 | 0.010488 | 0.011522 |  | 0.040825 | 0.020813 | 0.012111 | 0.01264 | 0.01472 |  | 6.0809e-17 | 0.9798
| Rel.Std.Dev. |  | 1.807% | 10.05% | 13.33% |  | 3.633% | 5.565% | 4.099% | 2.656% | 3.131% | 3.123% |  | 12.89% | 6.915% | 3.527% | 4.066% | 4.308% |  | 1.52e-14% | 90.44%
| PDM3 |  | 4.926% | -10.58% | 12.42% |  | -0.07079% | 5.925% | -0.07079% | -5.567% | 0.4289% | 10.6% |  | -5.067% | -9.764% | 2.927% | -6.816% | 2.427% |  | 19.92% | 224.8%
| Z-Score (Absolute) |  | 0.56 | 1.18 | NA | NA | 0 | 0.67 | 0 | 0.62 | 0.06 | 1.2 | 6.18 | NA | 1.09 | 0.34 | 0.76 | 0.28 | 2.81 | NA | 25.29
Results from laboratories 1.07, 1.08, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 211b (Certified Value 6.94 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.1429 | 7.0235 | 7.2578 | 6.77 | 7.05 | 6.86 | 6.74 | 6.92 | 7.02 | 7.04 | 7.0966 | 6.7598 | 7.07 | 7.4 | 6.6925 | 7.16 | 6.68 | 6.76 | 6.78 | 7.2409 | 1
| 1 | 2 | 7.3216 | 6.8501 | 7.2723 | 6.65 | 7.12 | 6.93 | 6.81 | 6.83 | 6.94 | 7 | 7.2967 | 6.7167 | 6.93 | 7.26 | 6.6594 | 7.31 | 7.13 | 6.78 | 6.66 | 7.3043 | #N/A
| 1 | 3 | 7.2546 | 6.751 | 7.2562 | 6.68 | 7.1 | 6.78 | 6.69 | 6.77 | 6.99 | 7.28 | 7.0803 | 6.7242 | 6.98 | 7.13 | 6.8168 | 7.15 | 6.27 | 6.78 | 6.79 | 7.1973 | 16
| 1 | 4 | 7.3777 | 6.6744 | 7.2752 | 6.76 | 7 | 6.75 | 6.72 | 6.53 | 7.13 | 7.08 | 7.3158 | 6.5984 | 6.71 | 7.12 | 6.7744 | 6.98 | 6.42 | 6.71 | 6.67 | 7.1278 | 6.9361
| 1 | 5 | 7.175 | 6.6522 | 7.2496 | 6.55 | 7.26 | 6.8 | 6.97 | 6.83 | 7.08 | 6.86 | 7.3787 | 6.7024 | 6.89 | 7.09 | 6.6813 | 7.21 | 6.48 | 6.79 | 6.69 | 7.1259 | 16
| 1 | 6 | 7.3649 | 6.777 | 7.2549 | 6.67 | 7.06 | 6.84 | 6.72 | 6.76 | 6.94 | 7.05 | 7.2573 | 6.7381 | 6.98 | 7.53 | 6.8105 | 7.05 | 6.46 | 6.76 | 6.75 | 7.2741
| Mean |  | 7.2728 | 6.788 | 7.261 | 6.68 | 7.0983 | 6.8267 | 6.775 | 6.7733 | 7.0167 | 7.0517 | 7.2376 | 6.7066 | 6.9267 | 7.255 | 6.7391 | 7.1433 | 6.5733 | 6.7633 | 6.7233 | 7.2117
| Median |  | 7.2881 | 6.764 | 7.257 | 6.675 | 7.08 | 6.82 | 6.73 | 6.8 | 7.005 | 7.045 | 7.277 | 6.7204 | 6.955 | 7.195 | 6.7334 | 7.155 | 6.47 | 6.77 | 6.72 | 7.2191
| Std Dev. |  | 0.098629 | 0.13575 | 0.010292 | 0.080498 | 0.089536 | 0.064395 | 0.10368 | 0.13216 | 0.076594 | 0.136 | 0.12209 | 0.056496 | 0.12209 | 0.17762 | 0.069636 | 0.11656 | 0.30277 | 0.028752 | 0.057155 | 0.074725
| Rel.Std.Dev. |  | 1.356% | 2% | 0.1417% | 1.205% | 1.261% | 0.9433% | 1.53% | 1.951% | 1.092% | 1.929% | 1.687% | 0.8424% | 1.763% | 2.448% | 1.033% | 1.632% | 4.606% | 0.4251% | 0.8501% | 1.036%
| PDM3 |  | 4.854% | -2.135% | 4.684% | -3.692% | 2.339% | -1.578% | -2.323% | -2.347% | 1.162% | 1.666% | 4.347% | -3.309% | -0.1359% | 4.598% | -2.839% | 2.988% | -5.23% | -2.491% | -3.067% | 3.974%
| Z-Score (Absolute) |  | 1.55 | 0.35 | 1.5 | 0.77 | 0.87 | 0.2 | 0.4 | 0.4 | 0.55 | 0.68 | 1.41 | 0.66 | 0.2 | 1.48 | 0.54 | 1.04 | 1.19 | 0.44 | 0.6 | 1.31
| Analytical results for As in OREAS 211b (Certified Value 33.7 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 35.1 | 34.5323 | 34.3 | 34.2 | 35.1 | 32 | 34.8 | 33.1 | 34 | 34.3893 | 27.2 | 33 | 31.2 | 32.3 | 28.4 | 33.1 | 22.13 | 33 | 81.882 |  | 1
| 1 | 2 | 34.6 | 35.9039 | 36.4 | 34.7 | 34.5 | 32.3 | 35.7 | 33.6 | 32.5 | 32.92 | 27.4 | 34 | 32.3 | 32.2 | 29.8 | 32.8 | 25.21 | 34 | 69.999 |  | 10
| 1 | 3 | 32.7 | 34.3999 | 34.9 | 34.1 | 35.2 | 31.1 | 34.5 | 32.7 | 36.6 | 33.2829 | 26 | 35 | 30.7 | 33 | 29 | 33.1 | 23.14 | 35 | 79.889 |  | 16
| 1 | 4 | 34.7 | 34.5135 | 35.3 | 32.6 | 33.2 | 32.5 | 33.6 | 32.8 | 35.1 | 33.0852 | 27.4 | 33 | 30.2 | 32.7 | 28.6 | 31.9 | 20.6 | 33 | 72.422 |  | 33.7495
| 1 | 5 | 33.6 | 34.6149 | 36.3 | 35.3 | 35 | 34 | 34.8 | 34.4 | 36.3 | 34.2242 | 27.2 | 35 | 31.6 | 32.3 | 31.1 | 32.6 | 23.01 | 35 | 76.588 |  | 17
| 1 | 6 | 34 | 33.7106 | 35.4 | 34.1 | 34.8 | 31.9 | 34.5 | 33.5 | 34.7 | 33.6092 | 26.9 | 32 | 33.4 | 33.2 | 28.5 | 32.4 | 20.35 | 34 | 74.947
| Mean |  | 34.1167 | 34.6125 | 35.4333 | 34.1667 | 34.6333 | 32.3 | 34.65 | 33.35 | 34.8667 | 33.5851 | 27.0167 | 33.6667 | 31.5667 | 32.6167 | 29.2333 | 32.65 | 22.4067 | 34 | 75.9545
| Median |  | 34.3 | 34.5229 | 35.35 | 34.15 | 34.9 | 32.15 | 34.65 | 33.3 | 34.9 | 33.446 | 27.2 | 33.5 | 31.4 | 32.5 | 28.8 | 32.7 | 22.57 | 34 | 75.7675
| Std Dev. |  | 0.87502 | 0.71311 | 0.80911 | 0.89815 | 0.74476 | 0.96125 | 0.6775 | 0.62849 | 1.5161 | 0.60651 | 0.53072 | 1.2111 | 1.1535 | 0.41673 | 1.0482 | 0.45935 | 1.8073 | 0.89443 | 4.4717
| Rel.Std.Dev. |  | 2.565% | 2.06% | 2.283% | 2.629% | 2.15% | 2.976% | 1.955% | 1.885% | 4.348% | 1.806% | 1.964% | 3.597% | 3.654% | 1.278% | 3.586% | 1.407% | 8.066% | 2.631% | 5.887%
| PDM3 |  | 1.088% | 2.557% | 4.989% | 1.236% | 2.619% | -4.295% | 2.668% | -1.184% | 3.31% | -0.4871% | -19.95% | -0.2455% | -6.468% | -3.357% | -13.38% | -3.258% | -33.61% | 0.7422% | 125.1%
| Z-Score (Absolute) |  | 0.3 | 0.63 | 1.17 | 0.33 | 0.64 | 0.91 | 0.65 | 0.21 | 0.8 | 0.05 | 4.41 | 0 | 1.39 | 0.7 | 2.94 | 0.67 | 7.47 | 0.22 | 28.05
| Analytical results for B in OREAS 211b (Indicative Value 44.7 ppm) | [21 empty] | Indicative
B | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.08 | [19 empty] | ppm
|  |  | 4A*OES | [19 empty] | 1
1
| 1 | 1 | 46 | [19 empty] | 1
| 1 | 2 | 45 | [19 empty] | 1
| 1 | 3 | 46 | [19 empty] | 16
| 1 | 4 | 43 | [19 empty] | 44.6667
| 1 | 5 | 43 | [19 empty] | 7
| 1 | 6 | 45
| Mean |  | 44.6667
| Median |  | 45
| Std Dev. |  | 1.3663
| Rel.Std.Dev. |  | 3.059%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 211b (Certified Value 151 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 158.5 | 151.7638 | 143.3876 | 151 | 153 | 150 | 150 | 150 | 150 | 150 | 155.1528 | 139 | 150 | 158 | 162.7 | 144 | 187 | 153.2 | 142 | 147.445 | 1
| 1 | 2 | 164.4 | 157.2742 | 142.3283 | 147 | 154 | 150 | 160 | 150 | 150 | 150 | 146.8897 | 142 | 153 | 155 | 161.8 | 148 | 158 | 154.6 | 134 | 148.486 | 24
| 1 | 3 | 172.4 | 151.5725 | 142.981 | 151 | 154 | 150 | 150 | 150 | 150 | 160 | 152.6512 | 140 | 152 | 153 | 163.5 | 147 | 155 | 152.3 | 142 | 148.036 | 16
| 1 | 4 | 163.7 | 152.4869 | 139.8245 | 151 | 152 | 150 | 150 | 140 | 150 | 150 | 149.7929 | 135 | 156 | 152 | 162.9 | 145 | 158 | 150.4 | 134 | 146.768 | 151.4677
| 1 | 5 | 153.3 | 152.63 | 139.3537 | 150 | 155 | 150 | 160 | 150 | 150 | 150 | 147.9504 | 138 | 151 | 153 | 158 | 149 | 162 | 148.7 | 140 | 147.216 | 18
| 1 | 6 | 162.2 | 150.1465 | 139.3109 | 151 | 152 | 150 | 150 | 150 | 150 | 150 | 157.1607 | 137 | 150 | 156 | 161.2 | 151 | 151 | 153.2 | 144 | 147.027
| Mean |  | 162.4167 | 152.6456 | 141.1977 | 150.1667 | 153.3333 | 150 | 153.3333 | 148.3333 | 150 | 151.6667 | 151.5996 | 138.5 | 152 | 154.5 | 161.6833 | 147.3333 | 161.8333 | 152.0667 | 139.3333 | 147.4963
| Median |  | 162.95 | 152.1253 | 141.0764 | 151 | 153.5 | 150 | 150 | 150 | 150 | 150 | 151.2221 | 138.5 | 151.5 | 154 | 162.25 | 147.5 | 158 | 152.75 | 141 | 147.3305
| Std Dev. |  | 6.3835 | 2.4341 | 1.9026 | 1.6021 | 1.2111 | 0 | 5.164 | 4.0825 | 0 | 4.0825 | 4.0846 | 2.429 | 2.2804 | 2.2583 | 1.9813 | 2.582 | 12.8595 | 2.1519 | 4.3205 | 0.64848
| Rel.Std.Dev. |  | 3.93% | 1.595% | 1.347% | 1.067% | 0.7898% | 0% | 3.368% | 2.752% | 0% | 2.692% | 2.694% | 1.754% | 1.5% | 1.462% | 1.225% | 1.752% | 7.946% | 1.415% | 3.101% | 0.4397%
| PDM3 |  | 7.229% | 0.7777% | -6.78% | -0.859% | 1.232% | -0.969% | 1.232% | -2.069% | -0.969% | 0.1313% | 0.08708% | -8.561% | 0.3514% | 2.002% | 6.744% | -2.73% | 6.843% | 0.3954% | -8.011% | -2.622%
| Z-Score (Absolute) |  | 3.18 | 0.3 | 3.08 | 0.43 | 0.5 | 0.48 | 0.5 | 0.97 | 0.48 | 0.01 | 0.01 | 3.88 | 0.11 | 0.85 | 2.97 | 1.27 | 3.01 | 0.13 | 3.63 | 1.22
| Analytical results for Be in OREAS 211b (Certified Value 0.34 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 0.35 | < 0.5 | 0.33624 | 0.4 | < 5 | 0.37 | 0.33 | 0.28 | 0.35 | 0.32 | < 0.5 | 0.33 | < 0.5 | < 1 | 0.36 | < 1 | 0.6 | 9.21 | < 0.5 |  | 1
| 1 | 2 | 0.41 | < 0.5 | 0.35676 | 0.4 | < 5 | 0.37 | 0.34 | 0.31 | 0.34 | 0.32 | < 0.5 | < 0.01 | < 0.5 | < 1 | 0.37 | < 1 | 0.57 | 9.03 | < 0.5 |  | 25
| 1 | 3 | 0.35 | < 0.5 | 0.33219 | 0.4 | < 5 | 0.37 | 0.33 | 0.3 | 0.34 | 0.34 | < 0.5 | 0.33 | < 0.5 | < 1 | 0.38 | < 1 | 0.63 | 9.08 | < 0.5 | 0.095 | 16
| 1 | 4 | 0.4 | < 0.5 | 0.30194 | 0.4 | < 5 | 0.36 | 0.34 | 0.28 | 0.35 | 0.32 | < 0.5 | < 0.01 | < 0.5 | < 1 | 0.36 | < 1 | 0.53 | 9.15 | < 0.5 | 0.127 | 0.34181
| 1 | 5 | 0.36 | < 0.5 | 0.36244 | 0.4 | < 5 | 0.36 | 0.35 | 0.3 | 0.34 | 0.32 | < 0.5 | 0.31 | < 0.5 | < 1 | 0.35 | < 1 | 0.42 | 8.98 | < 0.5 | 0.093 | 19
| 1 | 6 | 0.41 | < 0.5 | 0.32805 | 0.4 | < 5 | 0.36 | 0.35 | 0.31 | 0.34 | 0.32 | < 0.5 | 0.31 | < 0.5 | < 1 | 0.41 | < 1 | 0.43 | 9.12 | < 0.5
| Mean |  | 0.38 |  | 0.33627 | 0.4 |  | 0.365 | 0.34 | 0.29667 | 0.34333 | 0.32333 |  | 0.32 |  |  | 0.37167 |  | 0.53 | 9.095 |  | 0.105
| Median |  | 0.38 |  | 0.33421 | 0.4 |  | 0.365 | 0.34 | 0.3 | 0.34 | 0.32 |  | 0.32 |  |  | 0.365 |  | 0.55 | 9.1 |  | 0.095
| Std Dev. |  | 0.029665 |  | 0.02176 | 6.0809e-17 |  | 0.0054772 | 0.0089443 | 0.013663 | 0.005164 | 0.008165 |  | 0.011547 |  |  | 0.02137 |  | 0.087864 | 0.083126 |  | 0.019079
| Rel.Std.Dev. |  | 7.807% |  | 6.471% | 1.52e-14% |  | 1.501% | 2.631% | 4.605% | 1.504% | 2.525% |  | 3.608% |  |  | 5.75% |  | 16.58% | 0.914% |  | 18.17%
| PDM3 |  | 11.17% |  | -1.62% | 17.02% |  | 6.785% | -0.5289% | -13.21% | 0.4463% | -5.405% |  | -6.38% |  |  | 8.736% |  | 55.06% | 2561% |  | -69.28%
| Z-Score (Absolute) |  | 0.32 | 0.67 | 0.01 | NA | NA | 0.21 | 0.01 | 0.32 | 0.04 | 0.11 | 0.67 | 0.94 | 0.67 | 1.24 | 0.26 | 1.24 | 1.47 | 67 | 0.67 | 1.78
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.08 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 211b (Certified Value 0.041 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 3
3
| 1 | 1 | 0.06 | 0.041406 | 0.03 | 0.22 | 0.03 | 0.06 | 0.04 | 0.03 | 0.04 | < 0.05 | 0.04 | 0.05 | < 0.04 | 0.05 | < 0.04 | 0.03 | < 5 | < 0.1 |  |  | 1
| 1 | 2 | 0.04 | 0.049834 | 0.03 | 0.24 | 0.03 | 0.05 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.06 | < 0.05 | 0.04 | 0.05 | < 0.04 | 0.03 | < 5 | < 0.1 |  |  | 26
| 1 | 3 | 0.04 | 0.038746 | 0.03 | 0.21 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | < 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | < 0.04 | 0.03 | < 5 | < 0.1 |  |  | 16
| 1 | 4 | 0.05 | 0.043769 | 0.03 | 0.25 | 0.04 | 0.05 | 0.03 | 0.03 | 0.04 | < 0.05 | 0.03 | < 0.05 | 0.05 | 0.05 | < 0.04 | 0.03 | < 5 | < 0.1 |  |  | 0.040847
| 1 | 5 | 0.04 | 0.060988 | 0.03 | 0.25 | 0.03 | 0.05 | 0.03 | 0.03 | 0.04 | < 0.05 | 0.1 | 0.05 | 0.04 | 0.05 | 0.04 | 0.03 | < 5 | < 0.1 |  |  | 20
| 1 | 6 | 0.05 | 0.046406 | 0.03 | 0.23 | 0.04 | 0.05 | 0.04 | 0.03 | 0.04 | < 0.05 | 0.07 | < 0.05 | 0.04 | 0.06 | 0.04 | 0.03 | < 5 | < 0.1
| Mean |  | 0.046667 | 0.046858 | 0.03 | 0.23333 | 0.033333 | 0.05 | 0.035 | 0.03 | 0.038333 |  | 0.056667 | 0.05 | 0.042 | 0.051667 | 0.04 | 0.03
| Median |  | 0.045 | 0.045088 | 0.03 | 0.235 | 0.03 | 0.05 | 0.035 | 0.03 | 0.04 |  | 0.05 | 0.05 | 0.04 | 0.05 | 0.04 | 0.03
| Std Dev. |  | 0.008165 | 0.0079218 | 0 | 0.01633 | 0.005164 | 0.0063246 | 0.0054772 | 0 | 0.0040825 |  | 0.02582 | 8.4984e-18 | 0.0044721 | 0.0040825 | 0 | 0
| Rel.Std.Dev. |  | 17.5% | 16.91% | 0% | 6.999% | 15.49% | 12.65% | 15.65% | 0% | 10.65% |  | 45.56% | 1.7e-14% | 10.65% | 7.902% | 0% | 0%
| PDM3 |  | 14.25% | 14.72% | -26.56% | 471.2% | -18.39% | 22.41% | -14.31% | -26.56% | -6.154% |  | 38.73% | 22.41% | 2.823% | 26.49% | -2.074% | -26.56%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 0.67 | 15.6 | 0.4 | 0.93 | 0.27 | 0.67 | 0 | 1.07 | 1.47 | 0.07 | 0 | 1.07 | 0.93 | 0.67 | 196.93 | 0.93
| Analytical results for Ca in OREAS 211b (Certified Value 6.93 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.1038 | 6.8441 | 7.1327 | 6.99 | 6.66 | 7.07 | 7.02 | 6.94 | 6.82 | 6.61 | 6.9205 | 7.5406 | 7.11 | 6.93 | 6.9369 | 6.77 | 6.75 | 6.8 | 6.81 | 7.2111 | 1
| 1 | 2 | 7.3033 | 7.1273 | 7.1428 | 6.46 | 6.73 | 7.06 | 7.18 | 6.76 | 6.7 | 6.61 | 6.9287 | 7.5258 | 7.2 | 6.93 | 6.9139 | 6.73 | 7.05 | 7.02 | 6.86 | 7.2648 | #N/A
| 1 | 3 | 7.1914 | 6.8906 | 7.1249 | 6.73 | 6.71 | 6.96 | 7.01 | 6.6 | 6.67 | 6.82 | 7.0727 | 7.5289 | 7.01 | 6.7 | 7.0561 | 6.65 | 6.07 | 6.96 | 6.82 | 7.1156 | 16
| 1 | 4 | 7.291 | 6.92 | 7.1139 | 6.92 | 6.62 | 6.9 | 7.02 | 6.44 | 6.82 | 6.62 | 7.0643 | 7.4015 | 7.26 | 6.87 | 7.0975 | 6.69 | 6.22 | 6.85 | 6.86 | 7.1298 | 6.9344
| 1 | 5 | 7.1343 | 7.0955 | 7.1186 | 7.53 | 6.86 | 6.94 | 7.33 | 6.65 | 6.84 | 6.58 | 7.1784 | 7.4978 | 6.84 | 6.73 | 6.9577 | 6.78 | 6.48 | 7.02 | 6.9 | 7.0811 | 21
| 1 | 6 | 7.318 | 6.7514 | 7.112 | 6.65 | 6.69 | 7 | 7.1 | 6.62 | 6.75 | 6.63 | 7.0842 | 7.4387 | 6.94 | 7.06 | 7.0567 | 6.88 | 6.2 | 6.8 | 6.82 | 7.1619
| Mean |  | 7.2236 | 6.9382 | 7.1242 | 6.88 | 6.7117 | 6.9883 | 7.11 | 6.6683 | 6.7667 | 6.645 | 7.0415 | 7.4889 | 7.06 | 6.87 | 7.0031 | 6.75 | 6.4617 | 6.9083 | 6.845 | 7.1607
| Median |  | 7.2412 | 6.9053 | 7.1218 | 6.825 | 6.7 | 6.98 | 7.06 | 6.635 | 6.785 | 6.615 | 7.0685 | 7.5118 | 7.06 | 6.9 | 7.0069 | 6.75 | 6.35 | 6.905 | 6.84 | 7.1458
| Std Dev. |  | 0.092919 | 0.1462 | 0.011872 | 0.37094 | 0.082321 | 0.067651 | 0.12617 | 0.16833 | 0.070899 | 0.08735 | 0.099413 | 0.056343 | 0.15962 | 0.1355 | 0.076149 | 0.08025 | 0.3768 | 0.10439 | 0.034496 | 0.067393
| Rel.Std.Dev. |  | 1.286% | 2.107% | 0.1667% | 5.392% | 1.227% | 0.9681% | 1.775% | 2.524% | 1.048% | 1.315% | 1.412% | 0.7524% | 2.261% | 1.972% | 1.087% | 1.189% | 5.831% | 1.511% | 0.504% | 0.9411%
| PDM3 |  | 4.171% | 0.05384% | 2.736% | -0.7849% | -3.212% | 0.7774% | 2.532% | -3.837% | -2.419% | -4.174% | 1.544% | 7.996% | 1.811% | -0.9291% | 0.9908% | -2.66% | -6.818% | -0.3763% | -1.29% | 3.263%
| Z-Score (Absolute) |  | 1.23 | 0.06 | 0.82 | 0.18 | 0.87 | 0.27 | 0.76 | 1.04 | 0.64 | 1.14 | 0.48 | 2.31 | 0.56 | 0.22 | 0.33 | 0.71 | 1.89 | 0.06 | 0.32 | 0.97
| Analytical results for Cd in OREAS 211b (Certified Value 0.45 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 0.45 | 0.50415 | 0.4 | 0.54 | 0.46 | 0.46 | 0.5 | 0.44 | 0.47 | 0.34953 | 0.4 | 0.48 | 0.37 | 0.49 | 0.43 | 0.37 | < 1 | < 0.5 | 2.227 |  | 1
| 1 | 2 | 0.49 | 0.47121 | 0.4 | 0.47 | 0.48 | 0.48 | 0.46 | 0.45 | 0.48 | 0.42994 | 0.43 | 0.47 | 0.35 | 0.48 | 0.42 | 0.41 | < 1 | < 0.5 | 2.359 |  | 27
| 1 | 3 | 0.46 | 0.45214 | 0.5 | 0.45 | 0.48 | 0.44 | 0.47 | 0.41 | 0.46 | 0.43419 | 0.41 | 0.49 | 0.47 | 0.41 | 0.41 | 0.4 | < 1 | < 0.5 | 2.331 |  | 16
| 1 | 4 | 0.45 | 0.43305 | 0.5 | 0.51 | 0.48 | 0.47 | 0.45 | 0.42 | 0.51 | 0.46901 | 0.42 | 0.49 | 0.52 | 0.45 | 0.41 | 0.37 | < 1 | < 0.5 | 2.37 |  | 0.45264
| 1 | 5 | 0.46 | 0.44716 | 0.4 | 0.49 | 0.47 | 0.48 | 0.46 | 0.45 | 0.48 | 0.3832 | 0.43 | 0.46 | 0.41 | 0.46 | 0.42 | 0.33 | < 1 | < 0.5 | 2.359 |  | 22
| 1 | 6 | 0.49 | 0.4329 | 0.4 | 0.47 | 0.47 | 0.48 | 0.48 | 0.45 | 0.48 | 0.37495 | 0.43 | 0.46 | 0.43 | 0.41 | 0.43 | 0.38 | < 1 | < 0.5 | 2.294
| Mean |  | 0.46667 | 0.45677 | 0.43333 | 0.48833 | 0.47333 | 0.46833 | 0.47 | 0.43667 | 0.48 | 0.4068 | 0.42 | 0.475 | 0.425 | 0.45 | 0.42 | 0.37667 |  |  | 2.3233
| Median |  | 0.46 | 0.44965 | 0.4 | 0.48 | 0.475 | 0.475 | 0.465 | 0.445 | 0.48 | 0.40657 | 0.425 | 0.475 | 0.42 | 0.455 | 0.42 | 0.375 |  |  | 2.345
| Std Dev. |  | 0.018619 | 0.0272 | 0.05164 | 0.032506 | 0.008165 | 0.016021 | 0.017889 | 0.017512 | 0.016733 | 0.044737 | 0.012649 | 0.013784 | 0.063166 | 0.034059 | 0.0089443 | 0.028048 |  |  | 0.054628
| Rel.Std.Dev. |  | 3.99% | 5.955% | 11.92% | 6.657% | 1.725% | 3.421% | 3.806% | 4.01% | 3.486% | 11% | 3.012% | 2.902% | 14.86% | 7.569% | 2.13% | 7.446% |  |  | 2.351%
| PDM3 |  | 3.1% | 0.9133% | -4.265% | 7.887% | 4.573% | 3.468% | 3.836% | -3.528% | 6.045% | -10.13% | -7.21% | 4.941% | -6.106% | -0.5824% | -7.21% | -16.78% |  |  | 413.3%
| Z-Score (Absolute) |  | 0.13 | 0.13 | NA | 0.69 | 0.3 | 0.17 | 0.22 | 0.65 | 0.48 | 1.43 | 1.09 | 0.35 | 0.96 | 0.31 | 1.09 | 2.22 | 1 | 5.53 | 48.59
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 211b (Certified Value 11.7 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 11.76 | 11.8501 | 12.3111 | 11.7 | 12.65 | 11.6 | 11.6 | 11.1 | 11.85 | 11.9 | 10.39 | 11.37 | 11.79 | 10.7 | 11.6 | 43.922 |  |  |  |  | 1
| 1 | 2 | 11.57 | 12.0246 | 12.2224 | 11.1 | 12.6 | 11.6 | 12.05 | 11.4 | 12.05 | 12.3 | 10.41 | 11.34 | 11.7 | 11.6 | 11.2 | 43.08 |  |  |  |  | 28
| 1 | 3 | 11.25 | 12.1467 | 12.309 | 11.3 | 12.9 | 10.95 | 12.3 | 11.2 | 12.45 | 12 | 10.63 | 11.56 | 11.69 | 11.7 | 11.2 | 44.03 |  |  |  |  | 16
| 1 | 4 | 11.43 | 11.5833 | 12.1252 | 11.6 | 12.25 | 11.75 | 11.55 | 11.05 | 12 | 11.5 | 11.01 | 11.28 | 11.43 | 11.5 | 11.2 | 43.142 |  |  |  |  | 11.7322
| 1 | 5 | 11.28 | 12.3181 | 12.3072 | 12.1 | 12.6 | 12.1 | 12.15 | 11.4 | 12.45 | 12.1 | 10.45 | 11.32 | 11.75 | 11 | 11.9 | 43.204 |  |  |  |  | 23
| 1 | 6 | 11.4 | 11.3578 | 12.2657 | 11.4 | 12.45 | 11.55 | 11.4 | 11.05 | 12.65 | 11.7 | 12.4 | 11.59 | 11 | 11.8 | 11.3 | 42.457
| Mean |  | 11.4483 | 11.8801 | 12.2568 | 11.5333 | 12.575 | 11.5917 | 11.8417 | 11.2 | 12.2417 | 11.9167 | 10.8817 | 11.41 | 11.56 | 11.3833 | 11.4 | 43.3058
| Median |  | 11.415 | 11.9374 | 12.2864 | 11.5 | 12.6 | 11.6 | 11.825 | 11.15 | 12.25 | 11.95 | 10.54 | 11.355 | 11.695 | 11.55 | 11.25 | 43.173
| Std Dev. |  | 0.19094 | 0.35933 | 0.0732 | 0.35024 | 0.21622 | 0.37339 | 0.3707 | 0.16432 | 0.31689 | 0.28577 | 0.77919 | 0.13145 | 0.30199 | 0.43551 | 0.28983 | 0.58529
| Rel.Std.Dev. |  | 1.668% | 3.025% | 0.5972% | 3.037% | 1.719% | 3.221% | 3.13% | 1.467% | 2.589% | 2.398% | 7.161% | 1.152% | 2.612% | 3.826% | 2.542% | 1.352%
| PDM3 |  | -2.419% | 1.261% | 4.471% | -1.695% | 7.184% | -1.198% | 0.9332% | -4.536% | 4.343% | 1.572% | -7.249% | -2.746% | -1.468% | -2.973% | -2.831% | 269.1%
| Z-Score (Absolute) |  | 0.3 | 0.72 | 1.61 | 0.1 | 2.36 | 0.04 | 0.63 | 0.89 | 1.58 | 0.81 | 1.64 | 0.39 | 0.04 | 0.46 | 0.42 | 75.06
| Analytical results for Co in OREAS 211b (Certified Value 44.4 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 46.3 | 43.681 | 43.1568 | 42 | 43 | 46.3 | 43.3 | 45 | 41.3 | 45.3 | 43.2699 | 46.3 | 43 | 44 | 44.2 | 45.9 | 46.6 | 44.72 | 46 | 44.271 | 1
| 1 | 2 | 45.9 | 45.3932 | 43.5145 | 41 | 43 | 45.4 | 44.1 | 45.2 | 43 | 43 | 43.1976 | 47.4 | 44 | 42.4 | 43.6 | 45.8 | 46.8 | 43.79 | 44 | 44.645 | 14
| 1 | 3 | 45.6 | 44.088 | 43.2136 | 41 | 43 | 46.3 | 41.4 | 45.1 | 41 | 43.7 | 44.7794 | 46 | 44 | 43.5 | 44.1 | 46 | 44.4 | 44.66 | 44 | 44.461 | 16
| 1 | 4 | 46 | 44.3498 | 42.2574 | 41 | 43 | 43.9 | 44.5 | 43.4 | 41.5 | 43.5 | 44.3048 | 45.3 | 45 | 45.6 | 44.2 | 45.5 | 45 | 44.88 | 44 | 43.145 | 44.3712
| 1 | 5 | 45.7 | 44.2186 | 41.6828 | 40 | 44 | 45.8 | 45.4 | 44.7 | 44.2 | 44.4 | 44.9823 | 45.4 | 43 | 44.6 | 43.7 | 46.2 | 46.9 | 43.71 | 45 | 44.908 | 24
| 1 | 6 | 45.7 | 42.96 | 41.3173 | 41 | 43 | 45.7 | 44.5 | 43.6 | 42.8 | 43.3 | 44.1644 | 44.9 | 41 | 45.7 | 43.7 | 46.8 | 44.9 | 43.03 | 45 | 44.401
| Mean |  | 45.8667 | 44.1151 | 42.5237 | 41 | 43.1667 | 45.5667 | 43.8667 | 44.5 | 42.3 | 43.8667 | 44.1164 | 45.8833 | 43.3333 | 44.3 | 43.9167 | 46.0333 | 45.7667 | 44.1317 | 44.6667 | 44.3052
| Median |  | 45.8 | 44.1533 | 42.7071 | 41 | 43 | 45.75 | 44.3 | 44.85 | 42.15 | 43.6 | 44.2346 | 45.7 | 43.5 | 44.3 | 43.9 | 45.95 | 45.8 | 44.225 | 44.5 | 44.431
| Std Dev. |  | 0.2582 | 0.80285 | 0.90463 | 0.63246 | 0.40825 | 0.88919 | 1.3866 | 0.79498 | 1.2394 | 0.84538 | 0.7467 | 0.89759 | 1.3663 | 1.2712 | 0.27869 | 0.44121 | 1.1183 | 0.73396 | 0.8165 | 0.60984
| Rel.Std.Dev. |  | 0.5629% | 1.82% | 2.127% | 1.543% | 0.9457% | 1.951% | 3.161% | 1.786% | 2.93% | 1.927% | 1.693% | 1.956% | 3.153% | 2.87% | 0.6346% | 0.9585% | 2.444% | 1.663% | 1.828% | 1.376%
| PDM3 |  | 3.37% | -0.5771% | -4.164% | -7.598% | -2.715% | 2.694% | -1.137% | 0.2904% | -4.668% | -1.137% | -0.5742% | 3.408% | -2.339% | -0.1604% | -1.024% | 3.746% | 3.145% | -0.5398% | 0.666% | -0.1487%
| Z-Score (Absolute) |  | 1.76 | 0.01 | 1.62 | 3.16 | 0.97 | 1.46 | 0.26 | 0.38 | 1.84 | 0.26 | 0.01 | 1.78 | 0.8 | 0.18 | 0.21 | 1.93 | 1.66 | 0.01 | 0.55 | 0.18
| Analytical results for Cr in OREAS 211b (Certified Value 121 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 138 | 117.0403 | 70 | 139 | 112 | 98 | 118 | 119 | 113 | 113.0422 | 130.3 | 124 | 123 | 121 | 123 | 125 | 137.3 | 140 | 148.256 |  | 1
| 1 | 2 | 133 | 123.8029 | 105 | 135 | 113 | 102 | 116 | 120 | 119 | 108.8612 | 130.2 | 113 | 121 | 114 | 126 | 122 | 132.8 | 140 | 148.601 |  | 30
| 1 | 3 | 144 | 122.2529 | 89 | 141 | 111 | 98 | 110 | 118 | 118 | 115.3231 | 129.9 | 116 | 119 | 116 | 122 | 110 | 127 | 140 | 146.857 |  | 16
| 1 | 4 | 144 | 120.6354 | 100 | 135 | 109 | 172 | 107 | 117 | 115 | 110.5629 | 128.3 | 134 | 115 | 126 | 123 | 111 | 125.1 | 140 | 144.589 |  | 121.3588
| 1 | 5 | 135 | 114.3153 | 76 | 141 | 109 | 105 | 115 | 119 | 119 | 113.2713 | 129.5 | 114 | 118 | 122 | 125 | 126 | 127.1 | 140 | 146.655 |  | 25
| 1 | 6 | 140 | 116.1548 | 97 | 134 | 113 | 98 | 112 | 118 | 119 | 106.9316 | 127.4 | 114 | 112 | 126 | 124 | 112 | 133.1 | 140 | 147.145
| Mean |  | 139 | 119.0336 | 89.5 | 137.5 | 111.1667 | 112.1667 | 113 | 118.5 | 117.1667 | 111.3321 | 129.2667 | 119.1667 | 118 | 120.8333 | 123.8333 | 117.6667 | 130.4 | 140 | 147.0172
| Median |  | 139 | 118.8379 | 93 | 137 | 111.5 | 100 | 113.5 | 118.5 | 118.5 | 111.8025 | 129.7 | 115 | 118.5 | 121.5 | 123.5 | 117 | 129.95 | 140 | 147.001
| Std Dev. |  | 4.5607 | 3.747 | 13.9248 | 3.2094 | 1.8348 | 29.4511 | 4.0988 | 1.0488 | 2.5626 | 3.12 | 1.1673 | 8.3046 | 4 | 4.9967 | 1.472 | 7.4476 | 4.7159 | 0 | 1.4221
| Rel.Std.Dev. |  | 3.281% | 3.148% | 15.56% | 2.334% | 1.651% | 26.26% | 3.627% | 0.8851% | 2.187% | 2.802% | 0.903% | 6.969% | 3.39% | 4.135% | 1.189% | 6.329% | 3.617% | 0% | 0.9673%
| PDM3 |  | 14.54% | -1.916% | -26.25% | 13.3% | -8.398% | -7.574% | -6.888% | -2.356% | -3.454% | -8.262% | 6.516% | -1.806% | -2.768% | -0.433% | 2.039% | -3.042% | 7.45% | 15.36% | 21.14%
| Z-Score (Absolute) |  | 1.96 | 0 | 2.9 | 1.81 | 0.77 | 0.67 | 0.59 | 0.05 | 0.18 | 0.76 | 1 | 0.01 | 0.1 | 0.18 | 0.47 | 0.13 | 1.12 | 2.06 | 2.75
| Analytical results for Cs in OREAS 211b (Certified Value 0.69 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.73 | 0.66043 | 0.66 | 0.72 | 0.68 | 0.64 | 0.62 | 0.72 | 0.69362 | 0.6 | 0.7 | 0.7 | 0.71 | 0.7 | 0.8 | 0.8 |  |  |  |  | 1
| 1 | 2 | 0.7 | 0.66626 | 0.62 | 0.69 | 0.69 | 0.66 | 0.64 | 0.7 | 0.68951 | 0.6 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 |  |  |  |  | 31
| 1 | 3 | 0.68 | 0.68841 | 0.65 | 0.68 | 0.64 | 0.66 | 0.57 | 0.72 | 0.69162 | 0.6 | 0.7 | 0.7 | 0.69 | 0.6 | 0.7 | 0.7 |  |  |  |  | 16
| 1 | 4 | 0.71 | 0.66288 | 0.67 | 0.68 | 0.68 | 0.63 | 0.58 | 0.72 | 0.67786 | 0.6 | 0.7 | 0.7 | 0.74 | 0.6 | 0.7 | 0.7 |  |  |  |  | 0.68548
| 1 | 5 | 0.67 | 0.65727 | 0.71 | 0.71 | 0.7 | 0.65 | 0.63 | 0.74 | 0.67496 | 0.6 | 0.7 | 0.7 | 0.72 | 0.7 | 0.7 | 0.7 |  |  |  |  | 26
| 1 | 6 | 0.69 | 0.68547 | 0.66 | 0.69 | 0.68 | 0.62 | 0.62 | 0.73 | 0.69188 | 0.6 | 0.7 | 0.7 | 0.69 | 0.7 | 0.7 | 0.7
| Mean |  | 0.69667 | 0.67012 | 0.66167 | 0.695 | 0.67833 | 0.64333 | 0.61 | 0.72167 | 0.68657 | 0.6 | 0.7 | 0.7 | 0.70833 | 0.66667 | 0.71667 | 0.71667
| Median |  | 0.695 | 0.66457 | 0.66 | 0.69 | 0.68 | 0.645 | 0.62 | 0.72 | 0.69056 | 0.6 | 0.7 | 0.7 | 0.705 | 0.7 | 0.7 | 0.7
| Std Dev. |  | 0.021602 | 0.013389 | 0.029269 | 0.016432 | 0.020412 | 0.01633 | 0.028284 | 0.013292 | 0.0080341 | 0 | 1.2162e-16 | 1.2162e-16 | 0.019408 | 0.05164 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 3.101% | 1.998% | 4.424% | 2.364% | 3.009% | 2.538% | 4.637% | 1.842% | 1.17% | 0% | 1.737e-14% | 1.737e-14% | 2.74% | 7.746% | 5.696% | 5.696%
| PDM3 |  | 1.631% | -2.241% | -3.475% | 1.388% | -1.043% | -6.149% | -11.01% | 5.278% | 0.159% | -12.47% | 2.118% | 2.118% | 3.333% | -2.745% | 4.549% | 4.549%
| Z-Score (Absolute) |  | 0.55 | 0.48 | 0.8 | 0.48 | 0.16 | 1.51 | 2.79 | 1.51 | 0.16 | NA | NA | NA | 1 | NA | NA | NA
Results from laboratories 1.15, 1.16, 1.17, 1.19, 1.21 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 211b (Certified Value 177 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 182.2 | 172.2225 | 181.585 | 178 | 172 | 178 | 176.5 | 184.5 | 178.5 | 187 | 175.908 | 160.3 | 180 | 187.6 | 172 | 170.6 | 217.4 | 162.2 | 164 | 182.672 | 1
| 1 | 2 | 182.1 | 179.3884 | 181.648 | 177 | 174 | 178.5 | 178.5 | 183.5 | 179.5 | 180 | 173.8735 | 159.3 | 185 | 183.4 | 171.6 | 171.1 | 200.1 | 159.4 | 160 | 184.975 | 16
| 1 | 3 | 174.3 | 177.0588 | 183.0205 | 176 | 181 | 176.5 | 173 | 182.5 | 177 | 186.5 | 173.0573 | 158.7 | 178 | 180 | 172.7 | 166.4 | 198.8 | 163.4 | 164 | 182.696 | 16
| 1 | 4 | 177.7 | 177.0634 | 180.0175 | 177 | 173 | 172 | 177.5 | 174 | 180 | 183.5 | 171.9268 | 153.7 | 183 | 180.4 | 175.8 | 167 | 191.5 | 162.4 | 164 | 177.854 | 177.0683
| 1 | 5 | 178.7 | 177.708 | 179.9965 | 174 | 181 | 176 | 184 | 183 | 180 | 179 | 173.0063 | 157 | 172 | 180.6 | 171.3 | 168.6 | 207.8 | 161 | 166 | 184.186 | 27
| 1 | 6 | 175.8 | 171.4549 | 178.033 | 174 | 170 | 178 | 175.5 | 180 | 176 | 180.5 | 175.0105 | 156.2 | 176 | 187.2 | 173.2 | 174.3 | 195.7 | 160.3 | 168 | 182.102
| Mean |  | 178.4667 | 175.816 | 180.7168 | 176 | 175.1667 | 176.5 | 177.5 | 181.25 | 178.5 | 182.75 | 173.7971 | 157.5333 | 179 | 183.2 | 172.7667 | 169.6667 | 201.8833 | 161.45 | 164.3333 | 182.4142
| Median |  | 178.2 | 177.0611 | 180.8013 | 176.5 | 173.5 | 177.25 | 177 | 182.75 | 179 | 182 | 173.4654 | 157.85 | 179 | 182 | 172.35 | 169.6 | 199.45 | 161.6 | 164 | 182.684
| Std Dev. |  | 3.2327 | 3.2054 | 1.74 | 1.6733 | 4.7081 | 2.4083 | 3.7014 | 3.8568 | 1.6733 | 3.446 | 1.4563 | 2.4039 | 4.7329 | 3.471 | 1.6428 | 2.9432 | 9.3221 | 1.4829 | 2.6583 | 2.4793
| Rel.Std.Dev. |  | 1.811% | 1.823% | 0.9629% | 0.9508% | 2.688% | 1.364% | 2.085% | 2.128% | 0.9374% | 1.886% | 0.838% | 1.526% | 2.644% | 1.895% | 0.9509% | 1.735% | 4.618% | 0.9185% | 1.618% | 1.359%
| PDM3 |  | 0.7897% | -0.7072% | 2.06% | -0.6033% | -1.074% | -0.3209% | 0.2438% | 2.362% | 0.8086% | 3.209% | -1.847% | -11.03% | 1.091% | 3.463% | -2.429% | -4.18% | 14.01% | -8.82% | -7.192% | 3.019%
| Z-Score (Absolute) |  | 0.25 | 0.2 | 0.63 | 0.17 | 0.31 | 0.08 | 0.08 | 0.72 | 0.25 | 0.98 | 0.54 | 3.3 | 0.34 | 1.05 | 0.72 | 1.24 | 4.22 | 2.64 | 2.15 | 0.92
| Analytical results for Dy in OREAS 211b (Certified Value 3.9 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 4.05 | 3.6259 | 4.5032 | 4 | 3.7 | 3.92 | 3.7 | 3.61 | 4.05 | 4.209 | [10 empty] | 1
| 1 | 2 | 3.96 | 3.8524 | 4.4874 | 3.9 | 3.8 | 4 | 3.7 | 3.9 | 3.9 | 4.194 | [10 empty] | 4
| 1 | 3 | 3.98 | 3.9476 | 4.4475 | 3.7 | 3.8 | 3.91 | 3.7 | 3.89 | 3.9 | 4.155 | [10 empty] | 16
| 1 | 4 | 4 | 3.9372 | 4.3888 | 3.9 | 3.8 | 3.95 | 3.6 | 3.93 | 3.9 | 4.153 | [10 empty] | 3.9042
| 1 | 5 | 3.84 | 3.8525 | 4.505 | 4.1 | 3.6 | 3.94 | 3.8 | 3.69 | 3.8 | 4.245 | [10 empty] | 28
| 1 | 6 | 3.99 | 3.7229 | 4.4928 | 3.8 | 4.1 | 3.95 | 3.8 | 3.98 | 3.95 | 4.286
| Mean |  | 3.97 | 3.8231 | 4.4708 | 3.9 | 3.8 | 3.945 | 3.7167 | 3.8333 | 3.9167 | 4.207
| Median |  | 3.985 | 3.8524 | 4.4901 | 3.9 | 3.8 | 3.945 | 3.7 | 3.895 | 3.9 | 4.2015
| Std Dev. |  | 0.070427 | 0.12585 | 0.045255 | 0.14142 | 0.16733 | 0.031464 | 0.075277 | 0.1476 | 0.08165 | 0.051919
| Rel.Std.Dev. |  | 1.774% | 3.292% | 1.012% | 3.626% | 4.403% | 0.7976% | 2.025% | 3.851% | 2.085% | 1.234%
| PDM3 |  | 1.686% | -2.078% | 14.51% | -0.1074% | -2.669% | 1.045% | -4.803% | -1.815% | 0.3195% | 7.756%
| Z-Score (Absolute) |  | 0.52 | 0.72 | 4.73 | 0.07 | 0.91 | 0.31 | 1.61 | 0.63 | 0.07 | 2.51
| Analytical results for Er in OREAS 211b (Certified Value 2.34 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.48 | 2.1757 | 2.7683 | 2.4 | 2.3 | 2.29 | 2.2 | 2.34 | 2.35 | [11 empty] | 1
| 1 | 2 | 2.41 | 2.3047 | 2.837 | 2.5 | 2.3 | 2.22 | 2.1 | 2.28 | 2.3 | [11 empty] | 5
| 1 | 3 | 2.6 | 2.3105 | 2.8253 | 2.5 | 2.3 | 2.3 | 2.2 | 2.35 | 2.35 | [11 empty] | 16
| 1 | 4 | 2.48 | 2.3911 | 2.8647 | 2.5 | 2.5 | 2.31 | 2.2 | 2.33 | 2.25 | [11 empty] | 2.339
| 1 | 5 | 2.43 | 2.304 | 2.8224 | 2.3 | 2.4 | 2.33 | 2.2 | 2.41 | 2.4 | [11 empty] | 29
| 1 | 6 | 2.48 | 2.2204 | 2.7794 | 2.3 | 2.7 | 2.29 | 2.3 | 2.34 | 2.3
| Mean |  | 2.48 | 2.2844 | 2.8162 | 2.4167 | 2.4167 | 2.29 | 2.2 | 2.3417 | 2.325
| Median |  | 2.48 | 2.3043 | 2.8238 | 2.45 | 2.35 | 2.295 | 2.2 | 2.34 | 2.325
| Std Dev. |  | 0.06603 | 0.07587 | 0.036206 | 0.098319 | 0.16021 | 0.037417 | 0.063246 | 0.041673 | 0.05244
| Rel.Std.Dev. |  | 2.663% | 3.321% | 1.286% | 4.068% | 6.629% | 1.634% | 2.875% | 1.78% | 2.256%
| PDM3 |  | 6.03% | -2.333% | 20.4% | 3.322% | 3.322% | -2.094% | -5.941% | 0.1154% | -0.5971%
| Z-Score (Absolute) |  | 1.24 | 0.51 | 4.27 | 0.67 | 0.67 | 0.46 | 1.27 | 0 | 0.15
| Analytical results for Eu in OREAS 211b (Certified Value 0.98 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
2
| 1 | 1 | 0.99 | 0.90454 | 1.1482 | 0.96 | 0.8 | 0.92 | 0.9 | 0.83 | 1.15 | 1.437 | [10 empty] | 1
| 1 | 2 | 0.95 | 0.94128 | 1.1393 | 0.91 | 0.9 | 0.88 | 0.9 | 0.87 | 1.1 | 1.475 | [10 empty] | 35
| 1 | 3 | 0.92 | 0.95317 | 1.0984 | 0.95 | 0.9 | 0.93 | 0.9 | 0.87 | 1.1 | 1.433 | [10 empty] | 16
| 1 | 4 | 0.95 | 0.9941 | 1.1708 | 1.02 | 0.9 | 0.9 | 0.9 | 0.86 | 1.1 | 1.378 | [10 empty] | 0.97963
| 1 | 5 | 0.93 | 0.96063 | 1.1102 | 0.94 | 0.9 | 0.9 | 0.9 | 0.84 | 1.1 | 1.422 | [10 empty] | 30
| 1 | 6 | 0.96 | 0.89931 | 1.1846 | 0.94 | 1.2 | 0.93 | 0.9 | 0.92 | 1.1 | 1.428
| Mean |  | 0.95 | 0.94217 | 1.1419 | 0.95333 | 0.93333 | 0.91 | 0.9 | 0.865 | 1.1083 | 1.4288
| Median |  | 0.95 | 0.94722 | 1.1437 | 0.945 | 0.9 | 0.91 | 0.9 | 0.865 | 1.1 | 1.4305
| Std Dev. |  | 0.024495 | 0.035811 | 0.033449 | 0.036697 | 0.13663 | 0.02 | 0 | 0.031464 | 0.020412 | 0.031135
| Rel.Std.Dev. |  | 2.578% | 3.801% | 2.929% | 3.849% | 14.64% | 2.198% | 0% | 3.637% | 1.842% | 2.179%
| PDM3 |  | -3.025% | -3.824% | 16.57% | -2.684% | -4.726% | -7.108% | -8.129% | -11.7% | 13.14% | 45.85%
| Z-Score (Absolute) |  | 0.02 | 0.1 | 2 | 0.02 | NA | 0.44 | NA | 0.91 | 1.65 | 5.01
Results from laboratories 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 211b (Certified Value 7.97 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.87 | 7.9198 | 7.8637 | 7.83 | 8.04 | 7.88 | 7.94 | 8.21 | 7.93 | 7.83 | 8.5437 | 8.2904 | 8.14 | 7.85 | 7.7756 | 7.66 | 8.16 | 7.94 | 8.19 | 8.6567 | 1
| 1 | 2 | 8.12 | 7.8765 | 7.8636 | 7.79 | 8.18 | 7.92 | 8.09 | 8.06 | 7.82 | 7.82 | 8.4129 | 8.2636 | 7.9 | 7.75 | 7.764 | 7.61 | 8.32 | 8.02 | 8.15 | 8.7514 | #N/A
| 1 | 3 | 8.02 | 7.7546 | 7.8662 | 7.72 | 8.4 | 7.75 | 7.91 | 8.04 | 7.9 | 8.08 | 8.5135 | 8.2457 | 7.98 | 7.6 | 7.9343 | 7.5 | 8.06 | 7.98 | 8.24 | 8.5847 | 16
| 1 | 4 | 8.16 | 7.742 | 7.8755 | 7.75 | 7.9 | 7.71 | 7.98 | 7.73 | 8.01 | 7.89 | 8.475 | 8.1334 | 8.48 | 7.6 | 7.8155 | 7.53 | 8.05 | 7.97 | 8.16 | 8.5607 | 7.9712
| 1 | 5 | 7.95 | 7.7443 | 7.8693 | 7.96 | 8.28 | 7.76 | 8.27 | 8.11 | 8.01 | 7.63 | 8.3963 | 8.2446 | 7.84 | 7.54 | 7.7689 | 7.64 | 8.2 | 8.05 | 8.22 | 8.5383 | 31
| 1 | 6 | 8.14 | 7.8512 | 7.8577 | 7.7 | 8.03 | 7.81 | 7.97 | 8.01 | 7.91 | 7.87 | 8.5463 | 8.1525 | 8.05 | 7.64 | 7.8168 | 7.73 | 7.91 | 7.98 | 8.15 | 8.6463
| Mean |  | 8.0433 | 7.8147 | 7.866 | 7.7917 | 8.1383 | 7.805 | 8.0267 | 8.0267 | 7.93 | 7.8533 | 8.4813 | 8.2217 | 8.065 | 7.6633 | 7.8125 | 7.6117 | 8.1167 | 7.99 | 8.185 | 8.623
| Median |  | 8.07 | 7.8029 | 7.8649 | 7.77 | 8.11 | 7.785 | 7.975 | 8.05 | 7.92 | 7.85 | 8.4943 | 8.2451 | 8.015 | 7.62 | 7.7956 | 7.625 | 8.11 | 7.98 | 8.175 | 8.6155
| Std Dev. |  | 0.11673 | 0.077528 | 0.0060406 | 0.094956 | 0.18357 | 0.081179 | 0.13397 | 0.16133 | 0.072388 | 0.14459 | 0.064978 | 0.0635 | 0.22941 | 0.11501 | 0.063996 | 0.085186 | 0.1418 | 0.038987 | 0.038341 | 0.078358
| Rel.Std.Dev. |  | 1.451% | 0.9921% | 0.07679% | 1.219% | 2.256% | 1.04% | 1.669% | 2.01% | 0.9128% | 1.841% | 0.7661% | 0.7723% | 2.845% | 1.501% | 0.8191% | 1.119% | 1.747% | 0.4879% | 0.4684% | 0.9087%
| PDM3 |  | 0.9044% | -1.963% | -1.32% | -2.253% | 2.096% | -2.086% | 0.6953% | 0.6953% | -0.5174% | -1.479% | 6.399% | 3.142% | 1.176% | -3.863% | -1.991% | -4.511% | 1.824% | 0.2353% | 2.682% | 8.177%
| Z-Score (Absolute) |  | 0.14 | 0.79 | 0.58 | 0.88 | 0.53 | 0.83 | 0.07 | 0.07 | 0.32 | 0.63 | 1.92 | 0.87 | 0.23 | 1.4 | 0.8 | 1.61 | 0.44 | 0.07 | 0.72 | 2.5
| Analytical results for Ga in OREAS 211b (Certified Value 16 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 16.56 | 15.5019 | 18.326 | 16.1 | 16.1 | 16.2 | 16.1 | 16.15 | 15.55 | 16.1 | 16.5805 | 16.72 | 17 | 14.69 | 17.05 | 14.95 | 16.7 | 16.2 |  |  | 1
| 1 | 2 | 16.8 | 16.0833 | 18.387 | 15.4 | 15.9 | 16.15 | 16.4 | 16 | 15.85 | 15.5 | 16.1328 | 16.98 | 17.2 | 14.39 | 16.92 | 14.87 | 14.6 | 16.4 |  |  | 36
| 1 | 3 | 16.08 | 15.3991 | 18.364 | 15.7 | 16.2 | 16.3 | 15.35 | 15.8 | 15.45 | 16.1 | 16.0555 | 16.53 | 17.1 | 14.96 | 17.07 | 14.59 | 15.2 | 16.2 |  |  | 16
| 1 | 4 | 16.38 | 15.6784 | 18.354 | 15.6 | 16 | 15.5 | 16.35 | 15.15 | 14.95 | 16 | 15.9143 | 16.24 | 16.6 | 15.17 | 16.76 | 14.92 | 15.4 | 16 |  |  | 15.9611
| 1 | 5 | 15.96 | 15.7325 | 18.393 | 16.4 | 16.3 | 16.05 | 16.75 | 15.65 | 16.1 | 16.5 | 16.2694 | 16.48 | 16.9 | 14.56 | 16.42 | 15 | 15.2 | 16.2 |  |  | 32
| 1 | 6 | 16.54 | 15.5544 | 18.372 | 16 | 16.4 | 16.1 | 16.3 | 15.75 | 15.85 | 15.65 | 16.4592 | 16.22 | 16.7 | 17.14 | 16.41 | 15.32 | 14.8 | 16.2
| Mean |  | 16.3867 | 15.6583 | 18.366 | 15.8667 | 16.15 | 16.05 | 16.2083 | 15.75 | 15.625 | 15.975 | 16.2353 | 16.5283 | 16.9167 | 15.1517 | 16.7717 | 14.9417 | 15.3167 | 16.2
| Median |  | 16.46 | 15.6164 | 18.368 | 15.85 | 16.15 | 16.125 | 16.325 | 15.775 | 15.7 | 16.05 | 16.2011 | 16.505 | 16.95 | 14.825 | 16.84 | 14.935 | 15.2 | 16.2
| Std Dev. |  | 0.3164 | 0.24039 | 0.024306 | 0.36697 | 0.18708 | 0.28284 | 0.47055 | 0.34496 | 0.40466 | 0.35742 | 0.25154 | 0.29027 | 0.23166 | 1.0133 | 0.29769 | 0.2349 | 0.73869 | 0.12649
| Rel.Std.Dev. |  | 1.931% | 1.535% | 0.1323% | 2.313% | 1.158% | 1.762% | 2.903% | 2.19% | 2.59% | 2.237% | 1.549% | 1.756% | 1.369% | 6.688% | 1.775% | 1.572% | 4.823% | 0.7808%
| PDM3 |  | 2.666% | -1.897% | 15.07% | -0.5918% | 1.183% | 0.5568% | 1.549% | -1.323% | -2.106% | 0.08689% | 1.718% | 3.554% | 5.987% | -5.071% | 5.078% | -6.387% | -4.038% | 1.497%
| Z-Score (Absolute) |  | 0.5 | 0.77 | 3.93 | 0.4 | 0.09 | 0.09 | 0.19 | 0.61 | 0.82 | 0.22 | 0.23 | 0.74 | 1.42 | 1.64 | 1.16 | 2.01 | 1.36 | 0.17
| Analytical results for Gd in OREAS 211b (Certified Value 3.28 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 3.46 | 3.0788 | 3.2 | 3.3 | 3.29 | 3.2 | 3.2 | 3.4 | [12 empty] | 1
| 1 | 2 | 3.37 | 3.2314 | 3 | 3.3 | 3.29 | 3.2 | 3.62 | 3.4 | [12 empty] | 6
| 1 | 3 | 3.3 | 3.3079 | 3.2 | 3.2 | 3.32 | 3.1 | 3.44 | 3.4 | [12 empty] | 16
| 1 | 4 | 3.44 | 3.2318 | 3.1 | 3.5 | 3.27 | 3.1 | 3.47 | 3.4 | [12 empty] | 3.2839
| 1 | 5 | 3.35 | 3.296 | 3 | 3.3 | 3.41 | 3.2 | 3.29 | 3.4 | [12 empty] | 33
| 1 | 6 | 3.34 | 3.0294 | 3 | 3.7 | 3.26 | 3.2 | 3.51 | 3.4
| Mean |  | 3.3767 | 3.1959 | 3.0833 | 3.3833 | 3.3067 | 3.1667 | 3.4217 | 3.4
| Median |  | 3.36 | 3.2316 | 3.05 | 3.3 | 3.29 | 3.2 | 3.455 | 3.4
| Std Dev. |  | 0.061536 | 0.1154 | 0.098319 | 0.18348 | 0.05465 | 0.05164 | 0.1525 | 0
| Rel.Std.Dev. |  | 1.822% | 3.611% | 3.189% | 5.423% | 1.653% | 1.631% | 4.457% | 0%
| PDM3 |  | 2.826% | -2.679% | -6.106% | 3.029% | 0.6945% | -3.569% | 4.196% | 3.537%
| Z-Score (Absolute) |  | 0.34 | 1.42 | 2.52 | 0.41 | 0.34 | 1.71 | 0.78 | 0.57
| Analytical results for Ge in OREAS 211b (Indicative Value 0.19 ppm) | [21 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.6 | < 0.2 | 0.1838 | 0.4 | 0.08 | 0.13 | 0.09 | 0.08 | 0.21 | 0.25258 | NA | 1.3 | 0.26 | 0.9 | 1.54 | [5 empty] | 1
| 1 | 2 | 1.6 | < 0.2 | 0.1455 | 0.3 | 0.08 | 0.11 | 0.13 | 0.1 | 0.18 | 0.44877 | NA | 1.3 | 0.3 | 0.9 | 1.36 | [5 empty] | 2
| 1 | 3 | 1.6 | < 0.2 | 0.2109 | 0.5 | 0.09 | 0.11 | 0.13 | 0.1 | 0.21 | 0.42999 | NA | 1.3 | 0.19 | 0.8 | 1.41 | [5 empty] | 16
| 1 | 4 | 1.7 | < 0.2 | 0.2009 | 0.2 | 0.09 | 0.1 | 0.11 | 0.09 | 0.22 | 0.59196 | NA | 1.4 | 0.24 | 1 | 1.41 | [5 empty] | 0.19376
| 1 | 5 | 1.5 | < 0.2 | 0.1845 | 0.2 | 0.08 | 0.11 | 0.1 | 0.1 | 0.22 | 0.53445 | NA | 1.3 | 0.15 | 0.7 | 1.39 | [5 empty] | 8
| 1 | 6 | 1.6 | < 0.2 | 0.1714 | 0.4 | 0.08 | 0.11 | 0.1 | 0.11 | 0.16 | 0.33093 | NA | 1.3 | 0.22 | 1 | 1.36
| Mean |  | 1.6 |  | 0.18283 | 0.33333 | 0.083333 | 0.11167 | 0.11 | 0.096667 | 0.2 | 0.43145 |  | 1.3167 | 0.22667 | 0.88333 | 1.4117
| Median |  | 1.6 |  | 0.18415 | 0.35 | 0.08 | 0.11 | 0.105 | 0.1 | 0.21 | 0.43938 |  | 1.3 | 0.23 | 0.9 | 1.4
| Std Dev. |  | 0.063246 |  | 0.022988 | 0.12111 | 0.005164 | 0.0098319 | 0.016733 | 0.010328 | 0.024495 | 0.12552 |  | 0.040825 | 0.052789 | 0.1169 | 0.066758
| Rel.Std.Dev. |  | 3.953% |  | 12.57% | 36.33% | 6.197% | 8.805% | 15.21% | 10.68% | 12.25% | 29.09% |  | 3.101% | 23.29% | 13.23% | 4.729%
| PDM3 |  | 725.8% |  | -5.64% | 72.03% | -56.99% | -42.37% | -43.23% | -50.11% | 3.22% | 122.7% |  | 579.5% | 16.98% | 355.9% | 628.6%
| Z-Score (Absolute) |  | 7.9 | 0.65 | 0.17 | 0.68 | 0.74 | 0.58 | 0.59 | 0.66 | 0.08 | 1.24 | NA | 6.29 | 0.08 | 3.82 | 6.83
| Analytical results for Hf in OREAS 211b (Certified Value 1.55 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 1.61 | 1.5549 | 1.2 | 1.6 | 1.5 | 1.5 | 1.5 | 1.6 | 1.6728 | 1.56 | 1.5 | 1.51 | 1.47 | 1.41 | 1.52 | 1.8 |  |  |  |  | 1
| 1 | 2 | 1.45 | 1.5416 | 1.6 | 1.6 | 1.7 | 1.5 | 1.6 | 1.5 | 1.5431 | 1.6 | 1.5 | 1.57 | 1.5 | 1.38 | 1.51 | 1.6 |  |  |  |  | 22
| 1 | 3 | 1.51 | 1.5352 | 1.2 | 1.6 | 1.6 | 1.5 | 1.5 | 1.6 | 1.6602 | 1.55 | 1.5 | 1.5 | 1.52 | 1.39 | 1.51 | 1.8 |  |  |  |  | 16
| 1 | 4 | 1.51 | 1.4964 | 0.8 | 1.5 | 1.6 | 1.6 | 1.5 | 1.6 | 1.604 | 1.54 | 1.6 | 1.77 | 1.42 | 1.55 | 1.56 | 1.6 |  |  |  |  | 1.5476
| 1 | 5 | 1.48 | 1.4874 | 0.7 | 1.6 | 1.6 | 1.5 | 1.7 | 1.5 | 1.5825 | 1.76 | 1.5 | 1.55 | 1.55 | 1.46 | 1.55 | 1.6 |  |  |  |  | 34
| 1 | 6 | 1.52 | 1.5203 | 1.3 | 1.6 | 1.8 | 1.5 | 1.6 | 1.6 | 1.55 | 1.51 | 1.5 | 1.69 | 1.48 | 1.53 | 1.54 | 1.6
| Mean |  | 1.5133 | 1.5226 | 1.1333 | 1.5833 | 1.6333 | 1.5167 | 1.5667 | 1.5667 | 1.6021 | 1.5867 | 1.5167 | 1.5983 | 1.49 | 1.4533 | 1.5317 | 1.6667
| Median |  | 1.51 | 1.5277 | 1.2 | 1.6 | 1.6 | 1.5 | 1.55 | 1.6 | 1.5932 | 1.555 | 1.5 | 1.56 | 1.49 | 1.435 | 1.53 | 1.6
| Std Dev. |  | 0.053914 | 0.026438 | 0.33267 | 0.040825 | 0.10328 | 0.040825 | 0.08165 | 0.05164 | 0.054733 | 0.089815 | 0.040825 | 0.10815 | 0.044721 | 0.072847 | 0.02137 | 0.10328
| Rel.Std.Dev. |  | 3.563% | 1.736% | 29.35% | 2.578% | 6.323% | 2.692% | 5.212% | 3.296% | 3.416% | 5.661% | 2.692% | 6.766% | 3.001% | 5.012% | 1.395% | 6.197%
| PDM3 |  | -2.212% | -1.611% | -26.77% | 2.312% | 5.543% | -1.996% | 1.235% | 1.235% | 3.524% | 2.527% | -1.996% | 3.281% | -3.719% | -6.089% | -1.027% | 7.697%
| Z-Score (Absolute) |  | 0.66 | 0.49 | 7.65 | 0.63 | 1.55 | 0.6 | 0.32 | 0.32 | 0.97 | 0.69 | 0.6 | 0.9 | 1.09 | 1.76 | 0.32 | 2.16
| Analytical results for Hg in OREAS 211b (Indicative Value 0.12 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.15 | 1.21 | [17 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 1 | 0.13 | 0.18 | [17 empty] | 1
| 1 | 2 | < 1 | 0.14 | 0.11 | [17 empty] | 1
| 1 | 3 | < 1 | 0.14 | 0.13 | [17 empty] | 16
| 1 | 4 | < 1 | 0.13 | 0.09 | [17 empty] | 0.12167
| 1 | 5 | < 1 | 0.13 | 0.11 | [17 empty] | 7
| 1 | 6 | < 1 | 0.13 | 0.11
| Mean |  |  | 0.13333 | 0.12167
| Median |  |  | 0.13 | 0.11
| Std Dev. |  |  | 0.005164 | 0.031252
| Rel.Std.Dev. |  |  | 3.873% | 25.69%
| PDM3 |  |  | 9.589% | -2.887e-13%
| Z-Score (Absolute) |  | 21.19 | 0 | 0.67
| Analytical results for Ho in OREAS 211b (Certified Value 0.79 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.84 | 0.74681 | 0.93077 | 0.8 | 0.7 | 0.74 | 0.7 | 0.78 | 0.8 | [11 empty] | 1
| 1 | 2 | 0.81 | 0.80951 | 0.90512 | 0.8 | 0.8 | 0.76 | 0.7 | 0.83 | 0.76 | [11 empty] | 23
| 1 | 3 | 0.82 | 0.78836 | 0.94148 | 0.8 | 0.8 | 0.75 | 0.7 | 0.82 | 0.78 | [11 empty] | 16
| 1 | 4 | 0.83 | 0.80406 | 0.93437 | 0.9 | 0.8 | 0.76 | 0.7 | 0.79 | 0.76 | [11 empty] | 0.78957
| 1 | 5 | 0.8 | 0.79343 | 0.96165 | 0.9 | 0.7 | 0.73 | 0.7 | 0.79 | 0.78 | [11 empty] | 35
| 1 | 6 | 0.82 | 0.78566 | 0.95642 | 0.8 | 0.8 | 0.78 | 0.8 | 0.82 | 0.76
| Mean |  | 0.82 | 0.78797 | 0.9383 | 0.83333 | 0.76667 | 0.75333 | 0.71667 | 0.805 | 0.77333
| Median |  | 0.82 | 0.79089 | 0.93792 | 0.8 | 0.8 | 0.755 | 0.7 | 0.805 | 0.77
| Std Dev. |  | 0.014142 | 0.022145 | 0.020282 | 0.05164 | 0.05164 | 0.017512 | 0.040825 | 0.020736 | 0.01633
| Rel.Std.Dev. |  | 1.725% | 2.81% | 2.162% | 6.197% | 6.736% | 2.325% | 5.696% | 2.576% | 2.112%
| PDM3 |  | 3.853% | -0.203% | 18.84% | 5.542% | -2.901% | -4.59% | -9.234% | 1.954% | -2.057%
| Z-Score (Absolute) |  | 0.68 | 0.25 | 4.1 | NA | NA | 1.25 | NA | 0.25 | 0.67
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 211b (Certified Value 0.073 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.08 | < 0.1 | 0.25 | 0.076 | 0.075 | 0.073 | 0.07 | 0.077 | 0.075574 | 0.073 | 0.07 | 0.07 | 0.07 | 0.07 | 0.074 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.08 | < 0.1 | 0.24 | 0.077 | 0.079 | 0.076 | 0.071 | 0.075 | 0.075246 | 0.071 | 0.08 | 0.07 | 0.07 | 0.07 | 0.071 | 0.05 |  |  |  |  | 7
| 1 | 3 | 0.08 | < 0.1 | 0.23 | 0.08 | 0.072 | 0.07 | 0.064 | 0.082 | 0.07431 | 0.069 | 0.08 | 0.06 | 0.07 | 0.07 | 0.074 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.07 | < 0.1 | 0.22 | 0.075 | 0.079 | 0.067 | 0.065 | 0.082 | 0.07702 | 0.07 | 0.08 | 0.07 | 0.08 | 0.07 | 0.078 | 0.05 |  |  |  |  | 0.073368
| 1 | 5 | 0.08 | < 0.1 | 0.23 | 0.075 | 0.078 | 0.067 | 0.072 | 0.077 | 0.076854 | 0.072 | 0.07 | 0.08 | 0.07 | 0.07 | 0.07 | 0.05 |  |  |  |  | 36
| 1 | 6 | 0.07 | < 0.1 | 0.25 | 0.074 | 0.077 | 0.073 | 0.069 | 0.077 | 0.073702 | 0.068 | 0.07 | 0.09 | 0.07 | 0.07 | 0.076 | 0.05
| Mean |  | 0.076667 |  | 0.23667 | 0.076167 | 0.076667 | 0.071 | 0.0685 | 0.078333 | 0.075451 | 0.0705 | 0.075 | 0.073333 | 0.071667 | 0.07 | 0.073833 | 0.05
| Median |  | 0.08 |  | 0.235 | 0.0755 | 0.0775 | 0.0715 | 0.0695 | 0.077 | 0.07541 | 0.0705 | 0.075 | 0.07 | 0.07 | 0.07 | 0.074 | 0.05
| Std Dev. |  | 0.005164 |  | 0.012111 | 0.002137 | 0.0027325 | 0.0036332 | 0.0032711 | 0.0029439 | 0.0013302 | 0.0018708 | 0.0054772 | 0.010328 | 0.0040825 | 0 | 0.0029944 | 7.6012e-18
| Rel.Std.Dev. |  | 6.736% |  | 5.117% | 2.806% | 3.564% | 5.117% | 4.775% | 3.758% | 1.763% | 2.654% | 7.303% | 14.08% | 5.696% | 0% | 4.056% | 1.52e-14%
| PDM3 |  | 4.496% |  | 222.6% | 3.815% | 4.496% | -3.228% | -6.635% | 6.768% | 2.839% | -3.909% | 2.224% | -0.04727% | -2.319% | -4.591% | 0.6342% | -31.85%
| Z-Score (Absolute) |  | 0.67 | 5.16 | 35.67 | 0.56 | 0.67 | 0.56 | 1.11 | 1.04 | 0.41 | 0.67 | 0.31 | 0.05 | 0.42 | 0.78 | 0.05 | 5.16
| Analytical results for K in OREAS 211b (Certified Value 0.362 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.3505 | 0.36971 | 0.33225 | 0.37 | 0.39 | 0.36 | 0.36 | 0.35 | 0.36 | 0.36 | 0.3777 | 0.4248 | 0.392 | 0.39 | 0.3554 | 0.35 | 0.348 | 0.34 | 0.37 | 0.36452 | 1
| 1 | 2 | 0.3604 | 0.38553 | 0.33484 | 0.36 | 0.39 | 0.36 | 0.37 | 0.35 | 0.36 | 0.35 | 0.37655 | 0.4221 | 0.37 | 0.38 | 0.3605 | 0.36 | 0.367 | 0.34 | 0.36 | 0.36529 | #N/A
| 1 | 3 | 0.3563 | 0.37489 | 0.33452 | 0.37 | 0.39 | 0.35 | 0.36 | 0.35 | 0.36 | 0.38 | 0.36999 | 0.4263 | 0.387 | 0.38 | 0.3626 | 0.35 | 0.312 | 0.34 | 0.36 | 0.36532 | 16
| 1 | 4 | 0.3649 | 0.3763 | 0.33512 | 0.38 | 0.38 | 0.35 | 0.36 | 0.33 | 0.36 | 0.36 | 0.36948 | 0.4137 | 0.402 | 0.37 | 0.3585 | 0.34 | 0.306 | 0.34 | 0.36 | 0.3656 | 0.36168
| 1 | 5 | 0.3548 | 0.38824 | 0.33244 | 0.38 | 0.4 | 0.35 | 0.38 | 0.35 | 0.36 | 0.36 | 0.38101 | 0.4239 | 0.379 | 0.38 | 0.3532 | 0.36 | 0.33 | 0.34 | 0.36 | 0.36485 | 37
| 1 | 6 | 0.3602 | 0.36629 | 0.33386 | 0.36 | 0.39 | 0.36 | 0.37 | 0.35 | 0.36 | 0.36 | 0.37252 | 0.4213 | 0.374 | 0.39 | 0.3661 | 0.36 | 0.324 | 0.33 | 0.36 | 0.35996
| Mean |  | 0.35785 | 0.37682 | 0.33384 | 0.37 | 0.39 | 0.355 | 0.36667 | 0.34667 | 0.36 | 0.36167 | 0.37454 | 0.42202 | 0.384 | 0.38167 | 0.35938 | 0.35333 | 0.33117 | 0.33833 | 0.36167 | 0.36426
| Median |  | 0.35825 | 0.37559 | 0.33419 | 0.37 | 0.39 | 0.355 | 0.365 | 0.35 | 0.36 | 0.36 | 0.37454 | 0.423 | 0.383 | 0.38 | 0.3595 | 0.355 | 0.327 | 0.34 | 0.36 | 0.36507
| Std Dev. |  | 0.0050489 | 0.0086228 | 0.0012331 | 0.0089443 | 0.0063246 | 0.0054772 | 0.008165 | 0.008165 | 6.0809e-17 | 0.0098319 | 0.0046095 | 0.0044571 | 0.011983 | 0.0075277 | 0.0047242 | 0.008165 | 0.022895 | 0.0040825 | 0.0040825 | 0.0021418
| Rel.Std.Dev. |  | 1.411% | 2.288% | 0.3694% | 2.417% | 1.622% | 1.543% | 2.227% | 2.355% | 1.689e-14% | 2.719% | 1.231% | 1.056% | 3.121% | 1.972% | 1.315% | 2.311% | 6.913% | 1.207% | 1.129% | 0.588%
| PDM3 |  | -1.059% | 4.188% | -7.697% | 2.301% | 7.831% | -1.847% | 1.379% | -4.151% | -0.4642% | -0.003336% | 3.556% | 16.68% | 6.172% | 5.526% | -0.6347% | -2.307% | -8.436% | -6.455% | -0.003336% | 0.7127%
| Z-Score (Absolute) |  | 0.24 | 0.96 | 1.77 | 0.53 | 1.8 | 0.42 | 0.32 | 0.95 | 0.11 | 0 | 0.82 | 3.84 | 1.42 | 1.27 | 0.15 | 0.53 | 1.94 | 1.48 | 0 | 0.16
| Analytical results for La in OREAS 211b (Certified Value 4.74 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  |  | 2
3
| 1 | 1 | 4.91 | 4.741 | 5.6742 | 5.1 | 4.6 | 5 | 4.7 | 4.6 | 4.5 | 4.9 | 4.7 | 5.1 | 4.63 | 4.6 | 4.19 | 4.16 | 5.3 | 6.888 |  |  | 1
| 1 | 2 | 4.78 | 4.862 | 5.658 | 4.9 | 4.9 | 4.9 | 4.7 | 4.6 | 4.6 | 4.9 | 4.7 | 5 | 4.58 | 4.6 | 4.42 | 4.26 | 5.1 | 6.93 |  |  | 25
| 1 | 3 | 4.65 | 4.9643 | 5.6862 | 4.9 | 5.1 | 4.9 | 4.5 | 4.5 | 4.6 | 5 | 4.7 | 4.9 | 4.72 | 4.5 | 4.43 | 4.17 | 5.6 | 6.871 |  |  | 16
| 1 | 4 | 4.81 | 4.714 | 5.6898 | 4.9 | 5.1 | 4.7 | 4.8 | 4.4 | 4.4 | 5.1 | 4.6 | 5 | 4.58 | 4.5 | 4.34 | 4.44 | 5.1 | 6.876 |  |  | 4.7405
| 1 | 5 | 4.66 | 4.9869 | 5.6124 | 5.1 | 4.8 | 5 | 4.9 | 4.6 | 4.6 | 5.2 | 4.7 | 4.9 | 4.62 | 4.6 | 4.22 | 4.27 | 5.1 | 6.564 |  |  | 38
| 1 | 6 | 4.81 | 4.5895 | 5.6764 | 5 | 4.7 | 4.9 | 4.7 | 4.6 | 4.5 | 5.2 | 4.7 | 5 | 4.72 | 4.7 | 4.46 | 4.22 | 5.3 | 6.921
| Mean |  | 4.77 | 4.8096 | 5.6662 | 4.9833 | 4.8667 | 4.9 | 4.7167 | 4.55 | 4.5333 | 5.05 | 4.6833 | 4.9833 | 4.6417 | 4.5833 | 4.3433 | 4.2533 | 5.25 | 6.8417
| Median |  | 4.795 | 4.8015 | 5.6753 | 4.95 | 4.85 | 4.9 | 4.7 | 4.6 | 4.55 | 5.05 | 4.7 | 5 | 4.625 | 4.6 | 4.38 | 4.24 | 5.2 | 6.882
| Std Dev. |  | 0.099398 | 0.15516 | 0.02859 | 0.098319 | 0.20656 | 0.10954 | 0.13292 | 0.083666 | 0.08165 | 0.13784 | 0.040825 | 0.075277 | 0.064005 | 0.075277 | 0.11466 | 0.10191 | 0.19748 | 0.13812
| Rel.Std.Dev. |  | 2.084% | 3.226% | 0.5046% | 1.973% | 4.244% | 2.236% | 2.818% | 1.839% | 1.801% | 2.73% | 0.8717% | 1.511% | 1.379% | 1.642% | 2.64% | 2.396% | 3.762% | 2.019%
| PDM3 |  | 0.6224% | 1.458% | 19.53% | 5.123% | 2.662% | 3.365% | -0.5027% | -4.018% | -4.37% | 6.529% | -1.206% | 5.123% | -2.085% | -3.315% | -8.378% | -10.28% | 10.75% | 44.32%
| Z-Score (Absolute) |  | 0.07 | 0.07 | 2.95 | 0.65 | 0.26 | 0.37 | 0.25 | 0.81 | 0.86 | 0.88 | 0.36 | 0.65 | 0.5 | 0.7 | 1.51 | 1.81 | 1.55 | 6.92
| Analytical results for Li in OREAS 211b (Certified Value 11.6 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 11.5 | 11.2825 | 6.856 | 12 | 11 | 12.6 | 12.3 | 13 | 12.7 | 12.6 | 11.9405 | 11 | 12.1 | 9.5 | 11.3 | 10.6 | 12.7 | 12.05 | 10 | 11.982 | 1
| 1 | 2 | 11.4 | 11.8687 | 6.8869 | 12 | 10.9 | 12.9 | 12.4 | 13 | 12.5 | 12.7 | 11.6962 | 11 | 12.1 | 9.5 | 11.6 | 10.5 | 9.4 | 13.15 | < 10 | 12.127 | #N/A
| 1 | 3 | 11 | 11.6236 | 6.8643 | 12 | 11 | 12.4 | 12.4 | 12.9 | 12.7 | 13.1 | 11.7998 | 11 | 12.1 | 9.4 | 11.5 | 10.4 | 9.4 | 13.31 | 10 | 12.022 | 16
| 1 | 4 | 11.3 | 11.6897 | 6.7962 | 12 | 10.7 | 12.4 | 12.5 | 12.6 | 12.7 | 12.8 | 11.6057 | 11 | 12.1 | 10.2 | 11.5 | 10.4 | 9.9 | 13.01 | 10 | 11.994 | 11.5743
| 1 | 5 | 11.3 | 11.6267 | 6.7019 | 12 | 10.6 | 12.4 | 12.7 | 13.4 | 12.5 | 12.4 | 11.994 | 11 | 12.1 | 9.5 | 11.6 | 10.7 | 11.9 | 12.98 | < 10 | 12.035 | 39
| 1 | 6 | 11.5 | 11.1211 | 6.719 | 12 | 11.1 | 12.7 | 12.5 | 13.2 | 12.6 | 12.6 | 11.8303 | 11 | 12.1 | 12.2 | 11.6 | 11 | 9.5 | 14.22 | 10 | 12.001
| Mean |  | 11.3333 | 11.5354 | 6.804 | 12 | 10.8833 | 12.5667 | 12.4667 | 13.0167 | 12.6167 | 12.7 | 11.8111 | 11 | 12.1 | 10.05 | 11.5167 | 10.6 | 10.4667 | 13.12 | 10 | 12.0268
| Median |  | 11.35 | 11.6252 | 6.8261 | 12 | 10.95 | 12.5 | 12.45 | 13 | 12.65 | 12.65 | 11.815 | 11 | 12.1 | 9.5 | 11.55 | 10.55 | 9.7 | 13.08 | 10 | 12.0115
| Std Dev. |  | 0.18619 | 0.27808 | 0.078651 | 0 | 0.19408 | 0.20656 | 0.13663 | 0.27142 | 0.098319 | 0.23664 | 0.14563 | 0 | 0 | 1.0932 | 0.1169 | 0.22804 | 1.4542 | 0.69622 | 0 | 0.052686
| Rel.Std.Dev. |  | 1.643% | 2.411% | 1.156% | 0% | 1.783% | 1.644% | 1.096% | 2.085% | 0.7793% | 1.863% | 1.233% | 0% | 0% | 10.88% | 1.015% | 2.151% | 13.89% | 5.307% | 0% | 0.4381%
| PDM3 |  | -2.082% | -0.3362% | -41.21% | 3.678% | -5.97% | 8.574% | 7.71% | 12.46% | 9.006% | 9.726% | 2.046% | -4.962% | 4.542% | -13.17% | -0.4981% | -8.418% | -9.57% | 13.35% | -13.6% | 3.91%
| Z-Score (Absolute) |  | 0.14 | 0.01 | 3.41 | NA | 0.46 | 0.75 | 0.68 | 1.08 | NA | 0.85 | 0.21 | 0.38 | 0.41 | 1.07 | 0.01 | 0.67 | 0.77 | 1.15 | 2.31 | 0.36
Results from laboratory 1.12 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 211b (Certified Value 0.34 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.34 | 0.30668 | 0.4391 | 0.3 | 0.3 | 0.34 | 0.3 | 0.33 | 0.36 | [11 empty] | 1
| 1 | 2 | 0.35 | 0.36112 | 0.39401 | 0.4 | 0.3 | 0.34 | 0.3 | 0.33 | 0.34 | [11 empty] | 4
| 1 | 3 | 0.33 | 0.32648 | 0.40166 | 0.3 | 0.3 | 0.34 | 0.3 | 0.34 | 0.38 | [11 empty] | 16
| 1 | 4 | 0.33 | 0.34979 | 0.40409 | 0.3 | 0.3 | 0.35 | 0.3 | 0.32 | 0.36 | [11 empty] | 0.34045
| 1 | 5 | 0.35 | 0.37383 | 0.38897 | 0.3 | 0.3 | 0.33 | 0.3 | 0.29 | 0.34 | [11 empty] | 40
| 1 | 6 | 0.34 | 0.31548 | 0.38573 | 0.3 | 0.4 | 0.34 | 0.3 | 0.33 | 0.34
| Mean |  | 0.34 | 0.3389 | 0.40226 | 0.31667 | 0.31667 | 0.34 | 0.3 | 0.32333 | 0.35333
| Median |  | 0.34 | 0.33813 | 0.39783 | 0.3 | 0.3 | 0.34 | 0.3 | 0.33 | 0.35
| Std Dev. |  | 0.0089443 | 0.026734 | 0.019385 | 0.040825 | 0.040825 | 0.0063246 | 0 | 0.017512 | 0.01633
| Rel.Std.Dev. |  | 2.631% | 7.889% | 4.819% | 12.89% | 12.89% | 1.86% | 0% | 5.416% | 4.622%
| PDM3 |  | -0.131% | -0.455% | 18.16% | -6.985% | -6.985% | -0.131% | -11.88% | -5.027% | 3.785%
| Z-Score (Absolute) |  | 0 | 0.1 | 5.82 | NA | NA | 0 | NA | 1.56 | 1.25
| Analytical results for Mg in OREAS 211b (Certified Value 3.88 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.8296 | 3.7604 | 3.8051 | 4.17 | 3.86 | 3.83 | 3.86 | 3.96 | 3.8 | 3.72 | 3.9998 | 4.0083 | 3.98 | 3.95 | 3.7059 | 3.95 | 3.72 | 3.88 | 3.92 | 4.1018 | 1
| 1 | 2 | 3.9347 | 3.8757 | 3.8023 | 3.83 | 3.87 | 3.83 | 3.91 | 3.9 | 3.74 | 3.7 | 3.9771 | 4.0211 | 4.02 | 3.9 | 3.7257 | 3.93 | 3.95 | 3.91 | 3.88 | 4.1458 | #N/A
| 1 | 3 | 3.8737 | 3.7974 | 3.8097 | 3.82 | 3.82 | 3.78 | 3.83 | 3.86 | 3.77 | 3.83 | 4.0042 | 4.013 | 3.83 | 3.85 | 3.7804 | 3.88 | 3.46 | 3.9 | 3.95 | 4.0626 | 16
| 1 | 4 | 3.9324 | 3.8335 | 3.8023 | 4.08 | 3.82 | 3.74 | 3.85 | 3.72 | 3.84 | 3.74 | 3.9896 | 3.9149 | 4.03 | 3.84 | 3.7692 | 3.89 | 3.53 | 3.85 | 3.89 | 4.0623 | 3.8772
| 1 | 5 | 3.8447 | 3.809 | 3.805 | 3.88 | 3.88 | 3.77 | 3.99 | 3.89 | 3.82 | 3.65 | 3.9748 | 3.9941 | 3.79 | 3.79 | 3.7122 | 3.94 | 3.58 | 3.89 | 3.88 | 4.0458 | 41
| 1 | 6 | 3.9315 | 3.7105 | 3.8032 | 3.9 | 3.83 | 3.8 | 3.86 | 3.86 | 3.77 | 3.74 | 4.0016 | 3.9678 | 3.84 | 3.99 | 3.7696 | 3.99 | 3.55 | 3.76 | 3.99 | 4.0866
| Mean |  | 3.8911 | 3.7978 | 3.8046 | 3.9467 | 3.8467 | 3.7917 | 3.8833 | 3.865 | 3.79 | 3.73 | 3.9912 | 3.9865 | 3.915 | 3.8867 | 3.7438 | 3.93 | 3.6317 | 3.865 | 3.9183 | 4.0842
| Median |  | 3.9026 | 3.8032 | 3.8041 | 3.89 | 3.845 | 3.79 | 3.86 | 3.875 | 3.785 | 3.73 | 3.9947 | 4.0012 | 3.91 | 3.875 | 3.7475 | 3.935 | 3.565 | 3.885 | 3.905 | 4.0746
| Std Dev. |  | 0.04791 | 0.057391 | 0.002799 | 0.14418 | 0.026583 | 0.035449 | 0.058538 | 0.079937 | 0.036878 | 0.05933 | 0.012844 | 0.039771 | 0.10672 | 0.074476 | 0.032903 | 0.040497 | 0.17792 | 0.055408 | 0.04446 | 0.036131
| Rel.Std.Dev. |  | 1.231% | 1.511% | 0.07357% | 3.653% | 0.6911% | 0.9349% | 1.507% | 2.068% | 0.973% | 1.591% | 0.3218% | 0.9976% | 2.726% | 1.916% | 0.8789% | 1.03% | 4.899% | 1.434% | 1.135% | 0.8847%
| PDM3 |  | 0.358% | -2.049% | -1.873% | 1.791% | -0.788% | -2.207% | 0.1577% | -0.3152% | -2.25% | -3.797% | 2.939% | 2.819% | 0.9744% | 0.2437% | -3.44% | 1.361% | -6.333% | -0.3152% | 1.06% | 5.337%
| Z-Score (Absolute) |  | 0.16 | 0.73 | 0.66 | 0.69 | 0.26 | 0.78 | 0.09 | 0.09 | 0.8 | 1.37 | 1.11 | 1.07 | 0.39 | 0.12 | 1.24 | 0.53 | 2.3 | 0.09 | 0.42 | 1.99
| Analytical results for Mn in OREAS 211b (Certified Value 0.135 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.1379 | 0.13413 | 0.14395 | 0.145 | 0.1466 | 0.1305 | 0.1295 | 0.1405 | 0.133 | 0.133 | 0.14097 | 0.1207 | 0.128 | 0.1378 | 0.1376 | 0.1407 | 0.134 | 0.1357 | 0.134 | 0.13966 | 1
| 1 | 2 | 0.1416 | 0.13988 | 0.14461 | 0.137 | 0.1467 | 0.131 | 0.132 | 0.1385 | 0.1295 | 0.132 | 0.14162 | 0.121 | 0.134 | 0.1354 | 0.1365 | 0.1368 | 0.141 | 0.1375 | 0.131 | 0.13741 | 27
| 1 | 3 | 0.1398 | 0.13651 | 0.14453 | 0.139 | 0.1487 | 0.1285 | 0.1285 | 0.135 | 0.131 | 0.137 | 0.14207 | 0.1209 | 0.128 | 0.1337 | 0.1394 | 0.1351 | 0.122 | 0.1368 | 0.13 | 0.13959 | 16
| 1 | 4 | 0.1418 | 0.13593 | 0.1427 | 0.141 | 0.1469 | 0.128 | 0.131 | 0.131 | 0.133 | 0.133 | 0.14066 | 0.1187 | 0.137 | 0.1335 | 0.1385 | 0.139 | 0.125 | 0.1351 | 0.13 | 0.13679 | 0.13547
| 1 | 5 | 0.1387 | 0.13815 | 0.14104 | 0.142 | 0.1477 | 0.129 | 0.134 | 0.136 | 0.1345 | 0.131 | 0.14117 | 0.1206 | 0.13 | 0.1327 | 0.1369 | 0.1413 | 0.13 | 0.1365 | 0.13 | 0.13935 | 42
| 1 | 6 | 0.1417 | 0.13132 | 0.14338 | 0.137 | 0.1471 | 0.131 | 0.13 | 0.1345 | 0.1315 | 0.133 | 0.14213 | 0.1202 | 0.133 | 0.1385 | 0.1394 | 0.1426 | 0.126 | 0.1339 | 0.135 | 0.13639
| Mean |  | 0.14025 | 0.13599 | 0.14337 | 0.14017 | 0.14728 | 0.12967 | 0.13083 | 0.13592 | 0.13208 | 0.13317 | 0.14144 | 0.12035 | 0.13167 | 0.13527 | 0.13805 | 0.13925 | 0.12967 | 0.13592 | 0.13167 | 0.1382
| Median |  | 0.1407 | 0.13622 | 0.14366 | 0.14 | 0.147 | 0.12975 | 0.1305 | 0.1355 | 0.13225 | 0.133 | 0.14139 | 0.12065 | 0.1315 | 0.13455 | 0.13805 | 0.13985 | 0.128 | 0.1361 | 0.1305 | 0.13838
| Std Dev. |  | 0.0017003 | 0.0030122 | 0.0013489 | 0.0031252 | 0.00079603 | 0.0013292 | 0.0019664 | 0.0033078 | 0.0017725 | 0.0020412 | 0.00060277 | 0.00085499 | 0.0036148 | 0.0024105 | 0.001247 | 0.0028571 | 0.0069474 | 0.0012968 | 0.0022509 | 0.0015029
| Rel.Std.Dev. |  | 1.212% | 2.215% | 0.9408% | 2.23% | 0.5405% | 1.025% | 1.503% | 2.434% | 1.342% | 1.533% | 0.4262% | 0.7104% | 2.745% | 1.782% | 0.9033% | 2.052% | 5.358% | 0.9541% | 1.71% | 1.087%
| PDM3 |  | 3.528% | 0.3794% | 5.83% | 3.466% | 8.719% | -4.285% | -3.424% | 0.3288% | -2.501% | -1.701% | 4.404% | -11.16% | -2.808% | -0.151% | 1.904% | 2.789% | -4.285% | 0.3288% | -2.808% | 2.014%
| Z-Score (Absolute) |  | 0.69 | 0.01 | 1.18 | 0.67 | 1.8 | 0.99 | 0.81 | 0 | 0.61 | 0.44 | 0.88 | 2.47 | 0.67 | 0.1 | 0.34 | 0.53 | 0.99 | 0 | 0.67 | 0.36
| Analytical results for Mo in OREAS 211b (Certified Value 1.52 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 1.5 | 1.6431 | < 1 | < 2 | 1.58 | 1.61 | 1.61 | 1.53 | 1.48 | 1.4425 | 1.5 | 1.6 | 1.38 | 1.5 | 1.41 | 1.4 | < 1 | 1.5 | 1.908 |  | 1
| 1 | 2 | 1.6 | 1.5554 | < 1 | < 2 | 1.55 | 1.57 | 1.68 | 1.55 | 1.82 | 1.3988 | 1.6 | 1.6 | 1.36 | 1.6 | 1.39 | 1 | < 1 | 1.5 | 1.439 |  | 28
| 1 | 3 | 1.5 | 1.6135 | < 1 | < 2 | 1.58 | 1.42 | 1.72 | 1.48 | 1.5 | 1.3861 | 1.5 | 1.5 | 1.35 | 1.6 | 1.35 | 0.9 | < 1 | 1.5 | 1.78 |  | 16
| 1 | 4 | 1.5 | 1.5861 | < 1 | < 2 | 1.53 | 1.77 | 1.54 | 1.43 | 1.51 | 1.3905 | 1.5 | 1.6 | 1.48 | 1.6 | 1.36 | 1 | < 1 | 1.5 | 1.378 |  | 1.5214
| 1 | 5 | 1.6 | 1.5839 | < 1 | < 2 | 1.52 | 1.61 | 1.66 | 1.6 | 1.44 | 1.4461 | 1.5 | 1.5 | 1.33 | 1.5 | 1.4 | 1.1 | < 1 | 1.5 | 1.912 |  | 43
| 1 | 6 | 1.5 | 1.5401 | < 1 | 2 | 1.55 | 1.5 | 1.64 | 1.5 | 1.52 | 1.4262 | 1.4 | 1.5 | 1.54 | 1.6 | 1.39 | 1 | < 1 | 1.5 | 1.803
| Mean |  | 1.5333 | 1.587 |  | 2 | 1.5517 | 1.58 | 1.6417 | 1.515 | 1.545 | 1.415 | 1.5 | 1.55 | 1.4067 | 1.5667 | 1.3833 | 1.0667 |  | 1.5 | 1.7033
| Median |  | 1.5 | 1.585 |  | 2 | 1.55 | 1.59 | 1.65 | 1.515 | 1.505 | 1.4125 | 1.5 | 1.55 | 1.37 | 1.6 | 1.39 | 1 |  | 1.5 | 1.7915
| Std Dev. |  | 0.05164 | 0.03757 |  |  | 0.024833 | 0.11832 | 0.062102 | 0.058907 | 0.13766 | 0.02666 | 0.063246 | 0.054772 | 0.083825 | 0.05164 | 0.023381 | 0.17512 |  | 0 | 0.23535
| Rel.Std.Dev. |  | 3.368% | 2.367% |  |  | 1.6% | 7.489% | 3.783% | 3.888% | 8.91% | 1.884% | 4.216% | 3.534% | 5.959% | 3.296% | 1.69% | 16.42% |  | 0% | 13.82%
| PDM3 |  | 0.7871% | 4.316% |  | 31.46% | 1.992% | 3.854% | 7.908% | -0.418% | 1.554% | -6.989% | -1.404% | 1.883% | -7.539% | 2.978% | -9.073% | -29.89% |  | -1.404% | 11.96%
| Z-Score (Absolute) |  | 0.1 | 0.71 | 11.64 | NA | 0.31 | 0.63 | 1.34 | 0.1 | 0.24 | 1.24 | 0.27 | 0.29 | 1.34 | 0.48 | 1.6 | 5.2 | 11.64 | 0.27 | 2.04
Results from laboratory 1.08 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 211b (Certified Value 1.9 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.9041 | 1.8748 | 1.7409 | 1.9 | 1.95 | 1.85 | 1.87 | 1.93 | 1.84 | 1.92 | 1.8945 | 2.0048 | 1.96 | 1.984 | 1.8633 | 1.921 | 1.79 | 1.87 | 1.9 | 1.9086 | 1
| 1 | 2 | 1.9526 | 1.934 | 1.7391 | 1.79 | 1.96 | 1.87 | 1.89 | 1.9 | 1.83 | 1.91 | 1.8893 | 2.026 | 2.03 | 1.936 | 1.8662 | 1.906 | 1.91 | 1.91 | 1.86 | 1.9331 | #N/A
| 1 | 3 | 1.9336 | 1.8848 | 1.735 | 1.81 | 1.97 | 1.83 | 1.86 | 1.9 | 1.88 | 1.96 | 1.8875 | 2.0031 | 1.95 | 1.915 | 1.8916 | 1.88 | 1.67 | 1.92 | 1.86 | 1.9177 | 16
| 1 | 4 | 1.9683 | 1.8986 | 1.7432 | 1.86 | 1.92 | 1.81 | 1.87 | 1.82 | 1.86 | 1.94 | 1.8928 | 1.9661 | 2 | 1.903 | 1.8949 | 1.895 | 1.71 | 1.88 | 1.85 | 1.8901 | 1.9039
| 1 | 5 | 1.9144 | 1.9252 | 1.7438 | 1.91 | 1.98 | 1.83 | 1.94 | 1.9 | 1.86 | 1.87 | 1.886 | 1.988 | 1.91 | 1.914 | 1.8579 | 1.921 | 1.73 | 1.9 | 1.81 | 1.9065 | 44
| 1 | 6 | 1.965 | 1.8328 | 1.7484 | 1.77 | 1.94 | 1.85 | 1.87 | 1.89 | 1.84 | 1.92 | 1.8942 | 2.0008 | 1.97 | 1.966 | 1.8903 | 1.94 | 1.72 | 1.96 | 1.88 | 1.8976
| Mean |  | 1.9397 | 1.8917 | 1.7418 | 1.84 | 1.9533 | 1.84 | 1.8833 | 1.89 | 1.8517 | 1.92 | 1.8907 | 1.9981 | 1.97 | 1.9363 | 1.8774 | 1.9105 | 1.755 | 1.9067 | 1.86 | 1.9089
| Median |  | 1.9431 | 1.8917 | 1.7421 | 1.835 | 1.955 | 1.84 | 1.87 | 1.9 | 1.85 | 1.92 | 1.8911 | 2.0019 | 1.965 | 1.9255 | 1.8782 | 1.9135 | 1.725 | 1.905 | 1.86 | 1.9075
| Std Dev. |  | 0.026717 | 0.036764 | 0.0045734 | 0.058652 | 0.021602 | 0.020976 | 0.029439 | 0.036878 | 0.018348 | 0.030332 | 0.0036159 | 0.019908 | 0.041473 | 0.032303 | 0.016606 | 0.021361 | 0.085264 | 0.032042 | 0.030332 | 0.015151
| Rel.Std.Dev. |  | 1.377% | 1.943% | 0.2626% | 3.188% | 1.106% | 1.14% | 1.563% | 1.951% | 0.9909% | 1.58% | 0.1912% | 0.9963% | 2.105% | 1.668% | 0.8845% | 1.118% | 4.858% | 1.681% | 1.631% | 0.7937%
| PDM3 |  | 1.876% | -0.6438% | -8.519% | -3.359% | 2.594% | -3.359% | -1.083% | -0.7325% | -2.746% | 0.8432% | -0.6944% | 4.947% | 3.469% | 1.701% | -1.396% | 0.3443% | -7.823% | 0.1429% | -2.308% | 0.2623%
| Z-Score (Absolute) |  | 0.92 | 0.01 | 2.85 | 0.98 | 1.18 | 0.98 | 0.15 | 0.02 | 0.75 | 0.55 | 0.01 | 2.04 | 1.5 | 0.86 | 0.26 | 0.37 | 2.6 | 0.29 | 0.59 | 0.34
| Analytical results for Nb in OREAS 211b (Certified Value 3.43 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 3.53 | 3.4051 | 3.1 | < 5 | 3.7 | 3.4 | 3.5 | 3.3 | 3.4 | 3.6203 | NA | 3.6 | 3.23 | 3.44 | 3.17 | 3.7 | 4 |  |  |  | 1
| 1 | 2 | 3.47 | 3.4478 | 3.1 | < 5 | 3.6 | 3.5 | 3.5 | 3.4 | 3.3 | 3.4922 | NA | 3.8 | 3.46 | 3.37 | 3.11 | 3.4 | 3.5 |  |  |  | 29
| 1 | 3 | 3.34 | 3.4021 | 3.1 | < 5 | 3.6 | 3.2 | 3.5 | 3.2 | 3.4 | 3.5888 | NA | 3.8 | 3.26 | 3.44 | 3.06 | 3.4 | 3.5 |  |  |  | 16
| 1 | 4 | 3.46 | 3.2155 | 3.2 | < 5 | 3.5 | 3.5 | 3.3 | 3.2 | 3.3 | 3.5244 | NA | 3.8 | 3.31 | 3.44 | 3.07 | 3.6 | 3.5 |  |  |  | 3.4307
| 1 | 5 | 3.37 | 3.5934 | 3.2 | < 5 | 3.6 | 3.6 | 3.5 | 3.3 | 3.4 | 3.6348 | NA | 3.8 | 3.28 | 3.43 | 3.13 | 3.4 | 3.5 |  |  |  | 45
| 1 | 6 | 3.42 | 3.2383 | 3.2 | < 5 | 3.6 | 3.4 | 3.4 | 3.3 | 3.5 | 3.6963 | NA | 3.6 | 3.79 | 3.58 | 3.18 | 3.5 | 4
| Mean |  | 3.4317 | 3.3837 | 3.15 |  | 3.6 | 3.4333 | 3.45 | 3.2833 | 3.3833 | 3.5928 |  | 3.7333 | 3.3883 | 3.45 | 3.12 | 3.5 | 3.6667
| Median |  | 3.44 | 3.4036 | 3.15 |  | 3.6 | 3.45 | 3.5 | 3.3 | 3.4 | 3.6046 |  | 3.8 | 3.295 | 3.44 | 3.12 | 3.45 | 3.5
| Std Dev. |  | 0.06969 | 0.1402 | 0.054772 |  | 0.063246 | 0.13663 | 0.083666 | 0.075277 | 0.075277 | 0.074889 |  | 0.10328 | 0.21255 | 0.069282 | 0.0498 | 0.12649 | 0.2582
| Rel.Std.Dev. |  | 2.031% | 4.144% | 1.739% |  | 1.757% | 3.979% | 2.425% | 2.293% | 2.225% | 2.084% |  | 2.766% | 6.273% | 2.008% | 1.596% | 3.614% | 7.042%
| PDM3 |  | 0.02883% | -1.37% | -8.181% |  | 4.936% | 0.07741% | 0.5632% | -4.295% | -1.38% | 4.726% |  | 8.822% | -1.234% | 0.5632% | -9.056% | 2.021% | 6.879%
| Z-Score (Absolute) |  | 0.01 | 0.3 | 1.76 | 5.8 | 1.04 | 0.01 | 0.11 | 0.93 | 0.31 | 1 | NA | 1.87 | 0.27 | 0.11 | 1.95 | 0.42 | 1.46
| Analytical results for Nd in OREAS 211b (Certified Value 8.13 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 8.29 | 7.5701 | 10.3697 | 8.4 | 7.6 | 8.28 | 7.7 | 7.92 | 8.4 | [11 empty] | 1
| 1 | 2 | 8.06 | 8.1967 | 10.4201 | 8.4 | 7.5 | 8.2 | 7.8 | 8.69 | 8.4 | [11 empty] | 5
| 1 | 3 | 7.92 | 8.0234 | 10.4647 | 8 | 8 | 8.25 | 7.7 | 8.51 | 8.35 | [11 empty] | 16
| 1 | 4 | 8.22 | 8.1889 | 10.4576 | 8.3 | 8 | 8.28 | 7.6 | 8.47 | 8.35 | [11 empty] | 8.1287
| 1 | 5 | 8.12 | 8.4733 | 10.3935 | 8.2 | 7.6 | 8.28 | 8 | 7.75 | 8.6 | [11 empty] | 46
| 1 | 6 | 8.32 | 7.6565 | 10.4841 | 8.4 | 9.1 | 8.36 | 8 | 8.61 | 8.5
| Mean |  | 8.155 | 8.0182 | 10.4316 | 8.2833 | 7.9667 | 8.275 | 7.8 | 8.325 | 8.4333
| Median |  | 8.17 | 8.1061 | 10.4388 | 8.35 | 7.8 | 8.28 | 7.75 | 8.49 | 8.4
| Std Dev. |  | 0.15176 | 0.34634 | 0.044606 | 0.16021 | 0.59554 | 0.052058 | 0.16733 | 0.39099 | 0.098319
| Rel.Std.Dev. |  | 1.861% | 4.319% | 0.4276% | 1.934% | 7.475% | 0.6291% | 2.145% | 4.697% | 1.166%
| PDM3 |  | 0.3232% | -1.36% | 28.33% | 1.902% | -1.994% | 1.799% | -4.044% | 2.415% | 3.747%
| Z-Score (Absolute) |  | 0.51 | 1.09 | 9.18 | 0.04 | 1.31 | 0 | 2.02 | 0.21 | 0.67
| Analytical results for Ni in OREAS 211b (Certified Value 94 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 98.8 | 92.7319 | 102.6519 | 91 | 95 | 97.6 | 95.5 | 98.2 | 95.9 | 101 | 93.2644 | 83.9 | 91 | 90 | 94.7 | 91.5 | 93.2 | 88.89 | 114 | 90.281 | 1
| 1 | 2 | 98.8 | 96.9627 | 99.9363 | 88 | 96 | 96.9 | 96.6 | 99.6 | 97.7 | 99.1 | 96.691 | 90.6 | 94 | 88.9 | 93.6 | 92.3 | 96.3 | 86.45 | 112 | 93.967 | 14
| 1 | 3 | 95 | 95.1568 | 101.5195 | 88 | 96 | 96.5 | 91.1 | 100.5 | 93.5 | 102 | 96.5394 | 84.8 | 91 | 87.3 | 94.6 | 90.9 | 93.2 | 88.49 | 114 | 90.306 | 16
| 1 | 4 | 97.4 | 93.8739 | 100.2945 | 91 | 94 | 94 | 109.5 | 96.5 | 93.6 | 100 | 94.611 | 82 | 94 | 87.9 | 95.1 | 90.9 | 93.1 | 87.39 | 114 | 90.263 | 94.2933
| 1 | 5 | 96.7 | 96.2836 | 97.3015 | 86 | 96 | 96.5 | 99.7 | 99.4 | 97.4 | 100 | 94.6166 | 83.9 | 91 | 87.5 | 94.2 | 92.8 | 95.1 | 85.31 | 110 | 92.97 | 47
| 1 | 6 | 94.7 | 91.8594 | 98.5496 | 89 | 94 | 95.3 | 96.7 | 98 | 95.9 | 98.7 | 94.3934 | 83.3 | 91 | 88.8 | 94.5 | 94.2 | 94 | 87.41 | 110 | 93.387
| Mean |  | 96.9 | 94.478 | 100.0422 | 88.8333 | 95.1667 | 96.1333 | 98.1833 | 98.7 | 95.6667 | 100.1333 | 95.0193 | 84.75 | 92 | 88.4 | 94.45 | 92.1 | 94.15 | 87.3233 | 112.3333 | 91.8623
| Median |  | 97.05 | 94.5153 | 100.1154 | 88.5 | 95.5 | 96.5 | 96.65 | 98.8 | 95.9 | 100 | 94.6138 | 83.9 | 91 | 88.35 | 94.55 | 91.9 | 93.6 | 87.4 | 113 | 91.638
| Std Dev. |  | 1.7866 | 2.0083 | 1.9394 | 1.9408 | 0.98319 | 1.2848 | 6.204 | 1.4227 | 1.8007 | 1.216 | 1.3346 | 3.0111 | 1.5492 | 1.0237 | 0.50892 | 1.2791 | 1.3004 | 1.3148 | 1.9664 | 1.7585
| Rel.Std.Dev. |  | 1.844% | 2.126% | 1.939% | 2.185% | 1.033% | 1.336% | 6.319% | 1.441% | 1.882% | 1.214% | 1.405% | 3.553% | 1.684% | 1.158% | 0.5388% | 1.389% | 1.381% | 1.506% | 1.75% | 1.914%
| PDM3 |  | 2.765% | 0.196% | 6.097% | -5.79% | 0.9263% | 1.951% | 4.126% | 4.673% | 1.457% | 6.194% | 0.77% | -10.12% | -2.432% | -6.25% | 0.1662% | -2.326% | -0.1519% | -7.392% | 19.13% | -2.578%
| Z-Score (Absolute) |  | 0.51 | 0.06 | 1.27 | 1.42 | 0.1 | 0.33 | 0.82 | 0.95 | 0.22 | 1.29 | 0.06 | 2.39 | 0.66 | 1.52 | 0.07 | 0.63 | 0.14 | 1.78 | 4.21 | 0.69
| Analytical results for P in OREAS 211b (Certified Value 0.041 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0406 | 0.041212 | 0.038 | 0.0423 | 0.04 | 0.041 | 0.04 | 0.043 | 0.043 | 0.044774 | 0.0419 | 0.043 | 0.043 | 0.0437 | 0.038 | 0.0265 | 0.04 | 0.04 | 0.041628 |  | 1
| 1 | 2 | 0.0414 | 0.041957 | 0.041 | 0.0428 | 0.04 | 0.042 | 0.041 | 0.042 | 0.043 | 0.042339 | 0.0418 | 0.044 | 0.042 | 0.0413 | 0.039 | 0.0304 | 0.04 | 0.04 | 0.042079 |  | #N/A
| 1 | 3 | 0.0414 | 0.041202 | 0.037 | 0.0426 | 0.04 | 0.041 | 0.039 | 0.042 | 0.044 | 0.043513 | 0.0422 | 0.042 | 0.041 | 0.0394 | 0.038 | 0.0302 | 0.04 | 0.04 | 0.042074 |  | 16
| 1 | 4 | 0.0419 | 0.041824 | 0.037 | 0.042 | 0.039 | 0.041 | 0.039 | 0.044 | 0.043 | 0.043024 | 0.0407 | 0.043 | 0.041 | 0.0442 | 0.038 | 0.0239 | 0.04 | 0.04 | 0.04148 |  | 0.041496
| 1 | 5 | 0.041 | 0.041997 | 0.037 | 0.0428 | 0.04 | 0.043 | 0.041 | 0.043 | 0.042 | 0.044589 | 0.0412 | 0.042 | 0.041 | 0.0418 | 0.039 | 0.0293 | 0.05 | 0.04 | 0.042062 |  | 48
| 1 | 6 | 0.0419 | 0.040544 | 0.038 | 0.0428 | 0.04 | 0.042 | 0.041 | 0.042 | 0.043 | 0.043685 | 0.0414 | 0.043 | 0.042 | 0.0449 | 0.04 | 0.0244 | 0.04 | 0.04 | 0.041612
| Mean |  | 0.041367 | 0.041456 | 0.038 | 0.04255 | 0.039833 | 0.041667 | 0.040167 | 0.042667 | 0.043 | 0.043654 | 0.041533 | 0.042833 | 0.041667 | 0.04255 | 0.038667 | 0.02745 | 0.041667 | 0.04 | 0.041823
| Median |  | 0.0414 | 0.041518 | 0.0375 | 0.0427 | 0.04 | 0.0415 | 0.0405 | 0.0425 | 0.043 | 0.043599 | 0.0416 | 0.043 | 0.0415 | 0.04275 | 0.0385 | 0.0279 | 0.04 | 0.04 | 0.041845
| Std Dev. |  | 0.00050859 | 0.00057173 | 0.0015492 | 0.00033317 | 0.00040825 | 0.0008165 | 0.00098319 | 0.0008165 | 0.00063246 | 0.00092461 | 0.00054283 | 0.00075277 | 0.0008165 | 0.0020792 | 0.0008165 | 0.0029153 | 0.0040825 | 0 | 0.00027792
| Rel.Std.Dev. |  | 1.229% | 1.379% | 4.077% | 0.783% | 1.025% | 1.96% | 2.448% | 1.914% | 1.471% | 2.118% | 1.307% | 1.757% | 1.96% | 4.886% | 2.112% | 10.62% | 9.798% | 0% | 0.6645%
| PDM3 |  | -0.3118% | -0.0965% | -8.425% | 2.54% | -4.007% | 0.4112% | -3.204% | 2.821% | 3.624% | 5.201% | 0.08987% | 3.223% | 0.4112% | 2.54% | -6.818% | -33.85% | 0.4112% | -3.605% | 0.787%
| Z-Score (Absolute) |  | 0.2 | 0.14 | 2.47 | 0.6 | 1.24 | 0 | 1.01 | 0.67 | 0.9 | 1.34 | 0.09 | 0.79 | 0 | 0.6 | 2.02 | 9.59 | 0 | 1.12 | 0.11
| Analytical results for Pb in OREAS 211b (Certified Value 14.8 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 15.5 | 16.2852 | 12 | 14.4 | 14.7 | 14.5 | 14.7 | 13.8 | 16.9 | 14.9025 | 12 | 15 | 12.76 | 15.1 | 14.29 | 16.1 | 18.77 | 16 | 14.83 |  | 1
| 1 | 2 | 15 | 16.7334 | 8 | 15 | 14.5 | 14.9 | 14.8 | 15 | 14.5 | 14.4042 | 12 | 15 | 12.81 | 16.3 | 14.47 | 14.6 | 19.18 | 14 | 17.638 |  | 16
| 1 | 3 | 15.2 | 16.6309 | 11 | 15 | 14.3 | 14.1 | 14.6 | 13.3 | 16.1 | 14.9534 | 12 | 16 | 13.24 | 14.9 | 14.33 | 14.3 | 18.11 | 15 | 14.551 |  | 16
| 1 | 4 | 15.5 | 15.469 | 13 | 13.8 | 13.9 | 14.2 | 13.8 | 12.8 | 15.7 | 14.4102 | 13 | 15 | 14.19 | 15.1 | 13.94 | 14.6 | 18.65 | 14 | 13.841 |  | 14.8071
| 1 | 5 | 15.2 | 17.6581 | 12 | 14.6 | 14.8 | 15.4 | 14.8 | 14.1 | 15.2 | 14.7122 | 13 | 15 | 12.83 | 15.4 | 14.7 | 14.4 | 20.03 | 16 | 18.479 |  | 49
| 1 | 6 | 15.2 | 16.922 | 11 | 13.9 | 14.4 | 14.7 | 14.4 | 13.7 | 16.5 | 14.7267 | 13 | 15 | 13.3 | 16.1 | 14.78 | 17 | 18.01 | 15 | 14.463
| Mean |  | 15.2667 | 16.6164 | 11.1667 | 14.45 | 14.4333 | 14.6333 | 14.5167 | 13.7833 | 15.8167 | 14.6849 | 12.5 | 15.1667 | 13.1883 | 15.4833 | 14.4183 | 15.1667 | 18.7917 | 15 | 15.6337
| Median |  | 15.2 | 16.6821 | 11.5 | 14.5 | 14.45 | 14.6 | 14.65 | 13.75 | 15.9 | 14.7195 | 12.5 | 15 | 13.035 | 15.25 | 14.4 | 14.6 | 18.71 | 15 | 14.6905
| Std Dev. |  | 0.19664 | 0.72367 | 1.7224 | 0.52058 | 0.32042 | 0.48028 | 0.38166 | 0.74677 | 0.87731 | 0.23502 | 0.54772 | 0.40825 | 0.54286 | 0.58109 | 0.30525 | 1.1147 | 0.74572 | 0.89443 | 1.9243
| Rel.Std.Dev. |  | 1.288% | 4.355% | 15.42% | 3.603% | 2.22% | 3.282% | 2.629% | 5.418% | 5.547% | 1.6% | 4.382% | 2.692% | 4.116% | 3.753% | 2.117% | 7.35% | 3.968% | 5.963% | 12.31%
| PDM3 |  | 3.104% | 12.22% | -24.59% | -2.412% | -2.524% | -1.173% | -1.961% | -6.914% | 6.818% | -0.8254% | -15.58% | 2.428% | -10.93% | 4.567% | -2.625% | 2.428% | 26.91% | 1.303% | 5.582%
| Z-Score (Absolute) |  | 0.67 | 2.24 | NA | 0.27 | 0.29 | 0.06 | 0.19 | 1.04 | 1.31 | 0 | NA | NA | 1.73 | 0.93 | 0.31 | 0.56 | 4.76 | NA | 1.1
Results from laboratories 1.07, 1.15, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 211b (Certified Value 1.67 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.71 | 1.6202 | 2.143 | 1.7 | 1.6 | 1.64 | 1.5 | 1.6 | 1.7 | [11 empty] | 1
| 1 | 2 | 1.68 | 1.6954 | 2.1492 | 1.7 | 1.6 | 1.68 | 1.5 | 1.72 | 1.6 | [11 empty] | 33
| 1 | 3 | 1.63 | 1.7211 | 2.1546 | 1.6 | 1.6 | 1.66 | 1.5 | 1.73 | 1.7 | [11 empty] | 16
| 1 | 4 | 1.65 | 1.7049 | 2.1665 | 1.7 | 1.6 | 1.67 | 1.5 | 1.68 | 1.65 | [11 empty] | 1.666
| 1 | 5 | 1.61 | 1.6944 | 2.1846 | 1.7 | 1.6 | 1.65 | 1.6 | 1.61 | 1.65 | [11 empty] | 50
| 1 | 6 | 1.71 | 1.5734 | 2.1248 | 1.7 | 1.8 | 1.68 | 1.6 | 1.69 | 1.65
| Mean |  | 1.665 | 1.6682 | 2.1538 | 1.6833 | 1.6333 | 1.6633 | 1.5333 | 1.6717 | 1.6583
| Median |  | 1.665 | 1.6949 | 2.1519 | 1.7 | 1.6 | 1.665 | 1.5 | 1.685 | 1.65
| Std Dev. |  | 0.041833 | 0.058078 | 0.020449 | 0.040825 | 0.08165 | 0.01633 | 0.05164 | 0.054924 | 0.037639
| Rel.Std.Dev. |  | 2.512% | 3.481% | 0.9494% | 2.425% | 4.999% | 0.9818% | 3.368% | 3.286% | 2.27%
| PDM3 |  | -0.0618% | 0.1324% | 29.28% | 1.039% | -1.963% | -0.1618% | -7.965% | 0.3384% | -0.4619%
| Z-Score (Absolute) |  | 0 | 0.33 | 49.44 | 1.85 | 3.2 | 0.17 | 13.32 | 0.67 | 0.67
| Analytical results for Pt in OREAS 211b (Indicative Value 19.2 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.21 | [19 empty] | ppb
|  |  | 4A*OES/MS | [19 empty] | 1
1
| 1 | 1 | 21 | [19 empty] | 1
| 1 | 2 | 20 | [19 empty] | 2
| 1 | 3 | 13 | [19 empty] | 16
| 1 | 4 | 21 | [19 empty] | 19.1667
| 1 | 5 | 19 | [19 empty] | 8
| 1 | 6 | 21
| Mean |  | 19.1667
| Median |  | 20.5
| Std Dev. |  | 3.1252
| Rel.Std.Dev. |  | 16.31%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 211b (Certified Value 10.1 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 10.57 | 9.8075 | 11.59 | 9.6 | 10.7 | 10.6 | 10.4 | 10.3 | 10.4 | 10.1246 | 10.3 | 11 | 9 | 9.58 | 9.4 | 10.2 | 8.8 |  |  |  | 1
| 1 | 2 | 10.55 | 9.9595 | 11.5525 | 9.4 | 11 | 10.6 | 10.4 | 10.2 | 10.1 | 9.9609 | 10.5 | 11.1 | 8.8 | 9.78 | 9.6 | 10.5 | 8.8 |  |  |  | 34
| 1 | 3 | 9.95 | 10.7515 | 11.4032 | 9.5 | 10.8 | 10 | 10.4 | 9.7 | 10.5 | 10.2572 | 10.2 | 11.3 | 9.4 | 9.94 | 9.5 | 10.6 | 9.2 |  |  |  | 16
| 1 | 4 | 10.36 | 9.7586 | 11.5449 | 9.4 | 10.3 | 10.6 | 9.8 | 9.5 | 10.2 | 9.9212 | 10.1 | 11.1 | 9.5 | 9.72 | 9.5 | 9.9 | 8.8 |  |  |  | 10.1312
| 1 | 5 | 10.04 | 10.0598 | 11.5188 | 10.7 | 11 | 11 | 10.3 | 10.4 | 10.6 | 10.1942 | 10.2 | 11.3 | 8.9 | 9.69 | 9.6 | 10 | 9 |  |  |  | 51
| 1 | 6 | 10.19 | 9.492 | 11.5781 | 9.5 | 10.6 | 10.7 | 10.2 | 10.2 | 10.8 | 10.0132 | 10 | 10.5 | 11.1 | 9.92 | 9.7 | 10 | 9
| Mean |  | 10.2767 | 9.9715 | 11.5313 | 9.6833 | 10.7333 | 10.5833 | 10.25 | 10.05 | 10.4333 | 10.0786 | 10.2167 | 11.05 | 9.45 | 9.7717 | 9.55 | 10.2 | 8.9333
| Median |  | 10.275 | 9.8835 | 11.5487 | 9.5 | 10.75 | 10.6 | 10.35 | 10.2 | 10.45 | 10.0689 | 10.2 | 11.1 | 9.2 | 9.75 | 9.55 | 10.1 | 8.9
| Std Dev. |  | 0.25997 | 0.42861 | 0.067569 | 0.50365 | 0.26583 | 0.32506 | 0.23452 | 0.36194 | 0.2582 | 0.13435 | 0.17224 | 0.29496 | 0.85499 | 0.13891 | 0.10488 | 0.28983 | 0.1633
| Rel.Std.Dev. |  | 2.53% | 4.298% | 0.586% | 5.201% | 2.477% | 3.071% | 2.288% | 3.601% | 2.475% | 1.333% | 1.686% | 2.669% | 9.047% | 1.422% | 1.098% | 2.841% | 1.828%
| PDM3 |  | 1.436% | -1.576% | 13.82% | -4.42% | 5.944% | 4.463% | 1.173% | -0.801% | 2.983% | -0.5192% | 0.844% | 9.07% | -6.723% | -3.548% | -5.736% | 0.6795% | -11.82%
| Z-Score (Absolute) |  | 0.13 | 0.4 | 2.34 | 0.91 | 0.94 | 0.67 | 0.09 | 0.26 | 0.41 | 0.21 | 0.03 | 1.5 | 1.32 | 0.75 | 1.14 | 0 | 2.23
| Analytical results for Re in OREAS 211b (Certified Value 0.002 ppm) | [21 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 |  | 0.013307 | 0.002 | < 0.002 | 0.002 | 0.003 | 0.002 | 0.003 | < 0.05 | 0.009 | < 0.05 | < 0.002 | 0.003 | < 0.002 | 0.0019 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.002 | 0.010611 | 0.002 | 0.002 | 0.002 | 0.005 | 0.003 | 0.002 | < 0.05 | 0.009 | < 0.05 | 0.003 | 0.004 | < 0.002 | 0.0017 | < 0.1 |  |  |  |  | 35
| 1 | 3 | 0.002 | 0.010734 | 0.002 | 0.002 | 0.003 | 0.003 | 0.002 | 0.003 | < 0.05 | 0.008 | < 0.05 | 0.002 |  | < 0.002 | 0.0015 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.002 | 0.0105 | 0.002 | 0.002 | 0.002 | 0.003 | 0.003 | 0.002 | < 0.05 | 0.009 | < 0.05 | 0.002 |  | < 0.002 | 0.0018 | < 0.1 |  |  |  |  | 0.0022867
| 1 | 5 | 0.003 | 0.011376 | 0.003 | 0.003 | 0.002 | 0.002 | 0.003 | 0.002 | < 0.05 | 0.009 | < 0.05 | 0.004 |  | 0.002 | 0.002 | < 0.1 |  |  |  |  | 52
| 1 | 6 |  | 0.012976 | 0.002 | 0.002 | 0.002 | 0.002 | 0.003 | 0.002 | < 0.05 | 0.008 | < 0.05 | 0.005 |  | 0.002 | 0.0015 | < 0.1
| Mean |  | 0.00225 | 0.011584 | 0.0021667 | 0.0022 | 0.0021667 | 0.003 | 0.0026667 | 0.0023333 |  | 0.0086667 |  | 0.0032 | 0.0035 | 0.002 | 0.0017333
| Median |  | 0.002 | 0.011055 | 0.002 | 0.002 | 0.002 | 0.003 | 0.003 | 0.002 |  | 0.009 |  | 0.003 | 0.0035 | 0.002 | 0.00175
| Std Dev. |  | 0.0005 | 0.0012486 | 0.00040825 | 0.00044721 | 0.00040825 | 0.0010954 | 0.0005164 | 0.0005164 |  | 0.0005164 |  | 0.0013038 | 0.00070711 | 0 | 0.00020656
| Rel.Std.Dev. |  | 22.22% | 10.78% | 18.84% | 20.33% | 18.84% | 36.51% | 19.36% | 22.13% |  | 5.958% |  | 40.75% | 20.2% | 0% | 11.92%
| PDM3 |  | -1.603% | 406.6% | -5.248% | -3.79% | -5.248% | 31.2% | 16.62% | 2.041% |  | 279% |  | 39.94% | 53.06% | -12.54% | -24.2%
| Z-Score (Absolute) |  | 0.45 | 7.94 | 0.52 | 0.67 | 0.52 | 0.22 | 0.07 | 0.37 | 20 | 5.32 | 20 | 0.07 | 0.67 | 1.27 | 0.91 | 42.48
| Analytical results for S in OREAS 211b (Certified Value 0.267 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2665 | 0.25739 | 0.27893 | 0.25 | 0.28 | 0.25 | 0.26 | 0.28 | 0.27 | 0.27 | 0.27264 | 0.2549 | 0.29 | 0.27 | 0.22 | 0.25 | 0.231 | 0.12 | 0.27 | 0.32431 | 1
| 1 | 2 | 0.2728 | 0.27881 | 0.27918 | 0.26 | 0.28 | 0.26 | 0.27 | 0.27 | 0.26 | 0.27 | 0.27936 | 0.256 | 0.3 | 0.27 | 0.28 | 0.25 | 0.231 | 0.12 | 0.26 | 0.32539 | 20
| 1 | 3 | 0.2686 | 0.26116 | 0.27824 | 0.25 | 0.28 | 0.25 | 0.27 | 0.27 | 0.26 | 0.28 | 0.27497 | 0.2529 | 0.284 | 0.27 | 0.26 | 0.24 | 0.208 | 0.12 | 0.26 | 0.33178 | 16
| 1 | 4 | 0.2729 | 0.27109 | 0.2772 | 0.25 | 0.28 | 0.25 | 0.26 | 0.26 | 0.27 | 0.27 | 0.27283 | 0.2467 | 0.306 | 0.26 | 0.21 | 0.25 | 0.206 | 0.12 | 0.285 | 0.3209 | 0.26704
| 1 | 5 | 0.2687 | 0.26917 | 0.27898 | 0.25 | 0.28 | 0.25 | 0.28 | 0.27 | 0.27 | 0.27 | 0.28046 | 0.2529 | 0.274 | 0.26 | 0.24 | 0.25 | 0.231 | 0.11 | 0.27 | 0.33106 | 53
| 1 | 6 | 0.2713 | 0.26393 | 0.27868 | 0.26 | 0.28 | 0.25 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27757 | 0.2524 | 0.277 | 0.28 | 0.26 | 0.25 | 0.209 | 0.11 | 0.265 | 0.3188
| Mean |  | 0.27013 | 0.26692 | 0.27853 | 0.25333 | 0.28 | 0.25167 | 0.26833 | 0.27 | 0.26667 | 0.27167 | 0.2763 | 0.25263 | 0.2885 | 0.26833 | 0.245 | 0.24833 | 0.21933 | 0.11667 | 0.26833 | 0.32537
| Median |  | 0.27 | 0.26655 | 0.2788 | 0.25 | 0.28 | 0.25 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27627 | 0.2529 | 0.287 | 0.27 | 0.25 | 0.25 | 0.22 | 0.12 | 0.2675 | 0.32485
| Std Dev. |  | 0.0025974 | 0.0077023 | 0.00072952 | 0.005164 | 0 | 0.0040825 | 0.0075277 | 0.0063246 | 0.005164 | 0.0040825 | 0.0033315 | 0.003221 | 0.012677 | 0.0075277 | 0.026646 | 0.0040825 | 0.012817 | 0.005164 | 0.0093095 | 0.0052483
| Rel.Std.Dev. |  | 0.9615% | 2.886% | 0.2619% | 2.038% | 0% | 1.622% | 2.805% | 2.342% | 1.936% | 1.503% | 1.206% | 1.275% | 4.394% | 2.805% | 10.88% | 1.644% | 5.843% | 4.426% | 3.469% | 1.613%
| PDM3 |  | 1.158% | -0.04366% | 4.303% | -5.133% | 4.853% | -5.757% | 0.484% | 1.108% | -0.1401% | 1.732% | 3.469% | -5.395% | 8.036% | 0.484% | -8.254% | -7.006% | -17.87% | -56.31% | 0.484% | 21.84%
| Z-Score (Absolute) |  | 0.11 | 0.09 | 0.63 | 0.93 | 0.72 | 1.03 | 0 | 0.1 | 0.1 | 0.21 | 0.49 | 0.97 | 1.24 | 0 | 1.44 | 1.23 | 3.02 | 9.35 | 0 | 3.52
| Analytical results for Sb in OREAS 211b (Certified Value 1.38 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 1.33 | 1.3555 | 1.3 | < 5 | 1.39 | 2.05 | 1.32 | 1.4 | 1.45 | 1.4874 | 1.36 | 1.3 | 1.16 | 1.36 | 1.18 | 1.4 | < 5 | 1.5 | 1.091 |  | 1
| 1 | 2 | 1.34 | 1.4017 | 1.4 | < 5 | 1.39 | 1.63 | 1.34 | 1.32 | 1.47 | 1.4356 | 1.39 | 1.3 | 1.24 | 1.36 | 1.21 | 1.3 | < 5 | 1.6 | 2.007 |  | 36
| 1 | 3 | 1.32 | 1.3432 | 1.3 | < 5 | 1.42 | 1.49 | 1.38 | 1.32 | 1.48 | 1.4352 | 1.35 | 1.3 | 1.24 | 1.38 | 1.18 | 1.3 | < 5 | 1.6 |  |  | 16
| 1 | 4 | 1.33 | 1.3605 | 1.4 | < 5 | 1.33 | 1.62 | 1.3 | 1.36 | 1.58 | 1.4881 | 1.32 | 1.4 | 1.32 | 1.39 | 1.17 | 1.3 | < 5 | 1.5 | 4.062 |  | 1.3757
| 1 | 5 | 1.3 | 1.3933 | 1.4 | < 5 | 1.4 | 1.7 | 1.32 | 1.49 | 1.55 | 1.5083 | 1.37 | 1.3 | 1.22 | 1.48 | 1.18 | 1.3 | < 5 | 1.7 | 0.1 |  | 54
| 1 | 6 | 1.33 | 1.3175 | 1.4 | < 5 | 1.39 | 1.52 | 1.32 | 1.31 | 1.51 | 1.4994 | 1.31 | 1.3 | 1.33 | 1.44 | 1.22 | 1.3 | < 5 | 1.7 | 4.725
| Mean |  | 1.325 | 1.3619 | 1.3667 |  | 1.3867 | 1.6683 | 1.33 | 1.3667 | 1.5067 | 1.4757 | 1.35 | 1.3167 | 1.2517 | 1.4017 | 1.19 | 1.3167 |  | 1.6 | 2.397
| Median |  | 1.33 | 1.358 | 1.4 |  | 1.39 | 1.625 | 1.32 | 1.34 | 1.495 | 1.4878 | 1.355 | 1.3 | 1.24 | 1.385 | 1.18 | 1.3 |  | 1.6 | 2.007
| Std Dev. |  | 0.013784 | 0.031403 | 0.05164 |  | 0.030111 | 0.20213 | 0.027568 | 0.069186 | 0.050067 | 0.032143 | 0.030332 | 0.040825 | 0.064005 | 0.048339 | 0.02 | 0.040825 |  | 0.089443 | 1.9574
| Rel.Std.Dev. |  | 1.04% | 2.306% | 3.779% |  | 2.171% | 12.12% | 2.073% | 5.062% | 3.323% | 2.178% | 2.247% | 3.101% | 5.114% | 3.449% | 1.681% | 3.101% |  | 5.59% | 81.66%
| PDM3 |  | -3.687% | -1.002% | -0.6585% |  | 0.7953% | 21.27% | -3.324% | -0.6585% | 9.518% | 7.264% | -1.87% | -4.293% | -9.018% | 1.886% | -13.5% | -4.293% |  | 16.3% | 74.24%
| Z-Score (Absolute) |  | 0.47 | 0.19 | NA | 8.44 | 0 | 2.13 | 0.43 | 0.15 | 0.91 | 0.67 | 0.28 | NA | 1.02 | 0.11 | 1.49 | NA | 8.44 | NA | 7.66
Results from laboratories 1.07, 1.16, 1.21 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 211b (Certified Value 41.1 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 43.3 | 40.909 | 44.554 | 41 | 41 | 44 | 39.9 | 43.2 | 38.8 | 39.1 | 42.9105 | 38.26 | 41 | 41.8 | 39.8 | 40.5 | 45 | 41 | 39.608 |  | 1
| 1 | 2 | 43.8 | 42.318 | 44.665 | 40 | 41 | 43.2 | 40 | 42.7 | 40.9 | 38.9 | 42.358 | 38.65 | 43 | 41.1 | 39.3 | 40.8 | 47 | 39 | 39.577 |  | 37
| 1 | 3 | 41.4 | 41.2256 | 44.679 | 40 | 42 | 44.4 | 37.6 | 43.1 | 38 | 39.7 | 43.8272 | 38.39 | 41 | 40.5 | 39.7 | 39.3 | 40 | 41 | 39.613 |  | 16
| 1 | 4 | 42 | 41.3974 | 44.615 | 40 | 41 | 42.3 | 40.4 | 41.2 | 38 | 39.6 | 43.4421 | 37.61 | 43 | 40.6 | 40.2 | 39.7 | 41 | 40 | 38.817 |  | 41.1073
| 1 | 5 | 41.8 | 41.6529 | 44.573 | 39 | 42 | 43.8 | 41.6 | 42.3 | 40.5 | 41.2 | 44.2344 | 38.38 | 42 | 40.2 | 39.3 | 39.7 | 43 | 40 | 39.227 |  | 55
| 1 | 6 | 42.9 | 40.2177 | 44.585 | 39 | 41 | 43.4 | 40.2 | 42.1 | 39.5 | 38.7 | 44.8324 | 37.73 | 40 | 41.5 | 40.1 | 41.5 | 42 | 41 | 39.258
| Mean |  | 42.5333 | 41.2868 | 44.6118 | 39.8333 | 41.3333 | 43.5167 | 39.95 | 42.4333 | 39.2833 | 39.5333 | 43.6007 | 38.17 | 41.6667 | 40.95 | 39.7333 | 40.25 | 43 | 40.3333 | 39.35
| Median |  | 42.45 | 41.3115 | 44.6 | 40 | 41 | 43.6 | 40.1 | 42.5 | 39.15 | 39.35 | 43.6346 | 38.32 | 41.5 | 40.85 | 39.75 | 40.1 | 42.5 | 40.5 | 39.4175
| Std Dev. |  | 0.94163 | 0.707 | 0.050835 | 0.75277 | 0.5164 | 0.73326 | 1.305 | 0.74207 | 1.2384 | 0.9048 | 0.89618 | 0.40944 | 1.2111 | 0.62209 | 0.38297 | 0.82885 | 2.6077 | 0.8165 | 0.31462
| Rel.Std.Dev. |  | 2.214% | 1.712% | 0.1139% | 1.89% | 1.249% | 1.685% | 3.267% | 1.749% | 3.153% | 2.289% | 2.055% | 1.073% | 2.907% | 1.519% | 0.9639% | 2.059% | 6.064% | 2.024% | 0.7996%
| PDM3 |  | 3.469% | 0.4365% | 8.525% | -3.099% | 0.5498% | 5.861% | -2.815% | 3.226% | -4.437% | -3.829% | 6.066% | -7.146% | 1.361% | -0.3827% | -3.342% | -2.086% | 4.604% | -1.883% | -4.275%
| Z-Score (Absolute) |  | 0.75 | 0.16 | 1.74 | 0.53 | 0.18 | 1.22 | 0.48 | 0.71 | 0.79 | 0.67 | 1.26 | 1.32 | 0.34 | 0 | 0.58 | 0.33 | 0.98 | 0.29 | 0.76
| Analytical results for Se in OREAS 211b (Indicative Value 0.84 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.5 | < 1.5 | 0.4 | < 1 | 1 | 1 | < 1 | 1 | 1 | < 2 | 0.47 | < 1 | 1.2 |  | 0.9 | 2.06 | < 5 | 6.032 |  |  | 1
| 1 | 2 | 0.5 | < 1.5 | 0.5 | < 1 | 1 | 1 | < 1 | 1 | < 1 | < 2 | 0.52 | < 1 | 1 |  | 1 | 1.57 | < 5 | 4.411 |  |  | 38
| 1 | 3 | 0.5 | < 1.5 | 0.5 | < 1 | 1 | 1 | 1 | < 1 | 1 | < 2 | 0.47 | < 1 | 0.9 |  | 0.9 | 1.82 | < 5 | 7.209 |  |  | 16
| 1 | 4 |  | < 1.5 | 0.5 | < 1 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.48 | < 1 | 1.4 |  | 0.9 | 2.14 | < 5 | 1.75 |  |  | 0.8363
| 1 | 5 | 0.7 | < 1.5 | 0.5 | < 1 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.49 | < 1 | 1.6 | 0.7 | 0.5 | 1.82 | < 5 | 5.13 |  |  | 7
| 1 | 6 | 0.6 | < 1.5 | 0.5 | < 1 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.47 | < 1 | 0.8 | 0.6 | 0.7 | 1.75 | < 5 | 5.071
| Mean |  | 0.56 |  | 0.48333 |  | 1 | 1 | 1 | 1 | 1 |  | 0.48333 |  | 1.15 | 0.65 | 0.81667 | 1.86 |  | 4.9338
| Median |  | 0.5 |  | 0.5 |  | 1 | 1 | 1 | 1 | 1 |  | 0.475 |  | 1.1 | 0.65 | 0.9 | 1.82 |  | 5.1005
| Std Dev. |  | 0.089443 |  | 0.040825 |  | 0 | 0 | 0 | 0 | 0 |  | 0.019664 |  | 0.30822 | 0.070711 | 0.18348 | 0.20871 |  | 1.8357
| Rel.Std.Dev. |  | 15.97% |  | 8.447% |  | 0% | 0% | 0% | 0% | 0% |  | 4.068% |  | 26.8% | 10.88% | 22.47% | 11.22% |  | 37.21%
| PDM3 |  | -33.04% |  | -42.21% |  | 19.57% | 19.57% | 19.57% | 19.57% | 19.57% |  | -42.21% |  | 37.51% | -22.28% | -2.348% | 122.4% |  | 490%
| Z-Score (Absolute) |  | 0.85 | 0.34 | 1.06 | 1.01 | 0.34 | 0.34 | 0.11 | 0.11 | 0.11 | 0.34 | 1.06 | 1.01 | 0.74 | 0.61 | 0.16 | 2.66 | 4.38 | 10.95
| Analytical results for Sm in OREAS 211b (Certified Value 2.52 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.59 | 2.4192 | 2.9916 | 2.5 | 2.4 | 2.52 | 2.4 | 2.39 | 2.7 | [11 empty] | 1
| 1 | 2 | 2.55 | 2.6474 | 2.9114 | 2.6 | 2.4 | 2.5 | 2.4 | 2.7 | 2.5 | [11 empty] | 4
| 1 | 3 | 2.57 | 2.5333 | 3.0674 | 2.4 | 2.4 | 2.57 | 2.4 | 2.61 | 2.5 | [11 empty] | 16
| 1 | 4 | 2.66 | 2.562 | 2.9726 | 2.6 | 2.6 | 2.55 | 2.3 | 2.51 | 2.5 | [11 empty] | 2.5163
| 1 | 5 | 2.57 | 2.5304 | 2.9963 | 2.3 | 2.6 | 2.39 | 2.4 | 2.41 | 2.5 | [11 empty] | 56
| 1 | 6 | 2.61 | 2.4124 | 2.964 | 2.3 | 2.6 | 2.47 | 2.5 | 2.59 | 2.6
| Mean |  | 2.5917 | 2.5175 | 2.9839 | 2.45 | 2.5 | 2.5 | 2.4 | 2.535 | 2.55
| Median |  | 2.58 | 2.5319 | 2.9821 | 2.45 | 2.5 | 2.51 | 2.4 | 2.55 | 2.5
| Std Dev. |  | 0.0392 | 0.089449 | 0.050918 | 0.13784 | 0.10954 | 0.064498 | 0.063246 | 0.12095 | 0.083666
| Rel.Std.Dev. |  | 1.513% | 3.553% | 1.706% | 5.626% | 4.382% | 2.58% | 2.635% | 4.771% | 3.281%
| PDM3 |  | 2.995% | 0.04574% | 18.58% | -2.635% | -0.6479% | -0.6479% | -4.622% | 0.743% | 1.339%
| Z-Score (Absolute) |  | 1.54 | 0 | 9.66 | 1.4 | 0.36 | 0.36 | 2.43 | 0.36 | 0.67
| Analytical results for Sn in OREAS 211b (Certified Value 0.92 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 0.9 | 0.88513 | < 1 | 1 | 0.9 | 0.9 | 0.8 | 0.9 | 1 | 0.83738 | 0.8 | 0.9 | 0.9 | 0.9 | 0.8 | 0.3 | < 5 | 1.1 | 1.201 |  | 1
| 1 | 2 | 0.9 | 1.0217 | 7 | 1 | 0.9 | 1 | 0.8 | 0.9 | 1 | 0.81819 | 0.4 | 0.9 | 1 | 0.9 | 0.8 | 0.2 | < 5 | 1.1 | 2.055 |  | 11
| 1 | 3 | 0.9 | 1.0233 | < 1 | 1 | 0.9 | 0.8 | 0.8 | 0.8 | 1 | 0.81772 | 0.7 | 1 | 1.3 | 1 | 0.7 | 0.3 | < 5 | 1.1 | 1.852 |  | 16
| 1 | 4 | 0.8 | 1.0038 | < 1 | 0.9 | 0.9 | 1 | 0.8 | 0.8 | 1 | 0.86813 | 0.6 | 1 | 1.1 | 1 | 0.8 | 0.4 | < 5 | 1 | 1.349 |  | 0.9161
| 1 | 5 | 0.9 | 0.99005 | < 1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.9 | 1 | 0.87081 | 0.4 | 0.9 | 0.9 | 0.9 | 0.8 | < 0.2 | < 5 | 1 | 1.552 |  | 57
| 1 | 6 | 0.9 | 0.88129 | < 1 | 1 | 0.9 | 0.9 | 0.8 | 0.9 | 0.9 | 0.85496 | 0.6 | 0.9 | 1 | 0.9 | 0.8 | 0.6 | < 5 | 1.1 | 1.479
| Mean |  | 0.88333 | 0.96754 | 7 | 0.96667 | 0.9 | 0.91667 | 0.8 | 0.86667 | 0.98333 | 0.84453 | 0.58333 | 0.93333 | 1.0333 | 0.93333 | 0.78333 | 0.36 |  | 1.0667 | 1.5813
| Median |  | 0.9 | 0.99694 | 7 | 1 | 0.9 | 0.9 | 0.8 | 0.9 | 1 | 0.84617 | 0.6 | 0.9 | 1 | 0.9 | 0.8 | 0.3 |  | 1.1 | 1.5155
| Std Dev. |  | 0.040825 | 0.06647 |  | 0.05164 | 0 | 0.075277 | 1.2162e-16 | 0.05164 | 0.040825 | 0.023749 | 0.16021 | 0.05164 | 0.15055 | 0.05164 | 0.040825 | 0.15166 |  | 0.05164 | 0.31862
| Rel.Std.Dev. |  | 4.622% | 6.87% |  | 5.342% | 0% | 8.212% | 1.52e-14% | 5.958% | 4.152% | 2.812% | 27.46% | 5.533% | 14.57% | 5.533% | 5.212% | 42.13% |  | 4.841% | 20.15%
| PDM3 |  | -3.577% | 5.615% | 664.1% | 5.52% | -1.758% | 0.0618% | -12.67% | -5.396% | 7.339% | -7.812% | -36.32% | 1.881% | 12.8% | 1.881% | -14.49% | -60.7% |  | 16.44% | 72.62%
| Z-Score (Absolute) |  | 0.38 | 0.26 | 4.94 | 0.25 | 0.25 | 0.13 | 1.01 | 0.51 | 0.38 | 0.67 | 2.66 | 0 | 0.76 | 0 | 1.14 | 4.68 | 11.9 | 1.01 | 4.92
| Analytical results for Sr in OREAS 211b (Certified Value 213 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 216.89 | 208.6623 | 214.127 | 220 | 210 | 215 | 218 | 227 | 226 | 217 | 206.815 | 197 | 215 | 219 | 208.41 | 203 | 182 | 208.6 | 220 | 198.218 | 1
| 1 | 2 | 215.7 | 213.5106 | 212.1978 | 217 | 213 | 217 | 220 | 224 | 224 | 217 | 212.4676 | 203 | 213 | 217 | 208.43 | 204 | 197 | 209.7 | 209 | 199.61 | 12
| 1 | 3 | 210.76 | 209.4598 | 214.4238 | 222 | 217 | 215 | 214 | 223 | 229 | 222 | 210.8627 | 196 | 214 | 214 | 209.2 | 206 | 172 | 208.2 | 213 | 199.241 | 16
| 1 | 4 | 216.86 | 211.11 | 208.5726 | 217 | 208 | 212 | 218 | 214 | 227 | 219 | 211.2067 | 192 | 217 | 213 | 209.86 | 202 | 176 | 205.7 | 208 | 194.547 | 213.3548
| 1 | 5 | 213.86 | 215.3745 | 209.0072 | 218 | 217 | 212 | 226 | 223 | 227 | 211 | 205.6616 | 194 | 212 | 213 | 206.77 | 206 | 176 | 211.6 | 213 | 199.8 | 58
| 1 | 6 | 212.83 | 206.1741 | 208.6786 | 220 | 212 | 216 | 216 | 222 | 223 | 218 | 214.6012 | 193 | 210 | 222 | 209.3 | 208 | 178 | 202.8 | 212 | 196.72
| Mean |  | 214.4833 | 210.7152 | 211.1678 | 219 | 212.8333 | 214.5 | 218.6667 | 222.1667 | 226 | 217.3333 | 210.2691 | 195.8333 | 213.5 | 216.3333 | 208.6617 | 204.8333 | 180.1667 | 207.7667 | 212.5 | 198.0227
| Median |  | 214.78 | 210.2849 | 210.6025 | 219 | 212.5 | 215 | 218 | 223 | 226.5 | 217.5 | 211.0347 | 195 | 213.5 | 215.5 | 208.815 | 205 | 177 | 208.4 | 212.5 | 198.7295
| Std Dev. |  | 2.4457 | 3.3494 | 2.7575 | 2 | 3.656 | 2.0736 | 4.1312 | 4.3551 | 2.1909 | 3.6148 | 3.4056 | 3.9707 | 2.429 | 3.6697 | 1.0797 | 2.2286 | 8.8638 | 3.1066 | 4.2308 | 2.0483
| Rel.Std.Dev. |  | 1.14% | 1.59% | 1.306% | 0.9132% | 1.718% | 0.9667% | 1.889% | 1.96% | 0.9694% | 1.663% | 1.62% | 2.028% | 1.138% | 1.696% | 0.5174% | 1.088% | 4.92% | 1.495% | 1.991% | 1.034%
| PDM3 |  | 0.5289% | -1.237% | -1.025% | 2.646% | -0.2444% | 0.5368% | 2.49% | 4.13% | 5.927% | 1.865% | -1.446% | -8.212% | 0.06805% | 1.396% | -2.2% | -3.994% | -15.56% | -2.619% | -0.4007% | -7.186%
| Z-Score (Absolute) |  | 0.28 | 0.3 | 0.23 | 0.98 | 0.03 | 0.29 | 0.93 | 1.48 | 2.07 | 0.73 | 0.37 | 2.62 | 0.13 | 0.57 | 0.62 | 1.22 | 5.05 | 0.76 | 0.03 | 2.28
| Analytical results for Ta in OREAS 211b (Certified Value 0.22 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.28 | 0.27683 | < 0.1 | 0.42 | 0.22 | 0.23 | 0.24 | 0.21 | 0.22 | 0.23454 | 0.3 | 0.23 | 0.4 | 0.23 | 0.2 | 0.26 | 0.3 |  |  |  | 1
| 1 | 2 | 0.21 | 0.27008 | 0.1 | 0.41 | 0.22 | 0.25 | 0.23 | 0.21 | 0.23 | 0.22546 | 0.29 | 0.23 | 0.5 | 0.22 | 0.2 | 0.23 | 0.3 |  |  |  | 25
| 1 | 3 | 0.21 | 0.25538 | < 0.1 | 0.48 | 0.22 | 0.23 | 0.23 | 0.21 | 0.23 | 0.23146 | 0.28 | 0.24 | 0.4 | 0.22 | 0.2 | 0.22 | 0.3 |  |  |  | 16
| 1 | 4 | 0.23 | 0.25911 | < 0.1 | 0.4 | 0.21 | 0.23 | 0.21 | 0.23 | 0.23 | 0.22564 | 0.28 | 0.24 | 0.5 | 0.22 | 0.2 | 0.24 | 0.3 |  |  |  | 0.2247
| 1 | 5 | 0.22 | 0.27088 | < 0.1 | 0.47 | 0.22 | 0.24 | 0.24 | 0.22 | 0.23 | 0.22957 | 0.29 | 0.24 | 0.4 | 0.23 | 0.2 | 0.22 | 0.3 |  |  |  | 59
| 1 | 6 | 0.23 | 0.26453 | < 0.1 | 0.43 | 0.22 | 0.23 | 0.22 | 0.22 | 0.25 | 0.23352 | 0.28 | 0.22 | 0.4 | 0.23 | 0.2 | 0.23 | 0.3
| Mean |  | 0.23 | 0.26614 | 0.1 | 0.435 | 0.21833 | 0.235 | 0.22833 | 0.21667 | 0.23167 | 0.23003 | 0.28667 | 0.23333 | 0.43333 | 0.225 | 0.2 | 0.23333 | 0.3
| Median |  | 0.225 | 0.2673 | 0.1 | 0.425 | 0.22 | 0.23 | 0.23 | 0.215 | 0.23 | 0.23051 | 0.285 | 0.235 | 0.4 | 0.225 | 0.2 | 0.23 | 0.3
| Std Dev. |  | 0.026077 | 0.0080014 |  | 0.032711 | 0.0040825 | 0.0083666 | 0.01169 | 0.008165 | 0.0098319 | 0.0038702 | 0.008165 | 0.008165 | 0.05164 | 0.0054772 | 3.0405e-17 | 0.015055 | 0
| Rel.Std.Dev. |  | 11.34% | 3.006% |  | 7.52% | 1.87% | 3.56% | 5.12% | 3.768% | 4.244% | 1.682% | 2.848% | 3.499% | 11.92% | 2.434% | 1.52e-14% | 6.452% | 0%
| PDM3 |  | 2.359% | 18.44% | -55.5% | 93.59% | -2.833% | 4.584% | 1.617% | -3.575% | 3.101% | 2.373% | 27.58% | 3.842% | 92.85% | 0.1336% | -10.99% | 3.842% | 33.51%
| Z-Score (Absolute) |  | 0.11 | 4.76 | NA | 27.53 | 1.69 | 0.56 | 0.34 | 1.91 | 0.11 | 0.11 | 7.53 | 0.34 | NA | 0.79 | 4.16 | 0.34 | NA
Results from laboratories 1.07, 1.17 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 211b (Certified Value 0.58 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 0.55 | 0.5549 | 0.68525 | 0.6 | 0.5 | 0.55 | 0.5 | 0.56 | 0.56 | 2.41 | [10 empty] | 1
| 1 | 2 | 0.53 | 0.58658 | 0.69488 | 0.6 | 0.5 | 0.54 | 0.5 | 0.61 | 0.56 | 2.937 | [10 empty] | 26
| 1 | 3 | 0.53 | 0.56977 | 0.71873 | 0.6 | 0.5 | 0.53 | 0.5 | 0.58 | 0.54 | 2.154 | [10 empty] | 16
| 1 | 4 | 0.56 | 0.58024 | 0.68777 | 0.6 | 0.6 | 0.53 | 0.5 | 0.61 | 0.56 | 2.03 | [10 empty] | 0.57997
| 1 | 5 | 0.53 | 0.58522 | 0.66203 | 0.6 | 0.5 | 0.52 | 0.5 | 0.58 | 0.56 | 2.226 | [10 empty] | 60
| 1 | 6 | 0.52 | 0.54487 | 0.69875 | 0.6 | 0.6 | 0.55 | 0.5 | 0.61 | 0.54 | 2.282
| Mean |  | 0.53667 | 0.57026 | 0.69124 | 0.6 | 0.53333 | 0.53667 | 0.5 | 0.59167 | 0.55333 | 2.3398
| Median |  | 0.53 | 0.575 | 0.69132 | 0.6 | 0.5 | 0.535 | 0.5 | 0.595 | 0.56 | 2.254
| Std Dev. |  | 0.015055 | 0.017148 | 0.018582 | 0 | 0.05164 | 0.012111 | 0 | 0.02137 | 0.010328 | 0.31886
| Rel.Std.Dev. |  | 2.805% | 3.007% | 2.688% | 0% | 9.682% | 2.257% | 0% | 3.612% | 1.866% | 13.63%
| PDM3 |  | -7.467% | -1.674% | 19.18% | 3.453% | -8.042% | -7.467% | -13.79% | 2.016% | -4.593% | 303.4%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.43 | NA | NA | 0.67 | NA | 0.43 | 0.34 | 35.52
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 211b (Certified Value 0.081 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.1 | 0.29 | 0.08 | 0.08 | < 0.05 | 0.1 | 0.05 | 0.10759 | 0.07 | 0.1 | < 0.05 | 0.12 | < 0.05 | < 0.2 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.29 | 0.07 | 0.06 | 0.06 | 0.12 | 0.06 | < 0.1 | 0.08 | 0.1 | 0.06 | 0.11 | < 0.05 | < 0.2 | [5 empty] | 27
| 1 | 3 | < 0.2 | < 0.1 | 0.3 | 0.08 | 0.06 | 0.06 | 0.1 | 0.06 | < 0.1 | 0.08 | 0.1 | 0.09 | 0.1 | 0.07 | < 0.2 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.3 | 0.08 | 0.06 | 0.06 | 0.12 | 0.07 | < 0.1 | 0.04 | 0.1 | 0.09 | 0.09 | 0.06 | < 0.2 | [5 empty] | 0.080686
| 1 | 5 | < 0.2 | < 0.1 | 0.3 | 0.07 | 0.11 | < 0.05 | 0.12 | 0.07 | < 0.1 | 0.09 | 0.1 | 0.06 | 0.09 | 0.05 | < 0.2 | [5 empty] | 61
| 1 | 6 | < 0.2 | < 0.1 | 0.32 | 0.08 | 0.09 | 0.05 | 0.11 | 0.08 | 0.09779 | 0.08 | 0.1 | < 0.05 | 0.1 | 0.06 | < 0.2
| Mean |  |  |  | 0.3 | 0.076667 | 0.076667 | 0.0575 | 0.11167 | 0.065 | 0.10269 | 0.073333 | 0.1 | 0.075 | 0.10167 | 0.06
| Median |  |  |  | 0.3 | 0.08 | 0.07 | 0.06 | 0.115 | 0.065 | 0.10269 | 0.08 | 0.1 | 0.075 | 0.1 | 0.06
| Std Dev. |  |  |  | 0.010954 | 0.005164 | 0.020656 | 0.005 | 0.0098319 | 0.010488 | 0.0069318 | 0.017512 | 1.5202e-17 | 0.017321 | 0.01169 | 0.008165
| Rel.Std.Dev. |  |  |  | 3.651% | 6.736% | 26.94% | 8.696% | 8.805% | 16.14% | 6.75% | 23.88% | 1.52e-14% | 23.09% | 11.5% | 13.61%
| PDM3 |  |  |  | 271.8% | -4.981% | -4.981% | -28.74% | 38.4% | -19.44% | 27.27% | -9.112% | 23.94% | -7.047% | 26% | -25.64%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 6.07 | 0.04 | 0.04 | 0.76 | 0.99 | 0.27 | 0.2 | 0.04 | NA | 0.45 | 0.72 | 0.72 | 0.67
Results from laboratory 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 211b (Certified Value 0.74 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.75 | 0.74528 | 1.3824 | 0.9 | 0.75 | 0.7 | 0.78 | 0.7 | 0.7 | 0.7702 | 0.75 | 0.79 | 0.7 | 0.76 | 0.7 | 0.74 | 0.7 |  |  |  | 1
| 1 | 2 | 0.75 | 0.77015 | 1.4733 | 0.9 | 0.72 | 0.69 | 0.79 | 0.73 | 0.69 | 0.76557 | 0.78 | 0.82 | 0.7 | 0.74 | 0.7 | 0.65 | 0.7 |  |  |  | 28
| 1 | 3 | 0.76 | 0.75642 | 1.4038 | 0.9 | 0.75 | 0.65 | 0.77 | 0.7 | 0.7 | 0.76008 | 0.77 | 0.83 | 0.7 | 0.79 | 0.7 | 0.63 | 0.8 |  |  |  | 16
| 1 | 4 | 0.77 | 0.72589 | 1.3905 | 0.9 | 0.72 | 0.69 | 0.73 | 0.68 | 0.69 | 0.74504 | 0.74 | 0.85 | 0.7 | 0.75 | 0.7 | 0.62 | 0.7 |  |  |  | 0.73983
| 1 | 5 | 0.72 | 0.78009 | 1.3958 | 0.8 | 0.75 | 0.73 | 0.78 | 0.77 | 0.69 | 0.75063 | 0.78 | 0.84 | 0.7 | 0.78 | 0.7 | 0.66 | 0.7 |  |  |  | 62
| 1 | 6 | 0.74 | 0.72344 | 1.3995 | 0.9 | 0.76 | 0.71 | 0.75 | 0.72 | 0.74 | 0.77281 | 0.76 | 0.78 | 0.8 | 0.75 | 0.7 | 0.62 | 0.8
| Mean |  | 0.74833 | 0.75021 | 1.4076 | 0.88333 | 0.74167 | 0.695 | 0.76667 | 0.71667 | 0.70167 | 0.76072 | 0.76333 | 0.81833 | 0.71667 | 0.76167 | 0.7 | 0.65333 | 0.73333
| Median |  | 0.75 | 0.75085 | 1.3977 | 0.9 | 0.75 | 0.695 | 0.775 | 0.71 | 0.695 | 0.76283 | 0.765 | 0.825 | 0.7 | 0.755 | 0.7 | 0.64 | 0.7
| Std Dev. |  | 0.017224 | 0.023073 | 0.033054 | 0.040825 | 0.017224 | 0.026646 | 0.022509 | 0.031411 | 0.019408 | 0.011022 | 0.01633 | 0.027869 | 0.040825 | 0.019408 | 1.2162e-16 | 0.045461 | 0.05164
| Rel.Std.Dev. |  | 2.302% | 3.076% | 2.348% | 4.622% | 2.322% | 3.834% | 2.936% | 4.383% | 2.766% | 1.449% | 2.139% | 3.406% | 5.696% | 2.548% | 1.737e-14% | 6.958% | 7.042%
| PDM3 |  | 1.15% | 1.404% | 90.25% | 19.4% | 0.2487% | -6.059% | 3.628% | -3.13% | -5.158% | 2.824% | 3.177% | 10.61% | -3.13% | 2.952% | -5.383% | -11.69% | -0.8777%
| Z-Score (Absolute) |  | 0.08 | 0 | 26.94 | NA | 0.35 | 2.26 | 0.67 | 1.37 | 1.99 | 0.43 | 0.54 | 2.79 | NA | 0.47 | NA | 3.97 | NA
Results from laboratories 1.07, 1.17, 1.19 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 211b (Certified Value 0.633 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.6167 | 0.62908 | 0.62758 | 0.639 | 0.68 | 0.597 | 0.606 | 0.648 | 0.613 | 0.616 | 0.67662 | 0.6873 | 0.674 | 0.61 | 0.6385 | 0.646 | 0.552 | 0.57 | 0.625 | 0.61374 | 1
| 1 | 2 | 0.6348 | 0.65909 | 0.62794 | 0.71 | 0.68 | 0.606 | 0.621 | 0.638 | 0.608 | 0.617 | 0.67444 | 0.6865 | 0.656 | 0.607 | 0.6415 | 0.643 | 0.626 | 0.57 | 0.605 | 0.61831 | 29
| 1 | 3 | 0.6291 | 0.62516 | 0.62819 | 0.685 | 0.68 | 0.591 | 0.604 | 0.626 | 0.608 | 0.634 | 0.67482 | 0.6847 | 0.668 | 0.593 | 0.6445 | 0.635 | 0.546 | 0.57 | 0.62 | 0.6156 | 16
| 1 | 4 | 0.6423 | 0.63801 | 0.62914 | 0.703 | 0.67 | 0.592 | 0.606 | 0.605 | 0.619 | 0.616 | 0.67803 | 0.672 | 0.665 | 0.59 | 0.651 | 0.639 | 0.56 | 0.57 | 0.61 | 0.61054 | 0.63328
| 1 | 5 | 0.6262 | 0.63541 | 0.62794 | 0.73 | 0.69 | 0.596 | 0.629 | 0.634 | 0.628 | 0.61 | 0.67577 | 0.6849 | 0.648 | 0.588 | 0.6313 | 0.645 | 0.539 | 0.56 | 0.61 | 0.61306 | 63
| 1 | 6 | 0.6373 | 0.6174 | 0.62971 | 0.68 | 0.67 | 0.598 | 0.608 | 0.625 | 0.609 | 0.618 | 0.67704 | 0.6824 | 0.653 | 0.624 | 0.6524 | 0.655 | 0.558 | 0.55 | 0.62 | 0.60582
| Mean |  | 0.63107 | 0.63403 | 0.62842 | 0.69117 | 0.67833 | 0.59667 | 0.61233 | 0.62933 | 0.61417 | 0.6185 | 0.67612 | 0.68297 | 0.66067 | 0.602 | 0.6432 | 0.64383 | 0.5635 | 0.565 | 0.615 | 0.61285
| Median |  | 0.63195 | 0.63225 | 0.62807 | 0.694 | 0.68 | 0.5965 | 0.607 | 0.63 | 0.611 | 0.6165 | 0.67619 | 0.6848 | 0.6605 | 0.6 | 0.643 | 0.644 | 0.555 | 0.57 | 0.615 | 0.6134
| Std Dev. |  | 0.0090877 | 0.01432 | 0.00082527 | 0.031263 | 0.0075277 | 0.0053541 | 0.010211 | 0.01461 | 0.0079854 | 0.0080932 | 0.00137 | 0.0056319 | 0.009913 | 0.0141 | 0.0079216 | 0.006824 | 0.031583 | 0.0083666 | 0.007746 | 0.0043092
| Rel.Std.Dev. |  | 1.44% | 2.259% | 0.1313% | 4.523% | 1.11% | 0.8973% | 1.668% | 2.322% | 1.3% | 1.309% | 0.2026% | 0.8246% | 1.5% | 2.342% | 1.232% | 1.06% | 5.605% | 1.481% | 1.26% | 0.7031%
| PDM3 |  | -0.3494% | 0.1179% | -0.7678% | 9.141% | 7.114% | -5.781% | -3.308% | -0.6231% | -3.018% | -2.334% | 6.765% | 7.846% | 4.325% | -4.939% | 1.567% | 1.667% | -11.02% | -10.78% | -2.886% | -3.227%
| Z-Score (Absolute) |  | 0.09 | 0.21 | 0.02 | 2.58 | 2.05 | 1.33 | 0.68 | 0.02 | 0.61 | 0.43 | 1.96 | 2.24 | 1.32 | 1.11 | 0.59 | 0.62 | 2.71 | 2.64 | 0.57 | 0.66
| Analytical results for Tl in OREAS 211b (Certified Value 0.15 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 0.16 | 0.14931 | 0.15 | 0.15 | 0.16 | 0.18 | 0.14 | 0.17 | 0.16121 | 0.14 | 0.15 | 0.15 | 0.15 | 0.15 | 0.17 | 0.2 |  |  |  |  | 1
| 1 | 2 | 0.15 | 0.14966 | 0.15 | 0.15 | 0.16 | 0.19 | 0.15 | 0.15 | 0.14604 | 0.141 | 0.16 | 0.14 | 0.15 | 0.16 | 0.15 | 0.1 | 0.257 |  |  |  | 30
| 1 | 3 | 0.15 | 0.15039 | 0.15 | 0.15 | 0.14 | 0.16 | 0.13 | 0.17 | 0.14935 | 0.14 | 0.15 | 0.15 | 0.16 | 0.15 | 0.15 | 0.2 |  |  |  |  | 16
| 1 | 4 | 0.16 | 0.14822 | 0.16 | 0.15 | 0.15 | 0.17 | 0.13 | 0.17 | 0.16229 | 0.133 | 0.16 | 0.15 | 0.15 | 0.15 | 0.15 | 0.1 |  |  |  |  | 0.15263
| 1 | 5 | 0.15 | 0.15405 | 0.15 | 0.16 | 0.16 | 0.17 | 0.14 | 0.16 | 0.15572 | 0.142 | 0.15 | 0.16 | 0.16 | 0.15 | 0.15 | 0.1 | 2.744 |  |  |  | 64
| 1 | 6 | 0.16 | 0.14146 | 0.15 | 0.15 | 0.16 | 0.17 | 0.14 | 0.17 | 0.15676 | 0.142 | 0.15 | 0.15 | 0.16 | 0.15 | 0.15 | 0.1
| Mean |  | 0.155 | 0.14885 | 0.15167 | 0.15167 | 0.155 | 0.17333 | 0.13833 | 0.165 | 0.15523 | 0.13967 | 0.15333 | 0.15 | 0.155 | 0.15167 | 0.15333 | 0.13333 | 1.5005
| Median |  | 0.155 | 0.14949 | 0.15 | 0.15 | 0.16 | 0.17 | 0.14 | 0.17 | 0.15624 | 0.1405 | 0.15 | 0.15 | 0.155 | 0.15 | 0.15 | 0.1 | 1.5005
| Std Dev. |  | 0.0054772 | 0.0041304 | 0.0040825 | 0.0040825 | 0.0083666 | 0.010328 | 0.0075277 | 0.0083666 | 0.006439 | 0.0033862 | 0.005164 | 0.0063246 | 0.0054772 | 0.0040825 | 0.008165 | 0.05164 | 1.7586
| Rel.Std.Dev. |  | 3.534% | 2.775% | 2.692% | 2.692% | 5.398% | 5.958% | 5.442% | 5.071% | 4.148% | 2.425% | 3.368% | 4.216% | 3.534% | 2.692% | 5.325% | 38.73% | 117.2%
| PDM3 |  | 1.555% | -2.475% | -0.6287% | -0.6287% | 1.555% | 13.57% | -9.365% | 8.107% | 1.705% | -8.491% | 0.4633% | -1.721% | 1.555% | -0.6287% | 0.4633% | -12.64% | 883.1%
| Z-Score (Absolute) |  | 0.63 | 1.7 | 0.63 | 0.63 | 0.63 | 7.57 | 5.68 | 4.42 | 0.72 | 5.17 | 0 | 1.26 | 0.63 | 0.63 | 0 | NA | 510.01
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 211b (Certified Value 0.33 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.35 | 0.30864 | 0.41967 | 0.3 | 0.3 | 0.35 | 0.3 | 0.32 | 0.3 | [11 empty] | 1
| 1 | 2 | 0.34 | 0.3843 | 0.39753 | 0.3 | 0.3 | 0.34 | 0.3 | 0.33 | 0.3 | [11 empty] | 5
| 1 | 3 | 0.33 | 0.34667 | 0.41283 | 0.3 | 0.3 | 0.34 | 0.3 | 0.33 | 0.3 | [11 empty] | 16
| 1 | 4 | 0.35 | 0.32206 | 0.40581 | 0.3 | 0.3 | 0.34 | 0.3 | 0.33 | 0.3 | [11 empty] | 0.32838
| 1 | 5 | 0.35 | 0.32603 | 0.39105 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3 | [11 empty] | 65
| 1 | 6 | 0.35 | 0.32206 | 0.39492 | 0.3 | 0.4 | 0.34 | 0.3 | 0.32 | 0.3
| Mean |  | 0.345 | 0.33496 | 0.40363 | 0.3 | 0.31667 | 0.34167 | 0.3 | 0.32333 | 0.3
| Median |  | 0.35 | 0.32404 | 0.40167 | 0.3 | 0.3 | 0.34 | 0.3 | 0.325 | 0.3
| Std Dev. |  | 0.0083666 | 0.027121 | 0.011118 | 0 | 0.040825 | 0.0040825 | 0 | 0.008165 | 0
| Rel.Std.Dev. |  | 2.425% | 8.097% | 2.754% | 0% | 12.89% | 1.195% | 0% | 2.525% | 0%
| PDM3 |  | 5.062% | 2.004% | 22.92% | -8.642% | -3.566% | 4.047% | -8.642% | -1.536% | -8.642%
| Z-Score (Absolute) |  | 0.67 | 0.36 | 2.5 | 0.73 | 0.21 | 0.57 | 0.73 | 0 | 0.73
| Analytical results for U in OREAS 211b (Certified Value 0.2 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.22 | 0.20929 | 0.23457 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.23504 | 0.21 | 0.22 | 0.2 | 0.23 | 0.2 | 0.26 | 0.2 |  |  |  | 1
| 1 | 2 | 0.2 | 0.21814 | 0.31857 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.22364 | 0.21 | 0.22 | 0.2 | 0.23 | 0.2 | 0.19 | 0.3 |  |  |  | 32
| 1 | 3 | 0.2 | 0.24282 | 0.3304 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.23144 | 0.2 | 0.22 | 0.2 | 0.25 | 0.2 | 0.18 | 0.2 |  |  |  | 16
| 1 | 4 | 0.22 | 0.20353 | 0.28644 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21334 | 0.2 | 0.23 | 0.2 | 0.24 | 0.2 | 0.18 | 0.3 |  |  |  | 0.20453
| 1 | 5 | 0.21 | 0.24478 | 0.33628 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.22022 | 0.2 | 0.22 | 0.2 | 0.25 | 0.2 | 0.23 | 0.3 |  |  |  | 66
| 1 | 6 | 0.21 | 0.21992 | 0.31241 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21873 | 0.2 | 0.22 | 0.2 | 0.23 | 0.2 | 0.18 | 0.2
| Mean |  | 0.21 | 0.22308 | 0.30311 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.22373 | 0.20333 | 0.22167 | 0.2 | 0.23833 | 0.2 | 0.20333 | 0.25
| Median |  | 0.21 | 0.21903 | 0.31549 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.22193 | 0.2 | 0.22 | 0.2 | 0.235 | 0.2 | 0.185 | 0.25
| Std Dev. |  | 0.0089443 | 0.017131 | 0.037805 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 0.0081576 | 0.005164 | 0.0040825 | 3.0405e-17 | 0.0098319 | 3.0405e-17 | 0.033862 | 0.054772
| Rel.Std.Dev. |  | 4.259% | 7.679% | 12.47% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 3.646% | 2.54% | 1.842% | 1.52e-14% | 4.125% | 1.52e-14% | 16.65% | 21.91%
| PDM3 |  | 2.673% | 9.068% | 48.2% | -2.216% | -2.216% | -2.216% | -2.216% | -2.216% | -2.216% | 9.388% | -0.5862% | 8.377% | -2.216% | 16.53% | -2.216% | -0.5862% | 22.23%
| Z-Score (Absolute) |  | 1.35 | 3.99 | 20.18 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 4.13 | 0 | 3.71 | 0.67 | 7.08 | 0.67 | 0 | 9.44
| Analytical results for V in OREAS 211b (Certified Value 282 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 279 | 279.1145 | 276.6395 | 217 | 296 | 272 | 272 | 289 | 274 | 269 | 305.7522 | 295 | 282 | 287 | 281 | 292 | 258 | 289.3 | 280 | 284.036 | 1
| 1 | 2 | 287 | 288.698 | 278.6135 | 281 | 298 | 273 | 270 | 283 | 269 | 267 | 294.8246 | 296 | 293 | 285 | 281 | 287 | 271 | 285.6 | 275 | 283.927 | 33
| 1 | 3 | 291 | 279.6606 | 278.603 | 177 | 299 | 269 | 271 | 280 | 273 | 279 | 296.7731 | 295 | 287 | 279 | 282 | 285 | 249 | 281.6 | 280 | 283.283 | 16
| 1 | 4 | 296 | 281.0581 | 285.264 | 149 | 292 | 268 | 274 | 270 | 278 | 272 | 292.6636 | 290 | 275 | 279 | 284 | 283 | 246 | 281.9 | 275 | 277.539 | 281.8985
| 1 | 5 | 279 | 285.4342 | 273.9975 | 163 | 292 | 269 | 284 | 282 | 275 | 266 | 304.0569 | 294 | 280 | 277 | 270 | 287 | 262 | 284.6 | 275 | 283.393 | 67
| 1 | 6 | 277 | 271.7971 | 280.1925 | 199 | 293 | 271 | 273 | 280 | 268 | 267 | 300.5576 | 293 | 275 | 288 | 280 | 292 | 253 | 285.9 | 275 | 281.463
| Mean |  | 284.8333 | 280.9604 | 278.885 | 197.6667 | 295 | 270.3333 | 274 | 280.6667 | 272.8333 | 270 | 299.1047 | 293.8333 | 282 | 282.5 | 279.6667 | 287.6667 | 256.5 | 284.8167 | 276.6667 | 282.2735
| Median |  | 283 | 280.3594 | 278.6082 | 188 | 294.5 | 270 | 272.5 | 281 | 273.5 | 268 | 298.6653 | 294.5 | 281 | 282 | 281 | 287 | 255.5 | 285.1 | 275 | 283.338
| Std Dev. |  | 7.705 | 5.8131 | 3.7811 | 47.5759 | 3.0984 | 1.9664 | 5.099 | 6.186 | 3.7639 | 4.899 | 5.2167 | 2.137 | 7.0427 | 4.7223 | 4.9261 | 3.6697 | 9.1815 | 2.8562 | 2.582 | 2.4974
| Rel.Std.Dev. |  | 2.705% | 2.069% | 1.356% | 24.07% | 1.05% | 0.7274% | 1.861% | 2.204% | 1.38% | 1.814% | 1.744% | 0.7273% | 2.497% | 1.672% | 1.761% | 1.276% | 3.58% | 1.003% | 0.9332% | 0.8847%
| PDM3 |  | 1.041% | -0.3328% | -1.069% | -29.88% | 4.648% | -4.103% | -2.802% | -0.437% | -3.216% | -4.221% | 6.104% | 4.234% | 0.03599% | 0.2134% | -0.7917% | 2.046% | -9.01% | 1.035% | -1.856% | 0.133%
| Z-Score (Absolute) |  | 0.49 | 0.02 | 0.24 | 10.23 | 1.75 | 1.29 | 0.84 | 0.02 | 0.98 | 1.33 | 2.25 | 1.6 | 0.15 | 0.21 | 0.14 | 0.84 | 2.99 | 0.49 | 0.51 | 0.18
| Analytical results for W in OREAS 211b (Certified Value 16.8 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 16.6 | 16.1126 | 16.2 | 15.8 | 17.6 | 16.4 | 17.8 | 15.8 | 16.9 | 17.219 | 18.12 | 16.6 | 15.9 | 17.4 | 15.9 | 19.6 | < 5 |  |  |  | 1
| 1 | 2 | 16.2 | 16.9346 | 16.7 | 16 | 17.1 | 16.8 | 17.8 | 16 | 17.2 | 16.5604 | 18.92 | 16.9 | 15.9 | 17.2 | 16.1 | 17.8 | < 5 |  |  |  | 34
| 1 | 3 | 16.3 | 16.1895 | 16.4 | 16.6 | 17.4 | 16.3 | 17.8 | 14.8 | 17.7 | 17.078 | 17.25 | 16.8 | 16 | 17.6 | 15.7 | 17.4 | < 5 |  |  |  | 16
| 1 | 4 | 17 | 16.3883 | 16.3 | 15.9 | 16.7 | 17.1 | 16.6 | 16.6 | 17.4 | 16.5954 | 18.33 | 17 | 16.6 | 17.7 | 15.7 | 17.4 | < 5 |  |  |  | 16.8373
| 1 | 5 | 16.3 | 16.0356 | 16.8 | 17.4 | 17.4 | 17.8 | 18 | 16.2 | 16.8 | 16.5737 | 19.32 | 17 | 16 | 17.3 | 16.4 | 17.4 | < 5 |  |  |  | 68
| 1 | 6 | 16.4 | 16.0352 | 16.6 | 15.8 | 17.2 | 16.5 | 17.8 | 15.9 | 18.6 | 16.8622 | 19.23 | 16.5 | 17.6 | 17.3 | 16.2 | 17.2 | < 5
| Mean |  | 16.4667 | 16.2826 | 16.5 | 16.25 | 17.2333 | 16.8167 | 17.6333 | 15.8833 | 17.4333 | 16.8148 | 18.5283 | 16.8 | 16.3333 | 17.4167 | 16 | 17.8
| Median |  | 16.35 | 16.1511 | 16.5 | 15.95 | 17.3 | 16.65 | 17.8 | 15.95 | 17.3 | 16.7288 | 18.625 | 16.85 | 16 | 17.35 | 16 | 17.4
| Std Dev. |  | 0.29439 | 0.34528 | 0.23664 | 0.63797 | 0.31411 | 0.56362 | 0.51251 | 0.60139 | 0.65929 | 0.28489 | 0.78901 | 0.20976 | 0.67429 | 0.19408 | 0.28284 | 0.90333
| Rel.Std.Dev. |  | 1.788% | 2.121% | 1.434% | 3.926% | 1.823% | 3.352% | 2.906% | 3.786% | 3.782% | 1.694% | 4.258% | 1.249% | 4.128% | 1.114% | 1.768% | 5.075%
| PDM3 |  | -2.201% | -3.294% | -2.004% | -3.488% | 2.352% | -0.1228% | 4.728% | -5.666% | 3.54% | -0.134% | 10.04% | -0.2218% | -2.993% | 3.441% | -4.973% | 5.717%
| Z-Score (Absolute) |  | 0.41 | 0.63 | 0.37 | 0.67 | 0.53 | 0.02 | 1.02 | 1.12 | 0.78 | 0.02 | 2.12 | 0 | 0.57 | 0.76 | 0.98 | 1.23 | 17.53
| Analytical results for Y in OREAS 211b (Certified Value 20.3 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 21.19 | 20.1978 | 22.7936 | 19.2 | 20.5 | 22.1 | 20.7 | 20.6 | 20.2 | 20.1 | 21.1261 | 20.13 | 19.6 | 19.8 | 20.22 | 20.3 | 20.7 | 19 | 19.174 |  | 1
| 1 | 2 | 21.21 | 20.7826 | 22.3535 | 18.8 | 20.3 | 21.6 | 20.8 | 20.8 | 20.1 | 19.8 | 20.9907 | 20.13 | 20.3 | 19.3 | 20.73 | 20.3 | 22.4 | 19 | 19.315 |  | 35
| 1 | 3 | 21.11 | 20.2193 | 23.1482 | 19.3 | 21.5 | 22 | 19.7 | 20.3 | 20.1 | 20.3 | 21.1191 | 20.16 | 20.3 | 19 | 20.35 | 19.9 | 22.1 | 19.7 | 19.177 |  | 16
| 1 | 4 | 21.08 | 20.4021 | 23.0825 | 18.9 | 19.3 | 20.9 | 20.7 | 19.5 | 19.7 | 20.4 | 21.0182 | 19.79 | 20.7 | 19.9 | 20.59 | 20 | 21.8 | 19 | 18.812 |  | 20.3193
| 1 | 5 | 21.03 | 20.9474 | 22.0232 | 21.1 | 20.9 | 22 | 21.5 | 20.4 | 21.5 | 20.4 | 21.0175 | 20.1 | 20.3 | 19.6 | 20.62 | 21.1 | 20.7 | 19.3 | 19.049 |  | 69
| 1 | 6 | 20.84 | 19.9034 | 23.1464 | 19.2 | 19.5 | 21.4 | 20.9 | 20.3 | 20 | 21 | 21.1902 | 19.95 | 19.8 | 20.4 | 20.41 | 21 | 21.9 | 19.2 | 19.204
| Mean |  | 21.0767 | 20.4088 | 22.7579 | 19.4167 | 20.3333 | 21.6667 | 20.7167 | 20.3167 | 20.2667 | 20.3333 | 21.077 | 20.0433 | 20.1667 | 19.6667 | 20.4867 | 20.4333 | 21.6 | 19.2 | 19.1218
| Median |  | 21.095 | 20.3107 | 22.9381 | 19.2 | 20.4 | 21.8 | 20.75 | 20.35 | 20.1 | 20.35 | 21.0687 | 20.115 | 20.3 | 19.7 | 20.5 | 20.3 | 21.85 | 19.1 | 19.1755
| Std Dev. |  | 0.13411 | 0.39133 | 0.47185 | 0.84715 | 0.83347 | 0.46332 | 0.58109 | 0.4446 | 0.62823 | 0.39833 | 0.079318 | 0.14473 | 0.39833 | 0.48854 | 0.19149 | 0.50465 | 0.72664 | 0.27568 | 0.17386
| Rel.Std.Dev. |  | 0.6363% | 1.917% | 2.073% | 4.363% | 4.099% | 2.138% | 2.805% | 2.188% | 3.1% | 1.959% | 0.3763% | 0.7221% | 1.975% | 2.484% | 0.9347% | 2.47% | 3.364% | 1.436% | 0.9092%
| PDM3 |  | 3.727% | 0.4402% | 12% | -4.442% | 0.06902% | 6.631% | 1.956% | -0.013% | -0.2591% | 0.06902% | 3.729% | -1.358% | -0.7512% | -3.212% | 0.8236% | 0.5612% | 6.303% | -5.509% | -5.893%
| Z-Score (Absolute) |  | 1.31 | 0.13 | 4.26 | 1.61 | 0 | 2.35 | 0.67 | 0.03 | 0.12 | 0 | 1.31 | 0.51 | 0.29 | 1.17 | 0.27 | 0.18 | 2.23 | 1.99 | 2.13
| Analytical results for Yb in OREAS 211b (Certified Value 2.18 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.27 | 2.0955 | 2.1481 | 2.3 | 2.1 | 2.21 | 2.2 | 2.06 | 2.35 | [11 empty] | 1
| 1 | 2 | 2.26 | 2.2588 | 2.1339 | 2.3 | 2 | 2.27 | 2.1 | 2.36 | 2.3 | [11 empty] | 6
| 1 | 3 | 2.24 | 2.2042 | 2.1915 | 2.3 | 2.1 | 2.29 | 2.1 | 2.22 | 2.35 | [11 empty] | 16
| 1 | 4 | 2.29 | 2.1081 | 2.115 | 2.4 | 2.2 | 2.22 | 2.1 | 2.15 | 2.3 | [11 empty] | 2.1796
| 1 | 5 | 2.22 | 2.2608 | 2.1159 | 2.2 | 2.2 | 2.25 | 2.2 | 2.07 | 2.3 | [11 empty] | 70
| 1 | 6 | 2.21 | 2.0126 | 2.1584 | 2.3 | 2.3 | 2.3 | 2.1 | 2.15 | 2.4
| Mean |  | 2.2483 | 2.1567 | 2.1438 | 2.3 | 2.15 | 2.2567 | 2.1333 | 2.1683 | 2.3333
| Median |  | 2.25 | 2.1561 | 2.141 | 2.3 | 2.15 | 2.26 | 2.1 | 2.15 | 2.325
| Std Dev. |  | 0.030605 | 0.10041 | 0.029038 | 0.063246 | 0.10488 | 0.036697 | 0.05164 | 0.11089 | 0.040825
| Rel.Std.Dev. |  | 1.361% | 4.656% | 1.355% | 2.75% | 4.878% | 1.626% | 2.421% | 5.114% | 1.75%
| PDM3 |  | 3.154% | -1.051% | -1.642% | 5.524% | -1.358% | 3.536% | -2.122% | -0.5165% | 7.054%
| Z-Score (Absolute) |  | 1.54 | 0.22 | 0.47 | 2.54 | 0.35 | 1.7 | 0.67 | 0 | 3.18
| Analytical results for Zn in OREAS 211b (Certified Value 126 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 130 | 124.7061 | 132.996 | 114 | 134 | 127 | 129 | 137 | 129 | 127 | 128.7384 | 112 | 122 | 121.7 | 125 | 117.1 | 106 | 148.9 | 124 | 119.914 | 1
| 1 | 2 | 131 | 131.3682 | 132.444 | 112 | 130 | 127 | 131 | 134 | 127 | 126 | 128.5954 | 113 | 128 | 120.3 | 124 | 119.2 | 108 | 147.6 | 126 | 119.888 | 15
| 1 | 3 | 126 | 126.8629 | 130.812 | 114 | 130 | 125 | 129 | 134 | 128 | 131 | 130.6354 | 113 | 122 | 118.8 | 127 | 115.7 | 107 | 150.5 | 126 | 119.779 | 16
| 1 | 4 | 126 | 126.988 | 125.388 | 115 | 127 | 124 | 128 | 129 | 130 | 126 | 130.4204 | 110 | 130 | 120.6 | 127 | 117.8 | 105 | 149.9 | 128 | 117.866 | 126.4197
| 1 | 5 | 127 | 127.714 | 130.788 | 113 | 126 | 126 | 134 | 134 | 129 | 125 | 131.0277 | 112 | 121 | 118.3 | 124 | 118.8 | 106 | 149.7 | 128 | 120.151 | 71
| 1 | 6 | 126 | 122.7389 | 127.584 | 111 | 128 | 128 | 129 | 134 | 127 | 128 | 131.9887 | 111 | 120 | 120.2 | 126 | 121.2 | 105 | 149.7 | 126 | 119.399
| Mean |  | 127.6667 | 126.7297 | 130.002 | 113.1667 | 129.1667 | 126.1667 | 130 | 133.6667 | 128.3333 | 127.1667 | 130.2343 | 111.8333 | 123.8333 | 119.9833 | 125.5 | 118.3 | 106.1667 | 149.3833 | 126.3333 | 119.4995
| Median |  | 126.5 | 126.9254 | 130.8 | 113.5 | 129 | 126.5 | 129 | 134 | 128.5 | 126.5 | 130.5279 | 112 | 122 | 120.25 | 125.5 | 118.3 | 106 | 149.7 | 126 | 119.8335
| Std Dev. |  | 2.2509 | 2.9186 | 2.944 | 1.472 | 2.8577 | 1.472 | 2.1909 | 2.582 | 1.2111 | 2.137 | 1.3286 | 1.169 | 4.1191 | 1.2416 | 1.3784 | 1.8931 | 1.169 | 1.0128 | 1.5055 | 0.83711
| Rel.Std.Dev. |  | 1.763% | 2.303% | 2.265% | 1.301% | 2.212% | 1.167% | 1.685% | 1.932% | 0.9437% | 1.68% | 1.02% | 1.045% | 3.326% | 1.035% | 1.098% | 1.6% | 1.101% | 0.678% | 1.192% | 0.7005%
| PDM3 |  | 0.9864% | 0.2452% | 2.834% | -10.48% | 2.173% | -0.2002% | 2.832% | 5.732% | 1.514% | 0.5908% | 3.017% | -11.54% | -2.046% | -5.091% | -0.7275% | -6.423% | -16.02% | 18.16% | -0.06833% | -5.474%
| Z-Score (Absolute) |  | 0.22 | 0.04 | 0.67 | 2.6 | 0.51 | 0.07 | 0.67 | 1.39 | 0.35 | 0.12 | 0.72 | 2.86 | 0.52 | 1.27 | 0.2 | 1.6 | 3.96 | 4.44 | 0.04 | 1.37
| Analytical results for Zr in OREAS 211b (Certified Value 47.1 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 47.1 | 48.5868 | 60.173 | 35 | 49.2 | 65.9 | 48.7 | 62.5 | 48.1 | 53 | 45.5624 | 45.9 | 45.6 | 45.1 | 48 | 59.2 | 54 | 44.64 | 35 |  | 1
| 1 | 2 | 45.4 | 50.4041 | 56.33 | 44 | 48.5 | 49.6 | 47.2 | 51.1 | 53.3 | 44.7 | 47.7555 | 45.2 | 45.2 | 41.9 | 46 | 46.8 | 50 | 43.61 | 30 |  | 16
| 1 | 3 | 49 | 47.9526 | 58.304 | 33 | 53 | 51.3 | 46.3 | 48.7 | 49.7 | 41.3 | 46.2652 | 45.5 | 45.7 | 45.5 | 47.3 | 45.4 | 49 | 43.54 | 35 |  | 16
| 1 | 4 | 45.6 | 50.7451 | 55.6055 | 21 | 49.9 | 53 | 49.4 | 44.1 | 48.5 | 42.8 | 45.3028 | 44.2 | 47.5 | 45 | 50.5 | 59.7 | 49 | 43.22 | 30 |  | 47.1279
| 1 | 5 | 46.7 | 49.6688 | 54.083 | 20 | 51 | 50.9 | 48.3 | 46.4 | 48.1 | 41.8 | 45.7657 | 45.5 | 46.5 | 41.7 | 48.6 | 52.7 | 50 | 41.31 | 30 |  | 72
| 1 | 6 | 45.1 | 46.9554 | 58.241 | 38 | 49.5 | 50.4 | 50.6 | 49.7 | 45 | 44.2 | 47.6066 | 44.3 | 45.5 | 52.8 | 48.7 | 44 | 49 | 41.7 | 30
| Mean |  | 46.4833 | 49.0521 | 57.1228 | 31.8333 | 50.1833 | 53.5167 | 48.4167 | 50.4167 | 48.7833 | 44.6333 | 46.3764 | 45.1 | 46 | 45.3333 | 48.1833 | 51.3 | 50.1667 | 43.0033 | 31.6667
| Median |  | 46.15 | 49.1278 | 57.2855 | 34 | 49.7 | 51.1 | 48.5 | 49.2 | 48.3 | 43.5 | 46.0155 | 45.35 | 45.65 | 45.05 | 48.3 | 49.75 | 49.5 | 43.38 | 30
| Std Dev. |  | 1.458 | 1.4762 | 2.1964 | 9.5376 | 1.6092 | 6.1714 | 1.5355 | 6.4182 | 2.7044 | 4.3057 | 1.0598 | 0.6957 | 0.85323 | 4.0233 | 1.5092 | 6.9754 | 1.9408 | 1.2608 | 2.582
| Rel.Std.Dev. |  | 3.137% | 3.01% | 3.845% | 29.96% | 3.207% | 11.53% | 3.171% | 12.73% | 5.544% | 9.647% | 2.285% | 1.543% | 1.855% | 8.875% | 3.132% | 13.6% | 3.869% | 2.932% | 8.154%
| PDM3 |  | -1.368% | 4.083% | 21.21% | -32.45% | 6.483% | 13.56% | 2.735% | 6.978% | 3.513% | -5.293% | -1.595% | -4.303% | -2.393% | -3.808% | 2.239% | 8.853% | 6.448% | -8.752% | -32.81%
| Z-Score (Absolute) |  | 0.51 | 0.26 | 2.7 | 4.94 | 0.6 | 1.61 | 0.07 | 0.67 | 0.18 | 1.07 | 0.55 | 0.93 | 0.66 | 0.86 | 0 | 0.94 | 0.6 | 1.56 | 4.99
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 211b (Certified Value 0.339 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.3 | 0.40437 | 0.13641 | 0.352 | 0.35 | 0.34 | 0.35 | 0.31 | 0.33 | 0.35 | 0.33642 | < 0.3 | 0.22 | 0.224 | 0.313 | 0.3 | < 0.5 | 0.4 | 1.712 |  |  | 1
| 1 | 2 | 0.34 | 0.37399 | 0.13872 | 0.392 | 0.39 | 0.35 | 0.35 | 0.31 | 0.35 | 0.34 | 0.33168 | < 0.3 | 0.22 | 0.322 | 0.318 | 0.35 | < 0.5 | 0.4 | 0.233 |  |  | 22
| 1 | 3 | 0.32 | 0.38871 | 0.15116 | 0.333 | 0.33 | 0.33 | 0.35 | 0.31 | 0.34 | 0.34 | 0.33547 | < 0.3 | 0.25 | 0.265 | 0.322 | 0.31 | < 0.5 | 0.4 | 0.386 |  |  | 16
| 1 | 4 | 0.32 | 0.40151 | 0.2141 | 0.34 | 0.37 | 0.35 | 0.36 | 0.31 | 0.36 | 0.33 | 0.34712 | < 0.3 | 0.24 | 0.295 | 0.326 | 0.34 | < 0.5 | 0.4 | 0.256 |  |  | 0.33853
| 1 | 5 | 0.31 | 0.34164 | 0.11502 | 0.349 | 0.36 | 0.36 | 0.34 | 0.31 | 0.33 | 0.35 | 0.32605 | < 0.3 | 0.25 | 0.295 | 0.322 | 0.31 | < 0.5 | 0.4 | 0.344 |  |  | 74
| 1 | 6 | 0.32 | 0.37437 | 0.11344 | 0.348 | 0.38 | 0.35 | 0.35 | 0.32 | 0.35 | 0.33 | 0.34845 | < 0.3 | 0.22 | 0.283 | 0.318 | 0.41 | < 0.5 | 0.35 | 0.677
| Mean |  | 0.31833 | 0.38077 | 0.14481 | 0.35233 | 0.36333 | 0.34667 | 0.35 | 0.31167 | 0.34333 | 0.34 | 0.33753 |  | 0.23333 | 0.28067 | 0.31983 | 0.33667 |  | 0.39167 | 0.60133
| Median |  | 0.32 | 0.38154 | 0.13756 | 0.3485 | 0.365 | 0.35 | 0.35 | 0.31 | 0.345 | 0.34 | 0.33595 |  | 0.23 | 0.289 | 0.32 | 0.325 |  | 0.4 | 0.365
| Std Dev. |  | 0.013292 | 0.023101 | 0.036937 | 0.020637 | 0.021602 | 0.010328 | 0.0063246 | 0.0040825 | 0.012111 | 0.0089443 | 0.0087467 |  | 0.015055 | 0.033411 | 0.0044907 | 0.040825 |  | 0.020412 | 0.56689
| Rel.Std.Dev. |  | 4.175% | 6.067% | 25.51% | 5.857% | 5.946% | 2.979% | 1.807% | 1.31% | 3.527% | 2.631% | 2.591% |  | 6.452% | 11.9% | 1.404% | 12.13% |  | 5.212% | 94.27%
| PDM3 |  | -5.966% | 12.48% | -57.22% | 4.077% | 7.327% | 2.404% | 3.388% | -7.935% | 1.419% | 0.4343% | -0.2948% |  | -31.07% | -17.09% | -5.523% | -0.5504% |  | 15.7% | 77.63%
| Z-Score (Absolute) |  | 0.5 | 1.13 | 5.04 | 0.39 | 0.67 | 0.24 | 0.33 | 0.68 | 0.15 | 0.06 | 0 | 4.9 | 2.72 | 1.49 | 0.46 | 0.02 | 2.29 | 1.41 | 6.89
| Analytical results for Al in OREAS 211b (Certified Value 3.28 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 3.1482 | 3.3411 | 3.21 | 3.23 | 3.24 | 3.55 | 3.25 | 3.23 | 3.44 | 3.496 | 4.1 | 3.13 | 3.517 | 2.89 | 3.86 | 3.11 | 3.36 | 3.2882 |  |  |  | 1
| 1 | 2 | 3.2717 | 3.3174 | 2.96 | 3.22 | 3.35 | 3.58 | 3.28 | 3.28 | 3.46 | 3.4804 | 4.1 | 3.15 | 3.4435 | 2.89 | 3.91 | 3.1 | 3.33 | 3.2815 |  |  |  | #N/A
| 1 | 3 | 3.2673 | 3.3102 | 2.99 | 3.24 | 3.29 | 3.56 | 3.26 | 3.23 | 3.37 | 3.436 | 4.15 | 3.16 | 3.5106 | 2.91 | 3.72 | 3.12 | 3.37 | 3.3805 |  |  |  | 16
| 1 | 4 | 3.1543 | 3.3734 | 3.04 | 3.24 | 3.51 | 3.55 | 3.3 | 3.33 | 3.39 | 3.4926 | 4.15 | 2.97 | 3.5908 | 3.05 | 3.7 | 3.11 | 3.35 | 3.2222 |  |  |  | 3.2816
| 1 | 5 | 3.2318 | 3.3953 | 3.13 | 3.25 | 3.34 | 3.5 | 3.25 | 3.3 | 3.42 | 3.441 | 4.22 | 3.04 | 3.5261 | 3 | 3.54 | 3.1 | 3.37 | 3.3136 |  |  |  | 75
| 1 | 6 | 3.2836 | 3.3088 | 2.99 | 3.24 | 3.36 | 3.54 | 3.31 | 3.23 | 3.37 | 3.4152 | 4.16 | 3.25 | 3.507 | 2.94 | 3.89 | 3.11 | 3.36 | 3.3086
| Mean |  | 3.2262 | 3.3411 | 3.0533 | 3.2367 | 3.3483 | 3.5467 | 3.275 | 3.2667 | 3.4083 | 3.4602 | 4.1467 | 3.1167 | 3.5158 | 2.9467 | 3.77 | 3.1083 | 3.3567 | 3.2991
| Median |  | 3.2496 | 3.3293 | 3.015 | 3.24 | 3.345 | 3.55 | 3.27 | 3.255 | 3.405 | 3.4607 | 4.15 | 3.14 | 3.5138 | 2.925 | 3.79 | 3.11 | 3.36 | 3.2984
| Std Dev. |  | 0.060563 | 0.036158 | 0.097297 | 0.010328 | 0.091086 | 0.026583 | 0.025884 | 0.043205 | 0.037639 | 0.033811 | 0.044572 | 0.098319 | 0.047041 | 0.06532 | 0.14311 | 0.0075277 | 0.015055 | 0.051515
| Rel.Std.Dev. |  | 1.877% | 1.082% | 3.187% | 0.3191% | 2.72% | 0.7495% | 0.7904% | 1.323% | 1.104% | 0.9772% | 1.075% | 3.155% | 1.338% | 2.217% | 3.796% | 0.2422% | 0.4485% | 1.561%
| PDM3 |  | -1.69% | 1.812% | -6.956% | -1.369% | 2.033% | 8.077% | -0.2013% | -0.4552% | 3.862% | 5.442% | 26.36% | -5.026% | 7.138% | -10.21% | 14.88% | -5.28% | 2.287% | 0.5334%
| Z-Score (Absolute) |  | 0.54 | 0.12 | 1.54 | 0.48 | 0.16 | 1.31 | 0.26 | 0.31 | 0.51 | 0.81 | 4.76 | 1.17 | 1.13 | 2.15 | 2.59 | 1.22 | 0.21 | 0.12
| Analytical results for As in OREAS 211b (Certified Value 32.6 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 2
| 1 | 1 | 32 | 32.416 | 26.476 | 28.462 | 33 | 33 | 32.5 | 34.2 | 31.8 | 34.8 | 33.6 | 33.6022 | 31.3 | 33 | 31.4 | 34 | 31.4 | 31.6 | 25.2 | 34 | 33.479 | 1
| 1 | 2 | 33 | 32.7908 | 28.318 | 30.453 | 32.6 | 32.9 | 33.9 | 35.3 | 31.3 | 34.6 | 35.1 | 33.4275 | 30.7 | 33 | 30.1 | 33 | 31.3 | 33.6 | 25.4 | 34 | 33.095 | 10
| 1 | 3 | 33 | 32.13 | 32.391 | 29.474 | 31.3 | 33.8 | 33.4 | 34.9 | 31.1 | 35 | 34.6 | 33.517 | 31 | 32 | 30.5 | 34 | 31.7 | 32.8 | 27.8 | 34 | 34.288 | 16
| 1 | 4 | 32 | 32.0264 | 34.206 | 31.102 | 32.6 | 31.9 | 35.7 | 34.9 | 32 | 35.2 | 33.6 | 33.4342 | 31.7 | 32 | 29.8 | 34 | 32.4 | 32.9 | 27.4 | 32 | 32.432 | 32.6363
| 1 | 5 | 32 | 33.1164 | 29.955 | 28.22 | 32.9 | 34.1 | 33.7 | 34.9 | 31.1 | 34.6 | 34 | 33.387 | 31.2 | 32 | 29.9 | 34 | 32.3 | 32.9 | 27.1 | 32.5 | 36.008 | 76
| 1 | 6 | 33 | 32.5594 | 27.905 | 28.638 | 32.1 | 33.2 | 34.1 | 35.2 | 32.5 | 35.6 | 35.2 | 33.5392 | 31 | 32 | 30.9 | 34 | 32 | 22.2 | 27.2 | 32 | 33.31
| Mean |  | 32.5 | 32.5065 | 29.8752 | 29.3915 | 32.4167 | 33.15 | 33.8833 | 34.9 | 31.6333 | 34.9667 | 34.35 | 33.4845 | 31.15 | 32.3333 | 30.4333 | 33.8333 | 31.85 | 31 | 26.6833 | 33.0833 | 33.7687
| Median |  | 32.5 | 32.4877 | 29.1365 | 29.056 | 32.6 | 33.1 | 33.8 | 34.9 | 31.55 | 34.9 | 34.3 | 33.4756 | 31.1 | 32 | 30.3 | 34 | 31.85 | 32.85 | 27.15 | 33.25 | 33.3945
| Std Dev. |  | 0.54772 | 0.40892 | 2.931 | 1.1716 | 0.63061 | 0.77136 | 1.0515 | 0.38471 | 0.56451 | 0.38816 | 0.72042 | 0.08145 | 0.33912 | 0.5164 | 0.62503 | 0.40825 | 0.45935 | 4.3594 | 1.0998 | 1.0206 | 1.2505
| Rel.Std.Dev. |  | 1.685% | 1.258% | 9.811% | 3.986% | 1.945% | 2.327% | 3.103% | 1.102% | 1.785% | 1.11% | 2.097% | 0.2432% | 1.089% | 1.597% | 2.054% | 1.207% | 1.442% | 14.06% | 4.122% | 3.085% | 3.703%
| PDM3 |  | -0.4175% | -0.3975% | -8.46% | -9.942% | -0.6728% | 1.574% | 3.821% | 6.936% | -3.073% | 7.141% | 5.251% | 2.599% | -4.554% | -0.9282% | -6.75% | 3.668% | -2.409% | -5.014% | -18.24% | 1.37% | 3.47%
| Z-Score (Absolute) |  | 0 | 0 | 1.33 | 1.57 | 0.04 | 0.33 | 0.7 | 1.21 | 0.44 | 1.25 | 0.94 | 0.5 | 0.68 | 0.08 | 1.05 | 0.67 | 0.33 | 0.76 | 2.94 | 0.3 | 0.64
| Analytical results for B in OREAS 211b (Certified Value 20 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [8 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 15g | [8 empty] | 2
| 1 | 1 | 14 | 19.7224 | 23 | 18 | 20 | 20 | 20 | 20 | 20 | 20.2343 | < 20 | 28 | 13 | [8 empty] | 1
| 1 | 2 | 15 | 20.1256 | 23 | 18 | 20 | 20 | 20 | 20 | 20 | 18.9526 | < 20 | 26 | 13 | [8 empty] | 1
| 1 | 3 | 15 | 19.6455 | 21 | 18 | 20 | 20 | 20 | 20 | 20 | 19.1338 | < 20 | 28 | 13 | [8 empty] | 16
| 1 | 4 | 15 | 19.7368 | 22 | 19 | 20 | 20 | 20 | 20 | 20 | 18.8317 | < 20 | 27 | 14 | [8 empty] | 20.0159
| 1 | 5 | 15 | 20.9496 | 23 | 18 | 20 | 20 | 20 | 20 | 20 | 20.3718 | < 20 | 26 | 14 | [8 empty] | 77
| 1 | 6 | 15 | 20.1643 | 22 | 19 | 20 | 20 | 20 | 20 | 20 | 18.9878 | < 20 | 26 | 13
| Mean |  | 14.8333 | 20.0574 | 22.3333 | 18.3333 | 20 | 20 | 20 | 20 | 20 | 19.4187 |  | 26.8333 | 13.3333
| Median |  | 15 | 19.9312 | 22.5 | 18 | 20 | 20 | 20 | 20 | 20 | 19.0608 |  | 26.5 | 13
| Std Dev. |  | 0.40825 | 0.48925 | 0.8165 | 0.5164 | 0 | 0 | 0 | 0 | 0 | 0.69314 |  | 0.98319 | 0.5164
| Rel.Std.Dev. |  | 2.752% | 2.439% | 3.656% | 2.817% | 0% | 0% | 0% | 0% | 0% | 3.569% |  | 3.664% | 3.873%
| PDM3 |  | -25.89% | 0.2075% | 11.58% | -8.406% | -0.07923% | -0.07923% | -0.07923% | -0.07923% | -0.07923% | -2.984% |  | 34.06% | -33.39%
| Z-Score (Absolute) |  | 5.99 | 0.07 | 2.71 | 1.93 | 0 | 0 | 0 | 0 | 0 | 0.67 | 11.6 | 7.93 | 7.73
| Analytical results for Ba in OREAS 211b (Certified Value 29.3 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 24 | 26.2063 | 29.975 | 30.1 | 29 | 30 | 30 | 30 | 30 | 30 | 27.4251 | 31 | 25.1 | 32 | 20.4 | 34.4 | 23.9 | 31.312 |  |  |  | 1
| 1 | 2 | 25 | 26.2388 | 30.102 | 30.6 | 29 | 30 | 30 | 30 | 30 | 30 | 24.7606 | 31 | 26.1 | 31 | 21.1 | 34.8 | 24.1 | 31.265 |  |  |  | 24
| 1 | 3 | 25 | 26.2445 | 30.428 | 28.6 | 29 | 30 | 30 | 30 | 30 | 20 | 24.577 | 31 | 26 | 32 | 20.6 | 31.6 | 25.4 | 30.984 |  |  |  | 16
| 1 | 4 | 24 | 26.6126 | 30.097 | 30.6 | 29 | 30 | 30 | 30 | 30 | 30 | 28.3383 | 31 | 25 | 33 | 21.5 | 32.8 | 24.5 | 31.132 |  |  |  | 29.3352
| 1 | 5 | 24 | 26.6475 | 30.769 | 29.9 | 29 | 30 | 30 | 30 | 30 | 30 | 27.0266 | 31 | 24.9 | 32 | 21.8 | 31.7 | 24.8 | 31.684 |  |  |  | 78
| 1 | 6 | 25 | 25.5932 | 28.758 | 28.7 | 29 | 30 | 30 | 30 | 30 | 30 | 27.0237 | 31 | 26.2 | 30 | 21.3 | 34.5 | 26.2 | 30.959
| Mean |  | 24.5 | 26.2572 | 30.0215 | 29.75 | 29 | 30 | 30 | 30 | 30 | 28.3333 | 26.5252 | 31 | 25.55 | 31.6667 | 21.1167 | 33.3 | 24.8167 | 31.2227
| Median |  | 24.5 | 26.2416 | 30.0995 | 30 | 29 | 30 | 30 | 30 | 30 | 30 | 27.0252 | 31 | 25.55 | 32 | 21.2 | 33.6 | 24.65 | 31.1985
| Std Dev. |  | 0.54772 | 0.38016 | 0.68334 | 0.8961 | 0 | 0 | 0 | 0 | 0 | 4.0825 | 1.517 | 0 | 0.6091 | 1.0328 | 0.53448 | 1.456 | 0.8612 | 0.26732
| Rel.Std.Dev. |  | 2.236% | 1.448% | 2.276% | 3.012% | 0% | 0% | 0% | 0% | 0% | 14.41% | 5.719% | 0% | 2.384% | 3.261% | 2.531% | 4.372% | 3.47% | 0.8562%
| PDM3 |  | -16.48% | -10.49% | 2.34% | 1.414% | -1.143% | 2.266% | 2.266% | 2.266% | 2.266% | -3.415% | -9.579% | 5.675% | -12.9% | 7.948% | -28.02% | 13.52% | -15.4% | 6.434%
| Z-Score (Absolute) |  | 2.51 | 1.69 | 0.07 | 0.06 | 0.41 | 0.06 | 0.06 | 0.06 | 0.06 | 0.72 | 1.56 | 0.53 | 2.02 | 0.84 | 4.09 | 1.6 | 2.36 | 0.63
| Analytical results for Be in OREAS 211b (Certified Value 0.19 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.19 | < 0.5 | 1.134 | 0.2 | 0.19 | 0.19 | 0.2 | 0.16 | 0.19 | 0.17 | < 0.5 | < 0.5 | 0.23 | < 2 | 0.018 | [6 empty] | 1
| 1 | 2 | 0.16 | < 0.5 | 1.122 | 0.3 | 0.21 | 0.2 | 0.2 | 0.16 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.23 | < 2 | 0.032 | [6 empty] | 25
| 1 | 3 | 0.19 | < 0.5 | 1.176 | 0.2 | 0.19 | 0.19 | 0.2 | 0.16 | 0.19 | 0.17 | < 0.5 | < 0.5 | 0.22 | < 2 | [7 empty] | 16
| 1 | 4 | 0.17 | < 0.5 | 1.006 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.24 | < 2 | 0.066 | [6 empty] | 0.19024
| 1 | 5 | 0.17 | < 0.5 | 1.151 | 0.2 | 0.21 | 0.19 | 0.2 | 0.15 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.23 | < 2 | 0.037 | [6 empty] | 79
| 1 | 6 | 0.21 | < 0.5 | 1.039 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16 | 0.19 | 0.17 | < 0.5 | < 0.5 | 0.24 | < 2 | 0.109
| Mean |  | 0.18167 |  | 1.1047 | 0.21667 | 0.2 | 0.195 | 0.2 | 0.15833 | 0.195 | 0.17 |  |  | 0.23167 |  | 0.0524
| Median |  | 0.18 |  | 1.128 | 0.2 | 0.2 | 0.195 | 0.2 | 0.16 | 0.195 | 0.17 |  |  | 0.23 |  | 0.037
| Std Dev. |  | 0.018348 |  | 0.066998 | 0.040825 | 0.0089443 | 0.0054772 | 3.0405e-17 | 0.0040825 | 0.0054772 | 0 |  |  | 0.0075277 |  | 0.036143
| Rel.Std.Dev. |  | 10.1% |  | 6.065% | 18.84% | 4.472% | 2.809% | 1.52e-14% | 2.578% | 2.809% | 0% |  |  | 3.249% |  | 68.97%
| PDM3 |  | -4.506% |  | 480.7% | 13.89% | 5.131% | 2.503% | 5.131% | -16.77% | 2.503% | -10.64% |  |  | 21.78% |  | -72.46%
| Z-Score (Absolute) |  | 0.3 | 0.81 | 14.64 | NA | 0 | 0.08 | NA | 0.67 | 0.08 | 0.49 | 0.81 | 0.81 | 0.51 | 12.95 | 2.39
Results from laboratory 1.10 were removed due to their 0.01 ppm reading resolution. Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 211b (Certified Value 0.033 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  |  | 3
| 1 | 1 | 0.03 | < 3.5 | 0.04 | 0.21 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | < 0.05 | < 0.02 | < 0.05 | 0.04 | 0.03 | 0.03 | < 5 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.03 | < 3.5 | 0.03 | 0.21 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | < 0.05 | < 0.02 | < 0.05 | 0.03 | 0.03 | 0.05 | < 5 | < 0.1 |  |  |  |  | 26
| 1 | 3 | 0.03 | < 3.5 | 0.03 | 0.2 | 0.04 | 0.03 | 0.03 | 0.04 | 0.05 | < 0.05 | < 0.02 | < 0.05 | 0.08 | 0.03 | 0.04 | < 5 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.03 | < 3.5 | 0.03 | 0.21 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | < 0.05 | < 0.02 | < 0.05 | 0.04 | 0.03 | 0.04 | < 5 | < 0.1 |  |  |  |  | 0.032967
| 1 | 5 | 0.03 | < 3.5 | 0.04 | 0.2 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | < 0.05 | < 0.02 | < 0.05 | 0.04 | 0.03 | 0.03 | < 5 | < 0.1 |  |  |  |  | 80
| 1 | 6 | 0.03 | < 3.5 | 0.03 | 0.21 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | < 0.05 | < 0.02 | < 0.05 | 0.04 | 0.03 | 0.04 | < 5 | < 0.1
| Mean |  | 0.03 |  | 0.033333 | 0.20667 | 0.033333 | 0.03 | 0.031667 | 0.033333 | 0.036667 |  |  |  | 0.045 | 0.03 | 0.038333
| Median |  | 0.03 |  | 0.03 | 0.21 | 0.03 | 0.03 | 0.03 | 0.03 | 0.035 |  |  |  | 0.04 | 0.03 | 0.04
| Std Dev. |  | 0 |  | 0.005164 | 0.005164 | 0.005164 | 0 | 0.0040825 | 0.005164 | 0.008165 |  |  |  | 0.017607 | 0 | 0.0075277
| Rel.Std.Dev. |  | 0% |  | 15.49% | 2.499% | 15.49% | 0% | 12.89% | 15.49% | 22.27% |  |  |  | 39.13% | 0% | 19.64%
| PDM3 |  | -8.999% |  | 1.112% | 526.9% | 1.112% | -8.999% | -3.943% | 1.112% | 11.22% |  |  |  | 36.5% | -8.999% | 16.28%
| Z-Score (Absolute) |  | 0.45 | 231.51 | 0 | 23.38 | 0 | 0.45 | 0.22 | 0 | 0.45 | 1.12 | 3.15 | 1.12 | 1.57 | 0.45 | 0.67 | 332.66 | 2.25
| Analytical results for Ca in OREAS 211b (Certified Value 2.36 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 2.24 | 2.5079 | 1.6426 | 3.39 | 2.5 | 2.33 | 2.68 | 2.27 | 2.19 | 2.39 | 2.7132 | 1.7158 | 3.34 | 1.84 | 3.57 | 1.88 | 3.45 | 2.27 | 1.98 | 2.4046 |  | 1
| 1 | 2 | 2.39 | 2.4678 | 1.6975 | 3.27 | 2.52 | 2.38 | 2.7 | 2.3 | 2.28 | 2.4 | 2.6912 | 1.6581 | 3.23 | 1.86 | 3.44 | 1.88 | 3.51 | 2.28 | 2.12 | 2.4212 |  | #N/A
| 1 | 3 | 2.39 | 2.4551 | 1.7451 | 3.19 | 2.45 | 2.32 | 2.7 | 2.27 | 2.19 | 2.31 | 2.6654 | 1.7094 | 3.31 | 1.91 | 3.55 | 1.92 | 3.37 | 2.31 | 2.05 | 2.564 |  | 16
| 1 | 4 | 2.28 | 2.5071 | 1.6223 | 3.27 | 2.55 | 2.51 | 2.71 | 2.33 | 2.28 | 2.35 | 2.6558 | 1.7281 | 3.27 | 1.83 | 3.63 | 1.98 | 3.56 | 2.3 | 2.03 | 2.2493 |  | 2.3592
| 1 | 5 | 2.33 | 2.5444 | 1.6795 | 3.41 | 2.45 | 2.39 | 2.64 | 2.31 | 2.27 | 2.39 | 2.7072 | 1.6676 | 3.24 | 1.82 | 3.54 | 1.98 | 3.46 | 2.35 | 2.02 | 2.3757 |  | 81
| 1 | 6 | 2.4 | 2.43 | 1.6742 | 3.22 | 2.53 | 2.41 | 2.65 | 2.32 | 2.18 | 2.35 | 2.7037 | 1.6989 | 3.23 | 1.95 | 3.55 | 1.92 | 3.45 | 2.29 | 2.07 | 2.4474
| Mean |  | 2.3383 | 2.4854 | 1.6769 | 3.2917 | 2.5 | 2.39 | 2.68 | 2.3 | 2.2317 | 2.365 | 2.6894 | 1.6963 | 3.27 | 1.8683 | 3.5467 | 1.9267 | 3.4667 | 2.3 | 2.045 | 2.4104
| Median |  | 2.36 | 2.4875 | 1.6769 | 3.27 | 2.51 | 2.385 | 2.69 | 2.305 | 2.23 | 2.37 | 2.6975 | 1.7042 | 3.255 | 1.85 | 3.55 | 1.92 | 3.455 | 2.295 | 2.04 | 2.4129
| Std Dev. |  | 0.066758 | 0.041855 | 0.043033 | 0.089536 | 0.041952 | 0.068411 | 0.028983 | 0.025298 | 0.049565 | 0.034496 | 0.023641 | 0.027759 | 0.046043 | 0.051153 | 0.061536 | 0.045019 | 0.064083 | 0.028284 | 0.047645 | 0.10225
| Rel.Std.Dev. |  | 2.855% | 1.684% | 2.566% | 2.72% | 1.678% | 2.862% | 1.081% | 1.1% | 2.221% | 1.459% | 0.8791% | 1.636% | 1.408% | 2.738% | 1.735% | 2.337% | 1.849% | 1.23% | 2.33% | 4.242%
| PDM3 |  | -0.8832% | 5.35% | -28.92% | 39.53% | 5.97% | 1.307% | 13.6% | -2.508% | -5.405% | 0.2472% | 14% | -28.1% | 38.61% | -20.81% | 50.34% | -18.33% | 46.94% | -2.508% | -13.32% | 2.171%
| Z-Score (Absolute) |  | 0.09 | 0.24 | 1.54 | 2.01 | 0.27 | 0.03 | 0.66 | 0.17 | 0.32 | 0.03 | 0.68 | 1.5 | 1.96 | 1.12 | 2.57 | 0.99 | 2.39 | 0.17 | 0.73 | 0.07
| Analytical results for Cd in OREAS 211b (Certified Value 0.42 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.39 | 0.45817 | 0.42465 | 0.4 | 0.43 | 0.49 | 0.42 | 0.4 | 0.44 | 0.45 | 0.42052 | 0.49 | 0.37 |  | 0.4 | 0.23 | < 1 | 0.5 | 0.488 |  |  | 1
| 1 | 2 | 0.4 | 0.45953 | 0.39103 | 0.4 | 0.47 | 0.42 | 0.42 | 0.4 | 0.45 | 0.44 | 0.42003 | 0.47 | 0.38 |  | 0.4 | 0.23 | < 1 | 0.5 | 0.576 |  |  | 27
| 1 | 3 | 0.43 | 0.47633 | 0.40211 | 0.37 | 0.42 | 0.45 | 0.42 | 0.38 | 0.47 | 0.45 | 0.38482 | 0.44 | 0.35 |  | 0.39 | 0.23 | < 1 | 0.5 | 0.671 |  |  | 16
| 1 | 4 | 0.41 | 0.47747 | 0.40062 | 0.47 | 0.43 | 0.44 | 0.43 | 0.39 | 0.47 | 0.45 | 0.39255 | 0.46 | 0.39 |  | 0.41 | 0.24 | < 1 | 0.5 | 0.687 |  |  | 0.42454
| 1 | 5 | 0.4 | 0.44797 | 0.3861 | 0.46 | 0.46 | 0.43 | 0.43 | 0.4 | 0.47 | 0.43 | 0.37531 | 0.46 | 0.39 | 0.5 | 0.4 | 0.22 | < 1 | 0.5 | 0.704 |  |  | 82
| 1 | 6 | 0.41 | 0.44788 | 0.41067 | 0.37 | 0.42 | 0.46 | 0.42 | 0.41 | 0.47 | 0.43 | 0.39942 | 0.47 | 0.43 |  | 0.39 | 0.29 | < 1 | 0.5 | 0.506
| Mean |  | 0.40667 | 0.46123 | 0.40253 | 0.41167 | 0.43833 | 0.44833 | 0.42333 | 0.39667 | 0.46167 | 0.44167 | 0.39877 | 0.465 | 0.385 | 0.5 | 0.39833 | 0.24 |  | 0.5 | 0.60533
| Median |  | 0.405 | 0.45885 | 0.40136 | 0.4 | 0.43 | 0.445 | 0.42 | 0.4 | 0.47 | 0.445 | 0.39599 | 0.465 | 0.385 | 0.5 | 0.4 | 0.23 |  | 0.5 | 0.6235
| Std Dev. |  | 0.013663 | 0.013101 | 0.013869 | 0.043551 | 0.02137 | 0.024833 | 0.005164 | 0.010328 | 0.013292 | 0.0098319 | 0.018487 | 0.016432 | 0.026646 |  | 0.0075277 | 0.025298 |  | 0 | 0.095091
| Rel.Std.Dev. |  | 3.36% | 2.841% | 3.445% | 10.58% | 4.875% | 5.539% | 1.22% | 2.604% | 2.879% | 2.226% | 4.636% | 3.534% | 6.921% |  | 1.89% | 10.54% |  | 0% | 15.71%
| PDM3 |  | -4.21% | 8.642% | -5.184% | -3.032% | 3.25% | 5.605% | -0.2837% | -6.565% | 8.746% | 4.035% | -6.069% | 9.531% | -9.313% | 17.78% | -6.172% | -43.47% |  | 17.78% | 42.59%
| Z-Score (Absolute) |  | 0.45 | 1.02 | 0.56 | 0.31 | 0.4 | 0.67 | 0 | 0.72 | 1.03 | 0.49 | 0.66 | 1.12 | 1.03 | NA | 0.67 | 4.94 | 2.07 | NA | 4.91
Results from laboratories 1.18 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 211b (Certified Value 8.66 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 7.597 | 7.8935 | 11.4338 | 8.95 | 8.71 | 8.48 | 8.62 | 8.21 | 8.08 | 8.71 | 8.413 | 9.97 | 10.7 | 10.6 | 21.734 | [6 empty] | 1
| 1 | 2 | 7.938 | 7.9734 | 11.4707 | 8.77 | 8.93 | 8.16 | 8.61 | 8.04 | 8.2 | 8.62 | 8.4259 | 9.63 | 12.5 | 10.8 | 22.394 | [6 empty] | 28
| 1 | 3 | 7.8 | 8.4707 | 11.4718 | 8.46 | 8.38 | 8.34 | 8.58 | 8.08 | 8.38 | 8.78 | 7.8863 | 9.67 | 11.2 | 10.6 | 23.519 | [6 empty] | 16
| 1 | 4 | 7.72 | 8.2677 | 11.4785 | 8.81 | 8.53 | 8.58 | 8.6 | 8.26 | 8 | 8.7 | 7.8349 | 9.88 | 12.2 | 10.2 | 21.602 | [6 empty] | 8.6565
| 1 | 5 | 7.861 | 8.5087 | 11.4111 | 8.75 | 8.79 | 8.45 | 8.76 | 8.28 | 8.17 | 8.51 | 8.2931 | 9.7 | 10.9 | 10.2 | 22.655 | [6 empty] | 83
| 1 | 6 | 8.068 | 7.8859 | 11.4403 | 8.51 | 8.62 | 8.32 | 8.64 | 8.4 | 8.06 | 9 | 8.3175 | 9.63 | 13.6 | 10.4 | 23.534
| Mean |  | 7.8307 | 8.1666 | 11.451 | 8.7083 | 8.66 | 8.3883 | 8.635 | 8.2117 | 8.1483 | 8.72 | 8.1951 | 9.7467 | 11.85 | 10.4667 | 22.573
| Median |  | 7.8305 | 8.1206 | 11.4555 | 8.76 | 8.665 | 8.395 | 8.615 | 8.235 | 8.125 | 8.705 | 8.3053 | 9.685 | 11.7 | 10.5 | 22.5245
| Std Dev. |  | 0.16526 | 0.28651 | 0.026747 | 0.18723 | 0.19453 | 0.14703 | 0.06442 | 0.13363 | 0.13512 | 0.16529 | 0.26477 | 0.14348 | 1.1185 | 0.24221 | 0.83713
| Rel.Std.Dev. |  | 2.11% | 3.508% | 0.2336% | 2.15% | 2.246% | 1.753% | 0.746% | 1.627% | 1.658% | 1.896% | 3.231% | 1.472% | 9.439% | 2.314% | 3.709%
| PDM3 |  | -9.54% | -5.658% | 32.28% | 0.5993% | 0.04099% | -3.097% | -0.2478% | -5.138% | -5.87% | 0.7341% | -5.329% | 12.59% | 36.89% | 20.91% | 160.8%
| Z-Score (Absolute) |  | 1.13 | 0.67 | 3.81 | 0.07 | 0 | 0.37 | 0.03 | 0.61 | 0.7 | 0.08 | 0.64 | 1.49 | 4.36 | 2.47 | 19.02
| Analytical results for Co in OREAS 211b (Certified Value 31.1 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 29.4 | 31.1539 | 30.788 | 32.5 | 31.5 | 30 | 32.6 | 30 | 31.8 | 29.5 | 31.8265 | 31.7 | 24 | 31 | 28.9 | 31.5 | 27.5 | 32.2 | 29.306 |  |  | 1
| 1 | 2 | 30.5 | 31.0594 | 30.921 | 31.8 | 32.3 | 29.9 | 33.2 | 29.5 | 31.4 | 31.5 | 32.6377 | 31 | 29.1 | 31 | 28.6 | 32.7 | 27.6 | 32.6 | 30.482 |  |  | 14
| 1 | 3 | 30.7 | 31.1152 | 30.928 | 30.8 | 30.4 | 30.5 | 32.7 | 29.6 | 31.9 | 30.2 | 33.3906 | 31.4 | 25.5 | 31 | 29.1 | 31.8 | 28.5 | 31.6 | 30.825 |  |  | 16
| 1 | 4 | 29.7 | 30.9711 | 30.815 | 31.9 | 31 | 31.5 | 32.3 | 30.4 | 31.3 | 29.7 | 30.9356 | 31.2 | 27.3 | 32 | 30 | 31.2 | 27.8 | 31 | 29.523 |  |  | 31.0953
| 1 | 5 | 29.8 | 31.2701 | 31.737 | 32.1 | 31.3 | 31.5 | 32.3 | 29.4 | 32.4 | 30.1 | 31.4602 | 31.1 | 29.7 | 31 | 30.2 | 32 | 28.1 | 31 | 30.451 |  |  | 84
| 1 | 6 | 30.8 | 31.4508 | 29.67 | 31.1 | 31.2 | 30.5 | 32.9 | 31.1 | 32.4 | 31.8 | 33.2307 | 31 | 28 | 31 | 29.1 | 38.4 | 28.2 | 31.4 | 30.568
| Mean |  | 30.15 | 31.1701 | 30.8098 | 31.7 | 31.2833 | 30.65 | 32.6667 | 30 | 31.8667 | 30.4667 | 32.2469 | 31.2333 | 27.2667 | 31.1667 | 29.3167 | 32.9333 | 27.95 | 31.6333 | 30.1925
| Median |  | 30.15 | 31.1346 | 30.868 | 31.85 | 31.25 | 30.5 | 32.65 | 29.8 | 31.85 | 30.15 | 32.2321 | 31.15 | 27.65 | 31 | 29.1 | 31.9 | 27.95 | 31.5 | 30.4665
| Std Dev. |  | 0.58907 | 0.16965 | 0.66116 | 0.63561 | 0.62423 | 0.70356 | 0.35024 | 0.65422 | 0.47188 | 0.95638 | 0.99438 | 0.27325 | 2.1732 | 0.40825 | 0.63692 | 2.7259 | 0.38341 | 0.65013 | 0.62061
| Rel.Std.Dev. |  | 1.954% | 0.5443% | 2.146% | 2.005% | 1.995% | 2.295% | 1.072% | 2.181% | 1.481% | 3.139% | 3.084% | 0.8749% | 7.97% | 1.31% | 2.173% | 8.277% | 1.372% | 2.055% | 2.056%
| PDM3 |  | -3.04% | 0.2404% | -0.9181% | 1.945% | 0.6046% | -1.432% | 5.053% | -3.522% | 2.481% | -2.022% | 3.703% | 0.4438% | -12.31% | 0.2294% | -5.72% | 5.911% | -10.12% | 1.73% | -2.903%
| Z-Score (Absolute) |  | 0.98 | 0 | 0.34 | 0.51 | 0.11 | 0.5 | 1.44 | 1.12 | 0.67 | 0.67 | 1.04 | 0.06 | 3.76 | 0 | 1.78 | 1.7 | 3.1 | 0.45 | 0.94
| Analytical results for Cr in OREAS 211b (Certified Value 19.3 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 40g | 0.5g |  |  | 2
| 1 | 1 | 18 | 19.4465 | 17.554 | 22.07 | 23 | 20 | 18 | 20 | 18 | 18 | 20 | 20.0156 | 22 | 18 | 21 | 18 | 20 | 20 | 15.852 |  |  | 1
| 1 | 2 | 19 | 19.3529 | 17.444 | 22.47 | 22 | 19 | 18 | 19 | 18 | 18 | 19 | 19.0195 | 21 | 17.1 | 21 | 18.5 | 21 | 20 | 15.966 |  |  | 30
| 1 | 3 | 19 | 19.3328 | 17.9 | 21.77 | 21 | 19 | 18 | 20 | 19 | 18 | 19 | 19.8267 | 21 | 17.5 | 21 | 19.3 | 18 | 20 | 18.142 |  |  | 16
| 1 | 4 | 18 | 19.2632 | 18.201 | 22.84 | 22 | 19 | 19 | 20 | 18 | 19 | 19 | 19.4794 | 21 | 16.7 | 22 | 19.1 | 20 | 20 | 16.66 |  |  | 19.2842
| 1 | 5 | 19 | 19.679 | 18.022 | 22.15 | 22 | 19 | 18 | 19 | 19 | 18 | 19 | 19.5149 | 21 | 17.2 | 21 | 19.4 | 21 | 20 | 16.685 |  |  | 85
| 1 | 6 | 19 | 19.2008 | 17.031 | 20.9 | 22 | 19 | 18 | 19 | 19 | 18 | 19 | 19.5694 | 21 | 17.6 | 21 | 19.2 | 14 | 20 | 15.736
| Mean |  | 18.6667 | 19.3792 | 17.692 | 22.0333 | 22 | 19.1667 | 18.1667 | 19.5 | 18.5 | 18.1667 | 19.1667 | 19.5709 | 21.1667 | 17.35 | 21.1667 | 18.9167 | 19 | 20 | 16.5068
| Median |  | 19 | 19.3428 | 17.727 | 22.11 | 22 | 19 | 18 | 19.5 | 18.5 | 18 | 19 | 19.5422 | 21 | 17.35 | 21 | 19.15 | 20 | 20 | 16.313
| Std Dev. |  | 0.5164 | 0.16885 | 0.43098 | 0.66485 | 0.63246 | 0.40825 | 0.40825 | 0.54772 | 0.54772 | 0.40825 | 0.40825 | 0.34023 | 0.40825 | 0.45056 | 0.40825 | 0.54924 | 2.6833 | 0 | 0.89938
| Rel.Std.Dev. |  | 2.766% | 0.8713% | 2.436% | 3.017% | 2.875% | 2.13% | 2.247% | 2.809% | 2.961% | 2.247% | 2.13% | 1.738% | 1.929% | 2.597% | 1.929% | 2.903% | 14.12% | 0% | 5.449%
| PDM3 |  | -3.202% | 0.4927% | -8.257% | 14.26% | 14.08% | -0.6095% | -5.795% | 1.119% | -4.067% | -5.795% | -0.6095% | 1.487% | 9.762% | -10.03% | 9.762% | -1.906% | -1.474% | 3.712% | -14.4%
| Z-Score (Absolute) |  | 0.37 | 0.27 | 1.25 | 2.65 | 2.62 | 0.07 | 0.82 | 0.37 | 0.52 | 0.82 | 0.07 | 0.44 | 1.87 | 1.56 | 1.87 | 0.15 | 0.07 | NA | 2.32
Results from laboratory 1.28 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 211b (Certified Value 0.49 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 3
| 1 | 1 | 0.5 | 0.48619 | 0.45 | 0.47 | 0.48 | 0.47 | 0.47 | 0.47 | 0.52 | 0.49954 | 0.52 | 0.5 | 0.5 | [8 empty] | 1
| 1 | 2 | 0.54 | 0.48376 | 0.44 | 0.5 | 0.48 | 0.48 | 0.47 | 0.49 | 0.53 | 0.50159 | 0.54 | 0.6 | 0.5 | [8 empty] | 31
| 1 | 3 | 0.52 | 0.49571 | 0.43 | 0.42 | 0.48 | 0.48 | 0.46 | 0.49 | 0.53 | 0.47512 | 0.54 | 0.5 | 0.6 | [8 empty] | 16
| 1 | 4 | 0.51 | 0.48968 | 0.46 | 0.47 | 0.51 | 0.5 | 0.48 | 0.48 | 0.54 | 0.49485 | 0.52 | 0.5 | 0.6 | [8 empty] | 0.48688
| 1 | 5 | 0.52 | 0.51086 | 0.43 | 0.44 | 0.48 | 0.49 | 0.47 | 0.48 | 0.52 | 0.48711 | 0.52 | 0.5 | 0.5 | [8 empty] | 86
| 1 | 6 | 0.53 | 0.47662 | 0.44 | 0.5 | 0.48 | 0.48 | 0.49 | 0.49 | 0.54 | 0.47343 | 0.52 | 0.7 | 0.6
| Mean |  | 0.52 | 0.49047 | 0.44167 | 0.46667 | 0.485 | 0.48333 | 0.47333 | 0.48333 | 0.53 | 0.4886 | 0.52667 | 0.55 | 0.55
| Median |  | 0.52 | 0.48794 | 0.44 | 0.47 | 0.48 | 0.48 | 0.47 | 0.485 | 0.53 | 0.49098 | 0.52 | 0.5 | 0.55
| Std Dev. |  | 0.014142 | 0.011824 | 0.01169 | 0.032042 | 0.012247 | 0.010328 | 0.010328 | 0.008165 | 0.0089443 | 0.012178 | 0.010328 | 0.083666 | 0.054772
| Rel.Std.Dev. |  | 2.72% | 2.411% | 2.647% | 6.866% | 2.525% | 2.137% | 2.182% | 1.689% | 1.688% | 2.492% | 1.961% | 15.21% | 9.959%
| PDM3 |  | 6.802% | 0.7365% | -9.287% | -4.152% | -0.387% | -0.7293% | -2.783% | -0.7293% | 8.855% | 0.3534% | 8.171% | 12.96% | 12.96%
| Z-Score (Absolute) |  | 2.02 | 0.32 | 2.5 | 1.06 | 0 | 0.1 | 0.67 | 0.1 | 2.6 | 0.21 | 2.41 | NA | NA
Results from laboratories 1.21 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 211b (Certified Value 176 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 170.6 | 175.2457 | 178.436 | 183.03 | 178 | 168 | 175 | 179 | 175 | 184 | 188 | 177.8089 | 182 | 172 | 179.61 | 172 | 162.36 | 172 | 175.9 | 164 | 172.284 | 1
| 1 | 2 | 175.7 | 175.9076 | 176.162 | 185.86 | 171 | 166 | 175 | 179 | 175 | 184 | 187 | 173.8234 | 178 | 173 | 179.75 | 169 | 159.94 | 175 | 171.2 | 167 | 178.871 | 16
| 1 | 3 | 171.9 | 174.866 | 183.232 | 189.39 | 168 | 167 | 175 | 179 | 173 | 184 | 186.5 | 178.086 | 181 | 174 | 183.48 | 173 | 162.28 | 174 | 170 | 165 | 180.599 | 16
| 1 | 4 | 169 | 172.5655 | 178.769 | 185.71 | 173 | 165 | 183.5 | 179.5 | 176.5 | 186 | 185.5 | 173.9603 | 182 | 174 | 175.32 | 177 | 163.19 | 176 | 170.8 | 165 | 178.377 | 176.0751
| 1 | 5 | 174.7 | 175.6008 | 184.345 | 186.17 | 175 | 172 | 176 | 177.5 | 169.5 | 184 | 188.5 | 176.7264 | 181 | 170 | 177.89 | 172 | 163.84 | 174 | 170.1 | 165 | 181.102 | 87
| 1 | 6 | 173.6 | 174.6799 | 172.501 | 169.98 | 168 | 163 | 175 | 180.5 | 176 | 183 | 185.5 | 177.2838 | 179 | 172 | 178.28 | 171 | 164.43 | 209 | 173.3 | 164 | 176.963
| Mean |  | 172.5833 | 174.8109 | 178.9075 | 183.3567 | 172.1667 | 166.8333 | 176.5833 | 179.0833 | 174.1667 | 184.1667 | 186.8333 | 176.2815 | 180.5 | 172.5 | 179.055 | 172.3333 | 162.6733 | 180 | 171.8833 | 165 | 178.0327
| Median |  | 172.75 | 175.0559 | 178.6025 | 185.785 | 172 | 166.5 | 175 | 179 | 175 | 184 | 186.75 | 177.0051 | 181 | 172.5 | 178.945 | 172 | 162.775 | 174.5 | 171 | 165 | 178.624
| Std Dev. |  | 2.5483 | 1.1898 | 4.4066 | 6.858 | 3.9707 | 3.0605 | 3.412 | 0.9704 | 2.582 | 0.98319 | 1.2517 | 1.909 | 1.6432 | 1.5166 | 2.6937 | 2.6583 | 1.5772 | 14.2688 | 2.3025 | 1.0954 | 3.1944
| Rel.Std.Dev. |  | 1.477% | 0.6806% | 2.463% | 3.74% | 2.306% | 1.834% | 1.932% | 0.5419% | 1.482% | 0.5339% | 0.6699% | 1.083% | 0.9103% | 0.8792% | 1.504% | 1.543% | 0.9696% | 7.927% | 1.34% | 0.6639% | 1.794%
| PDM3 |  | -1.983% | -0.718% | 1.609% | 4.135% | -2.22% | -5.249% | 0.2886% | 1.708% | -1.084% | 4.595% | 6.11% | 0.1172% | 2.513% | -2.03% | 1.692% | -2.125% | -7.611% | 2.229% | -2.381% | -6.29% | 1.112%
| Z-Score (Absolute) |  | 0.66 | 0.26 | 0.47 | 1.26 | 0.73 | 1.68 | 0.05 | 0.5 | 0.38 | 1.41 | 1.88 | 0 | 0.75 | 0.67 | 0.49 | 0.7 | 2.43 | 0.66 | 0.78 | 2.01 | 0.31
| Analytical results for Dy in OREAS 211b (Indicative Value 2.1 ppm) | [22 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 1.914 | 2.0564 | 3.697 | 2.4 | 2.06 | 0.239 | [15 empty] | 1
| 1 | 2 | 1.983 | 2.0294 | 3.666 | 2.5 | 1.98 | 0.645 | [15 empty] | 4
| 1 | 3 | 1.953 | 2.197 | 3.7012 | 2.4 | 1.95 | 1.238 | [15 empty] | 16
| 1 | 4 | 1.917 | 2.1061 | 3.7106 | 2.4 | 2.06 | 0.597 | [15 empty] | 2.1043
| 1 | 5 | 1.912 | 2.2386 | 3.6896 | 2.3 | 1.98 | 0.494 | [15 empty] | 10
| 1 | 6 | 1.966 | 2.0301 | 3.6536 | 2.2 | 1.97 | 0.907
| Mean |  | 1.9408 | 2.1096 | 3.6863 | 2.3667 | 2 | 0.68667
| Median |  | 1.935 | 2.0812 | 3.6933 | 2.4 | 1.98 | 0.621
| Std Dev. |  | 0.03059 | 0.089307 | 0.021992 | 0.10328 | 0.047749 | 0.34636
| Rel.Std.Dev. |  | 1.576% | 4.233% | 0.5966% | 4.364% | 2.387% | 50.44%
| PDM3 |  | -7.767% | 0.2532% | 75.18% | 12.47% | -4.955% | -67.37%
| Z-Score (Absolute) |  | 0.36 | 0.17 | 5.17 | 0.99 | 0.17 | 4.33
| Analytical results for Er in OREAS 211b (Indicative Value 1.3 ppm) | [22 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 1.16 | 1.2106 | 2.2117 | 1.5 | 1.25 | [16 empty] | 1
| 1 | 2 | 1.183 | 1.2092 | 2.1324 | 1.4 | 1.31 | [16 empty] | 5
| 1 | 3 | 1.2 | 1.3049 | 2.1478 | 1.5 | 1.31 | [16 empty] | 16
| 1 | 4 | 1.169 | 1.2729 | 2.2878 | 1.6 | 1.26 | [16 empty] | 1.2991
| 1 | 5 | 1.17 | 1.3191 | 2.1974 | 1.4 | 1.25 | [16 empty] | 11
| 1 | 6 | 1.208 | 1.2313 | 2.1543 | 1.5 | 1.26
| Mean |  | 1.1817 | 1.258 | 2.1886 | 1.4833 | 1.2733
| Median |  | 1.1765 | 1.2521 | 2.1758 | 1.5 | 1.26
| Std Dev. |  | 0.01896 | 0.047954 | 0.057369 | 0.075277 | 0.028752
| Rel.Std.Dev. |  | 1.604% | 3.812% | 2.621% | 5.075% | 2.258%
| PDM3 |  | -9.039% | -3.162% | 68.47% | 14.18% | -1.982%
| Z-Score (Absolute) |  | 0.67 | 0.11 | 6.73 | 1.54 | 0
| Analytical results for Eu in OREAS 211b (Certified Value 0.5 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 0.377 | 0.40216 | 0.73072 | 0.5 | 0.348 | 0.579 | [15 empty] | 1
| 1 | 2 | 0.38 | 0.42136 | 0.6814 | 0.5 | 0.366 | 0.588 | [15 empty] | 35
| 1 | 3 | 0.4 | 0.43688 | 0.68068 | 0.5 | 0.366 | 0.664 | [15 empty] | 16
| 1 | 4 | 0.381 | 0.42728 | 0.72772 | 0.5 | 0.35 | 0.577 | [15 empty] | 0.49515
| 1 | 5 | 0.384 | 0.43993 | 0.68996 | 0.5 | 0.356 | 0.649 | [15 empty] | 88
| 1 | 6 | 0.409 | 0.40763 | 0.66352 | 0.5 | 0.35 | 0.592
| Mean |  | 0.3885 | 0.42254 | 0.69567 | 0.5 | 0.356 | 0.60817
| Median |  | 0.3825 | 0.42432 | 0.68568 | 0.5 | 0.353 | 0.59
| Std Dev. |  | 0.012911 | 0.015297 | 0.027387 | 0 | 0.0081976 | 0.038144
| Rel.Std.Dev. |  | 3.323% | 3.62% | 3.937% | 0% | 2.303% | 6.272%
| PDM3 |  | -21.54% | -14.66% | 40.5% | 0.9804% | -28.1% | 22.83%
| Z-Score (Absolute) |  | 0.55 | 0.29 | 1.78 | 0.29 | 0.8 | 1.11
| Analytical results for Fe in OREAS 211b (Certified Value 5.63 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 5.62 | 5.6663 | 5.2161 | 6.01 | 5.66 | 5.43 | 5.7 | 5.53 | 5.58 | 5.55 | 5.8778 | 5.454 | 6.17 | 4.84 | 5.79 | 4.77 | 5.57 | 5.31 | 5.77 | 5.6671 |  | 1
| 1 | 2 | 5.81 | 5.6555 | 5.5007 | 5.94 | 5.59 | 5.46 | 5.71 | 5.59 | 5.69 | 5.55 | 5.8868 | 5.2616 | 6.1 | 4.76 | 5.73 | 4.76 | 5.84 | 5.42 | 5.75 | 5.6216 |  | #N/A
| 1 | 3 | 5.78 | 5.6024 | 5.4882 | 5.69 | 5.61 | 5.38 | 5.72 | 5.55 | 5.63 | 5.49 | 5.5975 | 5.4294 | 6.16 | 4.8 | 5.81 | 4.88 | 5.64 | 5.38 | 5.81 | 5.7294 |  | 16
| 1 | 4 | 5.58 | 5.5864 | 5.4415 | 5.94 | 5.6 | 5.67 | 5.69 | 5.57 | 5.7 | 5.51 | 5.7905 | 5.4112 | 6.17 | 4.77 | 5.96 | 4.94 | 5.86 | 5.36 | 5.75 | 5.645 |  | 5.6266
| 1 | 5 | 5.72 | 5.6926 | 5.4693 | 5.98 | 5.57 | 5.45 | 5.64 | 5.49 | 5.67 | 5.56 | 5.7611 | 5.3135 | 5.97 | 4.77 | 5.84 | 4.94 | 5.73 | 5.38 | 5.78 | 5.65 |  | 89
| 1 | 6 | 5.78 | 5.6169 | 5.4535 | 5.71 | 5.56 | 5.47 | 5.69 | 5.59 | 5.63 | 5.5 | 5.7498 | 5.3636 | 6.19 | 4.82 | 5.77 | 4.9 | 6.92 | 5.39 | 5.71 | 5.67
| Mean |  | 5.715 | 5.6367 | 5.4282 | 5.8783 | 5.5983 | 5.4767 | 5.6917 | 5.5533 | 5.65 | 5.5267 | 5.7772 | 5.3722 | 6.1267 | 4.7933 | 5.8167 | 4.865 | 5.9267 | 5.3733 | 5.7617 | 5.6639
| Median |  | 5.75 | 5.6362 | 5.4614 | 5.94 | 5.595 | 5.455 | 5.695 | 5.56 | 5.65 | 5.53 | 5.7758 | 5.3874 | 6.165 | 4.785 | 5.8 | 4.89 | 5.785 | 5.38 | 5.76 | 5.6585
| Std Dev. |  | 0.094604 | 0.041121 | 0.10617 | 0.14077 | 0.035449 | 0.099933 | 0.027869 | 0.038816 | 0.045166 | 0.030111 | 0.10546 | 0.073805 | 0.082624 | 0.032042 | 0.079415 | 0.080932 | 0.49935 | 0.036697 | 0.033714 | 0.036546
| Rel.Std.Dev. |  | 1.655% | 0.7295% | 1.956% | 2.395% | 0.6332% | 1.825% | 0.4896% | 0.699% | 0.7994% | 0.5448% | 1.825% | 1.374% | 1.349% | 0.6685% | 1.365% | 1.664% | 8.425% | 0.6829% | 0.5852% | 0.6453%
| PDM3 |  | 1.572% | 0.1795% | -3.525% | 4.475% | -0.5018% | -2.664% | 1.157% | -1.302% | 0.4165% | -1.776% | 2.678% | -4.521% | 8.888% | -14.81% | 3.379% | -13.54% | 5.334% | -4.501% | 2.401% | 0.6629%
| Z-Score (Absolute) |  | 0.32 | 0.03 | 0.97 | 1.05 | 0.2 | 0.75 | 0.22 | 0.4 | 0.03 | 0.52 | 0.6 | 1.22 | 2.17 | 3.81 | 0.78 | 3.49 | 1.27 | 1.21 | 0.53 | 0.09
| Analytical results for Ga in OREAS 211b (Certified Value 9.96 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | [6 empty] | 3
| 1 | 1 | 8.86 | 10.31 | 10.3 | 9.56 | 9.52 | 10.75 | 9.23 | 9.88 | 10.1 | 10.8783 | 12 | 7.5 | 8.1 | 10.4 | 16.2 | [6 empty] | 1
| 1 | 2 | 9.46 | 10.53 | 9.93 | 9.72 | 9.97 | 11.3 | 9.12 | 9.81 | 10.3 | 10.9142 | 11.7 | 9.7 | 8.1 | 10.5 | 16.4 | [6 empty] | 36
| 1 | 3 | 9.49 | 9.7111 | 9.7 | 9.42 | 10.15 | 11.05 | 8.96 | 9.62 | 9.93 | 10.6887 | 11.7 | 8.5 | 8.3 | 9.98 | 16.6 | [6 empty] | 16
| 1 | 4 | 9.13 | 9.967 | 10 | 9.55 | 10.6 | 11 | 9.28 | 10.2 | 10.05 | 10.7397 | 11.8 | 9.2 | 8.3 | 10.3 | 16.4 | [6 empty] | 9.9621
| 1 | 5 | 9.2 | 10.0306 | 10.2 | 9.63 | 10.3 | 10.8 | 9.19 | 9.73 | 10.1 | 10.4828 | 11.7 | 9.4 | 8.4 | 10 | 17.9 | [6 empty] | 90
| 1 | 6 | 9.44 | 9.8124 | 9.65 | 9.79 | 10.45 | 10.95 | 9.62 | 9.76 | 10.4 | 10.4999 | 11.8 | 9.5 | 8.3 | 10.6 | 16.1
| Mean |  | 9.2633 | 10.0602 | 9.9633 | 9.6117 | 10.165 | 10.975 | 9.2333 | 9.8333 | 10.1467 | 10.7006 | 11.7833 | 8.9667 | 8.25 | 10.2967 | 16.6
| Median |  | 9.32 | 9.9988 | 9.965 | 9.595 | 10.225 | 10.975 | 9.21 | 9.785 | 10.1 | 10.7142 | 11.75 | 9.3 | 8.3 | 10.35 | 16.4
| Std Dev. |  | 0.24728 | 0.30852 | 0.26051 | 0.13197 | 0.38547 | 0.19685 | 0.21961 | 0.19937 | 0.17224 | 0.18252 | 0.1169 | 0.82865 | 0.12247 | 0.25781 | 0.6603
| Rel.Std.Dev. |  | 2.669% | 3.067% | 2.615% | 1.373% | 3.792% | 1.794% | 2.378% | 2.027% | 1.698% | 1.706% | 0.9921% | 9.241% | 1.485% | 2.504% | 3.978%
| PDM3 |  | -7.014% | 0.9847% | 0.01257% | -3.517% | 2.037% | 10.17% | -7.315% | -1.292% | 1.853% | 7.414% | 18.28% | -9.992% | -17.19% | 3.359% | 66.63%
| Z-Score (Absolute) |  | 0.84 | 0 | 0.1 | 0.47 | 0.11 | 0.96 | 0.87 | 0.24 | 0.09 | 0.67 | 1.81 | 1.15 | 1.91 | 0.25 | 6.89
| Analytical results for Gd in OREAS 211b (Indicative Value 1.9 ppm) | [22 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 1.731 | 1.7672 | 2.1 | 1.95 | [17 empty] | 1
| 1 | 2 | 1.795 | 1.7843 | 2 | 1.91 | [17 empty] | 6
| 1 | 3 | 1.853 | 1.9374 | 2 | 1.96 | [17 empty] | 16
| 1 | 4 | 1.799 | 1.8615 | 2.1 | 1.97 | [17 empty] | 1.9022
| 1 | 5 | 1.77 | 1.9589 | 2 | 1.9 | [17 empty] | 12
| 1 | 6 | 1.834 | 1.7413 | 2 | 1.93
| Mean |  | 1.797 | 1.8418 | 2.0333 | 1.9367
| Median |  | 1.797 | 1.8229 | 2 | 1.94
| Std Dev. |  | 0.043813 | 0.091863 | 0.05164 | 0.028048
| Rel.Std.Dev. |  | 2.438% | 4.988% | 2.54% | 1.448%
| PDM3 |  | -5.53% | -3.177% | 6.894% | 1.812%
| Z-Score (Absolute) |  | 0.89 | 0.46 | 1.39 | 0.46
| Analytical results for Ge in OREAS 211b (Certified Value 0.14 ppm) | [22 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.21 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.1g | [12 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.11 | 0.12 | 0.14 | 0.17 | 0.12 | 0.24 | 0.17 | [12 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.1 | 0.13 | 0.14 | 0.18 | 0.12 | 0.24 | 0.18 | [12 empty] | 2
| 1 | 3 | < 0.2 | < 0.1 | 0.1 | 0.13 | 0.14 | 0.17 | 0.12 | 0.22 | 0.16 | [12 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.1 | 0.13 | 0.13 | 0.18 | 0.12 | 0.22 | 0.19 | [12 empty] | 0.13883
| 1 | 5 | < 0.2 | < 0.1 | 0.11 | 0.12 | 0.15 | 0.19 | 0.12 | 0.22 | 0.14 | [12 empty] | 91
| 1 | 6 | < 0.2 | < 0.1 | 0.11 | 0.12 | 0.13 | 0.17 | 0.12 | 0.25 | 0.06
| Mean |  |  |  | 0.105 | 0.125 | 0.13833 | 0.17667 | 0.12 | 0.23167 | 0.15
| Median |  |  |  | 0.105 | 0.125 | 0.14 | 0.175 | 0.12 | 0.23 | 0.165
| Std Dev. |  |  |  | 0.0054772 | 0.0054772 | 0.0075277 | 0.008165 | 0 | 0.013292 | 0.047329
| Rel.Std.Dev. |  |  |  | 5.216% | 4.382% | 5.442% | 4.622% | 0% | 5.737% | 31.55%
| PDM3 |  |  |  | -24.37% | -9.964% | -0.3601% | 27.25% | -13.57% | 66.87% | 8.043%
| Z-Score (Absolute) |  | 0.67 | 2.02 | 0.54 | 0 | 0.36 | 1.39 | 0.13 | 2.88 | 0.67
| Analytical results for Hf in OREAS 211b (Certified Value 0.44 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 3
| 1 | 1 | 0.38 | 0.42982 | 0.3 | 0.42 | 0.41 | 0.49 | 0.39 | 0.42 | 0.41 | 0.49991 | 0.57 | 0.56 | [9 empty] | 1
| 1 | 2 | 0.43 | 0.42958 | 0.4 | 0.44 | 0.44 | 0.48 | 0.39 | 0.42 | 0.41 | 0.53905 | 0.57 | 0.61 | [9 empty] | 22
| 1 | 3 | 0.44 | 0.46494 | 0.3 | 0.39 | 0.39 | 0.48 | 0.4 | 0.42 | 0.4 | 0.48726 | 0.57 | 0.56 | [9 empty] | 16
| 1 | 4 | 0.42 | 0.43146 | 0.3 | 0.44 | 0.41 | 0.5 | 0.43 | 0.42 | 0.41 | 0.4807 | 0.58 | 0.56 | [9 empty] | 0.43885
| 1 | 5 | 0.42 | 0.46478 | 0.3 | 0.42 | 0.42 | 0.48 | 0.43 | 0.41 | 0.43 | 0.52577 | 0.52 | 0.56 | [9 empty] | 92
| 1 | 6 | 0.46 | 0.43205 | 0.4 | 0.47 | 0.42 | 0.47 | 0.43 | 0.44 | 0.42 | 0.51249 | 0.59 | 0.14
| Mean |  | 0.425 | 0.44211 | 0.33333 | 0.43 | 0.415 | 0.48333 | 0.41167 | 0.42167 | 0.41333 | 0.50753 | 0.56667 | 0.49833
| Median |  | 0.425 | 0.43176 | 0.3 | 0.43 | 0.415 | 0.48 | 0.415 | 0.42 | 0.41 | 0.5062 | 0.57 | 0.56
| Std Dev. |  | 0.026646 | 0.017651 | 0.05164 | 0.026833 | 0.016432 | 0.010328 | 0.020412 | 0.0098319 | 0.010328 | 0.022535 | 0.024221 | 0.17668
| Rel.Std.Dev. |  | 6.27% | 3.992% | 15.49% | 6.24% | 3.959% | 2.137% | 4.958% | 2.332% | 2.499% | 4.44% | 4.274% | 35.45%
| PDM3 |  | -3.156% | 0.7422% | -24.04% | -2.016% | -5.434% | 10.14% | -6.194% | -3.915% | -5.814% | 15.65% | 29.13% | 13.55%
| Z-Score (Absolute) |  | 0.11 | 0.65 | 4.17 | 0.11 | 0.55 | 2.47 | 0.7 | 0.26 | 0.63 | 3.55 | 6.17 | 3.14
| Analytical results for Hg in OREAS 211b (Certified Value 0.052 ppm) | [22 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.1g | 01g | 40g | [6 empty] | 3
| 1 | 1 | < 1 | 0.1 | 0.3 | 0.03 | 0.05 | 0.04 | 0.05 | 0.06 | < 0.05 | 0.13 | 0.016 | 0.04 | 0.07 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 1 | 0.1 | 0.29 | 0.04 | 0.04 | 0.05 | 0.08 | 0.06 | < 0.05 | 0.12 | 0.042 | 0.038 | 0.06 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 3 | < 1 | 0.08 | 0.32 | 0.03 | 0.04 | 0.04 | 0.04 | 0.07 | < 0.05 | 0.13 | 0.063 | 0.038 | 0.06 | < 5 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 1 | 0.09 | 0.32 | 0.04 | 0.04 | 0.04 | 0.05 | 0.06 | < 0.05 | 0.13 | 0.068 | 0.035 | 0.06 | < 5 | < 0.5 | [6 empty] | 0.051685
| 1 | 5 | < 1 | 0.05 | 0.3 | 0.05 | 0.04 | 0.03 | 0.06 | 0.07 | < 0.05 | 0.12 | 0.067 | 0.042 | 0.07 | < 5 | < 0.5 | [6 empty] | 93
| 1 | 6 | < 1 | 0.06 | 0.3 | 0.03 | 0.03 | 0.06 | 0.05 | 0.07 | < 0.05 | 0.13 | 0.024 | 0.038 | 0.07 | < 5 | < 0.5
| Mean |  |  | 0.08 | 0.305 | 0.036667 | 0.04 | 0.043333 | 0.055 | 0.065 |  | 0.12667 | 0.046667 | 0.0385 | 0.065
| Median |  |  | 0.085 | 0.3 | 0.035 | 0.04 | 0.04 | 0.05 | 0.065 |  | 0.13 | 0.0525 | 0.038 | 0.065
| Std Dev. |  |  | 0.020976 | 0.012247 | 0.008165 | 0.0063246 | 0.010328 | 0.013784 | 0.0054772 |  | 0.005164 | 0.022853 | 0.0023452 | 0.0054772
| Rel.Std.Dev. |  |  | 26.22% | 4.016% | 22.27% | 15.81% | 23.83% | 25.06% | 8.427% |  | 4.077% | 48.97% | 6.091% | 8.427%
| PDM3 |  |  | 54.78% | 490.1% | -29.06% | -22.61% | -16.16% | 6.413% | 25.76% |  | 145.1% | -9.71% | -25.51% | 25.76%
| Z-Score (Absolute) |  | 11.07 | 0.38 | 6.11 | 0.72 | 0.64 | 0.55 | 0.25 | 0 | 1.02 | 1.57 | 0.47 | 0.67 | 0 | 61.96 | 4.71
| Analytical results for Ho in OREAS 211b (Certified Value 0.48 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.396 | 0.40175 | 0.66364 | 0.5 | 0.455 | [16 empty] | 1
| 1 | 2 | 0.417 | 0.40419 | 0.60472 | 0.5 | 0.435 | [16 empty] | 23
| 1 | 3 | 0.405 | 0.45022 | 0.60844 | 0.5 | 0.43 | [16 empty] | 16
| 1 | 4 | 0.402 | 0.42871 | 0.6364 | 0.5 | 0.455 | [16 empty] | 0.47861
| 1 | 5 | 0.406 | 0.45594 | 0.62732 | 0.5 | 0.435 | [16 empty] | 94
| 1 | 6 | 0.416 | 0.40627 | 0.58864 | 0.5 | 0.43
| Mean |  | 0.407 | 0.42451 | 0.62153 | 0.5 | 0.44
| Median |  | 0.4055 | 0.41749 | 0.61788 | 0.5 | 0.435
| Std Dev. |  | 0.0081486 | 0.024205 | 0.026697 | 0 | 0.011832
| Rel.Std.Dev. |  | 2.002% | 5.702% | 4.295% | 0% | 2.689%
| PDM3 |  | -14.96% | -11.3% | 29.86% | 4.47% | -8.067%
| Z-Score (Absolute) |  | 0.67 | 0.32 | 3.71 | 1.23 | 0
| Analytical results for In in OREAS 211b (Indicative Value 0.029 ppm) | [22 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [9 empty] | 3
| 1 | 1 | 0.03 | 0.04 | 0.088 | 0.028 | 0.029 | 0.028 | 0.028 | 0.028 | < 0.05 | < 0.05 | 0.05 | < 0.5 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.04 | 0.081 | 0.028 | 0.028 | 0.03 | 0.028 | 0.03 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 7
| 1 | 3 | 0.03 | 0.03 | 0.082 | 0.027 | 0.029 | 0.026 | 0.028 | 0.029 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.03 | 0.079 | 0.029 | 0.03 | 0.027 | 0.028 | 0.027 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 0.0285
| 1 | 5 | 0.03 | 0.04 | 0.08 | 0.031 | 0.029 | 0.026 | 0.026 | 0.029 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 13
| 1 | 6 | 0.02 | 0.03 | 0.081 | 0.027 | 0.026 | 0.027 | 0.028 | 0.029 | < 0.05 | < 0.05 | 0.04 | < 0.5
| Mean |  | 0.028333 | 0.035 | 0.081833 | 0.028333 | 0.0285 | 0.027333 | 0.027667 | 0.028667 |  |  | 0.041667
| Median |  | 0.03 | 0.035 | 0.081 | 0.028 | 0.029 | 0.027 | 0.028 | 0.029 |  |  | 0.04
| Std Dev. |  | 0.0040825 | 0.0054772 | 0.0031885 | 0.0015055 | 0.0013784 | 0.0015055 | 0.0008165 | 0.0010328 |  |  | 0.0040825
| Rel.Std.Dev. |  | 14.41% | 15.65% | 3.896% | 5.314% | 4.837% | 5.508% | 2.951% | 3.603% |  |  | 9.798%
| PDM3 |  | -0.5848% | 22.81% | 187.1% | -0.5848% | 0% | -4.094% | -2.924% | 0.5848% |  |  | 46.2%
| Z-Score (Absolute) |  | 0.02 | 1.97 | 16.01 | 0.02 | 0.02 | 0.32 | 0.22 | 0.07 | 1.02 | 1.02 | 3.97 | 66.41
| Analytical results for K in OREAS 211b (Certified Value 0.103 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0983 | 0.10875 | 0.095695 | 0.1 | 0.11 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.099972 | 0.128 | 0.1 | 0.1186 | 0.1 | 0.149 | 0.11 | 0.09 | 0.11002 |  |  | 1
| 1 | 2 | 0.1024 | 0.10719 | 0.097395 | 0.09 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.10038 | 0.129 | 0.1 | 0.1124 | 0.1 | 0.133 | 0.12 | 0.09 | 0.10633 |  |  | #N/A
| 1 | 3 | 0.1017 | 0.10972 | 0.10054 | 0.09 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.098473 | 0.129 | 0.1 | 0.1177 | 0.1 | 0.131 | 0.12 | 0.1 | 0.10816 |  |  | 16
| 1 | 4 | 0.0978 | 0.10834 | 0.096421 | 0.1 | 0.1 | 0.11 | 0.11 | 0.1 | 0.1 | 0.1 | 0.098885 | 0.131 | 0.09 | 0.1197 | 0.1 | 0.143 | 0.11 | 0.1 | 0.1077 |  |  | 0.10349
| 1 | 5 | 0.1003 | 0.11479 | 0.1022 | 0.1 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.0986 | 0.122 | 0.09 | 0.1158 | 0.1 | 0.125 | 0.12 | 0.1 | 0.10883 |  |  | 95
| 1 | 6 | 0.1021 | 0.10521 | 0.1004 | 0.09 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.098024 | 0.13 | 0.1 | 0.1145 | 0.1 | 0.135 | 0.11 | 0.1 | 0.10955
| Mean |  | 0.10043 | 0.109 | 0.098776 | 0.095 | 0.10167 | 0.10167 | 0.11 | 0.1 | 0.1 | 0.1 | 0.099055 | 0.12817 | 0.096667 | 0.11645 | 0.1 | 0.136 | 0.115 | 0.096667 | 0.10843
| Median |  | 0.101 | 0.10855 | 0.098899 | 0.095 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.098743 | 0.129 | 0.1 | 0.11675 | 0.1 | 0.134 | 0.115 | 0.1 | 0.10849
| Std Dev. |  | 0.0019876 | 0.0032305 | 0.0026245 | 0.0054772 | 0.0040825 | 0.0040825 | 0 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 0.00091947 | 0.0031885 | 0.005164 | 0.002734 | 1.5202e-17 | 0.0086487 | 0.0054772 | 0.005164 | 0.0013367
| Rel.Std.Dev. |  | 1.979% | 2.964% | 2.657% | 5.766% | 4.016% | 4.016% | 0% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 0.9282% | 2.488% | 5.342% | 2.348% | 1.52e-14% | 6.359% | 4.763% | 5.342% | 1.233%
| PDM3 |  | -2.95% | 5.327% | -4.552% | -8.201% | -1.759% | -1.759% | 6.294% | -3.369% | -3.369% | -3.369% | -4.282% | 23.85% | -6.59% | 12.53% | -3.369% | 31.42% | 11.13% | -6.59% | 4.779%
| Z-Score (Absolute) |  | 0.14 | 0.85 | 0.33 | 0.77 | 0 | 0 | 0.96 | 0.19 | NA | NA | 0.3 | 3.06 | 0.58 | 1.71 | NA | 3.97 | 1.54 | 0.58 | 0.78
Results from laboratories 1.12 and 1.19 were removed due to their 0.01 wt.% reading resolution. Results from laboratory 1.13 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La in OREAS 211b (Certified Value 3.54 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 3.174 | 3.3533 | 5.5687 | 3.8 | 3.6 | 3.3 | 3.7 | 3.3 | 3.6 | 3.7 | 3.4683 | 4.19 | 2.7 | 3 | 4.34 | 2.5 | 5.2 | 3.272 |  |  |  | 1
| 1 | 2 | 3.299 | 3.2726 | 5.6086 | 3.7 | 3.8 | 3.4 | 3.8 | 3.3 | 3.6 | 3.5 | 3.5937 | 4.1 | 3.2 | 3.1 | 4.72 | 2.3 | 5.2 | 3.274 |  |  |  | 25
| 1 | 3 | 3.288 | 3.5055 | 5.5578 | 3.6 | 3.5 | 3.4 | 3.7 | 3.2 | 3.6 | 3.7 | 3.3945 | 4.15 | 2.8 | 3 | 4.39 | 2.4 | 5.2 | 3.482 |  |  |  | 16
| 1 | 4 | 3.176 | 3.4766 | 5.6356 | 3.8 | 3.7 | 3.5 | 3.8 | 3.3 | 3.5 | 3.6 | 3.4622 | 4.19 | 3.1 | 3.1 | 4.38 | 2.5 | 5.4 | 3.147 |  |  |  | 3.5394
| 1 | 5 | 3.211 | 3.5223 | 5.5422 | 3.7 | 3.6 | 3.4 | 3.8 | 3.3 | 3.5 | 3.5 | 3.5036 | 4.26 | 3.2 | 3.2 | 4.2 | 2.4 | 5 | 3.235 |  |  |  | 96
| 1 | 6 | 3.316 | 3.3356 | 5.5458 | 3.7 | 3.7 | 3.4 | 3.8 | 3.4 | 3.6 | 3.8 | 3.5993 | 4.15 | 3.3 | 3.2 | 5.2 | 2.4 | 5.4 | 3.311
| Mean |  | 3.244 | 3.411 | 5.5764 | 3.7167 | 3.65 | 3.4 | 3.7667 | 3.3 | 3.5667 | 3.6333 | 3.5036 | 4.1733 | 3.05 | 3.1 | 4.5383 | 2.4167 | 5.2333 | 3.2868
| Median |  | 3.2495 | 3.415 | 5.5633 | 3.7 | 3.65 | 3.4 | 3.8 | 3.3 | 3.6 | 3.65 | 3.486 | 4.17 | 3.15 | 3.1 | 4.385 | 2.4 | 5.2 | 3.273
| Std Dev. |  | 0.064433 | 0.10372 | 0.037547 | 0.075277 | 0.10488 | 0.063246 | 0.05164 | 0.063246 | 0.05164 | 0.12111 | 0.08018 | 0.053914 | 0.2429 | 0.089443 | 0.36652 | 0.075277 | 0.15055 | 0.11072
| Rel.Std.Dev. |  | 1.986% | 3.041% | 0.6733% | 2.025% | 2.873% | 1.86% | 1.371% | 1.917% | 1.448% | 3.333% | 2.289% | 1.292% | 7.964% | 2.885% | 8.076% | 3.115% | 2.877% | 3.369%
| PDM3 |  | -8.345% | -3.628% | 57.55% | 5.009% | 3.126% | -3.938% | 6.422% | -6.763% | 0.7713% | 2.655% | -1.011% | 17.91% | -13.83% | -12.41% | 28.22% | -31.72% | 47.86% | -7.135%
| Z-Score (Absolute) |  | 0.81 | 0.35 | 5.69 | 0.51 | 0.32 | 0.38 | 0.65 | 0.66 | 0.09 | 0.27 | 0.09 | 1.78 | 1.35 | 1.21 | 2.8 | 3.12 | 4.74 | 0.69
| Analytical results for Li in OREAS 211b (Certified Value 10 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 2
| 1 | 1 | 9 | < 10 | 10.5817 | 9.4 | 8.8 | 10.4 | 10.4 | 10.2 | 10.4 | 10.4 | 9.3213 | 10.2 | 12.3 | < 10 | 9.471 | [6 empty] | 1
| 1 | 2 | 9.3 | < 10 | 10.6872 | 9.4 | 8.9 | 9.7 | 10.5 | 10.3 | 10.4 | 10.4 | 9.7698 | 10.1 | 10.7 | < 10 | 9.965 | [6 empty] | #N/A
| 1 | 3 | 9.1 | < 10 | 10.7705 | 9 | 8.7 | 9.7 | 10.5 | 10.2 | 10.2 | 10.3 | 9.8125 | 10.2 | 10.4 | 10 | 10.007 | [6 empty] | 16
| 1 | 4 | 8.8 | < 10 | 10.5955 | 9.2 | 8.5 | 9.9 | 10.5 | 10.3 | 10.4 | 10.4 | 9.038 | 10.1 | 11 | < 10 | 10.039 | [6 empty] | 10.0328
| 1 | 5 | 9.3 | < 10 | 10.7764 | 9.3 | 9 | 9.8 | 10.4 | 10.1 | 10.3 | 10.4 | 9.337 | 10.1 | 11.7 | < 10 | 10.275 | [6 empty] | 97
| 1 | 6 | 9.2 | < 10 | 10.4971 | 9.1 | 8.6 | 9.7 | 10.5 | 10.5 | 10.2 | 10.2 | 9.1356 | 10.2 | 10.3 | 10 | 9.599
| Mean |  | 9.1167 |  | 10.6514 | 9.2333 | 8.75 | 9.8667 | 10.4667 | 10.2667 | 10.3167 | 10.35 | 9.4024 | 10.15 | 11.0667 | 10 | 9.8927
| Median |  | 9.15 |  | 10.6413 | 9.25 | 8.75 | 9.75 | 10.5 | 10.25 | 10.35 | 10.4 | 9.3292 | 10.15 | 10.85 | 10 | 9.986
| Std Dev. |  | 0.19408 |  | 0.11214 | 0.1633 | 0.18708 | 0.27325 | 0.05164 | 0.13663 | 0.098319 | 0.083666 | 0.32184 | 0.054772 | 0.78655 | 0 | 0.29998
| Rel.Std.Dev. |  | 2.129% |  | 1.053% | 1.769% | 2.138% | 2.769% | 0.4934% | 1.331% | 0.953% | 0.8084% | 3.423% | 0.5396% | 7.107% | 0% | 3.032%
| PDM3 |  | -9.131% |  | 6.166% | -7.969% | -12.79% | -1.656% | 4.324% | 2.331% | 2.829% | 3.162% | -6.284% | 1.168% | 10.3% | -0.3269% | -1.397%
| Z-Score (Absolute) |  | 0.91 | 5.75 | 0.89 | 0.77 | 1.34 | 0.03 | 0.67 | 0.44 | 0.5 | 0.54 | 0.58 | 0.3 | 1.38 | 3.79 | 0
| Analytical results for Lu in OREAS 211b (Indicative Value 0.16 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 0.144 | 0.15351 | 0.2 | 0.16 | [17 empty] | 1
| 1 | 2 | 0.146 | 0.14679 | 0.2 | 0.16 | [17 empty] | 4
| 1 | 3 | 0.148 | 0.15392 | 0.2 | 0.165 | [17 empty] | 16
| 1 | 4 | 0.147 | 0.14823 | 0.2 | 0.155 | [17 empty] | 0.16465
| 1 | 5 | 0.145 | 0.15389 | 0.2 | 0.16 | [17 empty] | 10
| 1 | 6 | 0.151 | 0.15418 | 0.2 | 0.16
| Mean |  | 0.14683 | 0.15175 | 0.2 | 0.16
| Median |  | 0.1465 | 0.1537 | 0.2 | 0.16
| Std Dev. |  | 0.0024833 | 0.0033261 | 3.0405e-17 | 0.0031623
| Rel.Std.Dev. |  | 1.691% | 2.192% | 1.52e-14% | 1.976%
| PDM3 |  | -10.82% | -7.831% | 21.47% | -2.822%
| Z-Score (Absolute) |  | 0.93 | 0.42 | 4.52 | 0.42
| Analytical results for Mg in OREAS 211b (Certified Value 1.71 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 1.63 | 1.7242 | 1.6165 | 1.81 | 1.69 | 1.67 | 1.75 | 1.66 | 1.72 | 1.68 | 1.7541 | 1.75 | 1.64 | 1.72 | 1.61 | 1.73 | 1.76 | 1.81 | 1.7486 |  |  | 1
| 1 | 2 | 1.69 | 1.7177 | 1.696 | 1.71 | 1.67 | 1.68 | 1.75 | 1.68 | 1.75 | 1.68 | 1.7575 | 1.76 | 1.54 | 1.69 | 1.61 | 1.7 | 1.78 | 1.85 | 1.7567 |  |  | #N/A
| 1 | 3 | 1.68 | 1.708 | 1.7016 | 1.7 | 1.69 | 1.66 | 1.74 | 1.67 | 1.73 | 1.67 | 1.6774 | 1.77 | 1.55 | 1.71 | 1.62 | 1.63 | 1.79 | 1.81 | 1.7627 |  |  | 16
| 1 | 4 | 1.62 | 1.7124 | 1.7036 | 1.74 | 1.67 | 1.74 | 1.74 | 1.67 | 1.75 | 1.67 | 1.728 | 1.77 | 1.54 | 1.75 | 1.67 | 1.66 | 1.75 | 1.84 | 1.7388 |  |  | 1.7075
| 1 | 5 | 1.66 | 1.7316 | 1.7507 | 1.75 | 1.67 | 1.68 | 1.72 | 1.65 | 1.75 | 1.69 | 1.7103 | 1.74 | 1.55 | 1.73 | 1.66 | 1.62 | 1.78 | 1.85 | 1.7542 |  |  | 98
| 1 | 6 | 1.68 | 1.7292 | 1.7198 | 1.67 | 1.69 | 1.68 | 1.75 | 1.68 | 1.73 | 1.67 | 1.7019 | 1.79 | 1.63 | 1.73 | 1.64 | 1.92 | 1.79 | 1.81 | 1.7486
| Mean |  | 1.66 | 1.7205 | 1.698 | 1.73 | 1.68 | 1.685 | 1.7417 | 1.6683 | 1.7383 | 1.6767 | 1.7215 | 1.7633 | 1.575 | 1.7217 | 1.635 | 1.71 | 1.775 | 1.8283 | 1.7516
| Median |  | 1.67 | 1.7209 | 1.7026 | 1.725 | 1.68 | 1.68 | 1.745 | 1.67 | 1.74 | 1.675 | 1.7191 | 1.765 | 1.55 | 1.725 | 1.63 | 1.68 | 1.78 | 1.825 | 1.7514
| Std Dev. |  | 0.028983 | 0.0094239 | 0.044609 | 0.04858 | 0.010954 | 0.028107 | 0.01169 | 0.01169 | 0.013292 | 0.008165 | 0.031155 | 0.017512 | 0.046797 | 0.020412 | 0.025884 | 0.111 | 0.016432 | 0.020412 | 0.0082164
| Rel.Std.Dev. |  | 1.746% | 0.5477% | 2.627% | 2.808% | 0.6521% | 1.668% | 0.6712% | 0.7007% | 0.7646% | 0.487% | 1.81% | 0.9931% | 2.971% | 1.186% | 1.583% | 6.491% | 0.9257% | 1.116% | 0.4691%
| PDM3 |  | -2.784% | 0.7599% | -0.5563% | 1.315% | -1.613% | -1.32% | 1.998% | -2.296% | 1.803% | -1.808% | 0.8185% | 3.267% | -7.762% | 0.8271% | -4.248% | 0.1438% | 3.95% | 7.074% | 2.58%
| Z-Score (Absolute) |  | 1.15 | 0 | 0.43 | 0.18 | 0.77 | 0.67 | 0.4 | 0.99 | 0.34 | 0.83 | 0.02 | 0.81 | 2.76 | 0.02 | 1.62 | 0.2 | 1.03 | 2.05 | 0.59
| Analytical results for Mn in OREAS 211b (Certified Value 0.073 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0719 | 0.073668 | 0.068592 | 0.0804 | 0.0749 | 0.0687 | 0.0734 | 0.0693 | 0.0703 | 0.0709 | 0.075966 | 0.077 | 0.0709 | 0.0776 | 0.0699 | 0.07 | 0.07072 | 0.073 | 0.070426 |  |  | 1
| 1 | 2 | 0.0724 | 0.073211 | 0.069346 | 0.0782 | 0.0775 | 0.0693 | 0.074 | 0.0707 | 0.0707 | 0.0719 | 0.075885 | 0.0774 | 0.0692 | 0.0767 | 0.0698 | 0.0731 | 0.07127 | 0.074 | 0.070531 |  |  | 27
| 1 | 3 | 0.0718 | 0.073305 | 0.070343 | 0.0751 | 0.0725 | 0.0682 | 0.0736 | 0.0697 | 0.0697 | 0.0715 | 0.074652 | 0.0782 | 0.0696 | 0.078 | 0.0706 | 0.071 | 0.07193 | 0.072 | 0.071761 |  |  | 16
| 1 | 4 | 0.0714 | 0.073268 | 0.069387 | 0.0783 | 0.0742 | 0.0714 | 0.0738 | 0.0696 | 0.0702 | 0.0708 | 0.074667 | 0.078 | 0.0682 | 0.0799 | 0.0718 | 0.073 | 0.07034 | 0.074 | 0.070409 |  |  | 0.072608
| 1 | 5 | 0.0711 | 0.073641 | 0.069827 | 0.0784 | 0.075 | 0.0693 | 0.0732 | 0.0691 | 0.0702 | 0.0716 | 0.075019 | 0.0776 | 0.0693 | 0.0777 | 0.0722 | 0.0721 | 0.07194 | 0.074 | 0.070335 |  |  | 99
| 1 | 6 | 0.0718 | 0.073498 | 0.069467 | 0.0769 | 0.0741 | 0.0688 | 0.0734 | 0.0706 | 0.0701 | 0.0713 | 0.074239 | 0.0784 | 0.0711 | 0.0778 | 0.0713 | 0.0902 | 0.07183 | 0.072 | 0.070939
| Mean |  | 0.071733 | 0.073432 | 0.069494 | 0.077883 | 0.0747 | 0.069283 | 0.073567 | 0.069833 | 0.0702 | 0.071333 | 0.075071 | 0.077767 | 0.069717 | 0.07795 | 0.070933 | 0.0749 | 0.071338 | 0.073167 | 0.070734
| Median |  | 0.0718 | 0.073402 | 0.069427 | 0.07825 | 0.07455 | 0.06905 | 0.0735 | 0.06965 | 0.0702 | 0.0714 | 0.074843 | 0.0778 | 0.06945 | 0.07775 | 0.07095 | 0.07255 | 0.07155 | 0.0735 | 0.070479
| Std Dev. |  | 0.00044572 | 0.00019786 | 0.00057954 | 0.0017657 | 0.0016383 | 0.0011161 | 0.00029439 | 0.00066833 | 0.00032249 | 0.00042269 | 0.00070657 | 0.00052789 | 0.0011017 | 0.0010559 | 0.00099532 | 0.0075895 | 0.00068374 | 0.00098319 | 0.00054729
| Rel.Std.Dev. |  | 0.6214% | 0.2695% | 0.834% | 2.267% | 2.193% | 1.611% | 0.4002% | 0.957% | 0.4594% | 0.5926% | 0.9412% | 0.6788% | 1.58% | 1.355% | 1.403% | 10.13% | 0.9584% | 1.344% | 0.7737%
| PDM3 |  | -1.205% | 1.135% | -4.289% | 7.265% | 2.881% | -4.579% | 1.32% | -3.821% | -3.316% | -1.756% | 3.392% | 7.105% | -3.982% | 7.357% | -2.306% | 3.157% | -1.749% | 0.7694% | -2.581%
| Z-Score (Absolute) |  | 0 | 0.6 | 0.79 | 2.18 | 1.05 | 0.87 | 0.65 | 0.67 | 0.54 | 0.14 | 1.18 | 2.14 | 0.72 | 2.21 | 0.28 | 1.12 | 0.14 | 0.51 | 0.35
| Analytical results for Mo in OREAS 211b (Certified Value 1.48 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 1.5 | 1.4851 | 1.5465 | 1.49 | 1.51 | 1.43 | 1.47 | 1.36 | 1.43 | 1.67 | 0.92811 | 1.5 | 1.5 | 1.15 | 1 | 1.33 | 1.2 | 7.4 | 1.6 | 2.161 |  | 1
| 1 | 2 | 1.5 | 1.5113 | 1.4757 | 1.51 | 1.66 | 1.41 | 1.49 | 1.34 | 1.47 | 1.75 | 1.2031 | 1.5 | 1.4 | 1.28 | 1 | 1.33 | 1.2 | 7.7 | 1.6 | 2.057 |  | 28
| 1 | 3 | 1.6 | 1.4865 | 1.4505 | 1.48 | 1.46 | 1.48 | 1.49 | 1.35 | 1.4 | 1.69 | 0.75931 | 1.5 | 1.5 | 1.17 | 1 | 1.39 | 1.2 | 6.6 | 1.6 | 2.021 |  | 16
| 1 | 4 | 1.4 | 1.4375 | 1.495 | 1.46 | 1.49 | 1.5 | 1.49 | 1.4 | 1.44 | 1.61 | 0.89849 | 1.5 | 1.5 | 1.21 | 1 | 1.36 | 1.2 | 6.1 | 1.8 | 1.682 |  | 1.4846
| 1 | 5 | 1.5 | 1.6555 | 1.476 | 1.47 | 1.46 | 1.4 | 1.47 | 1.36 | 1.45 | 1.64 | 1.1285 | 1.5 | 1.5 | 1.22 | 1 | 1.38 | 1.1 | 6.5 | 1.6 | 1.902 |  | 100
| 1 | 6 | 1.5 | 1.482 | 1.4681 | 1.44 | 1.58 | 1.43 | 1.48 | 1.44 | 1.48 | 1.57 | 1.1148 | 1.6 | 1.4 | 1.29 | 1 | 1.37 | 1.2 | 8.6 | 1.8 | 1.749
| Mean |  | 1.5 | 1.5097 | 1.4853 | 1.475 | 1.5267 | 1.4417 | 1.4817 | 1.375 | 1.445 | 1.655 | 1.0054 | 1.5167 | 1.4667 | 1.22 | 1 | 1.36 | 1.1833 | 7.15 | 1.6667 | 1.9287
| Median |  | 1.5 | 1.4858 | 1.4758 | 1.475 | 1.5 | 1.43 | 1.485 | 1.36 | 1.445 | 1.655 | 1.0214 | 1.5 | 1.5 | 1.215 | 1 | 1.365 | 1.2 | 7 | 1.6 | 1.9615
| Std Dev. |  | 0.063246 | 0.075354 | 0.033267 | 0.02429 | 0.078909 | 0.039707 | 0.0098319 | 0.037815 | 0.02881 | 0.063166 | 0.1698 | 0.040825 | 0.05164 | 0.056569 | 0 | 0.025298 | 0.040825 | 0.92682 | 0.10328 | 0.18591
| Rel.Std.Dev. |  | 4.216% | 4.991% | 2.24% | 1.647% | 5.169% | 2.754% | 0.6636% | 2.75% | 1.994% | 3.817% | 16.89% | 2.692% | 3.521% | 4.637% | 0% | 1.86% | 3.45% | 12.96% | 6.197% | 9.639%
| PDM3 |  | 1.04% | 1.69% | 0.04985% | -0.6441% | 2.836% | -2.889% | -0.195% | -7.38% | -2.665% | 11.48% | -32.28% | 2.163% | -1.205% | -17.82% | -32.64% | -8.39% | -20.29% | 381.6% | 12.27% | 29.91%
| Z-Score (Absolute) |  | 0.27 | 0.42 | 0.05 | 0.1 | 0.67 | 0.6 | 0 | 1.6 | 0.55 | 2.6 | 7.14 | 0.52 | 0.22 | 3.92 | NA | 1.82 | 4.47 | 84.94 | 2.77 | 6.7
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 211b (Certified Value 0.137 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.14 | 0.14105 | 0.1422 | 0.118 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14 | 0.12 | 0.13437 | 0.15 | 0.137 | 0.16 | 0.138 | 0.117 | 0.1 | 0.13 | 0.11789 |  |  | 1
| 1 | 2 | 0.14 | 0.14026 | 0.14875 | 0.12 | 0.14 | 0.13 | 0.14 | 0.14 | 0.14 | 0.12 | 0.1349 | 0.16 | 0.138 | 0.15 | 0.136 | 0.119 | 0.11 | 0.13 | 0.12732 |  |  | #N/A
| 1 | 3 | 0.14 | 0.14047 | 0.15709 | 0.122 | 0.14 | 0.13 | 0.14 | 0.14 | 0.14 | 0.12 | 0.1349 | 0.16 | 0.138 | 0.16 | 0.139 | 0.11 | 0.11 | 0.15 | 0.10374 |  |  | 16
| 1 | 4 | 0.14 | 0.14378 | 0.15499 | 0.128 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.12 | 0.13519 | 0.16 | 0.132 | 0.16 | 0.14 | 0.112 | 0.1 | 0.14 | 0.10329 |  |  | 0.13728
| 1 | 5 | 0.14 | 0.14467 | 0.15625 | 0.126 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.13 | 0.1343 | 0.16 | 0.133 | 0.16 | 0.141 | 0.123 | 0.11 | 0.13 | 0.12434 |  |  | 101
| 1 | 6 | 0.14 | 0.14146 | 0.15123 | 0.119 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.12 | 0.13456 | 0.16 | 0.137 | 0.16 | 0.139 | 0.13 | 0.11 | 0.14 | 0.10808
| Mean |  | 0.14 | 0.14195 | 0.15175 | 0.12217 | 0.14 | 0.13667 | 0.14167 | 0.14 | 0.14 | 0.12167 | 0.1347 | 0.15833 | 0.13583 | 0.15833 | 0.13883 | 0.1185 | 0.10667 | 0.13667 | 0.11411
| Median |  | 0.14 | 0.14126 | 0.15311 | 0.121 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.12 | 0.13473 | 0.16 | 0.137 | 0.16 | 0.139 | 0.118 | 0.11 | 0.135 | 0.11298
| Std Dev. |  | 0 | 0.0018346 | 0.0056544 | 0.0040208 | 0 | 0.005164 | 0.0040825 | 0 | 0 | 0.0040825 | 0.00035073 | 0.0040825 | 0.0026394 | 0.0040825 | 0.0017224 | 0.0073417 | 0.005164 | 0.008165 | 0.010531
| Rel.Std.Dev. |  | 0% | 1.292% | 3.726% | 3.291% | 0% | 3.779% | 2.882% | 0% | 0% | 3.355% | 0.2604% | 2.578% | 1.943% | 2.578% | 1.241% | 6.195% | 4.841% | 5.974% | 9.229%
| PDM3 |  | 1.982% | 3.401% | 10.54% | -11.01% | 1.982% | -0.446% | 3.196% | 1.982% | 1.982% | -11.37% | -1.875% | 15.34% | -1.053% | 15.34% | 1.132% | -13.68% | -22.3% | -0.446% | -16.88%
| Z-Score (Absolute) |  | 0.25 | 0.67 | 2.8 | 3.61 | 0.25 | 0.47 | 0.61 | 0.25 | 0.25 | 3.72 | 0.89 | 4.22 | 0.65 | 4.22 | 0 | 4.4 | 6.96 | 0.47 | 5.35
| Analytical results for Nb in OREAS 211b (Certified Value 0.12 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 2
| 1 | 1 | 0.1 | 0.11345 | 0.3 | 1.09 | 0.12 | 0.17 | 0.1 | 0.08 | 0.2 | < 0.2 | < 0.2 | 0.39 | [9 empty] | 1
| 1 | 2 | 0.1 | 0.10613 | 0.2 | 1.06 | 0.13 | 0.16 | 0.11 | 0.08 | 0.24 | < 0.2 | < 0.2 | 0.46 | [9 empty] | 29
| 1 | 3 | 0.1 | 0.13304 | 0.3 | 1.11 | 0.11 | 0.16 | 0.11 | 0.08 | 0.2 | < 0.2 | < 0.2 | 0.31 | [9 empty] | 16
| 1 | 4 | 0.09 | 0.12322 | 0.2 | 1.06 | 0.14 | 0.17 | 0.12 | 0.09 | 0.21 | < 0.2 | < 0.2 | 0.38 | [9 empty] | 0.11658
| 1 | 5 | 0.08 | 0.13374 | 0.3 | 1.1 | 0.12 | 0.15 | 0.11 | 0.1 | 0.26 | < 0.2 | < 0.2 | 0.34 | [9 empty] | 102
| 1 | 6 | 0.1 | 0.12747 | 0.2 | 1.09 | 0.14 | 0.14 | 0.12 | 0.08 | 0.22 | < 0.2 | < 0.2 | 0.37
| Mean |  | 0.095 | 0.12284 | 0.25 | 1.085 | 0.12667 | 0.15833 | 0.11167 | 0.085 | 0.22167 |  |  | 0.375
| Median |  | 0.1 | 0.12535 | 0.25 | 1.09 | 0.125 | 0.16 | 0.11 | 0.08 | 0.215 |  |  | 0.375
| Std Dev. |  | 0.0083666 | 0.011059 | 0.054772 | 0.020736 | 0.012111 | 0.01169 | 0.0075277 | 0.0083666 | 0.024014 |  |  | 0.050892
| Rel.Std.Dev. |  | 8.807% | 9.003% | 21.91% | 1.911% | 9.561% | 7.383% | 6.741% | 9.843% | 10.83% |  |  | 13.57%
| PDM3 |  | -18.51% | 5.367% | 114.4% | 830.7% | 8.648% | 35.81% | -4.218% | -27.09% | 90.13% |  |  | 221.7%
| Z-Score (Absolute) |  | 0.67 | 0 | NA | 23.3 | 0.09 | 0.86 | 0.27 | 0.92 | 2.39 | 0.55 | 0.55 | 6.11
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 211b (Indicative Value 5.67 ppm) | [22 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 4.975 | 5.1438 | 8.5325 | 5.79 | 6.52 | [16 empty] | 1
| 1 | 2 | 5.183 | 5.1677 | 8.5525 | 5.65 | 6.51 | [16 empty] | 5
| 1 | 3 | 5.14 | 5.5797 | 8.4795 | 5.74 | 6.36 | [16 empty] | 16
| 1 | 4 | 5.023 | 5.3659 | 8.5103 | 5.82 | 6.45 | [16 empty] | 5.6705
| 1 | 5 | 5.017 | 5.5903 | 8.4776 | 5.8 | 6.57 | [16 empty] | 11
| 1 | 6 | 5.316 | 5.1908 | 8.5208 | 5.81 | 6.38
| Mean |  | 5.109 | 5.3397 | 8.5122 | 5.7683 | 6.465
| Median |  | 5.0815 | 5.2784 | 8.5155 | 5.795 | 6.48
| Std Dev. |  | 0.12892 | 0.20553 | 0.029603 | 0.064317 | 0.083126
| Rel.Std.Dev. |  | 2.523% | 3.849% | 0.3478% | 1.115% | 1.286%
| PDM3 |  | -9.902% | -5.834% | 50.11% | 1.725% | 14.01%
| Z-Score (Absolute) |  | 0.67 | 0.44 | 2.81 | 0 | 0.71
| Analytical results for Ni in OREAS 211b (Certified Value 65 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 1
| 1 | 1 | 62.5 | 64.8155 | 65.234 | 59.7 | 65 | 69 | 63.2 | 67.2 | 64.5 | 64.9 | 68 | 63.9867 | 66 | 59.7 | 63 | 62.1 | 62.8 | 57.28 | 70 | 62.526 |  | 1
| 1 | 2 | 65.1 | 64.917 | 65.608 | 59.82 | 62.9 | 70 | 71 | 67.5 | 64.2 | 65.8 | 71.5 | 66.4928 | 65 | 56.9 | 62 | 61.9 | 64.4 | 57.27 | 70 | 64.383 |  | 14
| 1 | 3 | 64.8 | 65.3153 | 65.968 | 58.49 | 60.6 | 66 | 70.5 | 67.5 | 64.5 | 67.5 | 69.3 | 60.306 | 66 | 57.5 | 63 | 62 | 62.8 | 58.82 | 69 | 64.591 |  | 16
| 1 | 4 | 63.3 | 64.5315 | 66.903 | 60.73 | 63.9 | 67 | 73.7 | 67.5 | 65.1 | 66.3 | 67.2 | 62.523 | 66 | 56.9 | 64 | 61.7 | 63.3 | 57.17 | 69 | 64.495 |  | 64.968
| 1 | 5 | 64.4 | 64.9644 | 67.718 | 60.29 | 63.8 | 68 | 70.9 | 67.8 | 63.9 | 65.5 | 68.5 | 65.3081 | 67 | 57.8 | 63 | 62.8 | 65.2 | 57.95 | 71 | 64.288 |  | 103
| 1 | 6 | 64.6 | 65.8312 | 63.063 | 56.42 | 62.8 | 68 | 69.1 | 67.5 | 67.1 | 66.3 | 70.3 | 62.3989 | 65 | 58.9 | 62 | 61.5 | 78 | 57.84 | 69 | 63.093
| Mean |  | 64.1167 | 65.0625 | 65.749 | 59.2417 | 63.1667 | 68 | 69.7333 | 67.5 | 64.8833 | 66.05 | 69.1333 | 63.5026 | 65.8333 | 57.95 | 62.8333 | 62 | 66.0833 | 57.7217 | 69.6667 | 63.896
| Median |  | 64.5 | 64.9407 | 65.788 | 59.76 | 63.35 | 68 | 70.7 | 67.5 | 64.5 | 66.05 | 68.9 | 63.2549 | 66 | 57.65 | 63 | 61.95 | 63.85 | 57.56 | 69.5 | 64.3355
| Std Dev. |  | 1.0028 | 0.45355 | 1.5973 | 1.5739 | 1.4895 | 1.4142 | 3.5319 | 0.18974 | 1.1566 | 0.88487 | 1.5756 | 2.2285 | 0.75277 | 1.1309 | 0.75277 | 0.44721 | 5.9149 | 0.62863 | 0.8165 | 0.86653
| Rel.Std.Dev. |  | 1.564% | 0.6971% | 2.429% | 2.657% | 2.358% | 2.08% | 5.065% | 0.2811% | 1.783% | 1.34% | 2.279% | 3.509% | 1.143% | 1.952% | 1.198% | 0.7213% | 8.951% | 1.089% | 1.172% | 1.356%
| PDM3 |  | -1.31% | 0.1454% | 1.202% | -8.814% | -2.773% | 4.667% | 7.335% | 3.897% | -0.1303% | 1.665% | 6.411% | -2.256% | 1.332% | -10.8% | -3.286% | -4.568% | 1.717% | -11.15% | 7.232% | -1.65%
| Z-Score (Absolute) |  | 0.29 | 0.03 | 0.27 | 1.96 | 0.62 | 1.03 | 1.63 | 0.86 | 0.03 | 0.37 | 1.42 | 0.5 | 0.29 | 2.4 | 0.73 | 1.02 | 0.38 | 2.48 | 1.6 | 0.37
| Analytical results for P in OREAS 211b (Certified Value 0.04 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0389 | 0.040472 | 0.042222 | 0.039 | 0.0401 | 0.04 | 0.041 | 0.039 | 0.041 | 0.042 | 0.038489 | 0.0427 | 0.038 | 0.0409 | 0.038 | 0.0396 | 0.04 | 0.042 | 0.040726 |  |  | 1
| 1 | 2 | 0.0403 | 0.040174 | 0.042369 | 0.039 | 0.0404 | 0.041 | 0.041 | 0.039 | 0.042 | 0.042 | 0.038893 | 0.0423 | 0.037 | 0.0391 | 0.039 | 0.0388 | 0.04 | 0.044 | 0.04202 |  |  | #N/A
| 1 | 3 | 0.0398 | 0.040318 | 0.042571 | 0.039 | 0.0404 | 0.04 | 0.041 | 0.039 | 0.041 | 0.041 | 0.036714 | 0.0426 | 0.037 | 0.0402 | 0.039 | 0.036 | 0.04 | 0.042 | 0.042367 |  |  | 16
| 1 | 4 | 0.0397 | 0.039385 | 0.04305 | 0.04 | 0.0394 | 0.042 | 0.041 | 0.039 | 0.042 | 0.041 | 0.03781 | 0.0427 | 0.036 | 0.0416 | 0.039 | 0.0383 | 0.04 | 0.044 | 0.041587 |  |  | 0.040239
| 1 | 5 | 0.0397 | 0.040131 | 0.0437 | 0.04 | 0.0399 | 0.041 | 0.04 | 0.038 | 0.042 | 0.042 | 0.037467 | 0.0418 | 0.037 | 0.0406 | 0.038 | 0.0337 | 0.04 | 0.046 | 0.042576 |  |  | 104
| 1 | 6 | 0.0399 | 0.039966 | 0.040362 | 0.039 | 0.0399 | 0.041 | 0.041 | 0.039 | 0.042 | 0.041 | 0.037479 | 0.0427 | 0.038 | 0.0389 | 0.038 | 0.0381 | 0.04 | 0.044 | 0.041708
| Mean |  | 0.039717 | 0.040074 | 0.042379 | 0.039333 | 0.040017 | 0.040833 | 0.040833 | 0.038833 | 0.041667 | 0.0415 | 0.037808 | 0.042467 | 0.037167 | 0.040217 | 0.0385 | 0.037417 | 0.04 | 0.043667 | 0.041831
| Median |  | 0.03975 | 0.040152 | 0.04247 | 0.039 | 0.04 | 0.041 | 0.041 | 0.039 | 0.042 | 0.0415 | 0.037644 | 0.04265 | 0.037 | 0.0404 | 0.0385 | 0.0382 | 0.04 | 0.044 | 0.041864
| Std Dev. |  | 0.00045789 | 0.00037887 | 0.0011246 | 0.0005164 | 0.00037639 | 0.00075277 | 0.00040825 | 0.00040825 | 0.0005164 | 0.00054772 | 0.00078271 | 0.00036148 | 0.00075277 | 0.0010496 | 0.00054772 | 0.0021794 | 0 | 0.0015055 | 0.0006596
| Rel.Std.Dev. |  | 1.153% | 0.9454% | 2.654% | 1.313% | 0.9406% | 1.844% | 0.9998% | 1.051% | 1.239% | 1.32% | 2.07% | 0.8512% | 2.025% | 2.61% | 1.423% | 5.825% | 0% | 3.448% | 1.577%
| PDM3 |  | -1.297% | -0.4089% | 5.319% | -2.25% | -0.5518% | 1.478% | 1.478% | -3.493% | 3.549% | 3.135% | -6.04% | 5.537% | -7.635% | -0.0548% | -4.321% | -7.013% | -0.5933% | 8.519% | 3.956%
| Z-Score (Absolute) |  | 0.17 | 0 | 1.09 | 0.35 | 0.03 | 0.36 | 0.36 | 0.59 | 0.75 | 0.67 | 1.07 | 1.13 | 1.38 | 0.07 | 0.74 | 1.26 | 0.04 | 1.7 | 0.83
| Analytical results for Pb in OREAS 211b (Certified Value 14.3 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 14 | 16.603 | 15.4972 | 14 | 13.9 | 26.1 | 14.9 | 13.5 | 14.4 | 14.4 | 15.4204 | 17 | 15 | 12.13 | 15 | 13.45 | 16.8 | 15.3 | 16 | 13.812 |  | 1
| 1 | 2 | 14.5 | 16.0293 | 16.1392 | 13.8 | 14.6 | 13.9 | 14.8 | 13.8 | 14.2 | 14.3 | 15.219 | 16 | 14.6 | 13.67 | 14 | 13.84 | 16.8 | 15.8 | 16 | 13.952 |  | 16
| 1 | 3 | 14.5 | 15.824 | 16.0794 | 13.4 | 14.2 | 13.4 | 14.7 | 13.6 | 14.5 | 14.7 | 15.0265 | 14 | 14.8 | 12.25 | 14 | 13.54 | 16.9 | 17.9 | 16 | 13.503 |  | 16
| 1 | 4 | 13.8 | 15.3448 | 16.7758 | 14 | 13.6 | 14.2 | 14.9 | 13.6 | 14.2 | 14.9 | 14.4604 | 15 | 15 | 12.56 | 15 | 13.85 | 16.6 | 17.1 | 15 | 15.388 |  | 14.3247
| 1 | 5 | 14.1 | 14.7107 | 15.5074 | 14.1 | 14 | 13.5 | 14.7 | 13.6 | 13.8 | 14.5 | 14.263 | 16 | 14.6 | 14.21 | 14 | 13.98 | 20.4 | 16.4 | 15 | 13.772 |  | 105
| 1 | 6 | 14.4 | 14.9171 | 15.7616 | 13.7 | 13.7 | 13.7 | 15.2 | 13.8 | 14.6 | 14.3 | 14.5441 | 16 | 14.8 | 13.59 | 14 | 14.11 | 21.1 | 18.9 | 15 | 13.692
| Mean |  | 14.2167 | 15.5715 | 15.9601 | 13.8333 | 14 | 15.8 | 14.8667 | 13.65 | 14.2833 | 14.5167 | 14.8222 | 15.6667 | 14.8 | 13.0683 | 14.3333 | 13.795 | 18.1 | 16.9 | 15.5 | 14.0198
| Median |  | 14.25 | 15.5844 | 15.9205 | 13.9 | 13.95 | 13.8 | 14.85 | 13.6 | 14.3 | 14.45 | 14.7853 | 16 | 14.8 | 13.075 | 14 | 13.845 | 16.85 | 16.75 | 15.5 | 13.792
| Std Dev. |  | 0.29269 | 0.71525 | 0.48381 | 0.2582 | 0.36332 | 5.0541 | 0.18619 | 0.12247 | 0.28577 | 0.24014 | 0.46434 | 1.0328 | 0.17889 | 0.86557 | 0.5164 | 0.25399 | 2.0669 | 1.3461 | 0.54772 | 0.68639
| Rel.Std.Dev. |  | 2.059% | 4.593% | 3.031% | 1.866% | 2.595% | 31.99% | 1.252% | 0.8972% | 2.001% | 1.654% | 3.133% | 6.592% | 1.209% | 6.623% | 3.603% | 1.841% | 11.42% | 7.965% | 3.534% | 4.896%
| PDM3 |  | -0.754% | 8.704% | 11.42% | -3.43% | -2.267% | 10.3% | 3.784% | -4.71% | -0.2886% | 1.34% | 3.474% | 9.368% | 3.318% | -8.77% | 0.06048% | -3.698% | 26.36% | 17.98% | 8.205% | -2.128%
| Z-Score (Absolute) |  | 0.3 | 1.04 | 1.42 | 0.67 | 0.51 | 1.27 | 0.35 | 0.86 | 0.23 | 0 | 0.3 | NA | 0.28 | 1.43 | NA | 0.71 | 3.54 | 2.35 | NA | 0.49
Results from laboratories 1.15, 1.18 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 211b (Indicative Value 16.7 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.14 | [19 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [19 empty] | 1
|  |  | 01g | 01g | [19 empty] | 1
| 1 | 1 | 19 | 15.7514 | [19 empty] | 1
| 1 | 2 | 16 | 14.9183 | [19 empty] | 6
| 1 | 3 | 16 | 16.2493 | [19 empty] | 16
| 1 | 4 | 21 | 14.0291 | [19 empty] | 16.7332
| 1 | 5 | 18 | 16.4002 | [19 empty] | 12
| 1 | 6 | 19 | 14.4503
| Mean |  | 18.1667 | 15.2998
| Median |  | 18.5 | 15.3349
| Std Dev. |  | 1.9408 | 0.97964
| Rel.Std.Dev. |  | 10.68% | 6.403%
| PDM3 |  | 8.566% | -8.566%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 211b (Certified Value 1.32 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 1.058 | 1.088 | 1.8437 | 1.3 | 1.37 | [16 empty] | 1
| 1 | 2 | 1.105 | 1.0936 | 1.7249 | 1.3 | 1.4 | [16 empty] | 33
| 1 | 3 | 1.081 | 1.1516 | 1.7252 | 1.2 | 1.38 | [16 empty] | 16
| 1 | 4 | 1.061 | 1.1216 | 1.818 | 1.3 | 1.42 | [16 empty] | 1.3244
| 1 | 5 | 1.09 | 1.1478 | 1.783 | 1.3 | 1.32 | [16 empty] | 106
| 1 | 6 | 1.121 | 1.0673 | 1.7716 | 1.2 | 1.39
| Mean |  | 1.086 | 1.1116 | 1.7777 | 1.2667 | 1.38
| Median |  | 1.0855 | 1.1076 | 1.7773 | 1.3 | 1.385
| Std Dev. |  | 0.024641 | 0.03423 | 0.048159 | 0.05164 | 0.034059
| Rel.Std.Dev. |  | 2.269% | 3.079% | 2.709% | 4.077% | 2.468%
| PDM3 |  | -18% | -16.07% | 34.23% | -4.36% | 4.197%
| Z-Score (Absolute) |  | 0.79 | 0.67 | 2.22 | 0 | 0.49
| Analytical results for Pt in OREAS 211b (Indicative Value 18.7 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.14 | 1.21 | 1.28 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [17 empty] | 1
|  |  | 01g | 01g | 0.1g | 40g | [17 empty] | 1
| 1 | 1 | 14 | 18.8566 | 22 | 20 | [17 empty] | 1
| 1 | 2 | 18 | 18.0419 | 20 | 20 | [17 empty] | 2
| 1 | 3 | 14 | 15.4558 | 24 | 20 | [17 empty] | 16
| 1 | 4 | 12 | 15.6842 | 22 | 20 | [17 empty] | 18.7059
| 1 | 5 | 16 | 17.2152 | 22 | 20 | [17 empty] | 13
| 1 | 6 | 16 | 17.6879 | 26 | 20
| Mean |  | 15 | 17.1569 | 22.6667 | 20
| Median |  | 15 | 17.4516 | 22 | 20
| Std Dev. |  | 2.0976 | 1.3431 | 2.0656 | 0
| Rel.Std.Dev. |  | 13.98% | 7.829% | 9.113% | 0%
| PDM3 |  | -19.81% | -8.281% | 21.17% | 6.918%
| Z-Score (Absolute) |  | 0.97 | 0.38 | 1.1 | 0.38
| Analytical results for Rb in OREAS 211b (Certified Value 4.36 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 3
| 1 | 1 | 4.1 | 4.4036 | 3.8 | 4.5 | 4.2 | 4.3 | 4.2 | 4.5 | 4.7 | 4.3905 | 5.1 | 5.2 | 6.3 | [8 empty] | 1
| 1 | 2 | 4.42 | 4.4339 | 3.7 | 4.5 | 4.2 | 4.3 | 4.2 | 4.4 | 4.7 | 4.4056 | 5.1 | 5.5 | 6.3 | [8 empty] | 34
| 1 | 3 | 4.37 | 4.3439 | 3.6 | 4.2 | 4.1 | 4.3 | 4.1 | 4.4 | 4.6 | 4.1038 | 5.21 | 5.1 | 5.9 | [8 empty] | 16
| 1 | 4 | 4.17 | 4.5211 | 3.8 | 4.4 | 4.4 | 4.3 | 4.3 | 4.5 | 4.6 | 4.2729 | 4.92 | 5.4 | 6.1 | [8 empty] | 4.3604
| 1 | 5 | 4.21 | 4.8353 | 3.8 | 4.4 | 4.2 | 4.4 | 4.3 | 4.4 | 4.7 | 4.2475 | 5.49 | 5.1 | 5.8 | [8 empty] | 107
| 1 | 6 | 4.29 | 4.4552 | 3.6 | 4.2 | 4.2 | 4.4 | 4.4 | 4.4 | 4.8 | 4.1895 | 5.26 | 5.9 | 6.1
| Mean |  | 4.26 | 4.4988 | 3.7167 | 4.3667 | 4.2167 | 4.3333 | 4.25 | 4.4333 | 4.6833 | 4.2683 | 5.18 | 5.3667 | 6.0833
| Median |  | 4.25 | 4.4446 | 3.75 | 4.4 | 4.2 | 4.3 | 4.25 | 4.4 | 4.7 | 4.2602 | 5.155 | 5.3 | 6.1
| Std Dev. |  | 0.12231 | 0.17488 | 0.098319 | 0.13663 | 0.098319 | 0.05164 | 0.10488 | 0.05164 | 0.075277 | 0.11621 | 0.19173 | 0.30768 | 0.20412
| Rel.Std.Dev. |  | 2.871% | 3.887% | 2.645% | 3.129% | 2.332% | 1.192% | 2.468% | 1.165% | 1.607% | 2.723% | 3.701% | 5.733% | 3.355%
| PDM3 |  | -2.301% | 3.176% | -14.76% | 0.1448% | -3.295% | -0.6196% | -2.531% | 1.674% | 7.407% | -2.111% | 18.8% | 23.08% | 39.51%
| Z-Score (Absolute) |  | 0.54 | 0.67 | 3.32 | 0 | 0.77 | 0.17 | 0.6 | 0.34 | 1.62 | 0.5 | 4.15 | 5.1 | 8.76
| Analytical results for Re in OREAS 211b (Certified Value 0.002 ppm) | [22 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 3
| 1 | 1 | 0.002 | < 0.01 | 0.001 | 0.002 | 0.002 | 0.002 | 0.001 | 0.002 | 0.001 | < 0.05 | < 0.05 | 0.002 | [9 empty] | 1
| 1 | 2 | 0.002 | < 0.01 | 0.001 | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002 | [9 empty] | 35
| 1 | 3 | 0.002 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002 | [9 empty] | 16
| 1 | 4 | 0.002 | < 0.01 | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002 | [9 empty] | 0.0018267
| 1 | 5 | 0.002 | < 0.01 | 0.001 | 0.003 | 0.001 | 0.002 | 0.002 | 0.001 | 0.002 | < 0.05 | < 0.05 | 0.001 | [9 empty] | 108
| 1 | 6 | 0.002 | < 0.01 | 0.001 | < 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002
| Mean |  | 0.002 |  | 0.001 | 0.0018 | 0.0016667 | 0.002 | 0.0018333 | 0.0018333 | 0.0018333 |  |  | 0.0018333
| Median |  | 0.002 |  | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 |  |  | 0.002
| Std Dev. |  | 0 |  | 0 | 0.00083666 | 0.0005164 | 0 | 0.00040825 | 0.00040825 | 0.00040825 |  |  | 0.00040825
| Rel.Std.Dev. |  | 0% |  | 0% | 46.48% | 30.98% | 0% | 22.27% | 22.27% | 22.27% |  |  | 22.27%
| PDM3 |  | 9.489% |  | -45.26% | -1.46% | -8.759% | 9.489% | 0.365% | 0.365% | 0.365% |  |  | 0.365%
| Z-Score (Absolute) |  | 0.67 | 12.81 | 3.37 | 1.01 | 0.67 | 0.67 | 0 | 0 | 0 | 93.73 | 93.73 | 0
| Analytical results for S in OREAS 211b (Certified Value 0.266 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 3
| 1 | 1 | 0.2599 | 0.2793 | 0.24174 | 0.25876 | 0.244 | 0.26 | 0.26 | 0.29 | 0.26 | 0.29 | 0.27 | 0.27885 | 0.2684 | 0.28 | 0.25 | 0.2593 | 0.25 | 0.251 | 0.19 | 0.27 | 0.296 | 1
| 1 | 2 | 0.2685 | 0.27824 | 0.2598 | 0.26918 | 0.238 | 0.25 | 0.26 | 0.29 | 0.26 | 0.3 | 0.28 | 0.28018 | 0.2586 | 0.28 | 0.24 | 0.2588 | 0.25 | 0.251 | 0.19 | 0.275 | 0.2936 | 20
| 1 | 3 | 0.2675 | 0.2716 | 0.25844 | 0.25268 | 0.237 | 0.25 | 0.25 | 0.29 | 0.26 | 0.3 | 0.27 | 0.26894 | 0.2665 | 0.28 | 0.24 | 0.2593 | 0.26 | 0.242 | 0.2 | 0.28 | 0.2881 | 16
| 1 | 4 | 0.2586 | 0.26616 | 0.26051 | 0.26345 | 0.245 | 0.25 | 1.22 | 0.29 | 0.26 | 0.29 | 0.27 | 0.27191 | 0.2641 | 0.28 | 0.23 | 0.2676 | 0.26 | 0.244 | 0.21 | 0.275 | 0.2863 | 0.2657
| 1 | 5 | 0.2639 | 0.26563 | 0.2576 | 0.26118 | 0.243 | 0.25 | 0.26 | 0.29 | 0.26 | 0.3 | 0.27 | 0.27423 | 0.2593 | 0.28 | 0.24 | 0.2728 | 0.26 | 0.256 | 0.2 | 0.285 | 0.2862 | 109
| 1 | 6 | 0.2669 | 0.26815 | 0.25342 | 0.24258 | 0.235 | 0.25 | 0.26 | 0.29 | 0.26 | 0.3 | 0.27 | 0.27078 | 0.2606 | 0.28 | 0.24 | 0.2524 | 0.26 | 0.27 | 0.2 | 0.27 | 0.2927
| Mean |  | 0.26422 | 0.27151 | 0.25525 | 0.25797 | 0.24033 | 0.25167 | 0.41833 | 0.29 | 0.26 | 0.29667 | 0.27167 | 0.27415 | 0.26292 | 0.28 | 0.24 | 0.2617 | 0.25667 | 0.25233 | 0.19833 | 0.27583 | 0.29048
| Median |  | 0.2654 | 0.26987 | 0.25802 | 0.25997 | 0.2405 | 0.25 | 0.26 | 0.29 | 0.26 | 0.3 | 0.27 | 0.27307 | 0.26235 | 0.28 | 0.24 | 0.2593 | 0.26 | 0.251 | 0.2 | 0.275 | 0.2904
| Std Dev. |  | 0.0041619 | 0.006009 | 0.0070685 | 0.0092885 | 0.0041793 | 0.0040825 | 0.39276 | 0 | 0 | 0.005164 | 0.0040825 | 0.0045171 | 0.0040346 | 0 | 0.0063246 | 0.0072724 | 0.005164 | 0.010053 | 0.0075277 | 0.0058452 | 0.0041615
| Rel.Std.Dev. |  | 1.575% | 2.213% | 2.769% | 3.601% | 1.739% | 1.622% | 93.89% | 0% | 0% | 1.741% | 1.503% | 1.648% | 1.535% | 0% | 2.635% | 2.779% | 2.012% | 3.984% | 3.795% | 2.119% | 1.433%
| PDM3 |  | -0.5595% | 2.186% | -3.934% | -2.91% | -9.548% | -5.283% | 57.44% | 9.144% | -2.147% | 11.65% | 2.244% | 3.179% | -1.049% | 5.381% | -9.674% | -1.507% | -3.401% | -5.032% | -25.36% | 3.813% | 9.326%
| Z-Score (Absolute) |  | 0.08 | 0.52 | 0.46 | 0.3 | 1.36 | 0.68 | 9.33 | 1.63 | 0.18 | 2.03 | 0.53 | 0.67 | 0 | 1.03 | 1.38 | 0.07 | 0.38 | 0.64 | 3.88 | 0.78 | 1.65
| Analytical results for Sb in OREAS 211b (Certified Value 0.88 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 0.94 | 0.77721 | 0.72071 | 0.85 | < 5 | 0.94 | 0.95 | 0.82 | 0.95 | 1.05 | 0.39734 | 0.38 | 0.46 | 0.81 | 0.78 | 1.15 | < 5 | 0.8 |  |  |  | 1
| 1 | 2 | 0.94 | 0.77832 | 0.6441 | 0.95 | < 5 | 0.88 | 0.95 | 0.83 | 0.96 | 1.07 | 0.40689 | 0.34 | 0.44 | 0.95 | 0.78 | 0.89 | < 5 | 0.8 |  |  |  | 36
| 1 | 3 | 1.02 | 0.8374 | 0.66274 | 0.82 | < 5 | 0.85 | 0.95 | 0.82 | 0.95 | 1.06 | 0.3636 | 0.33 | 0.43 | 0.89 | 0.78 | 0.84 | < 5 | 0.8 |  |  |  | 16
| 1 | 4 | 0.88 | 0.82253 | 0.75012 | 0.88 | < 5 | 0.86 | 0.95 | 0.86 | 0.95 | 1 | 0.28927 | 0.35 | 0.47 | 1.01 | 0.76 | 0.86 | < 5 | 0.8 | 0.511 |  |  | 0.88319
| 1 | 5 | 0.93 | 0.82932 | 0.70374 | 0.86 | < 5 | 0.85 | 0.89 | 0.83 | 0.94 | 1.07 | 0.4712 | 0.34 | 0.47 | 1.01 | 0.83 | 0.84 | < 5 | 0.8 |  |  |  | 110
| 1 | 6 | 0.95 | 0.79245 | 0.69482 | 0.95 | < 5 | 0.92 | 0.94 | 0.89 | 0.95 | 1.07 | 0.29892 | 0.35 | 0.43 | 0.88 | 0.83 | 0.93 | < 5 | 0.8
| Mean |  | 0.94333 | 0.8062 | 0.69604 | 0.885 |  | 0.88333 | 0.93833 | 0.84167 | 0.95 | 1.0533 | 0.3712 | 0.34833 | 0.45 | 0.925 | 0.79333 | 0.91833 |  | 0.8 | 0.511
| Median |  | 0.94 | 0.80749 | 0.69928 | 0.87 |  | 0.87 | 0.95 | 0.83 | 0.95 | 1.065 | 0.38047 | 0.345 | 0.45 | 0.92 | 0.78 | 0.875 |  | 0.8 | 0.511
| Std Dev. |  | 0.045019 | 0.026763 | 0.038474 | 0.053944 |  | 0.038297 | 0.024014 | 0.027869 | 0.0063246 | 0.027325 | 0.06921 | 0.017224 | 0.018974 | 0.079436 | 0.029439 | 0.11856 |  | 1.2162e-16
| Rel.Std.Dev. |  | 4.772% | 3.32% | 5.528% | 6.095% |  | 4.336% | 2.559% | 3.311% | 0.6657% | 2.594% | 18.64% | 4.945% | 4.216% | 8.588% | 3.711% | 12.91% |  | 1.52e-14%
| PDM3 |  | 6.81% | -8.717% | -21.19% | 0.2053% |  | 0.01658% | 6.244% | -4.701% | 7.565% | 19.27% | -57.97% | -60.56% | -49.05% | 4.734% | -10.17% | 3.979% |  | -9.419% | -42.14%
| Z-Score (Absolute) |  | 0.47 | 0.62 | 1.5 | 0.01 | 12.91 | 0.01 | 0.43 | 0.34 | 0.53 | 1.35 | 4.1 | 4.28 | 3.47 | 0.33 | 0.73 | 0.27 | 12.91 | NA | 2.98
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 211b (Certified Value 5.17 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 2
| 1 | 1 | 4.5 | 5.2319 | 6.8 | 5 | 5.1 | 6.3 | 5.5 | 5 | 5.7 | 5.5861 | 6 | 3.6 | 7 | 3.8 | 7 | 4.3 | 4 | 4.15 |  |  |  | 1
| 1 | 2 | 4.8 | 5.32 | 6.4 | 5 | 5 | 6.2 | 5.5 | 5.1 | 6 | 5.5502 | 6 | 4.9 | 7 | 3.8 | 6 | 4.5 | 4 | 4.334 |  |  |  | 37
| 1 | 3 | 4.8 | 5.2031 | 6.4 | 5 | 5.1 | 6.3 | 5.4 | 5.2 | 6 | 5.4495 | 6 | 4.4 | 7 | 3.9 | 6 | 4.6 | 4 | 4.656 |  |  |  | 16
| 1 | 4 | 4.6 | 5.3526 | 6.6 | 5 | 5.1 | 6.3 | 5.6 | 5.1 | 5.9 | 5.5626 | 6 | 4.6 | 7 | 4 | 7 | 4.5 | 4 | 4.135 |  |  |  | 5.1708
| 1 | 5 | 4.7 | 5.392 | 6.4 | 5 | 5.1 | 6.1 | 5.6 | 5.2 | 6.4 | 5.5392 | 6 | 4.8 | 7 | 4 | 6 | 4.7 | 4 | 4.373 |  |  |  | 111
| 1 | 6 | 4.8 | 5.2499 | 6.4 | 5 | 5.1 | 6.1 | 5.8 | 5.1 | 6.1 | 5.4283 | 6 | 4.7 | 7 | 3.9 | 6 | 4.5 | 4 | 4.41
| Mean |  | 4.7 | 5.2916 | 6.5 | 5 | 5.0833 | 6.2167 | 5.5667 | 5.1167 | 6.0167 | 5.5193 | 6 | 4.5 | 7 | 3.9 | 6.3333 | 4.5167 | 4 | 4.343
| Median |  | 4.75 | 5.2849 | 6.4 | 5 | 5.1 | 6.25 | 5.55 | 5.1 | 6 | 5.5447 | 6 | 4.65 | 7 | 3.9 | 6 | 4.5 | 4 | 4.3535
| Std Dev. |  | 0.12649 | 0.074471 | 0.16733 | 0 | 0.040825 | 0.098319 | 0.13663 | 0.075277 | 0.23166 | 0.06457 | 0 | 0.47329 | 0 | 0.089443 | 0.5164 | 0.13292 | 0 | 0.19181
| Rel.Std.Dev. |  | 2.691% | 1.407% | 2.574% | 0% | 0.8031% | 1.582% | 2.454% | 1.471% | 3.85% | 1.17% | 0% | 10.52% | 0% | 2.293% | 8.154% | 2.943% | 0% | 4.416%
| PDM3 |  | -9.104% | 2.337% | 25.71% | -3.302% | -1.691% | 20.23% | 7.657% | -1.046% | 16.36% | 6.741% | 16.04% | -12.97% | 35.38% | -24.58% | 22.48% | -12.65% | -22.64% | -16.01%
| Z-Score (Absolute) |  | 0.51 | 0.25 | 1.8 | 0.13 | 0.02 | 1.43 | 0.6 | 0.02 | 1.18 | 0.54 | NA | 0.77 | NA | 1.54 | NA | 0.75 | NA | 0.97
Results from laboratories 1.16, 1.18, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 211b (Certified Value 0.59 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1.5 | 0.74909 | 0.4 | < 1 | 0.5 | 0.6 | 0.6 | 0.7 | 0.7 | < 1 | 0.43 | < 1 | 0.6 | 0.6 | 1.43 | 1 | [5 empty] | 1
| 1 | 2 | < 1.5 | 0.80119 | 0.5 | < 1 | 0.6 | 0.6 | 0.5 | 0.6 | 0.8 | < 1 | 0.42 | < 1 | 0.8 | 0.5 | 1.24 | 1 | 0.479 |  |  |  |  | 38
| 1 | 3 | < 1.5 | 0.73172 | 0.4 | < 1 | 0.5 | 0.6 | 0.5 | 0.8 | 0.7 | < 1 | 0.44 | < 1 | 0.4 | 0.7 | 1.2 | 1 | [5 empty] | 16
| 1 | 4 | < 1.5 | 0.6705 | 0.5 | < 1 | 0.7 | 0.6 | 0.6 | 0.7 | 0.8 | < 1 | 0.43 | < 1 | 0.3 | 0.5 | 1.49 | 1 | [5 empty] | 0.58707
| 1 | 5 | < 1.5 | 0.68955 | 0.5 | < 1 | 0.6 | 0.6 | 0.6 | 0.7 | 0.7 | < 1 | 0.43 | < 1 | 0.2 | 0.6 | 1.28 | 1 | 1.872 |  |  |  |  | 112
| 1 | 6 | < 1.5 | 0.70233 | 0.4 | < 1 | 0.6 | 0.5 | 0.6 | 0.6 | 0.8 | < 1 | 0.46 | < 1 | 0.8 | 0.6 | 1.16 | 1
| Mean |  |  | 0.72406 | 0.45 |  | 0.58333 | 0.58333 | 0.56667 | 0.68333 | 0.75 |  | 0.435 |  | 0.51667 | 0.58333 | 1.3 | 1 | 1.1755
| Median |  |  | 0.71702 | 0.45 |  | 0.6 | 0.6 | 0.6 | 0.7 | 0.75 |  | 0.43 |  | 0.5 | 0.6 | 1.26 | 1 | 1.1755
| Std Dev. |  |  | 0.047257 | 0.054772 |  | 0.075277 | 0.040825 | 0.05164 | 0.075277 | 0.054772 |  | 0.013784 |  | 0.25626 | 0.075277 | 0.13161 | 0 | 0.985
| Rel.Std.Dev. |  |  | 6.527% | 12.17% |  | 12.9% | 6.999% | 9.113% | 11.02% | 7.303% |  | 3.169% |  | 49.6% | 12.9% | 10.12% | 0% | 83.79%
| PDM3 |  |  | 23.33% | -23.35% |  | -0.637% | -0.637% | -3.476% | 16.4% | 27.75% |  | -25.9% |  | -11.99% | -0.637% | 121.4% | 70.34% | 100.2%
| Z-Score (Absolute) |  | 1.23 | 1.04 | 0.98 | 0.61 | 0 | 0 | 0.12 | 0.74 | 1.23 | 0.61 | 1.09 | 0.61 | 0.49 | 0 | 5.27 | NA | 4.36
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sm in OREAS 211b (Indicative Value 1.45 ppm) | [22 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 1.378 | 1.4541 | 2.5928 | 1.9 | 1.52 | [16 empty] | 1
| 1 | 2 | 1.459 | 1.4791 | 2.4722 | 1.7 | 1.45 | [16 empty] | 4
| 1 | 3 | 1.447 | 1.5453 | 2.4776 | 1.6 | 1.44 | [16 empty] | 16
| 1 | 4 | 1.42 | 1.4894 | 2.5496 | 1.5 | 1.43 | [16 empty] | 1.4509
| 1 | 5 | 1.424 | 1.502 | 2.5374 | 1.9 | 1.47 | [16 empty] | 10
| 1 | 6 | 1.454 | 1.3767 | 2.5261 | 1.4 | 1.45
| Mean |  | 1.4303 | 1.4744 | 2.5259 | 1.6667 | 1.46
| Median |  | 1.4355 | 1.4842 | 2.5318 | 1.65 | 1.45
| Std Dev. |  | 0.030151 | 0.056543 | 0.045563 | 0.20656 | 0.032249
| Rel.Std.Dev. |  | 2.108% | 3.835% | 1.804% | 12.39% | 2.209%
| PDM3 |  | -1.419% | 1.62% | 74.09% | 14.87% | 0.6258%
| Z-Score (Absolute) |  | 0.67 | 0 | 16.08 | 2.94 | 0.22
| Analytical results for Sn in OREAS 211b (Certified Value 0.53 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | 0.49 | 0.49952 | 0.17571 | 0.58 | 0.6 | 0.5 | 0.4 | 0.4 | 0.5 | 0.5 | 0.54428 | 0.7 | 0.6 | < 5 | < 1 | [6 empty] | 1
| 1 | 2 | 0.52 | 0.53176 | 0.13689 | 0.57 | 0.6 | 0.5 | 0.4 | 0.4 | 0.5 | 0.5 | 0.50242 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 11
| 1 | 3 | 0.5 | 0.52083 | 0.11163 | 0.58 | 0.6 | 0.5 | 0.3 | 0.4 | 0.5 | 0.5 | 0.56099 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 16
| 1 | 4 | 0.47 | 0.5203 | 0.093723 | 0.59 | 0.7 | 0.5 | 0.4 | 0.4 | 0.5 | 0.5 | 0.4975 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 0.53273
| 1 | 5 | 0.5 | 0.58152 | 0.092629 | 0.59 | 0.6 | 0.5 | 0.3 | 0.5 | 0.5 | 0.5 | 0.59112 | 0.7 | 0.6 | < 5 | < 1 | [6 empty] | 113
| 1 | 6 | 0.52 | 0.51376 | 0.011619 | 0.57 | 0.7 | 0.5 | 0.4 | 0.5 | 0.5 | 0.5 | 0.57711 | 0.7 | 0.4 | < 5 | < 1
| Mean |  | 0.5 | 0.52795 | 0.1037 | 0.58 | 0.63333 | 0.5 | 0.36667 | 0.43333 | 0.5 | 0.5 | 0.54557 | 0.7 | 0.61667
| Median |  | 0.5 | 0.52057 | 0.10268 | 0.58 | 0.6 | 0.5 | 0.4 | 0.4 | 0.5 | 0.5 | 0.55263 | 0.7 | 0.65
| Std Dev. |  | 0.018974 | 0.028294 | 0.054865 | 0.0089443 | 0.05164 | 0 | 0.05164 | 0.05164 | 0 | 0 | 0.038681 | 1.2162e-16 | 0.1169
| Rel.Std.Dev. |  | 3.795% | 5.359% | 52.91% | 1.542% | 8.154% | 0% | 14.08% | 11.92% | 0% | 0% | 7.09% | 1.737e-14% | 18.96%
| PDM3 |  | -6.144% | -0.8981% | -80.53% | 8.872% | 18.88% | -6.144% | -31.17% | -18.66% | -6.144% | -6.144% | 2.41% | 31.4% | 15.76%
| Z-Score (Absolute) |  | 0 | 0.28 | 4.01 | 0.81 | 1.35 | 0 | 1.35 | 0.67 | 0 | 0 | 0.46 | 2.02 | 1.18 | 20.23 | 0
| Analytical results for Sr in OREAS 211b (Certified Value 35.3 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 36.25 | 37.5709 | 25.049 | 62.9 | 37.7 | 34.2 | 42.2 | 33.9 | 30.4 | 36.3 | 38.7994 | 63.5 | 32.2 | 68 | 26 | 55.8 | 37.6 | 31 | 24.863 |  |  | 1
| 1 | 2 | 36.66 | 38.0839 | 25.04 | 57.8 | 40.4 | 37 | 42.1 | 34 | 31.3 | 35.7 | 38.947 | 60.8 | 36.8 | 63 | 26.3 | 55.1 | 36.1 | 32 | 26.199 |  |  | 12
| 1 | 3 | 36.47 | 36.5583 | 26.061 | 57.4 | 35.8 | 37.1 | 42.3 | 32.7 | 30 | 33.8 | 37.9637 | 63.1 | 37.4 | 65 | 26.7 | 49.3 | 37.5 | 31 | 26.761 |  |  | 16
| 1 | 4 | 35.42 | 38.8394 | 24.608 | 59.8 | 38.2 | 39.3 | 42.5 | 34.1 | 31.7 | 34.1 | 39.5488 | 61 | 33.3 | 68 | 27.4 | 53.3 | 37.9 | 31 | 23.913 |  |  | 35.281
| 1 | 5 | 35.46 | 39.6618 | 25.029 | 59.6 | 38.7 | 37.6 | 42.4 | 33.8 | 30.3 | 35.6 | 39.0279 | 61.5 | 33.2 | 65 | 27.5 | 42.4 | 41.3 | 30 | 25.323 |  |  | 114
| 1 | 6 | 36.48 | 37.3435 | 24.959 | 56.9 | 40 | 37 | 41.6 | 34.7 | 29.9 | 35.9 | 38.0588 | 60.8 | 36.5 | 64 | 26.5 | 52.7 | 39.6 | 31 | 26.05
| Mean |  | 36.1233 | 38.0096 | 25.1243 | 59.0667 | 38.4667 | 37.0333 | 42.1833 | 33.8667 | 30.6 | 35.2333 | 38.7243 | 61.7833 | 34.9 | 65.5 | 26.7333 | 51.4333 | 38.3333 | 31 | 25.5182
| Median |  | 36.36 | 37.8274 | 25.0345 | 58.7 | 38.45 | 37.05 | 42.25 | 33.95 | 30.35 | 35.65 | 38.8732 | 61.25 | 34.9 | 65 | 26.6 | 53 | 37.75 | 31 | 25.6865
| Std Dev. |  | 0.54518 | 1.1105 | 0.4885 | 2.2178 | 1.6681 | 1.6428 | 0.31885 | 0.6532 | 0.73212 | 1.027 | 0.60805 | 1.209 | 2.2432 | 2.0736 | 0.60222 | 4.975 | 1.8338 | 0.63246 | 1.0329
| Rel.Std.Dev. |  | 1.509% | 2.922% | 1.944% | 3.755% | 4.337% | 4.436% | 0.7559% | 1.929% | 2.393% | 2.915% | 1.57% | 1.957% | 6.428% | 3.166% | 2.253% | 9.673% | 4.784% | 2.04% | 4.048%
| PDM3 |  | 2.388% | 7.734% | -28.79% | 67.42% | 9.029% | 4.967% | 19.56% | -4.009% | -13.27% | -0.1351% | 9.76% | 75.12% | -1.08% | 85.65% | -24.23% | 45.78% | 8.652% | -12.13% | -27.67%
| Z-Score (Absolute) |  | 0.12 | 0.13 | 1.56 | 2.88 | 0.19 | 0 | 0.67 | 0.41 | 0.84 | 0.24 | 0.22 | 3.24 | 0.28 | 3.73 | 1.35 | 1.89 | 0.17 | 0.79 | 1.51
| Analytical results for Ta in OREAS 211b (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 25
| 1 | 3 | < 0.01 | < 0.05 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 115
| 1 | 6 | < 0.01 | < 0.05 | < 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | 0.01
| Mean |  |  |  |  | 0.01 | 0.01 | 0.01 |  |  |  |  | 0.01
| Median |  |  |  |  | 0.01 | 0.01 | 0.01 |  |  |  |  | 0.01
| Std Dev. |  |  |  |  | 0 | 0 | 0
| Rel.Std.Dev. |  |  |  |  | 0% | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 7.42 | 0.67 | 0 | 0 | 1.01 | 0.67 | 0.67 | 7.42 | 7.42 | 0.34
| Analytical results for Tb in OREAS 211b (Certified Value 0.32 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.263 | 0.31593 | 0.42388 | 0.3 | 0.325 | [16 empty] | 1
| 1 | 2 | 0.268 | 0.32587 | 0.39136 | 0.3 | 0.325 | [16 empty] | 26
| 1 | 3 | 0.278 | 0.33192 | 0.39664 | 0.3 | 0.32 | [16 empty] | 16
| 1 | 4 | 0.266 | 0.32533 | 0.42076 | 0.3 | 0.335 | [16 empty] | 0.323
| 1 | 5 | 0.262 | 0.35722 | 0.37972 | 0.3 | 0.33 | [16 empty] | 116
| 1 | 6 | 0.274 | 0.31088 | 0.37972 | 0.3 | 0.32
| Mean |  | 0.2685 | 0.32786 | 0.39868 | 0.3 | 0.32583
| Median |  | 0.267 | 0.3256 | 0.394 | 0.3 | 0.325
| Std Dev. |  | 0.0063166 | 0.016244 | 0.01949 | 0 | 0.0058452
| Rel.Std.Dev. |  | 2.353% | 4.954% | 4.889% | 0% | 1.794%
| PDM3 |  | -16.87% | 1.504% | 23.43% | -7.121% | 0.8771%
| Z-Score (Absolute) |  | 1.5 | 0.05 | 1.9 | 0.67 | 0
| Analytical results for Te in OREAS 211b (Certified Value 0.061 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | [6 empty] | 3
| 1 | 1 | < 0.2 | 0.1 | 0.28 | 0.06 | 0.06 | 0.07 | 0.07 | 0.07 | 0.047365 | 0.07 | 0.06 | 0.06 | 0.05 | 0.06 | < 0.2 | [6 empty] | 1
| 1 | 2 | < 0.2 | 0.07 | 0.28 | 0.06 | 0.07 | 0.06 | 0.06 | 0.07 | 0.060786 | 0.05 | 0.06 | 0.07 | 0.07 | 0.05 | < 0.2 | [6 empty] | 27
| 1 | 3 | < 0.2 | 0.1 | 0.28 | 0.08 | 0.07 | 0.04 | 0.06 | 0.07 | 0.049938 | < 0.03 | 0.06 | 0.07 | 0.06 | 0.06 | < 0.2 | [6 empty] | 16
| 1 | 4 | < 0.2 | 0.09 | 0.29 | 0.05 | 0.06 | 0.07 | 0.06 | 0.06 | 0.064128 | 0.05 | 0.06 | 0.04 | 0.07 | 0.06 | < 0.2 | [6 empty] | 0.060855
| 1 | 5 | < 0.2 | 0.09 | 0.29 | 0.06 | 0.07 | 0.05 | 0.06 | 0.06 | 0.061183 | 0.05 | 0.06 | 0.09 | 0.06 | 0.06 | < 0.2 | [6 empty] | 117
| 1 | 6 | < 0.2 | 0.08 | 0.3 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.05703 | 0.05 | 0.06 | 0.07 | 0.06 | 0.04 | < 0.2
| Mean |  |  | 0.088333 | 0.28667 | 0.063333 | 0.066667 | 0.06 | 0.063333 | 0.066667 | 0.056738 | 0.054 | 0.06 | 0.066667 | 0.061667 | 0.055
| Median |  |  | 0.09 | 0.285 | 0.06 | 0.07 | 0.065 | 0.06 | 0.07 | 0.058908 | 0.05 | 0.06 | 0.07 | 0.06 | 0.06
| Std Dev. |  |  | 0.01169 | 0.008165 | 0.010328 | 0.005164 | 0.012649 | 0.005164 | 0.005164 | 0.0067071 | 0.0089443 | 0 | 0.01633 | 0.0075277 | 0.0083666
| Rel.Std.Dev. |  |  | 13.23% | 2.848% | 16.31% | 7.746% | 21.08% | 8.154% | 7.746% | 11.82% | 16.56% | 0% | 24.49% | 12.21% | 15.21%
| PDM3 |  |  | 45.15% | 371.1% | 4.073% | 9.55% | -1.405% | 4.073% | 9.55% | -6.765% | -11.26% | -1.405% | 9.55% | 1.334% | -9.621%
| Z-Score (Absolute) |  | 7.42 | 5.06 | 45.18 | 0 | 0.67 | 0.67 | 0 | 0.67 | 1.33 | 3.2 | 0.67 | 0.67 | 0.34 | 1.69 | 7.42
| Analytical results for Th in OREAS 211b (Certified Value 0.58 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 2
| 1 | 1 | 0.5 | 0.57549 | 3.557 | 0.74992 | 0.6 | 0.7 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.59406 | 0.64 | 0.4 | 0.5 | 0.52 | < 5 | 0.7 |  |  |  | 1
| 1 | 2 | 0.52 | 0.56806 | 3.792 | 0.66236 | 0.6 | 0.7 | 0.6 | 0.6 | 0.5 | 0.5 | 0.6 | 0.62988 | 0.62 | 0.6 | 0.5 | 0.57 | < 5 | 0.75 |  |  |  | 28
| 1 | 3 | 0.53 | 0.59054 | 4.043 | 0.60004 | 0.6 | 0.6 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.56791 | 0.63 | 0.5 | 0.5 | 0.56 | < 5 | 0.7 |  |  |  | 16
| 1 | 4 | 0.52 | 0.55606 | 3.275 | 0.6306 | 0.6 | 0.7 | 0.6 | 0.6 | 0.5 | 0.5 | 0.6 | 0.58962 | 0.64 | 0.5 | 0.5 | 0.54 | < 5 | 0.8 |  |  |  | 0.58064
| 1 | 5 | 0.53 | 0.59362 | 4.127 | 0.68908 | 0.6 | 0.7 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.59121 | 0.62 | 0.5 | 0.5 | 0.54 | < 5 | 0.7 |  |  |  | 118
| 1 | 6 | 0.52 | 0.53861 | 3.517 | 0.63496 | 0.6 | 0.7 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.63922 | 0.64 | 0.5 | 0.5 | 0.6 | < 5 | 0.75
| Mean |  | 0.52 | 0.5704 | 3.7185 | 0.66116 | 0.6 | 0.68333 | 0.53333 | 0.6 | 0.5 | 0.5 | 0.6 | 0.60199 | 0.63167 | 0.5 | 0.5 | 0.555 |  | 0.73333
| Median |  | 0.52 | 0.57178 | 3.6745 | 0.64866 | 0.6 | 0.7 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.59264 | 0.635 | 0.5 | 0.5 | 0.55 |  | 0.725
| Std Dev. |  | 0.010954 | 0.020938 | 0.32892 | 0.052944 | 0 | 0.040825 | 0.05164 | 0 | 0 | 0 | 0 | 0.027045 | 0.0098319 | 0.063246 | 0 | 0.028107 |  | 0.040825
| Rel.Std.Dev. |  | 2.107% | 3.671% | 8.846% | 8.008% | 0% | 5.974% | 9.682% | 0% | 0% | 0% | 0% | 4.493% | 1.557% | 12.65% | 0% | 5.064% |  | 5.567%
| PDM3 |  | -10.44% | -1.764% | 540.4% | 13.87% | 3.335% | 17.69% | -8.147% | 3.335% | -13.89% | -13.89% | 3.335% | 3.677% | 8.788% | -13.89% | -13.89% | -4.416% |  | 26.3%
| Z-Score (Absolute) |  | 0.74 | 0.27 | 28.67 | 0.56 | 0 | 0.77 | 0.61 | 0 | 0.92 | 0.92 | 0 | 0.02 | 0.29 | 0.92 | 0.92 | 0.41 | 17.47 | 1.23
| Analytical results for Ti in OREAS 211b (Certified Value 0.371 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.3147 | 0.35938 | 0.20236 | 0.41 | 0.349 | 0.338 | 0.417 | 0.321 | 0.315 | 0.338 | 0.44561 | 0.462 | 0.351 | 0.448 | 0.22 | 0.371 | 0.32 | 0.21 | 0.31735 |  |  | 1
| 1 | 2 | 0.3175 | 0.36109 | 0.19966 | 0.41 | 0.368 | 0.362 | 0.419 | 0.322 | 0.335 | 0.338 | 0.44268 | 0.464 | 0.393 | 0.4308 | 0.22 | 0.365 | 0.34 | 0.205 | 0.32503 |  |  | 29
| 1 | 3 | 0.3165 | 0.35103 | 0.21391 | 0.39 | 0.329 | 0.349 | 0.417 | 0.321 | 0.315 | 0.321 | 0.44166 | 0.467 | 0.4 | 0.4373 | 0.227 | 0.35 | 0.35 | 0.215 | 0.34999 |  |  | 16
| 1 | 4 | 0.306 | 0.3655 | 0.18554 | 0.41 | 0.35 | 0.384 | 0.421 | 0.335 | 0.325 | 0.329 | 0.44543 | 0.47 | 0.349 | 0.4548 | 0.248 | 0.351 | 0.35 | 0.21 | 0.28983 |  |  | 0.37124
| 1 | 5 | 0.311 | 0.36829 | 0.19908 | 0.4 | 0.353 | 0.36 | 0.409 | 0.332 | 0.33 | 0.337 | 0.44474 | 0.445 | 0.357 | 0.4397 | 0.249 | 0.348 | 0.36 | 0.205 | 0.31169 |  |  | 119
| 1 | 6 | 0.3209 | 0.34871 | 0.19818 | 0.39 | 0.365 | 0.369 | 0.406 | 0.338 | 0.315 | 0.328 | 0.44719 | 0.474 | 0.396 | 0.4355 | 0.229 | 0.474 | 0.35 | 0.22 | 0.3288
| Mean |  | 0.31443 | 0.359 | 0.19979 | 0.40167 | 0.35233 | 0.36033 | 0.41483 | 0.32817 | 0.3225 | 0.33183 | 0.44455 | 0.46367 | 0.37433 | 0.44102 | 0.23217 | 0.3765 | 0.345 | 0.21083 | 0.32045
| Median |  | 0.3156 | 0.36024 | 0.19937 | 0.405 | 0.3515 | 0.361 | 0.417 | 0.327 | 0.32 | 0.333 | 0.44508 | 0.4655 | 0.375 | 0.4385 | 0.228 | 0.358 | 0.35 | 0.21 | 0.32119
| Std Dev. |  | 0.0052626 | 0.0077793 | 0.0090823 | 0.0098319 | 0.013909 | 0.015908 | 0.005947 | 0.0077309 | 0.0088034 | 0.006969 | 0.0020375 | 0.010093 | 0.024345 | 0.0088217 | 0.013167 | 0.048649 | 0.013784 | 0.0058452 | 0.019935
| Rel.Std.Dev. |  | 1.674% | 2.167% | 4.546% | 2.448% | 3.948% | 4.415% | 1.434% | 2.356% | 2.73% | 2.1% | 0.4583% | 2.177% | 6.503% | 2% | 5.671% | 12.92% | 3.995% | 2.772% | 6.221%
| PDM3 |  | -15.3% | -3.297% | -46.18% | 8.196% | -5.093% | -2.938% | 11.74% | -11.6% | -13.13% | -10.61% | 19.75% | 24.9% | 0.8331% | 18.8% | -37.46% | 1.417% | -7.068% | -43.21% | -13.68%
| Z-Score (Absolute) |  | 0.8 | 0.14 | 3.23 | 1.04 | 0 | 0.17 | 1.32 | 0.51 | 0.63 | 0.43 | 1.95 | 2.35 | 0.47 | 1.88 | 2.54 | 0.51 | 0.16 | 2.99 | 0.67
| Analytical results for Tl in OREAS 211b (Certified Value 0.077 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 3
| 1 | 1 | 0.08 | 0.079443 | 0.069282 | 0.08 | 0.31 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.08241 | 0.075 | 0.08 | 0.07 | 0.07 | 0.09 | < 5 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.09 | 0.072979 | 0.062895 | 0.08 | 0.32 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.083248 | 0.076 | 0.08 | 0.09 | 0.08 | 0.09 | < 5 | < 0.1 |  |  |  | 30
| 1 | 3 | 0.08 | 0.079746 | 0.067919 | 0.07 | 0.32 | 0.08 | 0.09 | 0.07 | 0.07 | 0.08 | 0.079431 | 0.074 | 0.08 | 0.07 | 0.07 | 0.09 | < 5 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.08 | 0.07856 | 0.065879 | 0.08 | 0.32 | 0.08 | 0.09 | 0.08 | 0.08 | 0.08 | 0.078686 | 0.079 | 0.08 | 0.06 | 0.08 | 0.1 | < 5 | 0.1 |  |  |  | 0.077424
| 1 | 5 | 0.09 | 0.08267 | 0.070903 | 0.08 | 0.32 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.087505 | 0.078 | 0.08 | 0.07 | 0.08 | 0.09 | < 5 | < 0.1 |  |  |  | 120
| 1 | 6 | 0.08 | 0.072895 | 0.070527 | 0.08 | 0.32 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.085636 | 0.081 | 0.08 | 0.07 | 0.08 | 0.11 | < 5 | 0.1
| Mean |  | 0.083333 | 0.077715 | 0.067901 | 0.078333 | 0.31833 | 0.08 | 0.083333 | 0.073333 | 0.071667 | 0.08 | 0.082819 | 0.077167 | 0.08 | 0.071667 | 0.076667 | 0.095 |  | 0.1
| Median |  | 0.08 | 0.079001 | 0.068601 | 0.08 | 0.32 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.082829 | 0.077 | 0.08 | 0.07 | 0.08 | 0.09 |  | 0.1
| Std Dev. |  | 0.005164 | 0.0039506 | 0.0030648 | 0.0040825 | 0.0040825 | 0 | 0.005164 | 0.005164 | 0.0040825 | 0 | 0.0034295 | 0.0026394 | 0 | 0.0098319 | 0.005164 | 0.0083666 |  | 0
| Rel.Std.Dev. |  | 6.197% | 5.083% | 4.514% | 5.212% | 1.282% | 0% | 6.197% | 7.042% | 5.696% | 0% | 4.141% | 3.42% | 0% | 13.72% | 6.736% | 8.807% |  | 0%
| PDM3 |  | 7.632% | 0.3764% | -12.3% | 1.175% | 311.2% | 3.327% | 7.632% | -5.283% | -7.436% | 3.327% | 6.969% | -0.3323% | 3.327% | -7.436% | -0.9781% | 22.7% |  | 29.16%
| Z-Score (Absolute) |  | 0.67 | 0.23 | 1.82 | 0.13 | 38.71 | 0.13 | 0.67 | 0.94 | 1.21 | 0.13 | 0.59 | 0.32 | 0.13 | 1.21 | 0.4 | 2.56 | 391.77 | 2.02
| Analytical results for Tm in OREAS 211b (Indicative Value 0.18 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 0.158 | 0.17383 | 0.2 | 0.165 | [17 empty] | 1
| 1 | 2 | 0.166 | 0.17093 | 0.2 | 0.165 | [17 empty] | 5
| 1 | 3 | 0.165 | 0.18091 | 0.2 | 0.16 | [17 empty] | 16
| 1 | 4 | 0.168 | 0.17867 | 0.2 | 0.16 | [17 empty] | 0.17576
| 1 | 5 | 0.17 | 0.18033 | 0.2 | 0.165 | [17 empty] | 11
| 1 | 6 | 0.161 | 0.1655 | 0.2 | 0.165
| Mean |  | 0.16467 | 0.17503 | 0.2 | 0.16333
| Median |  | 0.1655 | 0.17625 | 0.2 | 0.165
| Std Dev. |  | 0.0044572 | 0.0060828 | 3.0405e-17 | 0.002582
| Rel.Std.Dev. |  | 2.707% | 3.475% | 1.52e-14% | 1.581%
| PDM3 |  | -6.31% | -0.4139% | 13.79% | -7.069%
| Z-Score (Absolute) |  | 0.6 | 0.6 | 3.48 | 0.75
| Analytical results for U in OREAS 211b (Certified Value 0.14 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g |  |  |  | 2
| 1 | 1 | 0.14 | 0.14745 | 0.12567 | 0.10372 | 0.2 | 0.15 | 0.13 | 0.14 | 0.12 | 0.14 | 0.13 | 0.16015 | 0.14 | 0.17 | 0.13 | 0.1 | 0.15 | 0.15 |  |  |  | 1
| 1 | 2 | 0.13 | 0.16372 | 0.12432 | < 0.1 | 0.2 | 0.15 | 0.14 | 0.14 | 0.13 | 0.14 | 0.14 | 0.15948 | 0.14 | 0.17 | 0.14 | 0.1 | 0.15 | 0.15 |  |  |  | 32
| 1 | 3 | 0.13 | 0.15713 | 0.12812 | 0.12196 | 0.1 | 0.15 | 0.14 | 0.15 | 0.12 | 0.14 | 0.14 | 0.17841 | 0.14 | 0.17 | 0.12 | 0.1 | 0.15 | 0.15 |  |  |  | 16
| 1 | 4 | 0.12 | 0.13779 | 0.12026 | < 0.1 | 0.2 | 0.15 | 0.15 | 0.14 | 0.13 | 0.14 | 0.14 | 0.16794 | 0.15 | 0.17 | 0.13 | 0.1 | 0.15 | 0.15 |  |  |  | 0.14338
| 1 | 5 | 0.13 | 0.15333 | 0.12139 | < 0.1 | 0.2 | 0.15 | 0.13 | 0.15 | 0.13 | 0.14 | 0.13 | 0.18282 | 0.14 | 0.17 | 0.13 | 0.1 | 0.14 | 0.15 |  |  |  | 121
| 1 | 6 | 0.13 | 0.13197 | 0.12048 | 0.1054 | 0.2 | 0.15 | 0.14 | 0.14 | 0.13 | 0.14 | 0.13 | 0.16351 | 0.15 | 0.17 | 0.15 | 0.1 | 0.16 | 0.2
| Mean |  | 0.13 | 0.14856 | 0.12337 | 0.11036 | 0.18333 | 0.15 | 0.13833 | 0.14333 | 0.12667 | 0.14 | 0.135 | 0.16872 | 0.14333 | 0.17 | 0.13333 | 0.1 | 0.15 | 0.15833
| Median |  | 0.13 | 0.15039 | 0.12286 | 0.1054 | 0.2 | 0.15 | 0.14 | 0.14 | 0.13 | 0.14 | 0.135 | 0.16572 | 0.14 | 0.17 | 0.13 | 0.1 | 0.15 | 0.15
| Std Dev. |  | 0.0063246 | 0.011988 | 0.0031857 | 0.010081 | 0.040825 | 0 | 0.0075277 | 0.005164 | 0.005164 | 0 | 0.0054772 | 0.0097904 | 0.005164 | 0 | 0.010328 | 1.5202e-17 | 0.0063246 | 0.020412
| Rel.Std.Dev. |  | 4.865% | 8.069% | 2.582% | 9.135% | 22.27% | 0% | 5.442% | 3.603% | 4.077% | 0% | 4.057% | 5.803% | 3.603% | 0% | 7.746% | 1.52e-14% | 4.216% | 12.89%
| PDM3 |  | -9.33% | 3.618% | -13.95% | -23.03% | 27.87% | 4.619% | -3.518% | -0.03049% | -11.65% | -2.355% | -5.843% | 17.67% | -0.03049% | 18.57% | -7.005% | -30.25% | 4.619% | 10.43%
| Z-Score (Absolute) |  | 0.94 | 0.56 | 1.48 | 4.98 | NA | 0.67 | 0.27 | 0.13 | 1.21 | 0.13 | 0.54 | 2.19 | 0.13 | 2.29 | 0.67 | NA | 0.67 | 1.35
Results from laboratories 1.07 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 211b (Certified Value 136 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 0
| 1 | 1 | 127 | 140.8163 | 95.687 | 167 | 141 | 131 | 152 | 131 | 132 | 129 | 157.0897 | 184 | 128 | 179 | 112 | 174 | 125.5 | 120 | 131.191 |  |  | 1
| 1 | 2 | 136 | 142.8829 | 93.139 | 161 | 142 | 135 | 153 | 132 | 138 | 130 | 151.1417 | 182 | 136 | 176 | 111 | 183 | 124.3 | 120 | 136.806 |  |  | 33
| 1 | 3 | 136 | 137.8413 | 99.582 | 157 | 138 | 132 | 152 | 130 | 132 | 125 | 150.3185 | 184 | 138 | 179 | 113 | 176 | 128.5 | 130 | 146.138 |  |  | 16
| 1 | 4 | 130 | 141.1895 | 91.172 | 162 | 141 | 142 | 153 | 133 | 136 | 127 | 158.1319 | 185 | 127 | 182 | 118 | 179 | 128.8 | 120 | 127.721 |  |  | 136.2931
| 1 | 5 | 133 | 143.7344 | 97.152 | 165 | 140 | 135 | 150 | 132 | 137 | 129 | 157.7119 | 180 | 128 | 177 | 117 | 175 | 125.7 | 120 | 134.771 |  |  | 122
| 1 | 6 | 136 | 139.6401 | 90.716 | 156 | 144 | 137 | 149 | 134 | 133 | 126 | 155.0122 | 184 | 138 | 175 | 114 | 125 | 124.2 | 130 | 138.242
| Mean |  | 133 | 141.0174 | 94.5747 | 161.3333 | 141 | 135.3333 | 151.5 | 132 | 134.6667 | 127.6667 | 154.901 | 183.1667 | 132.5 | 178 | 114.1667 | 168.6667 | 126.1667 | 123.3333 | 135.8115
| Median |  | 134.5 | 141.0029 | 94.413 | 161.5 | 141 | 135 | 152 | 132 | 134.5 | 128 | 156.051 | 184 | 132 | 178 | 113.5 | 175.5 | 125.6 | 120 | 135.7885
| Std Dev. |  | 3.7947 | 2.1409 | 3.5066 | 4.3205 | 2 | 3.9328 | 1.6432 | 1.4142 | 2.6583 | 1.9664 | 3.4139 | 1.8348 | 5.3572 | 2.5298 | 2.7869 | 21.6395 | 2.0196 | 5.164 | 6.347
| Rel.Std.Dev. |  | 2.853% | 1.518% | 3.708% | 2.678% | 1.418% | 2.906% | 1.085% | 1.071% | 1.974% | 1.54% | 2.204% | 1.002% | 4.043% | 1.421% | 2.441% | 12.83% | 1.601% | 4.187% | 4.673%
| PDM3 |  | -2.416% | 3.466% | -30.61% | 18.37% | 3.454% | -0.7042% | 11.16% | -3.15% | -1.193% | -6.329% | 13.65% | 34.39% | -2.783% | 30.6% | -16.23% | 23.75% | -7.43% | -9.509% | -0.3534%
| Z-Score (Absolute) |  | 0.17 | 0.42 | 3 | 1.91 | 0.42 | 0 | 1.19 | 0.25 | 0.05 | 0.56 | 1.44 | 3.52 | 0.21 | 3.14 | 1.56 | 2.45 | 0.67 | 0.88 | 0.04
| Analytical results for W in OREAS 211b (Certified Value 11 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g |  |  |  |  | 1
| 1 | 1 | 8.9 | 11.6794 | 18.95 | 10.6 | 11.73 | 12.15 | 12.3 | 11.1 | 11.5 | 11.6 | 11.7228 | 8.9 | 7.9 | 8 | 8.6 | 13.5 | < 5 |  |  |  |  | 1
| 1 | 2 | 8.98 | 11.7446 | 18.047 | 10.7 | 12.01 | 11.95 | 12.1 | 11.15 | 11.25 | 12.4 | 11.3774 | 8.3 | 9.14 | 8 | 8.8 | 13.7 | < 5 |  |  |  |  | 34
| 1 | 3 | 8.65 | 12.3618 | 17.506 | 10.1 | 10.79 | 12.05 | 12.75 | 11.05 | 11.35 | 11.7 | 10.4288 | 8.7 | 7.87 | 7 | 9 | 13.6 | < 5 |  |  |  |  | 16
| 1 | 4 | 8.91 | 11.6754 | 17.149 | 10.1 | 11.61 | 12.65 | 13 | 11 | 11.7 | 12.4 | 10.8311 | 9.1 | 9.09 | 8 | 8.7 | 13.7 | < 5 |  |  |  |  | 10.9694
| 1 | 5 | 8.61 | 12.3276 | 16.65 | 10.5 | 12.03 | 12.3 | 12.3 | 11.45 | 11.2 | 12.65 | 11.128 | 8.5 | 9.44 | 8 | 9 | 13.4 | < 5 |  |  |  |  | 123
| 1 | 6 | 9.11 | 11.5487 | 16.094 | 10.8 | 11.64 | 12.35 | 12.8 | 11.5 | 11.6 | 12.05 | 10.7499 | 9 | 8.78 | 7 | 9.2 | 13.8 | < 5
| Mean |  | 8.86 | 11.8896 | 17.3993 | 10.4667 | 11.635 | 12.2417 | 12.5417 | 11.2083 | 11.4333 | 12.1333 | 11.0397 | 8.75 | 8.7033 | 7.6667 | 8.8833 | 13.6167
| Median |  | 8.905 | 11.712 | 17.3275 | 10.55 | 11.685 | 12.225 | 12.525 | 11.125 | 11.425 | 12.225 | 10.9795 | 8.8 | 8.935 | 8 | 8.9 | 13.65
| Std Dev. |  | 0.1937 | 0.35836 | 1.0162 | 0.30111 | 0.45174 | 0.24983 | 0.35555 | 0.21311 | 0.19916 | 0.42151 | 0.46672 | 0.30822 | 0.66761 | 0.5164 | 0.22286 | 0.1472
| Rel.Std.Dev. |  | 2.186% | 3.014% | 5.841% | 2.877% | 3.883% | 2.041% | 2.835% | 1.901% | 1.742% | 3.474% | 4.228% | 3.523% | 7.671% | 6.736% | 2.509% | 1.081%
| PDM3 |  | -19.23% | 8.389% | 58.62% | -4.583% | 6.068% | 11.6% | 14.33% | 2.178% | 4.229% | 10.61% | 0.6406% | -20.23% | -20.66% | -30.11% | -19.02% | 24.13%
| Z-Score (Absolute) |  | 1.55 | 0.36 | 3.83 | 0.54 | 0.2 | 0.58 | 0.77 | 0.07 | 0.07 | 0.51 | 0.18 | 1.62 | 1.65 | NA | 1.53 | 1.45 | 5.55
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 211b (Certified Value 11.2 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 2
| 1 | 1 | 10.31 | 11.2946 | 10.018 | 12.1 | 10.93 | 11.1 | 11.5 | 11.1 | 11.3 | 11.45 | 11.8282 | 14.6 | 13.6 | 12.3 | 9.946 | [6 empty] | 1
| 1 | 2 | 10.93 | 11.3461 | 10.538 | 11.8 | 11.41 | 11.35 | 12.1 | 10.65 | 11.35 | 11.25 | 12.0333 | 14.4 | 14 | 11.8 | 10.44 | [6 empty] | 35
| 1 | 3 | 10.96 | 11.0439 | 10.982 | 11.7 | 10.55 | 11.6 | 11.7 | 10.85 | 11 | 11.05 | 11.1407 | 14.5 | 13.1 | 11.9 | 10.804 | [6 empty] | 16
| 1 | 4 | 10.39 | 11.2594 | 10.463 | 12.1 | 11.08 | 12 | 11.85 | 10.85 | 11.7 | 11.05 | 11.6175 | 14.4 | 13.6 | 12.2 | 10.051 | [6 empty] | 11.2365
| 1 | 5 | 10.74 | 11.6123 | 10.619 | 11.9 | 10.88 | 11.7 | 11.65 | 10.85 | 11.25 | 11.2 | 11.6322 | 14.2 | 12.9 | 11.7 | 10.374 | [6 empty] | 124
| 1 | 6 | 10.81 | 11.1562 | 10.689 | 11.6 | 11.34 | 11.8 | 11.5 | 11.4 | 11.25 | 11.6 | 11.3512 | 14.4 | 14.5 | 11.8 | 10.514
| Mean |  | 10.69 | 11.2854 | 10.5515 | 11.8667 | 11.0317 | 11.5917 | 11.7167 | 10.95 | 11.3083 | 11.2667 | 11.6005 | 14.4167 | 13.6167 | 11.95 | 10.3548
| Median |  | 10.775 | 11.277 | 10.5785 | 11.85 | 11.005 | 11.65 | 11.675 | 10.85 | 11.275 | 11.225 | 11.6249 | 14.4 | 13.6 | 11.85 | 10.407
| Std Dev. |  | 0.27633 | 0.19304 | 0.31668 | 0.20656 | 0.31808 | 0.32314 | 0.22949 | 0.26268 | 0.22675 | 0.22061 | 0.32062 | 0.13292 | 0.58452 | 0.2429 | 0.31438
| Rel.Std.Dev. |  | 2.585% | 1.711% | 3.001% | 1.741% | 2.883% | 2.788% | 1.959% | 2.399% | 2.005% | 1.958% | 2.764% | 0.922% | 4.293% | 2.033% | 3.036%
| PDM3 |  | -4.863% | 0.4356% | -6.096% | 5.609% | -1.823% | 3.161% | 4.274% | -2.549% | 0.6397% | 0.2689% | 3.24% | 28.3% | 21.18% | 6.35% | -7.846%
| Z-Score (Absolute) |  | 1.02 | 0.04 | 1.25 | 0.92 | 0.46 | 0.47 | 0.67 | 0.59 | 0 | 0.07 | 0.48 | 5.13 | 3.81 | 1.06 | 1.57
| Analytical results for Yb in OREAS 211b (Indicative Value 1.31 ppm) | [22 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.988 | 1.0264 | 2.0163 | 1.2 | [17 empty] | 1
| 1 | 2 | 1.057 | 1.0437 | 1.9131 | 1.2 | [17 empty] | 6
| 1 | 3 | 1.028 | 1.1231 | 1.9437 | 1.2 | [17 empty] | 16
| 1 | 4 | 0.997 | 1.1165 | 1.9985 | 1.2 | [17 empty] | 1.3132
| 1 | 5 | 0.988 | 1.0871 | 1.9841 | 1.2 | [17 empty] | 12
| 1 | 6 | 1.054 | 1.0375 | 1.9152 | 1.2
| Mean |  | 1.0187 | 1.0724 | 1.9618 | 1.2
| Median |  | 1.0125 | 1.0654 | 1.9639 | 1.2
| Std Dev. |  | 0.03211 | 0.042159 | 0.044024 | 0
| Rel.Std.Dev. |  | 3.152% | 3.931% | 2.244% | 0%
| PDM3 |  | -22.43% | -18.34% | 49.39% | -8.621%
| Z-Score (Absolute) |  | 0.87 | 0.47 | 6.14 | 0.47
| Analytical results for Zn in OREAS 211b (Certified Value 113 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 105 | 112.9464 | 106.862 | 109.504 | 112 | 116 | 141 | 115 | 110 | 116 | 117 | 115.5591 | 104 | 115 | 102.8 | 113 | 103 | 102 | 117.4 | 119 | 109.953 | 1
| 1 | 2 | 109 | 113.9266 | 116.317 | 113.5175 | 110 | 125 | 115 | 114 | 112 | 119 | 117 | 116.8395 | 102 | 116 | 97.1 | 110 | 104.2 | 107 | 118.2 | 119 | 112.377 | 15
| 1 | 3 | 107 | 112.1829 | 116.077 | 107.0545 | 108 | 110 | 115 | 115 | 110 | 117 | 117 | 115.49 | 102 | 117 | 97.3 | 113 | 107.5 | 104 | 120.4 | 124 | 112.141 | 16
| 1 | 4 | 105 | 111.9027 | 116.324 | 112.9425 | 109 | 114 | 120 | 115 | 111 | 119 | 117 | 112.91 | 106 | 116 | 98.8 | 116 | 106.4 | 107 | 119.2 | 120 | 114.215 | 112.8153
| 1 | 5 | 106 | 112.6707 | 115.971 | 111.3555 | 109 | 112 | 115 | 115 | 109 | 119 | 117 | 112.6449 | 105 | 113 | 99.3 | 113 | 107.1 | 103 | 119.9 | 122 | 113.174 | 125
| 1 | 6 | 108 | 112.8444 | 113.53 | 105.18 | 107 | 117 | 116 | 115 | 111 | 118 | 116 | 112.6803 | 111 | 118 | 101.5 | 111 | 103.9 | 104 | 117.2 | 125 | 111.357
| Mean |  | 106.6667 | 112.7456 | 114.1802 | 109.9257 | 109.1667 | 115.6667 | 120.3333 | 114.8333 | 110.5 | 118 | 116.8333 | 114.354 | 105 | 115.8333 | 99.4667 | 112.6667 | 105.35 | 104.5 | 118.7167 | 121.5 | 112.2028
| Median |  | 106.5 | 112.7576 | 116.024 | 110.4297 | 109 | 115 | 115.5 | 115 | 110.5 | 118.5 | 117 | 114.2 | 104.5 | 116 | 99.05 | 113 | 105.3 | 104 | 118.7 | 121 | 112.259
| Std Dev. |  | 1.633 | 0.70394 | 3.7402 | 3.3163 | 1.7224 | 5.2409 | 10.3086 | 0.40825 | 1.0488 | 1.2649 | 0.40825 | 1.8291 | 3.3466 | 1.7224 | 2.2809 | 2.0656 | 1.8833 | 2.0736 | 1.3243 | 2.5884 | 1.4699
| Rel.Std.Dev. |  | 1.531% | 0.6244% | 3.276% | 3.017% | 1.578% | 4.531% | 8.567% | 0.3555% | 0.9491% | 1.072% | 0.3494% | 1.599% | 3.187% | 1.487% | 2.293% | 1.833% | 1.788% | 1.984% | 1.115% | 2.13% | 1.31%
| PDM3 |  | -5.45% | -0.06174% | 1.21% | -2.561% | -3.234% | 2.528% | 6.664% | 1.789% | -2.052% | 4.596% | 3.562% | 1.364% | -6.927% | 2.675% | -11.83% | -0.1317% | -6.617% | -7.371% | 5.231% | 7.698% | -0.5429%
| Z-Score (Absolute) |  | 1.15 | 0 | 0.27 | 0.53 | 0.67 | 0.55 | 1.43 | 0.39 | 0.42 | 0.99 | 0.77 | 0.3 | 1.46 | 0.58 | 2.5 | 0.01 | 1.39 | 1.55 | 1.13 | 1.65 | 0.1
| Analytical results for Zr in OREAS 211b (Certified Value 14.7 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | 13 | 15.0913 | 22.6086 | 12.5 | 14.5 | 13.9 | 17.5 | 13.8 | 14.2 | 15.1 | 15.6324 | 16.4 | 19 | 12.9 | 2 | [6 empty] | 1
| 1 | 2 | 13.2 | 15.0221 | 23.2064 | 15.5 | 15.7 | 14.5 | 17 | 13.7 | 14.6 | 14.5 | 16.2097 | 16.2 | 19.5 | 12.2 | 2 | [6 empty] | 16
| 1 | 3 | 13.4 | 14.5954 | 21.3836 | 12.2 | 13.4 | 13.9 | 17.1 | 13.8 | 14.4 | 15.4 | 15.0801 | 16.6 | 18.2 | 13.2 | 2 | [6 empty] | 16
| 1 | 4 | 12.9 | 15.1031 | 23.4612 | 12.7 | 15 | 14.5 | 17.4 | 14.6 | 14.4 | 14.1 | 15.084 | 16.2 | 18.9 | 13.5 | 2 | [6 empty] | 14.6974
| 1 | 5 | 12.9 | 15.3511 | 22.638 | 13.7 | 14.9 | 13.8 | 17.4 | 14.1 | 14.7 | 14.4 | 16.2762 | 16.3 | 18.9 | 13.3 | 2 | [6 empty] | 126
| 1 | 6 | 13.6 | 14.6598 | 21.3738 | 14.8 | 15.9 | 14.5 | 16.8 | 14.9 | 14.4 | 15.1 | 15.808 | 16.7 | 19.3 | 12.5 | 2
| Mean |  | 13.1667 | 14.9705 | 22.4453 | 13.5667 | 14.9 | 14.1833 | 17.2 | 14.15 | 14.45 | 14.7667 | 15.6817 | 16.4 | 18.9667 | 12.9333 | 2
| Median |  | 13.1 | 15.0567 | 22.6233 | 13.2 | 14.95 | 14.2 | 17.25 | 13.95 | 14.4 | 14.8 | 15.7202 | 16.35 | 18.95 | 13.05 | 2
| Std Dev. |  | 0.28752 | 0.28876 | 0.88887 | 1.3441 | 0.90111 | 0.34881 | 0.27568 | 0.49295 | 0.17607 | 0.50465 | 0.52343 | 0.20976 | 0.44572 | 0.50067 | 0
| Rel.Std.Dev. |  | 2.184% | 1.929% | 3.96% | 9.908% | 6.048% | 2.459% | 1.603% | 3.484% | 1.218% | 3.417% | 3.338% | 1.279% | 2.35% | 3.871% | 0%
| PDM3 |  | -10.42% | 1.858% | 52.72% | -7.693% | 1.378% | -3.498% | 17.03% | -3.725% | -1.683% | 0.4712% | 6.697% | 11.58% | 29.05% | -12% | -86.39%
| Z-Score (Absolute) |  | 1.06 | 0.09 | 4.85 | 0.81 | 0.04 | 0.41 | 1.51 | 0.44 | 0.24 | 0.04 | 0.54 | 1 | 2.64 | 1.21 | NA
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 211b (Indicative Value 13.66 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.65 | 1
| 1 | 2 | 13.66 | 9
| Mean |  | 13.655 | 16
| Median |  | 13.655 | 13.655
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.05178%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 211b (Indicative Value 30 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 10
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 211b (Indicative Value 240 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 234 | 1
| 1 | 2 | 246 | 11
| Mean |  | 240 | 16
| Median |  | 240 | 240.0475
| Std Dev. |  | 8.4853 | 17
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | -0.01979%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 211b (Indicative Value 10.14 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.13 | 1
| 1 | 2 | 10.15 | 12
| Mean |  | 10.14 | 16
| Median |  | 10.14 | 10.14
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.1395%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 211b (Indicative Value 1510 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1480 | 1
| 1 | 2 | 1540 | 13
| Mean |  | 1510 | 16
| Median |  | 1510 | 1510
| Std Dev. |  | 42.4264 | 19
| Rel.Std.Dev. |  | 2.81%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 211b (Indicative Value 60 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 60 | 14
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 211b (Indicative Value 241 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 234 | 1
| 1 | 2 | 248 | 15
| Mean |  | 241 | 16
| Median |  | 241 | 241.1475
| Std Dev. |  | 9.8995 | 21
| Rel.Std.Dev. |  | 4.108%
| PDM3 |  | -0.06117%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 211b (Indicative Value 200 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 200 | 16
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 211b (Indicative Value 11.7 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.724 | 1
| 1 | 2 | 11.681 | 9
| Mean |  | 11.7025 | 16
| Median |  | 11.7025 | 11.7021
| Std Dev. |  | 0.030406 | 15
| Rel.Std.Dev. |  | 0.2598%
| PDM3 |  | 0.003465%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 211b (Indicative Value 0.446 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.447 | 1
| 1 | 2 | 0.445 | 10
| Mean |  | 0.446 | 16
| Median |  | 0.446 | 0.446
| Std Dev. |  | 0.0014142 | 16
| Rel.Std.Dev. |  | 0.3171%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 211b (Indicative Value 6.61 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.62 | 11
| Mean |  | 6.61 | 16
| Median |  | 6.61 | 6.61
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.214%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 211b (Indicative Value 0.186 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.186 | 1
| 1 | 2 | 0.186 | 12
| Mean |  | 0.186 | 16
| Median |  | 0.186 | 0.186
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 211b (Indicative Value 2.56 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.56 | 1
| 1 | 2 | 2.56 | 13
| Mean |  | 2.56 | 16
| Median |  | 2.56 | 2.56
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 211b (Indicative Value 155 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 150 | 14
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 7.0711 | 20
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 211b (Indicative Value 0.094 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.094 | 1
| 1 | 2 | 0.094 | 15
| Mean |  | 0.094 | 16
| Median |  | 0.094 | 0.093956
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.04726%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 211b (Indicative Value 85 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 80 | 16
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 211b (Indicative Value 0.252 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.254 | 20
| Mean |  | 0.252 | 16
| Median |  | 0.252 | 0.252
| Std Dev. |  | 0.0028284 | 15
| Rel.Std.Dev. |  | 1.122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 211b (Indicative Value 50.16 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 50.16 | 1
| 1 | 2 | 50.16 | 10
| Mean |  | 50.16 | 16
| Median |  | 50.16 | 50.16
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 211b (Indicative Value 45 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 40 | 11
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 7.0711 | 17
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 211b (Indicative Value 225 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 230 | 1
| 1 | 2 | 220 | 12
| Mean |  | 225 | 16
| Median |  | 225 | 225
| Std Dev. |  | 7.0711 | 18
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 211b (Indicative Value 1.11 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.107 | 1
| 1 | 2 | 1.11 | 13
| Mean |  | 1.1085 | 16
| Median |  | 1.1085 | 1.1085
| Std Dev. |  | 0.0021213 | 19
| Rel.Std.Dev. |  | 0.1914%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 211b (Indicative Value 527 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 518 | 1
| 1 | 2 | 536 | 14
| Mean |  | 527 | 16
| Median |  | 527 | 526.634
| Std Dev. |  | 12.7279 | 20
| Rel.Std.Dev. |  | 2.415%
| PDM3 |  | 0.0695%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 211b (Indicative Value 140 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 140 | 1
| 1 | 2 | 140 | 15
| Mean |  | 140 | 16
| Median |  | 140 | 140
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 211b (Indicative Value 100 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 16
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 211b (Indicative Value 2.96 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.96 | 1
| 1 | 2 | 2.95 | 18
| Mean |  | 2.955 | 16
| Median |  | 2.955 | 2.955
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.2393%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 211b (Indicative Value 0.115 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.12 | 1
| 1 | 2 | 0.11 | 20
| Mean |  | 0.115 | 16
| Median |  | 0.115 | 0.115
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 211b (Indicative Value 0.295 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.3 | 20
| Mean |  | 0.295 | 16
| Median |  | 0.295 | 0.295
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 211b (Indicative Value 0.35 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.3 | 22
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 28
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 211b (Indicative Value 33.2 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.6 | 1
| 1 | 2 | 32.8 | 10
| Mean |  | 33.2 | 16
| Median |  | 33.2 | 33.2
| Std Dev. |  | 0.56569 | 29
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 211b (Indicative Value 150 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 151 | 1
| 1 | 2 | 149 | 24
| Mean |  | 150 | 16
| Median |  | 150 | 150
| Std Dev. |  | 1.4142 | 30
| Rel.Std.Dev. |  | 0.9428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 211b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 25
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 211b (Indicative Value < 0.2 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 32
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 211b (Indicative Value 0.5 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 27
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 211b (Indicative Value 11.6 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.5 | 1
| 1 | 2 | 11.6 | 28
| Mean |  | 11.55 | 16
| Median |  | 11.55 | 11.55
| Std Dev. |  | 0.070711 | 34
| Rel.Std.Dev. |  | 0.6122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 211b (Indicative Value 45.7 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 45.4 | 14
| Mean |  | 45.7 | 16
| Median |  | 45.7 | 45.7
| Std Dev. |  | 0.42426 | 35
| Rel.Std.Dev. |  | 0.9284%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 211b (Indicative Value 143 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 143 | 1
| 1 | 2 | 142 | 30
| Mean |  | 142.5 | 16
| Median |  | 142.5 | 142.5
| Std Dev. |  | 0.70711 | 36
| Rel.Std.Dev. |  | 0.4962%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 211b (Indicative Value 0.74 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.72 | 1
| 1 | 2 | 0.75 | 31
| Mean |  | 0.735 | 16
| Median |  | 0.735 | 0.735
| Std Dev. |  | 0.021213 | 37
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 211b (Indicative Value 179 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 178 | 16
| Mean |  | 179 | 16
| Median |  | 179 | 179
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 0.7901%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 211b (Indicative Value 3.88 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.83 | 1
| 1 | 2 | 3.93 | 4
| Mean |  | 3.88 | 16
| Median |  | 3.88 | 3.88
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 1.822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 211b (Indicative Value 2.49 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.49 | 1
| 1 | 2 | 2.48 | 5
| Mean |  | 2.485 | 16
| Median |  | 2.485 | 2.485
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.2846%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 211b (Indicative Value 0.94 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.94 | 1
| 1 | 2 | 0.94 | 35
| Mean |  | 0.94 | 16
| Median |  | 0.94 | 0.94
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 211b (Indicative Value 15.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.7 | 1
| 1 | 2 | 15.4 | 36
| Mean |  | 15.55 | 16
| Median |  | 15.55 | 15.55
| Std Dev. |  | 0.21213 | 42
| Rel.Std.Dev. |  | 1.364%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 211b (Indicative Value 3.21 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 3.22 | 6
| Mean |  | 3.21 | 16
| Median |  | 3.21 | 3.21
| Std Dev. |  | 0.014142 | 43
| Rel.Std.Dev. |  | 0.4406%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 211b (Indicative Value 1.38 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.4 | 2
| Mean |  | 1.375 | 16
| Median |  | 1.375 | 1.375
| Std Dev. |  | 0.035355 | 44
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 211b (Indicative Value 1.85 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.86 | 1
| 1 | 2 | 1.83 | 22
| Mean |  | 1.845 | 16
| Median |  | 1.845 | 1.845
| Std Dev. |  | 0.021213 | 28
| Rel.Std.Dev. |  | 1.15%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 211b (Indicative Value 0.85 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.86 | 1
| 1 | 2 | 0.83 | 23
| Mean |  | 0.845 | 16
| Median |  | 0.845 | 0.845
| Std Dev. |  | 0.021213 | 29
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 211b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 7
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 211b (Indicative Value 4.77 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.82 | 1
| 1 | 2 | 4.72 | 25
| Mean |  | 4.77 | 16
| Median |  | 4.77 | 4.77
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 1.482%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 211b (Indicative Value 0.33 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.33 | 1
| 1 | 2 | 0.33 | 4
| Mean |  | 0.33 | 16
| Median |  | 0.33 | 0.33
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 211b (Indicative Value 0.145 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.144 | 1
| 1 | 2 | 0.145 | 27
| Mean |  | 0.1445 | 16
| Median |  | 0.1445 | 0.1445
| Std Dev. |  | 0.00070711 | 33
| Rel.Std.Dev. |  | 0.4893%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 211b (Indicative Value 1.8 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 28
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 211b (Indicative Value 3.43 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.34 | 1
| 1 | 2 | 3.52 | 29
| Mean |  | 3.43 | 16
| Median |  | 3.43 | 3.43
| Std Dev. |  | 0.12728 | 35
| Rel.Std.Dev. |  | 3.711%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 211b (Indicative Value 8.33 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.21 | 1
| 1 | 2 | 8.44 | 5
| Mean |  | 8.325 | 16
| Median |  | 8.325 | 8.325
| Std Dev. |  | 0.16263 | 36
| Rel.Std.Dev. |  | 1.954%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 211b (Indicative Value 98 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 98 | 1
| 1 | 2 | 98 | 14
| Mean |  | 98 | 16
| Median |  | 98 | 98
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 211b (Indicative Value 14.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 15 | 16
| Mean |  | 14.5 | 16
| Median |  | 14.5 | 14.5
| Std Dev. |  | 0.70711 | 38
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 211b (Indicative Value 1.71 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.7 | 1
| 1 | 2 | 1.72 | 33
| Mean |  | 1.71 | 16
| Median |  | 1.71 | 1.71
| Std Dev. |  | 0.014142 | 39
| Rel.Std.Dev. |  | 0.827%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 211b (Indicative Value 9.75 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.85 | 1
| 1 | 2 | 9.65 | 34
| Mean |  | 9.75 | 16
| Median |  | 9.75 | 9.75
| Std Dev. |  | 0.14142 | 40
| Rel.Std.Dev. |  | 1.45%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 211b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 35
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 211b (Indicative Value 1.45 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.5 | 36
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 211b (Indicative Value 41.2 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 41.3 | 1
| 1 | 2 | 41.1 | 37
| Mean |  | 41.2 | 16
| Median |  | 41.2 | 41.2
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 0.3433%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 211b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 211b (Indicative Value 2.62 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.57 | 1
| 1 | 2 | 2.66 | 4
| Mean |  | 2.615 | 16
| Median |  | 2.615 | 2.615
| Std Dev. |  | 0.06364 | 28
| Rel.Std.Dev. |  | 2.434%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 211b (Indicative Value 0.9 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 0.8 | 11
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 211b (Indicative Value 208 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 208 | 1
| 1 | 2 | 207 | 12
| Mean |  | 207.5 | 16
| Median |  | 207.5 | 207.5
| Std Dev. |  | 0.70711 | 30
| Rel.Std.Dev. |  | 0.3408%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 211b (Indicative Value 0.21 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.21 | 25
| Mean |  | 0.21 | 16
| Median |  | 0.21 | 0.21
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 211b (Indicative Value 0.59 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.58 | 26
| Mean |  | 0.585 | 16
| Median |  | 0.585 | 0.585
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 1.209%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 211b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 211b (Indicative Value 0.73 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.72 | 1
| 1 | 2 | 0.74 | 28
| Mean |  | 0.73 | 16
| Median |  | 0.73 | 0.73
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 211b (Indicative Value 0.649 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.65 | 1
| 1 | 2 | 0.647 | 29
| Mean |  | 0.6485 | 16
| Median |  | 0.6485 | 0.6485
| Std Dev. |  | 0.0021213 | 35
| Rel.Std.Dev. |  | 0.3271%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 211b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 30
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 211b (Indicative Value 0.37 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.36 | 5
| Mean |  | 0.365 | 16
| Median |  | 0.365 | 0.365
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 211b (Indicative Value 0.22 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.22 | 32
| Mean |  | 0.22 | 16
| Median |  | 0.22 | 0.22
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 211b (Indicative Value 299 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 299 | 1
| 1 | 2 | 299 | 33
| Mean |  | 299 | 16
| Median |  | 299 | 299
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 211b (Indicative Value 17.3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.5 | 1
| 1 | 2 | 17 | 34
| Mean |  | 17.25 | 16
| Median |  | 17.25 | 17.25
| Std Dev. |  | 0.35355 | 40
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 211b (Indicative Value 20.8 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.8 | 1
| 1 | 2 | 20.7 | 35
| Mean |  | 20.75 | 16
| Median |  | 20.75 | 20.75
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 0.3408%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 211b (Indicative Value 2.33 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.33 | 1
| 1 | 2 | 2.33 | 6
| Mean |  | 2.33 | 16
| Median |  | 2.33 | 2.33
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 211b (Indicative Value 125 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 125 | 1
| 1 | 2 | 125 | 15
| Mean |  | 125 | 16
| Median |  | 125 | 125
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 211b (Indicative Value 64 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 63.5 | 1
| 1 | 2 | 64 | 16
| Mean |  | 63.75 | 16
| Median |  | 63.75 | 63.75
| Std Dev. |  | 0.35355 | 44
| Rel.Std.Dev. |  | 0.5546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.