## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 266
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 7.9293 | 7.8217 | 8.0368 | 7.9013 | 7.9572
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 7.7831 | 7.6483 | 7.918 | 7.7531 | 7.8132
Cyanide Leach
Au, Gold (ppm) | 7.6364 | 7.4757 | 7.7972 | 7.626 | 7.6469
PhotonAssay (recommended gross mass* 326±15 g)
Au, Gold (ppm) | 7.962 | 7.8765 | 8.0475 | 7.9538 | 7.9702
4-Acid Digestion
Ag, Silver (ppm) | 4.471 | 4.2519 | 4.6901 | 4.3508 | 4.5912
Al, Aluminium (wt.%) | 6.4955 | 6.32 | 6.671 | 6.3829 | 6.608
As, Arsenic (ppm) | 186.7636 | 180.8051 | 192.7221 | 182.237 | 191.2902
Ba, Barium (ppm) | 459.0281 | 443.2223 | 474.8339 | 450.8914 | 467.1647
Be, Beryllium (ppm) | 2.2854 | 2.2015 | 2.3694 | 2.2268 | 2.344
Bi, Bismuth (ppm) | 3.107 | 2.8612 | 3.3527 | 2.978 | 3.2359
Ca, Calcium (wt.%) | 2.194 | 2.1395 | 2.2485 | 2.1463 | 2.2416
Cd, Cadmium (ppm) | 0.086449 | 0.05783 | 0.11507 | 0.08208 | 0.090819
Ce, Cerium (ppm) | 83.3672 | 78.6642 | 88.0703 | 81.1076 | 85.6269
Co, Cobalt (ppm) | 22.3984 | 21.3649 | 23.4319 | 21.8244 | 22.9724
Cr, Chromium (ppm) | 131.1252 | 124.4764 | 137.774 | 126.8898 | 135.3606
Cs, Caesium (ppm) | 4.3971 | 4.2454 | 4.5487 | 4.3039 | 4.4902
Cu, Copper (ppm) | 32.2112 | 30.1403 | 34.2821 | 30.8997 | 33.5227
Dy, Dysprosium (ppm) | 4.3673 | 4.0189 | 4.7157 | 4.1949 | 4.5397
Er, Erbium (ppm) | 2.0198 | 1.8098 | 2.2297 | 1.9331 | 2.1064
Eu, Europium (ppm) | 1.7111 | 1.5144 | 1.9079 | 1.6397 | 1.7825
Fe, Iron (wt.%) | 5.0488 | 4.9026 | 5.1951 | 4.9463 | 5.1514
Ga, Gallium (ppm) | 19.2499 | 18.1216 | 20.3783 | 18.7431 | 19.7568
Gd, Gadolinium (ppm) | 6.3214 | 5.7478 | 6.8949 | 6.1847 | 6.4581
Hf, Hafnium (ppm) | 4.9573 | 4.5931 | 5.3214 | 4.7683 | 5.1463
Ho, Holmium (ppm) | 0.76802 | 0.71399 | 0.82205 | 0.72501 | 0.81103
In, Indium (ppm) | 0.078564 | 0.070079 | 0.087048 | 0.074389 | 0.082738
K, Potassium (wt.%) | 1.508 | 1.4748 | 1.5412 | 1.4727 | 1.5433
La, Lanthanum (ppm) | 44.8686 | 43.3008 | 46.4364 | 43.7022 | 46.035
Li, Lithium (ppm) | 29.3825 | 28.2994 | 30.4656 | 28.6572 | 30.1079
Lu, Lutetium (ppm) | 0.2362 | 0.20301 | 0.26938 | 0.21298 | 0.25942
Mg, Magnesium (wt.%) | 1.7628 | 1.7155 | 1.8101 | 1.7303 | 1.7953
Mn, Manganese (wt.%) | 0.051686 | 0.050473 | 0.052899 | 0.050668 | 0.052704
Mo, Molybdenum (ppm) | 9.3148 | 8.8457 | 9.7839 | 9.0813 | 9.5484
Na, Sodium (wt.%) | 1.1464 | 1.1111 | 1.1818 | 1.1176 | 1.1753
Nb, Niobium (ppm) | 32.3713 | 30.2029 | 34.5397 | 30.8542 | 33.8884
Nd, Neodymium (ppm) | 37.3789 | 32.2727 | 42.4851 | 36.1156 | 38.6422
Ni, Nickel (ppm) | 95.4208 | 91.873 | 98.9687 | 93.822 | 97.0197
P, Phosphorus (wt.%) | 0.10349 | 0.10028 | 0.10669 | 0.10176 | 0.10521
Pb, Lead (ppm) | 15.1841 | 14.0011 | 16.367 | 14.5196 | 15.8485
Pr, Praseodymium (ppm) | 9.9176 | 9.2691 | 10.566 | 9.5794 | 10.2557
Rb, Rubidium (ppm) | 78.0318 | 73.5572 | 82.5065 | 76.1603 | 79.9034
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.031599 | 0.028123 | 0.035076 | 0.030386 | 0.032812
Sb, Antimony (ppm) | 10.3051 | 9.617 | 10.9931 | 9.8939 | 10.7163
Sc, Scandium (ppm) | 12.8293 | 12.1557 | 13.5028 | 12.3555 | 13.3031
Sm, Samarium (ppm) | 7.5228 | 6.8707 | 8.1749 | 7.3083 | 7.7374
Sn, Tin (ppm) | 6.0402 | 5.6632 | 6.4172 | 5.7365 | 6.3439
Sr, Strontium (ppm) | 334.5849 | 321.5058 | 347.6639 | 327.7139 | 341.4558
Ta, Tantalum (ppm) | 2.1211 | 1.9217 | 2.3205 | 2.0196 | 2.2226
Tb, Terbium (ppm) | 0.82788 | 0.7654 | 0.89036 | 0.77912 | 0.87663
Te, Tellurium (ppm) | 0.084937 | 0.067874 | 0.102 | IND | IND
Th, Thorium (ppm) | 10.9234 | 10.2408 | 11.6061 | 10.5174 | 11.3294
Ti, Titanium (wt.%) | 0.64996 | 0.62608 | 0.67383 | 0.63307 | 0.66684
Tl, Thallium (ppm) | 0.41794 | 0.3939 | 0.44197 | 0.40277 | 0.43311
Tm, Thulium (ppm) | 0.25595 | 0.23581 | 0.2761 | 0.24069 | 0.27122
U, Uranium (ppm) | 2.078 | 1.9447 | 2.2113 | 2.0124 | 2.1436
V, Vanadium (ppm) | 92.315 | 88.9766 | 95.6533 | 90.2702 | 94.3597
W, Tungsten (ppm) | 32.083 | 29.6504 | 34.5156 | 30.6121 | 33.5539
Y, Yttrium (ppm) | 19.5221 | 18.6469 | 20.3974 | 18.933 | 20.1113
Yb, Ytterbium (ppm) | 1.6734 | 1.5392 | 1.8075 | 1.5782 | 1.7685
Zn, Zinc (ppm) | 90.4702 | 86.4565 | 94.484 | 88.5469 | 92.3935
Zr, Zirconium (ppm) | 205.1445 | 196.2325 | 214.0564 | 200.3122 | 209.9767
Aqua Regia Digestion
Ag, Silver (ppm) | 4.3007 | 4.1352 | 4.4662 | 4.2013 | 4.4001
Al, Aluminium (wt.%) | 1.5465 | 1.418 | 1.675 | 1.5058 | 1.5873
As, Arsenic (ppm) | 168.0743 | 161.4414 | 174.7071 | 165.3066 | 170.8419
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 99.9199 | 93.9105 | 105.9294 | 96.526 | 103.3138
Be, Beryllium (ppm) | 1.0736 | 1.0065 | 1.1407 | 1.0337 | 1.1135
Bi, Bismuth (ppm) | 2.6475 | 2.3873 | 2.9077 | 2.536 | 2.759
Ca, Calcium (wt.%) | 0.63202 | 0.58391 | 0.68013 | 0.61801 | 0.64604
Cd, Cadmium (ppm) | 0.063413 | 0.04567 | 0.081156 | IND | IND
Ce, Cerium (ppm) | 47.7415 | 44.7727 | 50.7103 | 46.2627 | 49.2203
Co, Cobalt (ppm) | 16.1926 | 15.2542 | 17.1311 | 15.6809 | 16.7044
Cr, Chromium (ppm) | 49.0785 | 44.884 | 53.2731 | 47.4966 | 50.6605
Cs, Caesium (ppm) | 1.215 | 1.0852 | 1.3447 | 1.1636 | 1.2663
Cu, Copper (ppm) | 26.7025 | 24.8287 | 28.5763 | 25.9856 | 27.4194
Fe, Iron (wt.%) | 3.6153 | 3.4247 | 3.8059 | 3.5467 | 3.6839
Ga, Gallium (ppm) | 5.8248 | 5.3733 | 6.2763 | 5.5575 | 6.0921
Ge, Germanium (ppm) | 0.10867 | 0.067168 | 0.15017 | IND | IND
Hg, Mercury (ppm) | 0.039143 | 0.017334 | 0.060952 | IND | IND
In, Indium (ppm) | 0.030667 | 0.022284 | 0.03905 | 0.025218 | 0.036116
K, Potassium (wt.%) | 0.2747 | 0.25712 | 0.29227 | 0.26793 | 0.28146
La, Lanthanum (ppm) | 25.2001 | 23.8398 | 26.5604 | 24.334 | 26.0662
Li, Lithium (ppm) | 6.6225 | 6.0794 | 7.1655 | 6.323 | 6.9219
Mg, Magnesium (wt.%) | 0.96536 | 0.90582 | 1.0249 | 0.9452 | 0.98552
Mn, Manganese (wt.%) | 0.030743 | 0.028899 | 0.032586 | 0.030071 | 0.031414
Mo, Molybdenum (ppm) | 8.3065 | 7.7908 | 8.8222 | 8.055 | 8.558
Na, Sodium (wt.%) | 0.29105 | 0.27737 | 0.30472 | 0.28053 | 0.30156
Nb, Niobium (ppm) | 0.62243 | 0.46475 | 0.7801 | 0.55653 | 0.68832
Ni, Nickel (ppm) | 79.6198 | 75.4459 | 83.7937 | 77.8724 | 81.3673
P, Phosphorus (wt.%) | 0.063826 | 0.061019 | 0.066632 | 0.062692 | 0.06496
Pb, Lead (ppm) | 9.9482 | 9.178 | 10.7183 | 9.6046 | 10.2917
Rb, Rubidium (ppm) | 15.3084 | 14.4279 | 16.1889 | 14.6595 | 15.9572
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.029922 | 0.024175 | 0.035669 | 0.028528 | 0.031316
Sb, Antimony (ppm) | 7.0221 | 6.3365 | 7.7078 | 6.6981 | 7.3462
Sc, Scandium (ppm) | 3.8711 | 3.5798 | 4.1624 | 3.701 | 4.0412
Sn, Tin (ppm) | 1.558 | 1.3962 | 1.7197 | 1.4167 | 1.6992
Sr, Strontium (ppm) | 66.0081 | 62.6621 | 69.3541 | 63.3395 | 68.6767
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.50396 | 0.46503 | 0.54289 | IND | IND
Te, Tellurium (ppm) | 0.057689 | 0.035588 | 0.07979 | IND | IND
Th, Thorium (ppm) | 6.7478 | 6.2428 | 7.2528 | 6.5513 | 6.9443
Ti, Titanium (wt.%) | 0.14219 | 0.13001 | 0.15438 | 0.13746 | 0.14693
Tl, Thallium (ppm) | 0.11894 | 0.10755 | 0.13033 | IND | IND
U, Uranium (ppm) | 1.0492 | 0.98208 | 1.1164 | 1.0151 | 1.0834
V, Vanadium (ppm) | 34.5107 | 32.3807 | 36.6406 | 33.4946 | 35.5267
W, Tungsten (ppm) | 11.7105 | 9.8206 | 13.6005 | 11.0143 | 12.4068
Y, Yttrium (ppm) | 10.1425 | 9.7456 | 10.5395 | 9.8927 | 10.3924
Yb, Ytterbium (ppm) | 0.74358 | 0.66189 | 0.82526 | IND | IND
Zn, Zinc (ppm) | 58.2965 | 54.8709 | 61.7222 | 57.0896 | 59.5035
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 266
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.16682 | Hg | ppm | 0.17967 | Se | ppm | 0.96742
Aqua Regia Digestion
Dy | ppm | 2.4469 | Ho | ppm | 0.3995 | Pt | ppb | < 5
Er | ppm | 0.95058 | Lu | ppm | 0.10409 | Se | ppm | 0.55448
Eu | ppm | 1.0157 | Nd | ppm | 24.8433 | Sm | ppm | 4.4516
Gd | ppm | 3.7649 | Pd | ppb | < 10 | Tm | ppm | 0.11673
Hf | ppm | 0.67948 | Pr | ppm | 6.3606 | Zr | ppm | 32.3824
Borate Fusion XRF
Al2O3 | wt.% | 12.74 | MgO | wt.% | 3.06 | SiO2 | wt.% | 65.89
CaO | wt.% | 2.99 | MnO | wt.% | 0.0675 | SO3 | wt.% | 0.0805
Fe2O3 | wt.% | 7.45 | Na2O | wt.% | 1.565 | TiO2 | wt.% | 1.2
K2O | wt.% | 1.84 | P2O5 | wt.% | 0.2405
Thermogravimetry
LOI1000 | wt.% | 2.62
Infrared Combustion
C | wt.% | 0.11 | S | wt.% | 0.01
Laser Ablation ICP-MS
Ag | ppm | 4.45 | Hf | ppm | 7.14 | Sm | ppm | 7.57
As | ppm | 180.5 | Ho | ppm | 0.985 | Sn | ppm | 7.9
Ba | ppm | 459 | In | ppm | 0.05 | Sr | ppm | 326.5
Be | ppm | 2.8 | La | ppm | 45.1 | Ta | ppm | 2.29
Bi | ppm | 3.18 | Lu | ppm | 0.315 | Tb | ppm | 0.945
Cd | ppm | 0.125 | Mn | wt.% | 0.0554 | Te | ppm | < 0.2
Ce | ppm | 81.85 | Mo | ppm | 9.6 | Th | ppm | 11.2
Co | ppm | 23.05 | Nb | ppm | 34.2 | Ti | wt.% | 0.726
Cr | ppm | 150.5 | Nd | ppm | 38.9 | Tl | ppm | 0.3
Cs | ppm | 4.38 | Ni | ppm | 101 | Tm | ppm | 0.37
Cu | ppm | 35 | Pb | ppm | 16 | U | ppm | 2.28
Dy | ppm | 5.235 | Pr | ppm | 10.45 | V | ppm | 99.05
Er | ppm | 2.68 | Rb | ppm | 75.05 | W | ppm | 33.75
Eu | ppm | 1.835 | Re | ppm | < 0.01 | Y | ppm | 25.25
Ga | ppm | 18.8 | Sb | ppm | 11 | Yb | ppm | 2.255
Gd | ppm | 6.19 | Sc | ppm | 12.45 | Zn | ppm | 95
Ge | ppm | 1.275 | Se | ppm | < 5 | Zr | ppm | 269.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 266
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 7.9293 | 0.24268 | 7.4439 | 8.4146 | 7.2012 | 8.6573 | 3.061% | 6.121% | 9.182% | 7.5328 | 8.3257
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 7.7831 | 0.29788 | 7.1874 | 8.3789 | 6.8895 | 8.6768 | 3.827% | 7.655% | 11.48% | 7.394 | 8.1723
Cyanide Leach
Au, ppm | 7.6364 | 0.30214 | 7.0322 | 8.2407 | 6.73 | 8.5429 | 3.957% | 7.913% | 11.87% | 7.2546 | 8.0183
PhotonAssay
Au, ppm | 7.962 | 0.17994 | 7.6021 | 8.3219 | 7.4222 | 8.5018 | 2.26% | 4.52% | 6.78% | 7.5639 | 8.3601
4-Acid Digestion
Ag, ppm | 4.471 | 0.22311 | 4.0248 | 4.9172 | 3.8017 | 5.1403 | 4.99% | 9.98% | 14.97% | 4.2475 | 4.6946
Al, wt.% | 6.4955 | 0.17934 | 6.1368 | 6.8542 | 5.9575 | 7.0335 | 2.761% | 5.522% | 8.283% | 6.1707 | 6.8203
As, ppm | 186.7636 | 4.3 | 178.1636 | 195.3636 | 173.8635 | 199.6637 | 2.302% | 4.605% | 6.907% | 177.4254 | 196.1018
Ba, ppm | 459.0281 | 13.0219 | 432.9843 | 485.0719 | 419.9624 | 498.0938 | 2.837% | 5.674% | 8.511% | 436.0767 | 481.9795
Be, ppm | 2.2854 | 0.075148 | 2.1351 | 2.4357 | 2.06 | 2.5109 | 3.288% | 6.576% | 9.865% | 2.1711 | 2.3997
Bi, ppm | 3.107 | 0.21407 | 2.6788 | 3.5351 | 2.4647 | 3.7492 | 6.89% | 13.78% | 20.67% | 2.9516 | 3.2623
Ca, wt.% | 2.194 | 0.053124 | 2.0877 | 2.3002 | 2.0346 | 2.3533 | 2.421% | 4.843% | 7.264% | 2.0843 | 2.3037
Cd, ppm | 0.086449 | 0.019852 | 0.046746 | 0.12615 | 0.026894 | 0.14601 | 22.96% | 45.93% | 68.89% | 0.082127 | 0.090772
Ce, ppm | 83.3672 | 3.5018 | 76.3637 | 90.3707 | 72.862 | 93.8725 | 4.2% | 8.401% | 12.6% | 79.1989 | 87.5356
Co, ppm | 22.3984 | 1.4888 | 19.4208 | 25.3759 | 17.9321 | 26.8647 | 6.647% | 13.29% | 19.94% | 21.2785 | 23.5183
Cr, ppm | 131.1252 | 5.8842 | 119.3569 | 142.8935 | 113.4727 | 148.7777 | 4.487% | 8.975% | 13.46% | 124.5689 | 137.6815
Cs, ppm | 4.3971 | 0.17962 | 4.0378 | 4.7563 | 3.8582 | 4.9359 | 4.085% | 8.17% | 12.26% | 4.1772 | 4.6169
Cu, ppm | 32.2112 | 2.4827 | 27.2459 | 37.1765 | 24.7632 | 39.6592 | 7.707% | 15.41% | 23.12% | 30.6006 | 33.8217
Dy, ppm | 4.3673 | 0.20284 | 3.9616 | 4.773 | 3.7588 | 4.9758 | 4.645% | 9.289% | 13.93% | 4.1489 | 4.5857
Er, ppm | 2.0198 | 0.14463 | 1.7305 | 2.309 | 1.5859 | 2.4537 | 7.161% | 14.32% | 21.48% | 1.9188 | 2.1207
Eu, ppm | 1.7111 | 0.16326 | 1.3846 | 2.0376 | 1.2213 | 2.2009 | 9.541% | 19.08% | 28.62% | 1.6256 | 1.7967
Fe, wt.% | 5.0488 | 0.1388 | 4.7712 | 5.3264 | 4.6325 | 5.4652 | 2.749% | 5.498% | 8.247% | 4.7964 | 5.3013
Ga, ppm | 19.2499 | 1.5232 | 16.2036 | 22.2963 | 14.6804 | 23.8195 | 7.913% | 15.83% | 23.74% | 18.2874 | 20.2124
Gd, ppm | 6.3214 | 0.45845 | 5.4045 | 7.2383 | 4.946 | 7.6967 | 7.252% | 14.5% | 21.76% | 6.0053 | 6.6374
Hf, ppm | 4.9573 | 0.35543 | 4.2464 | 5.6681 | 3.891 | 6.0235 | 7.17% | 14.34% | 21.51% | 4.7094 | 5.2051
Ho, ppm | 0.76802 | 0.034242 | 0.69954 | 0.83651 | 0.66529 | 0.87075 | 4.459% | 8.917% | 13.38% | 0.72962 | 0.80642
In, ppm | 0.078564 | 0.0051279 | 0.068308 | 0.088819 | 0.06318 | 0.093947 | 6.527% | 13.05% | 19.58% | 0.074635 | 0.082492
K, wt.% | 1.508 | 0.026803 | 1.4544 | 1.5616 | 1.4276 | 1.5884 | 1.777% | 3.555% | 5.332% | 1.4326 | 1.5834
La, ppm | 44.8686 | 1.244 | 42.3806 | 47.3566 | 41.1366 | 48.6006 | 2.773% | 5.545% | 8.318% | 42.6252 | 47.112
Li, ppm | 29.3825 | 1.5044 | 26.3737 | 32.3913 | 24.8693 | 33.8957 | 5.12% | 10.24% | 15.36% | 27.9134 | 30.8516
Lu, ppm | 0.2362 | 0.030309 | 0.17558 | 0.29681 | 0.14527 | 0.32712 | 12.83% | 25.66% | 38.5% | 0.22439 | 0.24801
Mg, wt.% | 1.7628 | 0.068566 | 1.6257 | 1.8999 | 1.5571 | 1.9685 | 3.89% | 7.779% | 11.67% | 1.6747 | 1.8509
Mn, wt.% | 0.051686 | 0.0011387 | 0.049409 | 0.053963 | 0.04827 | 0.055102 | 2.203% | 4.406% | 6.609% | 0.049102 | 0.05427
Mo, ppm | 9.3148 | 0.54722 | 8.2204 | 10.4093 | 7.6732 | 10.9565 | 5.875% | 11.75% | 17.62% | 8.8491 | 9.7806
Na, wt.% | 1.1464 | 0.033445 | 1.0795 | 1.2133 | 1.0461 | 1.2468 | 2.917% | 5.835% | 8.752% | 1.0891 | 1.2038
Nb, ppm | 32.3713 | 2.7378 | 26.8957 | 37.8468 | 24.1579 | 40.5846 | 8.457% | 16.91% | 25.37% | 30.7527 | 33.9898
Nd, ppm | 37.3789 | 4.492 | 28.3948 | 46.363 | 23.9028 | 50.855 | 12.02% | 24.04% | 36.05% | 35.51 | 39.2479
Ni, ppm | 95.4208 | 3.5529 | 88.3151 | 102.5266 | 84.7622 | 106.0795 | 3.723% | 7.447% | 11.17% | 90.6498 | 100.1919
P, wt.% | 0.10349 | 0.0041096 | 0.095266 | 0.1117 | 0.091156 | 0.11581 | 3.971% | 7.942% | 11.91% | 0.098311 | 0.10866
Pb, ppm | 15.1841 | 1.1933 | 12.7975 | 17.5706 | 11.6042 | 18.7639 | 7.859% | 15.72% | 23.58% | 14.4248 | 15.9433
Pr, ppm | 9.9176 | 0.42597 | 9.0656 | 10.7695 | 8.6396 | 11.1955 | 4.295% | 8.59% | 12.89% | 9.4217 | 10.4134
Rb, ppm | 78.0318 | 5.6139 | 66.8041 | 89.2595 | 61.1903 | 94.8734 | 7.194% | 14.39% | 21.58% | 74.1302 | 81.9334
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.031599 | 0.0034545 | 0.02469 | 0.038508 | 0.021236 | 0.041963 | 10.93% | 21.86% | 32.8% | 0.030019 | 0.033179
Sb, ppm | 10.3051 | 0.6499 | 9.0053 | 11.6049 | 8.3554 | 12.2548 | 6.307% | 12.61% | 18.92% | 9.7898 | 10.8203
Sc, ppm | 12.8293 | 0.68291 | 11.4634 | 14.1951 | 10.7805 | 14.878 | 5.323% | 10.65% | 15.97% | 12.1878 | 13.4707
Sm, ppm | 7.5228 | 0.59227 | 6.3383 | 8.7073 | 5.746 | 9.2996 | 7.873% | 15.75% | 23.62% | 7.1467 | 7.899
Sn, ppm | 6.0402 | 0.36015 | 5.3199 | 6.7605 | 4.9597 | 7.1206 | 5.963% | 11.93% | 17.89% | 5.7382 | 6.3422
Sr, ppm | 334.5849 | 14.1711 | 306.2427 | 362.9271 | 292.0716 | 377.0982 | 4.235% | 8.471% | 12.71% | 317.8556 | 351.3141
Ta, ppm | 2.1211 | 0.25294 | 1.6152 | 2.627 | 1.3623 | 2.8799 | 11.92% | 23.85% | 35.77% | 2.0151 | 2.2272
Tb, ppm | 0.82788 | 0.042005 | 0.74387 | 0.91189 | 0.70186 | 0.95389 | 5.074% | 10.15% | 15.22% | 0.78648 | 0.86927
Te, ppm | 0.084937 | 0.012954 | 0.05903 | 0.11084 | 0.046076 | 0.1238 | 15.25% | 30.5% | 45.75% | 0.08069 | 0.089184
Th, ppm | 10.9234 | 0.75682 | 9.4098 | 12.4371 | 8.653 | 13.1939 | 6.928% | 13.86% | 20.79% | 10.3773 | 11.4696
Ti, wt.% | 0.64996 | 0.033759 | 0.58244 | 0.71747 | 0.54868 | 0.75123 | 5.194% | 10.39% | 15.58% | 0.61746 | 0.68245
Tl, ppm | 0.41794 | 0.024411 | 0.36912 | 0.46676 | 0.3447 | 0.49117 | 5.841% | 11.68% | 17.52% | 0.39704 | 0.43883
Tm, ppm | 0.25595 | 0.0079944 | 0.23997 | 0.27194 | 0.23197 | 0.27994 | 3.123% | 6.247% | 9.37% | 0.24316 | 0.26875
U, ppm | 2.078 | 0.13566 | 1.8067 | 2.3493 | 1.671 | 2.485 | 6.528% | 13.06% | 19.58% | 1.9741 | 2.1819
V, ppm | 92.315 | 3.8398 | 84.6354 | 99.9946 | 80.7956 | 103.8344 | 4.159% | 8.319% | 12.48% | 87.6992 | 96.9307
W, ppm | 32.083 | 3.0108 | 26.0615 | 38.1045 | 23.0508 | 41.1153 | 9.384% | 18.77% | 28.15% | 30.4789 | 33.6872
Y, ppm | 19.5221 | 0.9313 | 17.6595 | 21.3847 | 16.7282 | 22.316 | 4.77% | 9.541% | 14.31% | 18.546 | 20.4982
Yb, ppm | 1.6734 | 0.076618 | 1.5201 | 1.8266 | 1.4435 | 1.9032 | 4.579% | 9.157% | 13.74% | 1.5897 | 1.757
Zn, ppm | 90.4702 | 4.5842 | 81.3018 | 99.6387 | 76.7176 | 104.2229 | 5.067% | 10.13% | 15.2% | 85.9467 | 94.9938
Zr, ppm | 205.1445 | 8.168 | 188.8084 | 221.4805 | 180.6404 | 229.6486 | 3.982% | 7.963% | 11.94% | 194.8872 | 215.4017
Aqua Regia Digestion
Ag, ppm | 4.3007 | 0.10502 | 4.0907 | 4.5107 | 3.9856 | 4.6158 | 2.442% | 4.884% | 7.326% | 4.0857 | 4.5157
Al, wt.% | 1.5465 | 0.18639 | 1.1738 | 1.9193 | 0.98736 | 2.1057 | 12.05% | 24.1% | 36.16% | 1.4692 | 1.6239
As, ppm | 168.0743 | 6.9482 | 154.1778 | 181.9708 | 147.2295 | 188.919 | 4.134% | 8.268% | 12.4% | 159.6706 | 176.478
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 99.9199 | 5.368 | 89.1839 | 110.656 | 83.8159 | 116.024 | 5.372% | 10.74% | 16.12% | 94.924 | 104.9159
Be, ppm | 1.0736 | 0.050888 | 0.97184 | 1.1754 | 0.92095 | 1.2263 | 4.74% | 9.48% | 14.22% | 1.0199 | 1.1273
Bi, ppm | 2.6475 | 0.19302 | 2.2615 | 3.0335 | 2.0684 | 3.2265 | 7.291% | 14.58% | 21.87% | 2.5151 | 2.7799
Ca, wt.% | 0.63202 | 0.079949 | 0.47212 | 0.79192 | 0.39217 | 0.87187 | 12.65% | 25.3% | 37.95% | 0.60042 | 0.66362
Cd, ppm | 0.063413 | 0.0095651 | 0.044283 | 0.082543 | 0.034718 | 0.092108 | 15.08% | 30.17% | 45.25% | 0.060242 | 0.066584
Ce, ppm | 47.7415 | 3.1115 | 41.5185 | 53.9646 | 38.4069 | 57.0761 | 6.517% | 13.03% | 19.55% | 45.3544 | 50.1286
Co, ppm | 16.1926 | 1.2376 | 13.7175 | 18.6678 | 12.4799 | 19.9054 | 7.643% | 15.29% | 22.93% | 15.383 | 17.0023
Cr, ppm | 49.0785 | 7.302 | 34.4745 | 63.6826 | 27.1725 | 70.9846 | 14.88% | 29.76% | 44.63% | 46.6246 | 51.5325
Cs, ppm | 1.215 | 0.1844 | 0.84616 | 1.5837 | 0.66176 | 1.7681 | 15.18% | 30.35% | 45.53% | 1.1542 | 1.2757
Cu, ppm | 26.7025 | 2.9789 | 20.7448 | 32.6602 | 17.766 | 35.6391 | 11.16% | 22.31% | 33.47% | 25.3674 | 28.0376
Fe, wt.% | 3.6153 | 0.36144 | 2.8924 | 4.3382 | 2.531 | 4.6996 | 9.998% | 20% | 29.99% | 3.4345 | 3.7961
Ga, ppm | 5.8248 | 0.53826 | 4.7483 | 6.9013 | 4.21 | 7.4396 | 9.241% | 18.48% | 27.72% | 5.5336 | 6.116
Ge, ppm | 0.10867 | 0.020126 | 0.068415 | 0.14892 | 0.048289 | 0.16904 | 18.52% | 37.04% | 55.56% | 0.10323 | 0.1141
Hg, ppm | 0.039143 | 0.011559 | 0.016025 | 0.06226 | 0.0044668 | 0.073819 | 29.53% | 59.06% | 88.59% | 0.037186 | 0.0411
In, ppm | 0.030667 | 0.0048372 | 0.020992 | 0.040341 | 0.016155 | 0.045178 | 15.77% | 31.55% | 47.32% | 0.029133 | 0.0322
K, wt.% | 0.2747 | 0.029166 | 0.21636 | 0.33303 | 0.1872 | 0.36219 | 10.62% | 21.24% | 31.85% | 0.26096 | 0.28843
La, ppm | 25.2001 | 1.792 | 21.616 | 28.7841 | 19.824 | 30.5762 | 7.111% | 14.22% | 21.33% | 23.9401 | 26.4601
Li, ppm | 6.6225 | 0.62765 | 5.3672 | 7.8778 | 4.7395 | 8.5054 | 9.478% | 18.96% | 28.43% | 6.2913 | 6.9536
Mg, wt.% | 0.96536 | 0.097953 | 0.76945 | 1.1613 | 0.6715 | 1.2592 | 10.15% | 20.29% | 30.44% | 0.91709 | 1.0136
Mn, wt.% | 0.030743 | 0.0028609 | 0.025021 | 0.036464 | 0.02216 | 0.039325 | 9.306% | 18.61% | 27.92% | 0.029205 | 0.03228
Mo, ppm | 8.3065 | 0.64687 | 7.0128 | 9.6003 | 6.3659 | 10.2471 | 7.787% | 15.57% | 23.36% | 7.8912 | 8.7219
Na, wt.% | 0.29105 | 0.013333 | 0.26438 | 0.31771 | 0.25105 | 0.33104 | 4.581% | 9.162% | 13.74% | 0.27649 | 0.3056
Nb, ppm | 0.62243 | 0.15055 | 0.32132 | 0.92353 | 0.17077 | 1.0741 | 24.19% | 48.38% | 72.56% | 0.5913 | 0.65355
Ni, ppm | 79.6198 | 5.202 | 69.2159 | 90.0237 | 64.014 | 95.2257 | 6.533% | 13.07% | 19.6% | 75.6388 | 83.6008
P, wt.% | 0.063826 | 0.0038018 | 0.056222 | 0.071429 | 0.05242 | 0.075231 | 5.957% | 11.91% | 17.87% | 0.060634 | 0.067017
Pb, ppm | 9.9482 | 0.86135 | 8.2255 | 11.6709 | 7.3641 | 12.5322 | 8.658% | 17.32% | 25.98% | 9.4508 | 10.4456
Rb, ppm | 15.3084 | 0.86245 | 13.5835 | 17.0333 | 12.721 | 17.8957 | 5.634% | 11.27% | 16.9% | 14.5429 | 16.0738
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.029922 | 0.0029107 | 0.0241 | 0.035743 | 0.02119 | 0.038654 | 9.728% | 19.46% | 29.18% | 0.028426 | 0.031418
Sb, ppm | 7.0221 | 0.96578 | 5.0906 | 8.9537 | 4.1248 | 9.9195 | 13.75% | 27.51% | 41.26% | 6.671 | 7.3732
Sc, ppm | 3.8711 | 0.21897 | 3.4332 | 4.309 | 3.2142 | 4.528 | 5.656% | 11.31% | 16.97% | 3.6776 | 4.0647
Sn, ppm | 1.558 | 0.21402 | 1.1299 | 1.986 | 0.91589 | 2.2 | 13.74% | 27.47% | 41.21% | 1.4801 | 1.6359
Sr, ppm | 66.0081 | 2.4242 | 61.1597 | 70.8565 | 58.7355 | 73.2807 | 3.673% | 7.345% | 11.02% | 62.7077 | 69.3085
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.50396 | 0.032793 | 0.43838 | 0.56955 | 0.40558 | 0.60234 | 6.507% | 13.01% | 19.52% | 0.47876 | 0.52916
Te, ppm | 0.057689 | 0.011187 | 0.035316 | 0.080062 | 0.024129 | 0.091249 | 19.39% | 38.78% | 58.17% | 0.054804 | 0.060573
Th, ppm | 6.7478 | 0.42353 | 5.9008 | 7.5949 | 5.4772 | 8.0184 | 6.277% | 12.55% | 18.83% | 6.4104 | 7.0852
Ti, wt.% | 0.14219 | 0.015052 | 0.11209 | 0.1723 | 0.097039 | 0.18735 | 10.59% | 21.17% | 31.76% | 0.13508 | 0.1493
Tl, ppm | 0.11894 | 0.01228 | 0.094382 | 0.1435 | 0.082102 | 0.15578 | 10.32% | 20.65% | 30.97% | 0.11299 | 0.12489
U, ppm | 1.0492 | 0.046419 | 0.95641 | 1.1421 | 0.90999 | 1.1885 | 4.424% | 8.848% | 13.27% | 0.99678 | 1.1017
V, ppm | 34.5107 | 3.0061 | 28.4985 | 40.5228 | 25.4924 | 43.5289 | 8.711% | 17.42% | 26.13% | 32.7851 | 36.2362
W, ppm | 11.7105 | 2.2472 | 7.2161 | 16.2049 | 4.9689 | 18.4522 | 19.19% | 38.38% | 57.57% | 11.125 | 12.2961
Y, ppm | 10.1425 | 0.39282 | 9.3569 | 10.9282 | 8.9641 | 11.321 | 3.873% | 7.746% | 11.62% | 9.6354 | 10.6497
Yb, ppm | 0.74358 | 0.047569 | 0.64844 | 0.83871 | 0.60087 | 0.88628 | 6.397% | 12.79% | 19.19% | 0.7064 | 0.78076
Zn, ppm | 58.2965 | 5.6618 | 46.9729 | 69.6202 | 41.3111 | 75.282 | 9.712% | 19.42% | 29.14% | 55.3817 | 61.2114
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 266
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 266
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 266 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 18 | 9 | 245480 | 0.08406 | 8.0062 | 8.2136
2 | 1 | ANSLu | 1 | 4 | 2 | 245481 | 0.083616 | 8.3332 | 8.231
2 | 1 | ANSLu | 1 | 5 | 3 | 245482 | 0.085935 | 8.5349 | 8.2418
2 | 1 | ANSLu | 1 | 17 | 9 | 245483 | 0.086122 | 8.0992 | 8.2185
2 | 1 | ANSLu | 1 | 3 | 2 | 245484 | 0.087592 | 8.1846 | 8.2231
2 | 1 | ANSLu | 1 | 9 | 5 | 245485 | 0.087807 | 8.0989 | 8.2185
2 | 1 | ANSLu | 1 | 19 | 10 | 245486 | 0.085811 | 8.2449 | 8.2263
2 | 1 | ANSLu | 1 | 13 | 7 | 245487 | 0.083808 | 8.3988 | 8.2345
2 | 1 | ANSLu | 1 | 10 | 5 | 245488 | 0.087216 | 8.6026 | 8.2454
2 | 1 | ANSLu | 1 | 1 | 1 | 245489 | 0.086661 | 8.249 | 8.2265
2 | 1 | ANSLu | 1 | 6 | 3 | 245490 | 0.083481 | 8.1365 | 8.2205
2 | 1 | ANSLu | 1 | 20 | 10 | 245491 | 0.085527 | 8.0552 | 8.2162
2 | 1 | ANSLu | 1 | 15 | 8 | 245492 | 0.084511 | 8.0362 | 8.2152
2 | 1 | ANSLu | 1 | 14 | 7 | 245493 | 0.084815 | 8.3757 | 8.2333
2 | 1 | ANSLu | 1 | 11 | 6 | 245494 | 0.085384 | 8.0205 | 8.2143
2 | 1 | ANSLu | 1 | 2 | 1 | 245495 | 0.085561 | 8.3056 | 8.2295
2 | 1 | ANSLu | 1 | 7 | 4 | 245496 | 0.083518 | 8.1975 | 8.2238
2 | 1 | ANSLu | 1 | 16 | 8 | 245497 | 0.083245 | 8.2398 | 8.226
2 | 1 | ANSLu | 1 | 8 | 4 | 245498 | 0.087686 | 8.0352 | 8.2151
2 | 1 | ANSLu | 1 | 12 | 6 | 245499 | 0.087022 | 8.3504 | 8.2319
Mean | [6 empty] | 0.085469 | 8.2252 | 8.2252
Median | [7 empty] | 8.2187 | 8.2249
SD | [7 empty] | 0.1715 | 0.0091541
RSD | [7 empty] | 2.085% | 0.1113%
RSD Upscale Factor | [7 empty] | 0.053376
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 266 (Certified Value 7.93 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | 2
|  |  | 0.085g | 25g | 50g | 30g | 50g | 30g | 50g | 50g | 50g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 8.249 | 8.032 | 7.469 | 7.7454 | 7.91 | 8.13 | 7 | 8.04 | 8.128 | 8.07 | 7.66 | 8.168 | 7.77 | 8.387 | 8.098 | 8.05 | 8 | 8.027 | 7.79 | 7.95 | 7.97 | 7.44 | 1
| 1 | 2 | 8.3056 | 8.005 | 7.453 | 7.7928 | 8.3 | 8.09 | 7.245 | 8.1 | 7.868 | 7.77 | 7.5 | 8.139 | 7.73 | 8.203 | 8.081 | 8 | 7.8726 | 8.07 | 7.87 | 7.95 | 8 | 7.43 | #N/A
| 1 | 3 | 8.1846 | 8.046 | 7.433 | 7.75 | 8.7 | 7.96 | 7.14 | 8.13 | 8.002 | 7.71 | 7.64 | 8.25 | 7.75 | 8.153 | 8.086 | 8.05 | 8.1215 | 8.005 | 7.86 | 8.02 | 7.85 | 7.45 | 16
| 1 | 4 | 8.3332 | 7.959 | 7.442 | 7.8705 | 7.88 | 7.85 | 7.315 | 8.2 | 8.181 | 7.81 | 7.57 | 8.139 | 7.87 | 8.302 | 8.083 | 7.84 | 7.8814 | 7.921 | 7.86 | 7.83 | 7.91 | 7.45 | 7.9293
| 1 | 5 | 8.5349 | 8.317 | 7.427 | 7.8079 | 8.54 | 8.12 | 6.82 | 8.34 | 7.953 | 8.11 | 7.54 | 8.14 | 7.91 | 8.324 | 8.088 | 8.04 | 7.8576 | 8.023 | 7.86 | 7.8 | 7.92 | 7.53 | 7
| 1 | 6 | 8.1365 | 7.925 | 7.514 | 7.8038 | 8.08 | 7.92 | 7.155 | 8.42 | 8.202 | 8.02 | 7.51 | 8.26 | 7.95 | 8.17 | 8.084 | 7.99 | 8.1617 | 8.012 | 7.72 | 7.73 | 7.94 | 7.46
|  | 7 | 8.1975
|  | 8 | 8.0352
|  | 9 | 8.0989
|  | 10 | 8.6026
|  | 11 | 8.0205
|  | 12 | 8.3504
|  | 13 | 8.3988
|  | 14 | 8.3757
|  | 15 | 8.0362
|  | 16 | 8.2398
|  | 17 | 8.0992
|  | 18 | 8.0062
|  | 19 | 8.2449
|  | 20 | 8.0552
| Mean |  | 8.2252 | 8.0473 | 7.4563 | 7.7951 | 8.235 | 8.0117 | 7.1125 | 8.205 | 8.0557 | 7.915 | 7.57 | 8.1827 | 7.83 | 8.2565 | 8.0867 | 7.995 | 7.9825 | 8.0097 | 7.8267 | 7.88 | 7.9317 | 7.46
| Median |  | 8.2187 | 8.0185 | 7.4475 | 7.7983 | 8.19 | 8.025 | 7.1475 | 8.165 | 8.065 | 7.915 | 7.555 | 8.154 | 7.82 | 8.2525 | 8.085 | 8.02 | 7.9407 | 8.0175 | 7.86 | 7.89 | 7.93 | 7.45
| Std Dev. |  | 0.1715 | 0.13965 | 0.03196 | 0.045686 | 0.33739 | 0.11754 | 0.17851 | 0.14721 | 0.1349 | 0.17155 | 0.066933 | 0.057207 | 0.092087 | 0.094566 | 0.0060553 | 0.080187 | 0.1339 | 0.049005 | 0.059889 | 0.11027 | 0.051929 | 0.035777
| Rel.Std.Dev. |  | 2.085% | 1.735% | 0.4286% | 0.5861% | 4.097% | 1.467% | 2.51% | 1.794% | 1.675% | 2.167% | 0.8842% | 0.6991% | 1.176% | 1.145% | 0.07488% | 1.003% | 1.677% | 0.6118% | 0.7652% | 1.399% | 0.6547% | 0.4796%
| PDM3 |  | 3.733% | 1.489% | -5.964% | -1.693% | 3.856% | 1.039% | -10.3% | 3.477% | 1.594% | -0.18% | -4.531% | 3.196% | -1.252% | 4.127% | 1.985% | 0.8289% | 0.6709% | 1.014% | -1.294% | -0.6214% | 0.03021% | -5.918%
| Z-Score (Absolute) |  | NA | 0.29 | 2.33 | 0.83 | 1.12 | 0.13 | 3.85 | 0.98 | 0.32 | 0.3 | 1.82 | 0.89 | 0.67 | 1.21 | 0.46 | 0.06 | 0 | 0.12 | 0.69 | 0.45 | 0.22 | 2.31
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 266 (Certified Value 7.78 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | 2
|  |  | 0.085g | 25g | 30g | 50g | 30g | 50g | 30g | 30g | 25g | 15g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 3
| 1 | 1 | 8.249 | 7.86 | 7.5821 | 7.7 | 7.98 | 7.77 | 7.12 | 8.183 | 7.97 | 7.5602 | 7.655 | 7.38 | 7.9455 | 7.92 | 8.16 | 7.49 | 8.45 | 1
| 1 | 2 | 8.3056 | 8.01 | 7.5659 | 7.98 | 8 | 7.46 | 7.38 | 8.1983 | 7.93 | 7.6838 | 7.656 | 7.38 | 8.0444 | 7.73 | 7.53 | 7.69 | 8.13 | #N/A
| 1 | 3 | 8.1846 | 7.95 | 7.5456 | 8.32 | 8.08 | 7.82 | 7.11 | 8.142 | 7.56 | 7.4778 | 7.654 | 7.45 | 7.8308 | 7.44 | 7.52 | 7.79 | 8.16 | 16
| 1 | 4 | 8.3332 | 7.99 | 7.5547 | 7.78 | 7.93 | 8.04 | 8.04 | 8.162 | 7.72 | 7.1928 | 7.655 | 7.72 | 7.7037 | 7.68 | 7.58 | 7.74 | 8.34 | 7.7831
| 1 | 5 | 8.5349 | 7.93 | 7.5395 | 8.18 | 8.25 | 7.49 | 7.6 | 8.117 | 7.98 | 7.0658 | 7.66 | 7.53 | 7.9024 | 7.71 | 7.74 | 7.42 | 8.52 | 9
| 1 | 6 | 8.1365 | 7.82 | 7.6276 | 7.88 | 8.3 | 7.72 | 7.58 | 8.1796 | 7.97 | 7.7044 | 7.654 | 7.64 | 7.6887 | 8 | 7.51 | 7.48 | 8.02
|  | 7 | 8.1975
|  | 8 | 8.0352
|  | 9 | 8.0989
|  | 10 | 8.6026
|  | 11 | 8.0205
|  | 12 | 8.3504
|  | 13 | 8.3988
|  | 14 | 8.3757
|  | 15 | 8.0362
|  | 16 | 8.2398
|  | 17 | 8.0992
|  | 18 | 8.0062
|  | 19 | 8.2449
|  | 20 | 8.0552
| Mean |  | 8.2252 | 7.9267 | 7.5692 | 7.9733 | 8.09 | 7.7167 | 7.4717 | 8.1636 | 7.855 | 7.4475 | 7.6557 | 7.5167 | 7.8526 | 7.7467 | 7.6733 | 7.6017 | 8.27
| Median |  | 8.2187 | 7.94 | 7.5603 | 7.93 | 8.04 | 7.745 | 7.48 | 8.1708 | 7.95 | 7.519 | 7.655 | 7.49 | 7.8666 | 7.72 | 7.555 | 7.59 | 8.25
| Std Dev. |  | 0.1715 | 0.073937 | 0.032368 | 0.23822 | 0.15205 | 0.21695 | 0.35045 | 0.029927 | 0.1749 | 0.26307 | 0.0022509 | 0.14038 | 0.13957 | 0.19715 | 0.25327 | 0.15664 | 0.19698
| Rel.Std.Dev. |  | 2.085% | 0.9328% | 0.4276% | 2.988% | 1.88% | 2.811% | 4.69% | 0.3666% | 2.227% | 3.532% | 0.0294% | 1.868% | 1.777% | 2.545% | 3.301% | 2.061% | 2.382%
| PDM3 |  | 5.68% | 1.844% | -2.748% | 2.444% | 3.943% | -0.8541% | -4.002% | 4.889% | 0.9232% | -4.313% | -1.638% | -3.424% | 0.8924% | -0.4686% | -1.411% | -2.332% | 6.255%
| Z-Score (Absolute) |  | NA | 0.74 | 0.61 | 0.91 | 1.35 | 0.06 | 0.98 | 1.63 | 0.47 | 1.07 | 0.29 | 0.81 | 0.46 | 0.06 | 0.22 | 0.49 | 2.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 266 (Certified Value 7.64 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 05g | 50g | 50g | 200g | 50g | 30g | 30g | 30g | 30g | 20g | 05g | 3
| 1 | 1 | 8.249 | 7.83 | 7.3048 | 7.1888 | 7.28 | 7.5 | 7.6 | 7.74 | 7.2 | 7.55 | 7.924 | 6.377 | 7.72 | 8.0732 | 8.17 | 7.2686 | 7.886 | 1
| 1 | 2 | 8.3056 | 7.97 | 7.2892 | 7.2164 | 7.31 | 7.4 | 7.6 | 8 | 7.38 | 7.59 | 7.97 | 6.286 | 7.69 | 8.0731 | 7.66 | 7.2144 | 7.883 | #N/A
| 1 | 3 | 8.1846 | 7.99 | 7.2697 | 7.0626 | 7.76 | 7.37 | 7.5 | 8.04 | 7.38 | 7.51 | 7.914 | 6.487 | 7.7 | 8.074 | 8.13 | 7.451 | 7.769 | 16
| 1 | 4 | 8.3332 | 8 | 7.2785 | 6.9619 | 7.18 | 7.54 | 7.4 | 7.7902 | 7.49 | 7.62 | 7.9228 | 6.465 | 7.63 | 8.071 | 8.09 | 7.388 | 7.881 | 7.6364
| 1 | 5 | 8.5349 | 7.97 | 7.2638 | 7.3181 | 7.52 | 7.56 | 7.5 | 7.8 | 7.44 | 7.54 | 7.922 | 6.479 | 7.7 | 8.075 | 7.62 | 7.5958 | 7.839 | 11
| 1 | 6 | 8.1365 | 7.88 | 7.3487 | 7.0012 | 7.68 | 7.51 | 7.5 | 8.02 | 7.39 | 7.56 | 7.9158 | 6.414 | 7.67 | 8.075 | 8.18 | 7.4541 | 7.786
|  | 7 | 8.1975
|  | 8 | 8.0352
|  | 9 | 8.0989
|  | 10 | 8.6026
|  | 11 | 8.0205
|  | 12 | 8.3504
|  | 13 | 8.3988
|  | 14 | 8.3757
|  | 15 | 8.0362
|  | 16 | 8.2398
|  | 17 | 8.0992
|  | 18 | 8.0062
|  | 19 | 8.2449
|  | 20 | 8.0552
| Mean |  | 8.2252 | 7.94 | 7.2924 | 7.1248 | 7.455 | 7.48 | 7.5167 | 7.8984 | 7.38 | 7.5617 | 7.9281 | 6.418 | 7.685 | 8.0735 | 7.975 | 7.3953 | 7.8407
| Median |  | 8.2187 | 7.97 | 7.2838 | 7.1257 | 7.415 | 7.505 | 7.5 | 7.9 | 7.385 | 7.555 | 7.9224 | 6.4395 | 7.695 | 8.0736 | 8.11 | 7.4195 | 7.86
| Std Dev. |  | 0.1715 | 0.068702 | 0.031184 | 0.13823 | 0.23458 | 0.077201 | 0.075277 | 0.13538 | 0.098184 | 0.038687 | 0.020914 | 0.077217 | 0.031464 | 0.0014986 | 0.26174 | 0.13828 | 0.052156
| Rel.Std.Dev. |  | 2.085% | 0.8653% | 0.4276% | 1.94% | 3.147% | 1.032% | 1.001% | 1.714% | 1.33% | 0.5116% | 0.2638% | 1.203% | 0.4094% | 0.01856% | 3.282% | 1.87% | 0.6652%
| PDM3 |  | 7.71% | 3.975% | -4.505% | -6.7% | -2.376% | -2.049% | -1.568% | 3.43% | -3.358% | -0.9791% | 3.819% | -15.96% | 0.6359% | 5.724% | 4.433% | -3.158% | 2.674%
| Z-Score (Absolute) |  | NA | 0.99 | 0.61 | 1.02 | 0.21 | 0.15 | 0.06 | 0.88 | 0.39 | 0.06 | 0.96 | 2.76 | 0.36 | 1.31 | 1.07 | 0.35 | 0.74
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 266 (Certified Value 7.96 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 326±15 g | 3
| 1 | 1 | 8.249 | 8.014 | 8.089 | 8.27 | 7.84 | 8.11 | 7.95 | 8 | 7.892 | 7.912 | 7.943 | 7.823 | 1
| 1 | 2 | 8.3056 | 7.962 | 8.092 | 8.31 | 7.96 | 7.9 | 7.85 | 7.68 | 8.108 | 8.164 | 7.895 | 7.903 | #N/A
| 1 | 3 | 8.1846 | 8.273 | 7.927 | 8.03 | 7.86 | 8.26 | 7.91 | 7.8 | 7.65 | 7.842 | 7.655 | 7.984 | 16
| 1 | 4 | 8.3332 | 8.127 | 7.886 | 8.21 | 7.9 | 7.8 | 7.93 | 7.75 | 7.921 | 7.977 | 7.762 | 7.687 | 7.962
| 1 | 5 | 8.5349 | 8.053 | 8.152 | 8.27 | 8.01 | 8.31 | 7.73 | 8.28 | 7.962 | 7.734 | 7.871 | 8.054 | 13
| 1 | 6 | 8.1365 | 8.025 | 8.275 | 8.18 | 8.01 | 8.07 | 7.82 | 7.88 | 7.836 | 7.773 | 7.584 | 7.804
|  | 7 | 8.1975
|  | 8 | 8.0352
|  | 9 | 8.0989
|  | 10 | 8.6026
|  | 11 | 8.0205
|  | 12 | 8.3504
|  | 13 | 8.3988
|  | 14 | 8.3757
|  | 15 | 8.0362
|  | 16 | 8.2398
|  | 17 | 8.0992
|  | 18 | 8.0062
|  | 19 | 8.2449
|  | 20 | 8.0552
| Mean |  | 8.2252 | 8.0757 | 8.0702 | 8.2117 | 7.93 | 8.075 | 7.865 | 7.8983 | 7.8948 | 7.9003 | 7.785 | 7.8758
| Median |  | 8.2187 | 8.039 | 8.0905 | 8.24 | 7.93 | 8.09 | 7.88 | 7.84 | 7.9065 | 7.877 | 7.8165 | 7.863
| Std Dev. |  | 0.1715 | 0.1108 | 0.14415 | 0.10048 | 0.074297 | 0.19847 | 0.082401 | 0.21711 | 0.15103 | 0.15673 | 0.14303 | 0.13251
| Rel.Std.Dev. |  | 2.085% | 1.372% | 1.786% | 1.224% | 0.9369% | 2.458% | 1.048% | 2.749% | 1.913% | 1.984% | 1.837% | 1.682%
| PDM3 |  | 3.306% | 1.428% | 1.359% | 3.136% | -0.4017% | 1.419% | -1.218% | -0.7994% | -0.8434% | -0.7743% | -2.223% | -1.082%
| Z-Score (Absolute) |  | NA | 3.35 | 3.24 | 5.94 | 0.57 | 3.33 | 0.67 | 0.04 | 0.1 | 0 | 2.2 | 0.47
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 266 (Certified Value 4.47 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4.36 | 4.3776 | 3.5108 | 4.7 | 4.9 | 4.3 | 4.7754 | 4.31 | 4.7 | 4.2421 | 4.11 | 4.277 | 4.7 | 4.41 | 4.52 | 4.58 |  |  | 1
| 1 | 2 | 4.33 | 4.2868 | 3.5242 | 4.7 | 4.9 | 4.3 | 4.662 | 4.32 | 4.9 | 4.1473 | 4.14 | 4.2756 | 4.5 | 4.36 | 4.46 | 4.35 |  |  | 18
| 1 | 3 | 4.36 | 4.4571 | 3.5587 | 4.6 | 4.7 | 4.3 | 4.6561 | 4.4 | 4.6 | 4.2208 | 3.9 | 4.1502 | 4.7 | 4.28 | 4.53 | 4.34 |  |  | 16
| 1 | 4 | 4.4 | 4.3646 | 3.5643 | 4.7 | 5 | 4.4 | 4.6 | 4.54 | 4.8 | 4.2499 | 4.16 | 4.193 | 4.8 | 4.4 | 4.55 | 4.32 |  |  | 4.471
| 1 | 5 | 4.37 | 4.5701 | 3.5054 | 4.6 | 4.9 | 4.4 | 4.635 | 4.52 | 4.6 | 4.3339 | 4.31 | 4.13 | 4.7 | 4.37 | 4.6 | 4.22 |  |  | 15
| 1 | 6 | 4.38 | 4.4777 | 3.486 | 4.5 | 4.9 | 4.3 | 4.6681 | 4.34 | 4.8 | 4.2623 | 4.34 | 4.1885 | 4.7 | 4.42 | 4.69 | 4.6
| Mean |  | 4.3667 | 4.4223 | 3.5249 | 4.6333 | 4.8833 | 4.3333 | 4.6661 | 4.405 | 4.7333 | 4.2427 | 4.16 | 4.2024 | 4.6833 | 4.3733 | 4.5583 | 4.4017
| Median |  | 4.365 | 4.4174 | 3.5175 | 4.65 | 4.9 | 4.3 | 4.659 | 4.37 | 4.75 | 4.246 | 4.15 | 4.1907 | 4.7 | 4.385 | 4.54 | 4.345
| Std Dev. |  | 0.023381 | 0.099808 | 0.030944 | 0.08165 | 0.098319 | 0.05164 | 0.058999 | 0.10193 | 0.12111 | 0.060546 | 0.15837 | 0.061914 | 0.098319 | 0.051251 | 0.078846 | 0.15316
| Rel.Std.Dev. |  | 0.5354% | 2.257% | 0.8779% | 1.762% | 2.013% | 1.192% | 1.264% | 2.314% | 2.559% | 1.427% | 3.807% | 1.473% | 2.099% | 1.172% | 1.73% | 3.479%
| PDM3 |  | -2.334% | -1.089% | -21.16% | 3.631% | 9.222% | -3.079% | 4.363% | -1.476% | 5.867% | -5.106% | -6.956% | -6.008% | 4.749% | -2.185% | 1.953% | -1.551%
| Z-Score (Absolute) |  | 0.14 | 0.07 | 3.28 | 0.86 | 1.79 | 0.26 | 0.98 | 0.01 | 1.23 | 0.6 | 0.91 | 0.75 | 1.04 | 0.11 | 0.58 | 0.01
| Analytical results for Al in OREAS 266 (Certified Value 6.5 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 6.6046 | 6.5317 | 6.5574 | 5.9809 | 6.57 | 6.13 | 6.07 | 6.5 | 6.43 | 6.55 | 6.0216 | 6.49 | 6.5358 | 6.63 | 6.41 | 6.53 | 6.58 | 6.09 | 1
| 1 | 2 | 6.6689 | 6.535 | 6.479 | 5.8956 | 6.54 | 6.81 | 6.18 | 6.53 | 6.43 | 6.51 | 5.982 | 6.48 | 6.7512 | 6.5 | 6.28 | 6.44 | 6.48 | 6.48 | #N/A
| 1 | 3 | 6.6745 | 6.5035 | 6.6051 | 5.9399 | 6.58 | 6.88 | 6.11 | 6.55 | 6.44 | 6.73 | 6.0012 | 6.46 | 6.6192 | 6.55 | 6.22 | 6.33 | 6.45 | 6.33 | 16
| 1 | 4 | 6.7453 | 6.5398 | 6.6547 | 5.9465 | 6.58 | 6.63 | 6.17 | 6.41 | 6.45 | 6.34 | 6.0858 | 6.5 | 6.5699 | 6.77 | 6.5 | 6.38 | 6.3 | 6.14 | 6.4955
| 1 | 5 | 6.6359 | 6.4837 | 6.6601 | 5.8174 | 6.55 | 6.69 | 6.09 | 6.46 | 6.42 | 6.43 | 6.0277 | 6.49 | 6.7505 | 6.64 | 6.35 | 6.42 | 6.45 | 6.31 | 16
| 1 | 6 | 6.6956 | 6.569 | 6.6625 | 5.8999 | 6.7 | 6.83 | 6.07 | 6.55 | 6.37 | 6.46 | 6.0969 | 6.6 | 6.5849 | 6.64 | 6.36 | 6.44 | 6.58 | 5.97
| Mean |  | 6.6708 | 6.5271 | 6.6031 | 5.9134 | 6.5867 | 6.6617 | 6.115 | 6.5 | 6.4233 | 6.5033 | 6.0359 | 6.5033 | 6.6352 | 6.6217 | 6.3533 | 6.4233 | 6.4733 | 6.22
| Median |  | 6.6717 | 6.5333 | 6.6299 | 5.9199 | 6.575 | 6.75 | 6.1 | 6.515 | 6.43 | 6.485 | 6.0246 | 6.49 | 6.602 | 6.635 | 6.355 | 6.43 | 6.465 | 6.225
| Std Dev. |  | 0.048545 | 0.029786 | 0.073382 | 0.056706 | 0.05785 | 0.27658 | 0.048888 | 0.055857 | 0.028048 | 0.13231 | 0.046026 | 0.049261 | 0.093466 | 0.092394 | 0.097912 | 0.067132 | 0.10386 | 0.18633
| Rel.Std.Dev. |  | 0.7277% | 0.4563% | 1.111% | 0.959% | 0.8783% | 4.152% | 0.7995% | 0.8593% | 0.4367% | 2.035% | 0.7625% | 0.7575% | 1.409% | 1.395% | 1.541% | 1.045% | 1.604% | 2.996%
| PDM3 |  | 2.699% | 0.4873% | 1.657% | -8.962% | 1.404% | 2.559% | -5.858% | 0.06963% | -1.111% | 0.121% | -7.076% | 0.121% | 2.152% | 1.943% | -2.188% | -1.111% | -0.3409% | -4.241%
| Z-Score (Absolute) |  | 1.03 | 0.16 | 0.62 | 3.58 | 0.52 | 0.97 | 2.35 | 0.01 | 0.48 | 0.01 | 2.84 | 0.01 | 0.81 | 0.73 | 0.9 | 0.48 | 0.17 | 1.72
| Analytical results for As in OREAS 266 (Certified Value 187 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 188.2 | 191.0019 | 265.1377 | 186 | 165 | 188 | 184.4233 | 188 | 185 | 127.3819 | 183 | 181.2111 | 185 | 180 | 193 | 188.5 | 185 |  | 1
| 1 | 2 | 191.5 | 187.4749 | 288.3359 | 189 | 151 | 189 | 182.8767 | 190 | 189 | 128.0493 | 187.5 | 191.5543 | 178 | 176 | 188 | 186 | 183 |  | 19
| 1 | 3 | 191 | 197.6538 | 270.8519 | 187 | 149 | 190 | 186.66 | 183.5 | 183 | 127.7915 | 185.5 | 187.8774 | 182 | 175 | 189.5 | 183.5 | 183 |  | 16
| 1 | 4 | 195.4 | 192.8695 | 321.9415 | 193 | 155 | 188 | 182.435 | 186.5 | 183 | 126.4053 | 188 | 189.4006 | 188 | 180.5 | 190.5 | 183.5 | 184 |  | 186.7636
| 1 | 5 | 186.6 | 195.5788 | 259.7383 | 188 | 155 | 187 | 185.4 | 187 | 187 | 126.4574 | 189 | 187.5527 | 184 | 179.5 | 188.5 | 187.5 | 185 |  | 17
| 1 | 6 | 200.6 | 191.1764 | 271.8493 | 194 | 157 | 191 | 187.3167 | 183.5 | 183 | 127.6386 | 191 | 182.4392 | 182 | 180 | 189.5 | 192.5 | 187
| Mean |  | 192.2167 | 192.6259 | 279.6424 | 189.5 | 155.3333 | 188.8333 | 184.8519 | 186.4167 | 185 | 127.2873 | 187.3333 | 186.6726 | 183.1667 | 178.5 | 189.8333 | 186.9167 | 184.5
| Median |  | 191.25 | 192.0229 | 271.3506 | 188.5 | 155 | 188.5 | 184.9117 | 186.75 | 184 | 127.5103 | 187.75 | 187.715 | 183 | 179.75 | 189.5 | 186.75 | 184.5
| Std Dev. |  | 5.1008 | 3.6147 | 22.8445 | 3.2711 | 5.5737 | 1.472 | 1.9785 | 2.5577 | 2.5298 | 0.69769 | 2.7869 | 4.0311 | 3.3714 | 2.3664 | 1.7795 | 3.412 | 1.5166
| Rel.Std.Dev. |  | 2.654% | 1.877% | 8.169% | 1.726% | 3.588% | 0.7795% | 1.07% | 1.372% | 1.367% | 0.5481% | 1.488% | 2.159% | 1.841% | 1.326% | 0.9374% | 1.825% | 0.822%
| PDM3 |  | 2.92% | 3.139% | 49.73% | 1.465% | -16.83% | 1.108% | -1.024% | -0.1858% | -0.9443% | -31.85% | 0.3051% | -0.04874% | -1.926% | -4.425% | 1.644% | 0.08196% | -1.212%
| Z-Score (Absolute) |  | 1.32 | 1.42 | 22.17 | 0.67 | 7.47 | 0.52 | 0.43 | 0.06 | 0.4 | 14.16 | 0.16 | 0 | 0.84 | 1.95 | 0.75 | 0.06 | 0.52
| Analytical results for Ba in OREAS 266 (Certified Value 459 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 468 | 463.529 | 462.7809 | 420.3724 | 458 | 414 | 444 | 459 | 460 | 296 | 436.352 | 460 | 465.6326 | 467 | 440 | 480 | 470 |  | 1
| 1 | 2 | 478.8 | 466.785 | 447.6996 | 412.592 | 454 | 417 | 446 | 459 | 460 | 295 | 438.2022 | 460 | 485.7279 | 451 | 440 | 480 | 460 |  | 20
| 1 | 3 | 469.8 | 461.868 | 470.8537 | 417.4518 | 444 | 417 | 448 | 455 | 460 | 290 | 436.5071 | 460 | 473.8853 | 462 | 430 | 460 | 450 |  | 16
| 1 | 4 | 486.9 | 458.403 | 458.9693 | 415.3139 | 458 | 424 | 454 | 452 | 460 | 288 | 437.0634 | 460 | 479.2947 | 470 | 450 | 470 | 440 |  | 459.0281
| 1 | 5 | 466 | 460.614 | 470.936 | 411.2044 | 451 | 418 | 444 | 453 | 450 | 295 | 436.6091 | 460 | 474.1949 | 462 | 440 | 470 | 460 |  | 18
| 1 | 6 | 496 | 471.229 | 469.4285 | 414.9863 | 455 | 428 | 448 | 455 | 460 | 294 | 436.8393 | 470 | 462.4528 | 458 | 460 | 480 | 460
| Mean |  | 477.5833 | 463.738 | 463.4447 | 415.3201 | 453.3333 | 419.6667 | 447.3333 | 455.5 | 458.3333 | 293 | 436.9289 | 461.6667 | 473.5314 | 461.6667 | 443.3333 | 473.3333 | 456.6667
| Median |  | 474.3 | 462.6985 | 466.1047 | 415.1501 | 454.5 | 417.5 | 447 | 455 | 460 | 294.5 | 436.7242 | 460 | 474.0401 | 462 | 440 | 475 | 460
| Std Dev. |  | 11.9503 | 4.6295 | 9.1126 | 3.3034 | 5.2789 | 5.2409 | 3.7238 | 2.9496 | 4.0825 | 3.2249 | 0.67237 | 4.0825 | 8.5751 | 6.7132 | 10.328 | 8.165 | 10.328
| Rel.Std.Dev. |  | 2.502% | 0.9983% | 1.966% | 0.7954% | 1.164% | 1.249% | 0.8324% | 0.6475% | 0.8907% | 1.101% | 0.1539% | 0.8843% | 1.811% | 1.454% | 2.33% | 1.725% | 2.262%
| PDM3 |  | 4.042% | 1.026% | 0.9622% | -9.522% | -1.241% | -8.575% | -2.548% | -0.7686% | -0.1513% | -36.17% | -4.814% | 0.5748% | 3.16% | 0.5748% | -3.419% | 3.116% | -0.5144%
| Z-Score (Absolute) |  | 1.51 | 0.51 | 0.49 | 2.99 | 0.24 | 2.67 | 0.67 | 0.08 | 0.12 | 11.82 | 1.43 | 0.36 | 1.22 | 0.36 | 0.96 | 1.2 | 0
| Analytical results for Be in OREAS 266 (Certified Value 2.29 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.09 | 2.2451 | 1.6589 | 2.3 | 2.3 | 2.31 | < 5 | 2.32 | 2 | 1.8859 | 2.35 | 2.5192 | 2.3 | 2.15 | 2.36 | 2.29 | 2.69 |  | 1
| 1 | 2 | 2.13 | 2.2906 | 1.6167 | 2.3 | 2.3 | 2.27 | < 5 | 2.32 | 2 | 1.8821 | 2.33 | 2.4536 | 2.3 | 2.14 | 2.34 | 2.27 | 2.64 |  | 21
| 1 | 3 | 2.23 | 2.3197 | 1.692 | 2.2 | 2.6 | 2.31 | < 5 | 2.31 | 2 | 1.8955 | 2.31 | 2.4048 | 2.2 | 2.12 | 2.3 | 2.23 | 2.69 |  | 16
| 1 | 4 | 2.27 | 2.3527 | 1.6051 | 2.2 | 2.3 | 2.32 | < 5 | 2.3 | 2 | 1.8957 | 2.35 | 2.5915 | 2.3 | 2.21 | 2.32 | 2.16 | 2.61 |  | 2.2854
| 1 | 5 | 2.16 | 2.4065 | 1.6828 | 2.3 | 2.4 | 2.27 | < 5 | 2.29 | 2 | 1.9378 | 2.33 | 2.5275 | 2.3 | 2.15 | 2.33 | 2.26 | 2.62 |  | 19
| 1 | 6 | 2.34 | 2.4006 | 1.5739 | 2.3 | 2.5 | 2.3 | < 5 | 2.29 | 2 | 1.878 | 2.37 | 2.6082 | 2.3 | 2.16 | 2.34 | 2.26 | 2.57
| Mean |  | 2.2033 | 2.3359 | 1.6382 | 2.2667 | 2.4 | 2.2967 |  | 2.305 | 2 | 1.8958 | 2.34 | 2.5175 | 2.2833 | 2.155 | 2.3317 | 2.245 | 2.6367
| Median |  | 2.195 | 2.3362 | 1.6378 | 2.3 | 2.35 | 2.305 |  | 2.305 | 2 | 1.8907 | 2.34 | 2.5233 | 2.3 | 2.15 | 2.335 | 2.26 | 2.63
| Std Dev. |  | 0.093737 | 0.063263 | 0.046914 | 0.05164 | 0.12649 | 0.021602 |  | 0.013784 | 0 | 0.021743 | 0.020976 | 0.078219 | 0.040825 | 0.030166 | 0.020412 | 0.045935 | 0.047188
| Rel.Std.Dev. |  | 4.254% | 2.708% | 2.864% | 2.278% | 5.27% | 0.9406% |  | 0.598% | 0% | 1.147% | 0.8964% | 3.107% | 1.788% | 1.4% | 0.8754% | 2.046% | 1.79%
| PDM3 |  | -3.591% | 2.208% | -28.32% | -0.8202% | 5.014% | 0.4924% |  | 0.8571% | -12.49% | -17.05% | 2.389% | 10.15% | -0.09097% | -5.706% | 2.024% | -1.768% | 15.37%
| Z-Score (Absolute) |  | 0.86 | 0.31 | 5.83 | 0.3 | 0.87 | 0.04 | 1.75 | 0.04 | NA | 3.56 | 0.34 | 1.91 | 0.15 | 1.28 | 0.27 | 0.49 | 2.95
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 266 (Certified Value 3.11 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.23 | 2.6681 | 3.2403 | 1.8764 | 3.22 | 3.37 | 3.02 | < 5 | 3.34 | 3 | 3.2502 | 2.75 | 3.0697 | 3.17 | 2.85 | 2.78 | 3 | 3.17 | 1
| 1 | 2 | 3.08 | 2.6308 | 3.1249 | 2.4188 | 3.3 | 3.26 | 2.95 | < 5 | 3.32 | 3.5 | 3.2893 | 3.02 | 3.061 | 3.15 | 2.93 | 2.83 | 2.97 | 3.2 | 22
| 1 | 3 | 3.25 | 2.707 | 3.5256 | 2.1042 | 3.2 | 3.35 | 3.03 | < 5 | 2.89 | 3.2 | 3.2932 | 3.01 | 3.0861 | 3.51 | 2.9 | 2.93 | 3.09 | 3.26 | 16
| 1 | 4 | 3.25 | 2.5372 | 3.509 | 2.9149 | 3.27 | 3.55 | 3 | < 5 | 2.95 | 3.2 | 3.2815 | 2.9 | 3.0901 | 3.24 | 3.12 | 2.89 | 3.23 | 3.12 | 3.107
| 1 | 5 | 3.08 | 2.6417 | 3.1931 | 3.127 | 3.22 | 3.45 | 2.98 | < 5 | 2.9 | 3.9 | 3.4487 | 3.08 | 3.1489 | 3.13 | 2.9 | 2.99 | 2.86 | 3.19 | 20
| 1 | 6 | 3.29 | 2.7909 | 3.1709 | 1.8819 | 3.26 | 3.55 | 3 | < 5 | 3.21 | 3.3 | 3.2687 | 2.95 | 3.161 | 3.15 | 2.84 | 2.96 | 3.41 | 3.14
| Mean |  | 3.1967 | 2.6626 | 3.294 | 2.3872 | 3.245 | 3.4217 | 2.9967 |  | 3.1017 | 3.35 | 3.3053 | 2.9517 | 3.1028 | 3.225 | 2.9233 | 2.8967 | 3.0933 | 3.18
| Median |  | 3.24 | 2.6549 | 3.2167 | 2.2615 | 3.24 | 3.41 | 3 |  | 3.08 | 3.25 | 3.2854 | 2.98 | 3.0881 | 3.16 | 2.9 | 2.91 | 3.045 | 3.18
| Std Dev. |  | 0.092448 | 0.084414 | 0.17704 | 0.5335 | 0.037815 | 0.11635 | 0.028752 |  | 0.21198 | 0.31464 | 0.071983 | 0.11652 | 0.041958 | 0.14474 | 0.10211 | 0.079917 | 0.19866 | 0.049396
| Rel.Std.Dev. |  | 2.892% | 3.17% | 5.375% | 22.35% | 1.165% | 3.4% | 0.9595% |  | 6.834% | 9.392% | 2.178% | 3.948% | 1.352% | 4.488% | 3.493% | 2.759% | 6.422% | 1.553%
| PDM3 |  | 2.887% | -14.3% | 6.019% | -23.17% | 4.443% | 10.13% | -3.55% |  | -0.1702% | 7.823% | 6.383% | -4.998% | -0.1337% | 3.799% | -5.91% | -6.768% | -0.4385% | 2.351%
| Z-Score (Absolute) |  | 0.39 | 1.8 | 0.78 | 2.93 | 0.58 | 1.31 | 0.43 | 2.47 | 0 | 1.01 | 0.83 | 0.62 | 0 | 0.5 | 0.73 | 0.84 | 0.04 | 0.32
| Analytical results for Ca in OREAS 266 (Certified Value 2.19 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.1752 | 2.1804 | 2.1058 | 2.8925 | 2.16 | 2.13 | 2.16 | 2.2 | 2.16 | 2.2 | 2.2472 | 2.28 | 2.225 | 2.18 | 2.15 | 2.28 | 2.3 | 2.206 | 1
| 1 | 2 | 2.2017 | 2.185 | 2.0682 | 2.8362 | 2.16 | 2.15 | 2.15 | 2.2 | 2.17 | 2.24 | 2.24 | 2.27 | 2.2082 | 2.14 | 2.13 | 2.24 | 2.27 | 2.343 | #N/A
| 1 | 3 | 2.215 | 2.1871 | 2.1771 | 2.8023 | 2.14 | 2.2 | 2.19 | 2.12 | 2.16 | 2.21 | 2.2501 | 2.25 | 2.1683 | 2.14 | 2.09 | 2.22 | 2.17 | 2.266 | 16
| 1 | 4 | 2.2202 | 2.1525 | 2.1349 | 2.8686 | 2.15 | 2.14 | 2.2 | 2.15 | 2.15 | 2.16 | 2.272 | 2.26 | 2.1887 | 2.2 | 2.2 | 2.22 | 2.11 | 2.255 | 2.194
| 1 | 5 | 2.1945 | 2.1866 | 2.1943 | 2.851 | 2.15 | 2.14 | 2.14 | 2.17 | 2.15 | 2.23 | 2.3108 | 2.25 | 2.2448 | 2.15 | 2.14 | 2.24 | 2.27 | 2.28 | 21
| 1 | 6 | 2.2069 | 2.1246 | 2.1433 | 2.8688 | 2.19 | 2.25 | 2.19 | 2.21 | 2.14 | 2.22 | 2.2766 | 2.31 | 2.2161 | 2.19 | 2.14 | 2.25 | 2.18 | 2.182
| Mean |  | 2.2022 | 2.1694 | 2.1373 | 2.8533 | 2.1583 | 2.1683 | 2.1717 | 2.175 | 2.155 | 2.21 | 2.2661 | 2.27 | 2.2085 | 2.1667 | 2.1417 | 2.2417 | 2.2167 | 2.2553
| Median |  | 2.2043 | 2.1827 | 2.1391 | 2.8598 | 2.155 | 2.145 | 2.175 | 2.185 | 2.155 | 2.215 | 2.261 | 2.265 | 2.2122 | 2.165 | 2.14 | 2.24 | 2.225 | 2.2605
| Std Dev. |  | 0.016115 | 0.025555 | 0.046134 | 0.03136 | 0.017224 | 0.047081 | 0.024833 | 0.035071 | 0.010488 | 0.028284 | 0.026249 | 0.022804 | 0.027016 | 0.026583 | 0.035449 | 0.022286 | 0.074207 | 0.056955
| Rel.Std.Dev. |  | 0.7317% | 1.178% | 2.159% | 1.099% | 0.798% | 2.171% | 1.143% | 1.612% | 0.4867% | 1.28% | 1.158% | 1.005% | 1.223% | 1.227% | 1.655% | 0.9942% | 3.348% | 2.525%
| PDM3 |  | 0.3773% | -1.121% | -2.585% | 30.05% | -1.624% | -1.169% | -1.017% | -0.8647% | -1.776% | 0.7306% | 3.287% | 3.465% | 0.6632% | -1.245% | -2.384% | 2.174% | 1.034% | 2.797%
| Z-Score (Absolute) |  | 0.31 | 0.44 | 1.19 | 15.37 | 0.7 | 0.47 | 0.39 | 0.31 | 0.78 | 0.49 | 1.79 | 1.88 | 0.46 | 0.51 | 1.09 | 1.23 | 0.65 | 1.54
| Analytical results for Cd in OREAS 266 (Certified Value 0.086 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.1 | < 0.5 | 0.5225 | 0.1 | < 0.3 | 0.29 | 1.374 | 0.06 | 0.2 | 0.1133 | 0.07 | 0.079633 | 0.34 | 0.09 | 0.06 | 0.1 |  |  | 1
| 1 | 2 | 0.1 | < 0.5 | 0.6496 | 0.09 | < 0.3 | 0.3 | 1.3701 | 0.11 | 0.3 | 0.11238 | 0.07 | 0.11317 | 0.33 | 0.07 | 0.07 | 0.09 |  |  | 23
| 1 | 3 | 0.09 | < 0.5 | 0.6632 | 0.12 | < 0.3 | 0.3 | 1.369 | 0.07 | 0.3 | 0.11085 | 0.07 | 0.10895 | 0.33 | 0.07 | 0.09 | 0.08 |  |  | 16
| 1 | 4 | 0.08 | < 0.5 | 0.6999 | 0.11 | < 0.3 | 0.31 | 1.3116 | 0.06 | 0.3 | 0.10885 | 0.05 | 0.090839 | 0.33 | 0.07 | 0.07 | 0.11 |  |  | 0.086449
| 1 | 5 | 0.1 | < 0.5 | 0.6591 | 0.1 | < 0.3 | 0.3 | 1.3828 | 0.05 | 0.2 | 0.11028 | 0.07 | 0.086341 | 0.33 | 0.08 | 0.08 | 0.1 |  |  | 22
| 1 | 6 | 0.1 | < 0.5 | 0.5613 | 0.1 | < 0.3 | 0.3 | 1.3384 | 0.1 | 0.3 | 0.10975 | 0.07 | 0.093928 | 0.32 | 0.05 | 0.06 | 0.05
| Mean |  | 0.095 |  | 0.62593 | 0.10333 |  | 0.3 | 1.3576 | 0.075 | 0.26667 | 0.1109 | 0.066667 | 0.095477 | 0.33 | 0.071667 | 0.071667 | 0.088333
| Median |  | 0.1 |  | 0.65435 | 0.1 |  | 0.3 | 1.3696 | 0.065 | 0.3 | 0.11056 | 0.07 | 0.092383 | 0.33 | 0.07 | 0.07 | 0.095
| Std Dev. |  | 0.0083666 |  | 0.068405 | 0.010328 |  | 0.0063246 | 0.02711 | 0.02429 | 0.05164 | 0.0016651 | 0.008165 | 0.013061 | 0.0063246 | 0.013292 | 0.01169 | 0.02137
| Rel.Std.Dev. |  | 8.807% |  | 10.93% | 9.995% |  | 2.108% | 1.997% | 32.39% | 19.36% | 1.501% | 12.25% | 13.68% | 1.917% | 18.55% | 16.31% | 24.19%
| PDM3 |  | 9.891% |  | 624% | 19.53% |  | 247% | 1470% | -13.24% | 208.5% | 28.28% | -22.88% | 10.44% | 281.7% | -17.1% | -17.1% | 2.179%
| Z-Score (Absolute) |  | 0.18 | 3.12 | 11.13 | 0 | 0.99 | 4.19 | 26.71 | 0.6 | NA | 0.16 | 0.78 | 0.17 | 4.83 | 0.67 | 0.67 | 0.32
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 266 (Certified Value 83 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
1
| 1 | 1 | 81.15 | 81.513 | 85.3885 | 69.4 | 69 | 84.1 | 84 | 65.4884 | 77.3 | 87.2 | 82.3 | 87.2 | 85.3 | 106 |  |  |  |  | 1
| 1 | 2 | 80.8 | 81.5583 | 90.4922 | 77 | 68 | 84.6 | 94 | 65.8657 | 80.6 | 83 | 78.9 | 85.9 | 81.6 | 106 |  |  |  |  | 24
| 1 | 3 | 80.91 | 81.3821 | 84.369 | 78 | 71 | 79.2 | 100 | 66.0132 | 76.3 | 84.9 | 81.8 | 85.2 | 81.7 | 107 |  |  |  |  | 16
| 1 | 4 | 83.13 | 81.1433 | 88.2867 | 77.9 | 70 | 82.7 | 87 | 65.9327 | 82.6 | 87.1 | 83.5 | 86.2 | 83 | 104 |  |  |  |  | 83.3672
| 1 | 5 | 83.7 | 81.8837 | 82.6629 | 78 | 70 | 83.4 | 81 | 65.9925 | 85.6 | 86.7 | 80.9 | 83.9 | 79.2 | 113 |  |  |  |  | 23
| 1 | 6 | 83.65 | 81.7944 | 84.6631 | 79.7 | 70 | 83.9 | 93 | 66.3146 | 85.7 | 85.9 | 82.4 | 86.5 | 87.4 | 110
| Mean |  | 82.2233 | 81.5458 | 85.9771 | 76.6667 | 69.6667 | 82.9833 | 89.8333 | 65.9345 | 81.35 | 85.8 | 81.6333 | 85.8167 | 83.0333 | 107.6667
| Median |  | 82.14 | 81.5356 | 85.0258 | 77.95 | 70 | 83.65 | 90 | 65.9626 | 81.6 | 86.3 | 82.05 | 86.05 | 82.35 | 106.5
| Std Dev. |  | 1.41 | 0.27057 | 2.8762 | 3.6659 | 1.0328 | 1.9631 | 7.0828 | 0.2675 | 4.0253 | 1.6199 | 1.5845 | 1.1479 | 2.9235 | 3.266
| Rel.Std.Dev. |  | 1.715% | 0.3318% | 3.345% | 4.782% | 1.482% | 2.366% | 7.884% | 0.4057% | 4.948% | 1.888% | 1.941% | 1.338% | 3.521% | 3.033%
| PDM3 |  | -1.372% | -2.185% | 3.131% | -8.037% | -16.43% | -0.4605% | 7.756% | -20.91% | -2.42% | 2.918% | -2.08% | 2.938% | -0.4005% | 29.15%
| Z-Score (Absolute) |  | 0.08 | 0.22 | 0.71 | 1.25 | 2.72 | 0.08 | 1.52 | 3.51 | 0.26 | 0.67 | 0.2 | 0.68 | 0.09 | 5.27
| Analytical results for Co in OREAS 266 (Certified Value 22.4 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 22.4 | 24.102 | 22.282 | 23.791 | 22 | 21 | 21.7 | 25.8114 | 21.6 | 21 | 19.1597 | 19.6 | 22.689 | 22.44 | 23.2 | 24.2 | 22.4 | 23.4 | 1
| 1 | 2 | 23 | 24.345 | 21.7241 | 23.6221 | 22 | 21 | 22.1 | 26.1197 | 21.4 | 21.1 | 18.9335 | 20.1 | 22.7417 | 21.82 | 23 | 23.3 | 21.5 | 22.8 | 25
| 1 | 3 | 22.6 | 23.868 | 22.71 | 23.8667 | 22 | 22 | 21.8 | 25.1076 | 21.2 | 20.7 | 19.3556 | 19.8 | 22.1602 | 22.38 | 22.1 | 24.4 | 21 | 23.9 | 16
| 1 | 4 | 23.2 | 23.787 | 22.2876 | 25.2892 | 22 | 21 | 22.1 | 25.2228 | 21.4 | 20.7 | 19.3442 | 20.9 | 22.4699 | 22.77 | 23.1 | 22.8 | 21.1 | 23.1 | 22.3984
| 1 | 5 | 22.3 | 24.129 | 23.2658 | 23.4626 | 22 | 21 | 22.2 | 25.4668 | 21.7 | 21.4 | 19.7674 | 22.6 | 22.0735 | 22.32 | 22.9 | 24.1 | 21.2 | 23.3 | 24
| 1 | 6 | 23.4 | 25.02 | 22.7233 | 23.8832 | 22 | 21 | 22 | 26.0818 | 21.8 | 21.1 | 19.0644 | 22.6 | 22.0749 | 22.27 | 23.1 | 24.5 | 22.6 | 23.3
| Mean |  | 22.8167 | 24.2085 | 22.4988 | 23.9858 | 22 | 21.1667 | 21.9833 | 25.635 | 21.5167 | 21 | 19.2708 | 20.9333 | 22.3682 | 22.3333 | 22.9 | 23.8833 | 21.6333 | 23.3
| Median |  | 22.8 | 24.1155 | 22.4988 | 23.8289 | 22 | 21 | 22.05 | 25.6391 | 21.5 | 21.05 | 19.252 | 20.5 | 22.315 | 22.35 | 23.05 | 24.15 | 21.35 | 23.3
| Std Dev. |  | 0.44907 | 0.44462 | 0.52443 | 0.65841 | 0 | 0.40825 | 0.19408 | 0.43406 | 0.22286 | 0.26833 | 0.29251 | 1.3648 | 0.30624 | 0.30722 | 0.40497 | 0.67946 | 0.69474 | 0.36332
| Rel.Std.Dev. |  | 1.968% | 1.837% | 2.331% | 2.745% | 0% | 1.929% | 0.8828% | 1.693% | 1.036% | 1.278% | 1.518% | 6.52% | 1.369% | 1.376% | 1.768% | 2.845% | 3.211% | 1.559%
| PDM3 |  | 1.867% | 8.081% | 0.4483% | 7.087% | -1.779% | -5.499% | -1.853% | 14.45% | -3.936% | -6.243% | -13.96% | -6.541% | -0.1348% | -0.2904% | 2.24% | 6.63% | -3.416% | 4.025%
| Z-Score (Absolute) |  | 0.35 | 1.4 | 0.11 | 1.24 | 0.27 | 0.9 | 0.28 | 2.48 | 0.63 | 1.02 | 2.33 | 1.07 | 0.01 | 0.01 | 0.42 | 1.16 | 0.54 | 0.72
| Analytical results for Cr in OREAS 266 (Certified Value 131 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 150 | 127.4326 | 135.8811 | 133 | 156 | 133 | 119.1836 | 131 | 122 | 103.407 | 138 | 133.1368 | 135 | 134 | 135 | 136 | 112 |  | 1
| 1 | 2 | 151 | 125.1199 | 131.6857 | 129 | 182 | 133 | 121.4633 | 130 | 132 | 102.7079 | 137 | 130.8295 | 129 | 128 | 133 | 135 | 110 |  | 26
| 1 | 3 | 152 | 130.1856 | 146.3351 | 127 | 131 | 131 | 122.5093 | 129 | 122 | 104.7042 | 137 | 135.9586 | 132 | 127 | 133 | 123 | 112 |  | 16
| 1 | 4 | 152 | 130.1167 | 148.4778 | 124 | 138 | 133 | 119.3947 | 128 | 123 | 103.8067 | 138 | 137.9503 | 135 | 128 | 132 | 122 | 106 |  | 131.1252
| 1 | 5 | 146 | 129.9996 | 147.4911 | 130 | 137 | 134 | 121.4392 | 127 | 126 | 109.8952 | 135 | 139.5446 | 132 | 133 | 134 | 134 | 116 |  | 25
| 1 | 6 | 151 | 129.5995 | 131.8556 | 132 | 150 | 134 | 121.3895 | 129 | 145 | 100.4114 | 138 | 130.4388 | 131 | 131 | 136 | 128 | 109
| Mean |  | 150.3333 | 128.7423 | 140.2877 | 129.1667 | 149 | 133 | 120.8966 | 129 | 128.3333 | 104.1554 | 137.1667 | 134.6431 | 132.3333 | 130.1667 | 133.8333 | 129.6667 | 110.8333
| Median |  | 151 | 129.7995 | 141.1081 | 129.5 | 144 | 133 | 121.4143 | 129 | 124.5 | 103.6069 | 137.5 | 134.5477 | 132 | 129.5 | 133.5 | 131 | 111
| Std Dev. |  | 2.2509 | 2.0555 | 8.0006 | 3.3116 | 18.5903 | 1.0954 | 1.3153 | 1.4142 | 9.0037 | 3.1636 | 1.169 | 3.7731 | 2.3381 | 2.9269 | 1.472 | 6.2183 | 3.3714
| Rel.Std.Dev. |  | 1.497% | 1.597% | 5.703% | 2.564% | 12.48% | 0.8236% | 1.088% | 1.096% | 7.016% | 3.037% | 0.8523% | 2.802% | 1.767% | 2.249% | 1.1% | 4.796% | 3.042%
| PDM3 |  | 14.65% | -1.817% | 6.988% | -1.494% | 13.63% | 1.43% | -7.801% | -1.621% | -2.129% | -20.57% | 4.607% | 2.683% | 0.9214% | -0.731% | 2.065% | -1.112% | -15.48%
| Z-Score (Absolute) |  | 3.71 | 0.26 | 1.86 | 0.18 | 3.46 | 0.52 | 1.7 | 0.21 | 0.34 | 4.78 | 1.29 | 0.82 | 0.4 | 0 | 0.67 | 0.09 | 3.56
| Analytical results for Cs in OREAS 266 (Certified Value 4.4 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 4.59 | 4.4514 | 4.2 | 4.25 | 4.32 | 4.45 | 4.1417 | 4.15 | 4.5074 | 4.27 | 4.32 | 4.53 | 4.5 | 4.71 |  |  |  |  | 1
| 1 | 2 | 4.58 | 4.6284 | 4.2 | 4.22 | 4.35 | 4.43 | 4.1021 | 4.29 | 4.7306 | 4.14 | 4.21 | 4.4 | 4.32 | 4.69 |  |  |  |  | 27
| 1 | 3 | 4.62 | 4.5923 | 4.2 | 4.22 | 4.33 | 4.35 | 4.1862 | 4.1 | 4.5224 | 4.25 | 4.24 | 4.66 | 4.38 | 4.74 |  |  |  |  | 16
| 1 | 4 | 4.64 | 4.4653 | 4.3 | 4.29 | 4.34 | 4.41 | 4.172 | 4.29 | 4.7364 | 4.35 | 4.28 | 4.39 | 4.41 | 4.62 |  |  |  |  | 4.3971
| 1 | 5 | 4.53 | 4.2356 | 4.2 | 4.25 | 4.32 | 4.49 | 4.2829 | 4.48 | 4.4971 | 4.29 | 4.23 | 4.69 | 4.31 | 4.66 |  |  |  |  | 26
| 1 | 6 | 4.57 | 4.4688 | 4.2 | 4.24 | 4.34 | 4.54 | 4.1421 | 4.48 | 4.4796 | 4.27 | 4.33 | 4.76 | 4.74 | 4.77
| Mean |  | 4.5883 | 4.4736 | 4.2167 | 4.245 | 4.3333 | 4.445 | 4.1712 | 4.2983 | 4.5789 | 4.2617 | 4.2683 | 4.5717 | 4.4433 | 4.6983
| Median |  | 4.585 | 4.4671 | 4.2 | 4.245 | 4.335 | 4.44 | 4.157 | 4.29 | 4.5149 | 4.27 | 4.26 | 4.595 | 4.395 | 4.7
| Std Dev. |  | 0.038687 | 0.13804 | 0.040825 | 0.025884 | 0.012111 | 0.065651 | 0.061959 | 0.15968 | 0.12055 | 0.068823 | 0.049565 | 0.15587 | 0.16083 | 0.054191
| Rel.Std.Dev. |  | 0.8432% | 3.086% | 0.9682% | 0.6098% | 0.2795% | 1.477% | 1.485% | 3.715% | 2.633% | 1.615% | 1.161% | 3.41% | 3.62% | 1.153%
| PDM3 |  | 4.35% | 1.742% | -4.102% | -3.458% | -1.449% | 1.09% | -5.137% | -2.245% | 4.136% | -3.079% | -2.927% | 3.971% | 1.053% | 6.852%
| Z-Score (Absolute) |  | 1 | 0.43 | 0.86 | 0.72 | 0.27 | 0.28 | 1.08 | 0.45 | 0.95 | 0.63 | 0.6 | 0.92 | 0.27 | 1.55
| Analytical results for Cu in OREAS 266 (Certified Value 32.2 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 33.5 | 31.7312 | 29.5724 | 32 | 34 | 34.4 | 30.36 | 32 | 35.9 | 20.3064 | 28.4 | 26.6943 | 33.4 | 31.7 | 35.3 | 33.5 | 31.3 |  | 1
| 1 | 2 | 33.8 | 31.0224 | 28.8169 | 32 | 33 | 35 | 30.485 | 31.8 | 36.4 | 19.2466 | 29.4 | 27.8684 | 33.1 | 30.2 | 34.3 | 32 | 29.7 |  | 28
| 1 | 3 | 33.2 | 32.5137 | 30.2353 | 31 | 33 | 33.8 | 29.6533 | 33 | 35.4 | 21.2874 | 28.1 | 26.5047 | 33.7 | 30.2 | 36.7 | 33.8 | 29.9 |  | 16
| 1 | 4 | 34.7 | 32.1155 | 30.8658 | 32 | 32 | 35 | 29.405 | 33.4 | 35.7 | 19.6011 | 29.6 | 32.4891 | 34.2 | 31.4 | 34.2 | 33.2 | 28 |  | 32.2112
| 1 | 5 | 34.5 | 33.0778 | 28.7524 | 32 | 33 | 35.4 | 29.2233 | 33.4 | 40 | 20.239 | 31.2 | 33.7263 | 33.5 | 30.8 | 37.4 | 31.5 | 29.4 |  | 27
| 1 | 6 | 34.8 | 32.3593 | 29.4992 | 31 | 34 | 35.1 | 29.91 | 31.9 | 35.9 | 21.1249 | 31.2 | 25.4517 | 33.3 | 30.8 | 36.7 | 35.3 | 29.5
| Mean |  | 34.0833 | 32.1366 | 29.6237 | 31.6667 | 33.1667 | 34.7833 | 29.8394 | 32.5833 | 36.55 | 20.3009 | 29.65 | 28.7891 | 33.5333 | 30.85 | 35.7667 | 33.2167 | 29.6333
| Median |  | 34.15 | 32.2374 | 29.5358 | 32 | 33 | 35 | 29.7817 | 32.5 | 35.9 | 20.2727 | 29.5 | 27.2814 | 33.45 | 30.8 | 36 | 33.35 | 29.6
| Std Dev. |  | 0.67355 | 0.70484 | 0.81794 | 0.5164 | 0.75277 | 0.58109 | 0.50905 | 0.76529 | 1.7213 | 0.80713 | 1.3293 | 3.4542 | 0.38297 | 0.61237 | 1.3589 | 1.3556 | 1.0577
| Rel.Std.Dev. |  | 1.976% | 2.193% | 2.761% | 1.631% | 2.27% | 1.671% | 1.706% | 2.349% | 4.71% | 3.976% | 4.483% | 12% | 1.142% | 1.985% | 3.799% | 4.081% | 3.569%
| PDM3 |  | 5.812% | -0.2314% | -8.033% | -1.69% | 2.966% | 7.985% | -7.363% | 1.155% | 13.47% | -36.98% | -7.951% | -10.62% | 4.105% | -4.226% | 11.04% | 3.122% | -8.003%
| Z-Score (Absolute) |  | 0.57 | 0 | 0.74 | 0.14 | 0.3 | 0.78 | 0.67 | 0.13 | 1.3 | 3.47 | 0.73 | 0.98 | 0.41 | 0.38 | 1.07 | 0.32 | 0.73
| Analytical results for Dy in OREAS 266 (Certified Value 4.37 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 4.37 | 4.4057 | 4.7676 | 3.9 | 4.3 | 3.1852 | 5.99 | [11 empty] | 1
| 1 | 2 | 4.32 | 4.1164 | 4.8786 | 4.2 | 4.21 | 3.1946 | 6.06 | [11 empty] | 3
| 1 | 3 | 4.32 | 4.0707 | 4.7108 | 4.3 | 4.3 | 3.2224 | 5.94 | [11 empty] | 16
| 1 | 4 | 4.37 | 4.1274 | 4.6628 | 4.3 | 4.31 | 3.253 | 5.93 | [11 empty] | 4.3673
| 1 | 5 | 4.27 | 4.3662 | 4.4982 | 4.3 | 4.19 | 3.3601 | 6.03 | [11 empty] | 28
| 1 | 6 | 4.41 | 4.2293 | 4.7449 | 4.4 | 4.27 | 3.21 | 6.19
| Mean |  | 4.3433 | 4.2193 | 4.7105 | 4.2333 | 4.2633 | 3.2375 | 6.0233
| Median |  | 4.345 | 4.1784 | 4.7279 | 4.3 | 4.285 | 3.2162 | 6.01
| Std Dev. |  | 0.049666 | 0.13966 | 0.12654 | 0.17512 | 0.051251 | 0.064553 | 0.095847
| Rel.Std.Dev. |  | 1.143% | 3.31% | 2.686% | 4.137% | 1.202% | 1.994% | 1.591%
| PDM3 |  | -0.5485% | -3.389% | 7.859% | -3.067% | -2.38% | -25.87% | 37.92%
| Z-Score (Absolute) |  | 0.67 | 0.37 | 3.77 | 0.25 | 0 | 8.65 | 14.83
| Analytical results for Er in OREAS 266 (Certified Value 2.02 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.01 | 2.0216 | 2.1586 | 1.9 | 1.77 | 3.8157 | 2.25 | [11 empty] | 1
| 1 | 2 | 1.95 | 2.0548 | 2.0052 | 2.1 | 1.81 | 3.8009 | 2.19 | [11 empty] | 4
| 1 | 3 | 2.05 | 2.0554 | 1.9773 | 2.1 | 1.81 | 3.8695 | 2.27 | [11 empty] | 16
| 1 | 4 | 2 | 2.0146 | 1.9203 | 2 | 1.8 | 3.8368 | 2.23 | [11 empty] | 2.0198
| 1 | 5 | 2.01 | 1.9717 | 1.8563 | 2.1 | 1.78 | 3.9801 | 2.22 | [11 empty] | 29
| 1 | 6 | 2.07 | 2.103 | 1.9726 | 2.1 | 1.79 | 3.6989 | 2.29
| Mean |  | 2.015 | 2.0368 | 1.9817 | 2.05 | 1.7933 | 3.8336 | 2.2417
| Median |  | 2.01 | 2.0382 | 1.9749 | 2.1 | 1.795 | 3.8263 | 2.24
| Std Dev. |  | 0.041833 | 0.044774 | 0.10144 | 0.083666 | 0.01633 | 0.091957 | 0.036009
| Rel.Std.Dev. |  | 2.076% | 2.198% | 5.119% | 4.081% | 0.9106% | 2.399% | 1.606%
| PDM3 |  | -0.2356% | 0.8461% | -1.884% | 1.497% | -11.21% | 89.81% | 10.99%
| Z-Score (Absolute) |  | 0.27 | 0 | 0.67 | 0.16 | 2.98 | 21.97 | 2.5
| Analytical results for Eu in OREAS 266 (Certified Value 1.71 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.83 | 0.71349 | 1.7874 | 1.34 | 1.7 | 1.4805 | 1.95 | [11 empty] | 1
| 1 | 2 | 1.82 | 0.64632 | 1.8599 | 1.52 | 1.71 | 1.4804 | 1.93 | [11 empty] | 5
| 1 | 3 | 1.82 | 0.68655 | 1.7257 | 1.51 | 1.71 | 1.4997 | 1.9 | [11 empty] | 16
| 1 | 4 | 1.83 | 0.60646 | 1.6996 | 1.52 | 1.74 | 1.4979 | 1.87 | [11 empty] | 1.7111
| 1 | 5 | 1.89 | 0.72199 | 1.6415 | 1.52 | 1.7 | 1.5438 | 1.96 | [11 empty] | 30
| 1 | 6 | 1.86 | 0.66514 | 1.7246 | 1.58 | 1.73 | 1.4989 | 2.03
| Mean |  | 1.8417 | 0.67333 | 1.7398 | 1.4983 | 1.715 | 1.5002 | 1.94
| Median |  | 1.83 | 0.67584 | 1.7251 | 1.52 | 1.71 | 1.4984 | 1.94
| Std Dev. |  | 0.027869 | 0.043412 | 0.075352 | 0.081588 | 0.016432 | 0.023197 | 0.055136
| Rel.Std.Dev. |  | 1.513% | 6.447% | 4.331% | 5.445% | 0.9581% | 1.546% | 2.842%
| PDM3 |  | 7.63% | -60.65% | 1.675% | -12.43% | 0.2274% | -12.32% | 13.38%
| Z-Score (Absolute) |  | 0.4 | 3.27 | 0.08 | 0.68 | 0 | 0.67 | 0.71
| Analytical results for Fe in OREAS 266 (Certified Value 5.05 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.14 | 5.334 | 5.0356 | 3.9158 | 5.19 | 4.75 | 4.9 | 5.04 | 5.11 | 4.76 | 5.0059 | 5.13 | 5.2416 | 5.04 | 4.83 | 5.16 | 5.11 | 5.01 | 1
| 1 | 2 | 5.17 | 5.382 | 4.9697 | 3.874 | 5.21 | 4.83 | 4.95 | 5.04 | 5.09 | 4.79 | 5.0417 | 5.08 | 5.3604 | 4.93 | 4.74 | 5.12 | 5.04 | 5.25 | #N/A
| 1 | 3 | 5.22 | 5.3165 | 5.1483 | 3.9306 | 5.24 | 4.87 | 4.99 | 5 | 5 | 4.73 | 5.0318 | 5.08 | 5.2201 | 4.97 | 4.7 | 5 | 4.99 | 5.05 | 16
| 1 | 4 | 5.23 | 5.2826 | 5.1127 | 3.8979 | 5.19 | 4.78 | 5.03 | 4.91 | 5.02 | 4.69 | 5.0703 | 5.13 | 5.1462 | 5.1 | 4.86 | 5.02 | 4.87 | 5.05 | 5.0488
| 1 | 5 | 5.15 | 5.3104 | 5.1743 | 3.8581 | 5.13 | 4.79 | 4.99 | 5 | 5.01 | 4.79 | 5.1427 | 5.05 | 5.291 | 5.02 | 4.78 | 5.06 | 5.01 | 5.19 | 31
| 1 | 6 | 5.19 | 5.3433 | 5.1508 | 3.8461 | 5.23 | 5 | 4.96 | 5.07 | 4.96 | 4.77 | 4.9408 | 5.21 | 5.3721 | 5.06 | 4.81 | 5.15 | 5.06 | 4.9
| Mean |  | 5.1833 | 5.3281 | 5.0986 | 3.8871 | 5.1983 | 4.8367 | 4.97 | 5.01 | 5.0317 | 4.755 | 5.0388 | 5.1133 | 5.2719 | 5.02 | 4.7867 | 5.085 | 5.0133 | 5.075
| Median |  | 5.18 | 5.3253 | 5.1305 | 3.886 | 5.2 | 4.81 | 4.975 | 5.02 | 5.015 | 4.765 | 5.0367 | 5.105 | 5.2663 | 5.03 | 4.795 | 5.09 | 5.025 | 5.05
| Std Dev. |  | 0.036697 | 0.033723 | 0.07965 | 0.03324 | 0.0392 | 0.090259 | 0.044272 | 0.055857 | 0.057067 | 0.038859 | 0.067151 | 0.056804 | 0.086763 | 0.061644 | 0.059217 | 0.068044 | 0.08165 | 0.12645
| Rel.Std.Dev. |  | 0.708% | 0.6329% | 1.562% | 0.8551% | 0.7541% | 1.866% | 0.8908% | 1.115% | 1.134% | 0.8172% | 1.333% | 1.111% | 1.646% | 1.228% | 1.237% | 1.338% | 1.629% | 2.492%
| PDM3 |  | 2.664% | 5.532% | 0.9847% | -23.01% | 2.961% | -4.202% | -1.562% | -0.7694% | -0.3402% | -5.82% | -0.1981% | 1.277% | 4.418% | -0.5713% | -5.193% | 0.7161% | -0.7033% | 0.5181%
| Z-Score (Absolute) |  | 1.39 | 2.76 | 0.6 | 10.8 | 1.53 | 1.87 | 0.61 | 0.24 | 0.03 | 2.64 | 0.03 | 0.73 | 2.23 | 0.14 | 2.34 | 0.47 | 0.21 | 0.37
| Analytical results for Ga in OREAS 266 (Certified Value 19.2 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 20.13 | 17.78 | 18.3441 | 21.2 | 11.8 | 20.52 | 21.637 | 19.05 | 15.9358 | 16.75 | 19.7548 | 20.01 | 18.45 | 19.9 | 19.1 | 20.23 |  |  | 1
| 1 | 2 | 20.99 | 17.58 | 18.6483 | 21 | 12.7 | 20.26 | 21.541 | 19 | 15.8922 | 17.5 | 20.5219 | 19.36 | 17.6 | 19 | 18.75 | 19.98 |  |  | 32
| 1 | 3 | 20.52 | 18.13 | 19.2721 | 20.7 | 13.4 | 20.39 | 20.9361 | 18.5 | 15.9885 | 16.7 | 19.8199 | 20 | 17.65 | 20.3 | 18.5 | 20.84 |  |  | 16
| 1 | 4 | 21.24 | 18.06 | 18.8467 | 21.3 | 11.6 | 20.41 | 20.7903 | 18.65 | 16.1538 | 17.4 | 20.3799 | 20.37 | 17.65 | 18.95 | 18.3 | 20.33 |  |  | 19.2499
| 1 | 5 | 19.87 | 18.29 | 18.9634 | 21.1 | 12.5 | 20.35 | 21.0019 | 19.25 | 16.0205 | 18.45 | 19.8788 | 19.79 | 17.9 | 20 | 18.3 | 20.21 |  |  | 32
| 1 | 6 | 21.61 | 16.92 | 19.0268 | 21.2 | 13.9 | 20.46 | 21.3288 | 19.15 | 15.8022 | 18.6 | 19.3404 | 19.75 | 17.65 | 20.6 | 20.2 | 20.36
| Mean |  | 20.7267 | 17.7933 | 18.8502 | 21.0833 | 12.65 | 20.3983 | 21.2058 | 18.9333 | 15.9655 | 17.5667 | 19.9493 | 19.88 | 17.8167 | 19.7917 | 18.8583 | 20.325
| Median |  | 20.755 | 17.92 | 18.9051 | 21.15 | 12.6 | 20.4 | 21.1654 | 19.025 | 15.9622 | 17.45 | 19.8494 | 19.895 | 17.65 | 19.95 | 18.625 | 20.28
| Std Dev. |  | 0.67018 | 0.4979 | 0.32214 | 0.2137 | 0.89163 | 0.089759 | 0.3465 | 0.29439 | 0.11996 | 0.8122 | 0.43442 | 0.33669 | 0.32813 | 0.67854 | 0.72417 | 0.28557
| Rel.Std.Dev. |  | 3.233% | 2.798% | 1.709% | 1.014% | 7.048% | 0.44% | 1.634% | 1.555% | 0.7514% | 4.624% | 2.178% | 1.694% | 1.842% | 3.428% | 3.84% | 1.405%
| PDM3 |  | 7.671% | -7.567% | -2.076% | 9.524% | -34.29% | 5.966% | 10.16% | -1.645% | -17.06% | -8.744% | 3.633% | 3.273% | -7.446% | 2.814% | -2.034% | 5.585%
| Z-Score (Absolute) |  | 0.77 | 0.88 | 0.29 | 0.97 | 3.77 | 0.58 | 1.04 | 0.24 | 1.91 | 1.01 | 0.33 | 0.29 | 0.87 | 0.24 | 0.28 | 0.54
| Analytical results for Gd in OREAS 266 (Certified Value 6.32 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 5.97 | 6.8854 | 6.2485 | 4.9 | 5.95 | 3.0551 | 7.16 | [11 empty] | 1
| 1 | 2 | 6.07 | 6.7739 | 6.4682 | 5.8 | 6 | 2.8802 | 6.86 | [11 empty] | 6
| 1 | 3 | 5.89 | 6.739 | 6.4554 | 5.6 | 6.01 | 2.9883 | 6.91 | [11 empty] | 16
| 1 | 4 | 6.01 | 6.7214 | 6.2765 | 5.7 | 6.05 | 3.1164 | 6.69 | [11 empty] | 6.3214
| 1 | 5 | 5.99 | 7.0922 | 6.1694 | 5.9 | 6.01 | 3.159 | 6.87 | [11 empty] | 33
| 1 | 6 | 6.23 | 6.9357 | 6.424 | 5.8 | 6.02 | 3.0679 | 7.13
| Mean |  | 6.0267 | 6.8579 | 6.3403 | 5.6167 | 6.0067 | 3.0445 | 6.9367
| Median |  | 6 | 6.8296 | 6.3502 | 5.75 | 6.01 | 3.0615 | 6.89
| Std Dev. |  | 0.11553 | 0.14251 | 0.12515 | 0.3656 | 0.03266 | 0.09909 | 0.1784
| Rel.Std.Dev. |  | 1.917% | 2.078% | 1.974% | 6.509% | 0.5437% | 3.255% | 2.572%
| PDM3 |  | -4.662% | 8.488% | 0.2998% | -11.15% | -4.979% | -51.84% | 9.733%
| Z-Score (Absolute) |  | 0 | 1.37 | 0.52 | 0.67 | 0.03 | 4.9 | 1.5
| Analytical results for Ge in OREAS 266 (Indicative Value 0.17 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.4 | < 0.2 | 1.7 | 0.1 | 0.18 | 0.25 | 0.15 | 0.2614 | 0.12 | 0.18777 | 0.14 | 0.27 | 0.12 | [5 empty] | 1
| 1 | 2 | 1.5 | < 0.2 | 1.7 | < 0.1 | 0.18 | 0.23 | 0.23 | 0.3696 | 0.13 | 0.20693 | 0.14 | 0.27 | 0.12 | [5 empty] | 1
| 1 | 3 | 1.5 | < 0.2 | 1.7 | < 0.1 | 0.17 | 0.17 | 0.24 | 0.1995 | 0.12 | 0.19947 | 0.13 | 0.3 | 0.14 | [5 empty] | 16
| 1 | 4 | 1.5 | < 0.2 | 1.7 | < 0.1 | 0.16 | 0.15 | 0.11 | 0.2492 | 0.12 | 0.1897 | 0.14 | 0.27 | 0.21 | [5 empty] | 0.16682
| 1 | 5 | 1.3 | < 0.2 | 1.8 | < 0.1 | 0.18 | 0.17 | 0.06 | 0.1003 | 0.13 | 0.20211 | 0.13 | 0.29 | 0.11 | [5 empty] | 7
| 1 | 6 | 1.6 | < 0.2 | 1.8 | < 0.1 | 0.18 | 0.26 | 0.13 | 0.0683 | 0.13 | 0.20688 | 0.12 | 0.28 | 0.12
| Mean |  | 1.4667 |  | 1.7333 | 0.1 | 0.175 | 0.205 | 0.15333 | 0.20805 | 0.125 | 0.19881 | 0.13333 | 0.28 | 0.13667
| Median |  | 1.5 |  | 1.7 | 0.1 | 0.18 | 0.2 | 0.14 | 0.22435 | 0.125 | 0.20079 | 0.135 | 0.275 | 0.12
| Std Dev. |  | 0.10328 |  | 0.05164 |  | 0.0083666 | 0.047223 | 0.070048 | 0.11122 | 0.0054772 | 0.0083364 | 0.008165 | 0.012649 | 0.037238
| Rel.Std.Dev. |  | 7.042% |  | 2.979% |  | 4.781% | 23.04% | 45.68% | 53.46% | 4.382% | 4.193% | 6.124% | 4.518% | 27.25%
| PDM3 |  | 779.2% |  | 939% | -40.06% | 4.902% | 22.89% | -8.086% | 24.71% | -25.07% | 19.18% | -20.07% | 67.84% | -18.08%
| Z-Score (Absolute) |  | 20.9 | 1.21 | 25.22 | 1.89 | 0 | 0.49 | 0.35 | 0.53 | 0.81 | 0.39 | 0.67 | 1.7 | 0.62
| Analytical results for Hf in OREAS 266 (Certified Value 4.96 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 5.25 | 5.4228 | 4.961 | 5 | 4.8 | 4.5 | 4.9 | 4.9 | 5.2967 | 4.3 | 4.8331 | 4.74 | 4.7 | 5.3 | 4.8 | 4.8 |  |  | 1
| 1 | 2 | 5.22 | 5.5015 | 5.3203 | 5 | 4.1 | 4.6 | 5.1 | 5.7 | 5.2035 | 4.4 | 5.2261 | 4.81 | 4.7 | 5.3 | 4.5 | 4.65 |  |  | 7
| 1 | 3 | 5.22 | 5.2193 | 5.1112 | 5 | 0.7 | 4.7 | 5 | 5.8 | 5.3349 | 4.2 | 5.0013 | 4.67 | 4.7 | 5.3 | 4.9 | 4.86 |  |  | 16
| 1 | 4 | 5.39 | 5.4553 | 4.9535 | 4.9 | 4.6 | 4.7 | 5.1 | 5.4 | 5.2947 | 4.6 | 5.2093 | 4.72 | 4.8 | 5 | 4.7 | 4.64 |  |  | 4.9573
| 1 | 5 | 5.27 | 5.2526 | 4.7176 | 5 | 4.4 | 4.7 | 5.1 | 4.9 | 5.5707 | 4.8 | 5.0499 | 4.73 | 4.6 | 5.5 | 4.6 | 4.73 |  |  | 34
| 1 | 6 | 5.65 | 5.5938 | 4.9988 | 5 | 4.1 | 4.8 | 5 | 5.7 | 5.1505 | 4.8 | 4.8384 | 4.69 | 4.7 | 5.6 | 5.2 | 4.82
| Mean |  | 5.3333 | 5.4076 | 5.0104 | 4.9833 | 3.7833 | 4.6667 | 5.0333 | 5.4 | 5.3085 | 4.5167 | 5.0264 | 4.7267 | 4.7 | 5.3333 | 4.7833 | 4.75
| Median |  | 5.26 | 5.4391 | 4.9799 | 5 | 4.25 | 4.7 | 5.05 | 5.55 | 5.2957 | 4.5 | 5.0256 | 4.725 | 4.7 | 5.3 | 4.75 | 4.765
| Std Dev. |  | 0.16741 | 0.14524 | 0.19886 | 0.040825 | 1.5355 | 0.10328 | 0.08165 | 0.40988 | 0.14544 | 0.25626 | 0.17159 | 0.048442 | 0.063246 | 0.20656 | 0.24833 | 0.091652
| Rel.Std.Dev. |  | 3.139% | 2.686% | 3.969% | 0.8192% | 40.59% | 2.213% | 1.622% | 7.59% | 2.74% | 5.674% | 3.414% | 1.025% | 1.346% | 3.873% | 5.192% | 1.93%
| PDM3 |  | 7.586% | 9.084% | 1.072% | 0.5259% | -23.68% | -5.862% | 1.535% | 8.931% | 7.085% | -8.888% | 1.394% | -4.652% | -5.19% | 7.586% | -3.509% | -4.181%
| Z-Score (Absolute) |  | 0.75 | 0.91 | 0.03 | 0.03 | 2.69 | 0.73 | 0.08 | 0.89 | 0.69 | 1.06 | 0.07 | 0.6 | 0.66 | 0.75 | 0.47 | 0.55
| Analytical results for Hg in OREAS 266 (Indicative Value 0.18 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.10 | 1.14 | [15 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [15 empty] | 2
2
| 1 | 1 | < 1 | < 2 | 0.1817 | [15 empty] | 1
| 1 | 2 | < 1 | < 2 | 0.1808 | [15 empty] | 1
| 1 | 3 | < 1 | < 2 | 0.1794 | [15 empty] | 16
| 1 | 4 | < 1 | < 2 | 0.178 | [15 empty] | 0.17967
| 1 | 5 | < 1 | < 2 | 0.1806 | [15 empty] | 7
| 1 | 6 | < 1 | < 2 | 0.1775
| Mean |  |  |  | 0.17967
| Median |  |  |  | 0.18
| Std Dev. |  |  |  | 0.0016633
| Rel.Std.Dev. |  |  |  | 0.9258%
| PDM3 |  |  |  | -1.998e-13%
| Z-Score (Absolute) |  | 0 | 1.05 | 0.67
| Analytical results for Ho in OREAS 266 (Certified Value 0.77 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.78 | 0.7499 | 0.80666 | 0.7 | 0.75 | 0.7555 | 0.98 | [11 empty] | 1
| 1 | 2 | 0.77 | 0.89836 | 0.88315 | 0.8 | 0.75 | 0.7316 | 0.95 | [11 empty] | 3
| 1 | 3 | 0.75 | 0.70526 | 0.83508 | 0.8 | 0.76 | 0.76 | 0.98 | [11 empty] | 16
| 1 | 4 | 0.77 | 0.74734 | 0.78455 | 0.8 | 0.78 | 0.7658 | 0.92 | [11 empty] | 0.76802
| 1 | 5 | 0.75 | 0.74427 | 0.80728 | 0.8 | 0.74 | 0.7698 | 0.96 | [11 empty] | 35
| 1 | 6 | 0.78 | 0.73513 | 0.92444 | 0.8 | 0.76 | 0.7596 | 0.98
| Mean |  | 0.76667 | 0.76338 | 0.84019 | 0.78333 | 0.75667 | 0.75705 | 0.96167
| Median |  | 0.77 | 0.7458 | 0.82118 | 0.8 | 0.755 | 0.7598 | 0.97
| Std Dev. |  | 0.013663 | 0.068116 | 0.05343 | 0.040825 | 0.013663 | 0.013449 | 0.024014
| Rel.Std.Dev. |  | 1.782% | 8.923% | 6.359% | 5.212% | 1.806% | 1.777% | 2.497%
| PDM3 |  | -0.1764% | -0.6049% | 9.397% | 1.994% | -1.478% | -1.429% | 25.21%
| Z-Score (Absolute) |  | 0.14 | 0.14 | 6.21 | NA | 0.69 | 0.66 | 16.24
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 266 (Certified Value 0.079 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.09 | 0.08 | < 0.1 | < 2 | 0.08 | 0.07 | 0.09 | 0.08 | 0.082791 | < 0.1 | 0.076 | 0.077 | 0.081 | 0.078 |  |  |  |  | 1
| 1 | 2 | 0.09 | 0.08 | < 0.1 | < 2 | 0.074 | 0.07 | 0.0951 | 0.084 | 0.08602 | < 0.1 | 0.076 | 0.077 | 0.071 | 0.08 |  |  |  |  | 20
| 1 | 3 | 0.07 | 0.08 | < 0.1 | < 2 | 0.077 | 0.08 | 0.089 | 0.072 | 0.087616 | < 0.1 | 0.074 | 0.076 | 0.085 | 0.081 |  |  |  |  | 16
| 1 | 4 | 0.08 | 0.07 | < 0.1 | < 2 | 0.086 | 0.07 | 0.0908 | 0.079 | 0.083577 | < 0.1 | 0.073 | 0.077 | 0.077 | 0.083 |  |  |  |  | 0.078564
| 1 | 5 | 0.07 | 0.08 | < 0.1 | < 2 | 0.083 | 0.08 | 0.0959 | 0.085 | 0.080996 | < 0.1 | 0.077 | 0.076 | 0.078 | 0.079 |  |  |  |  | 36
| 1 | 6 | 0.07 | 0.08 | < 0.1 | < 2 | 0.076 | 0.08 | 0.0885 | 0.088 | 0.078812 | < 0.1 | 0.077 | 0.078 | 0.073 | 0.08
| Mean |  | 0.078333 | 0.078333 |  |  | 0.079333 | 0.075 | 0.09155 | 0.081333 | 0.083302 |  | 0.0755 | 0.076833 | 0.0775 | 0.080167
| Median |  | 0.075 | 0.08 |  |  | 0.0785 | 0.075 | 0.0904 | 0.082 | 0.083184 |  | 0.076 | 0.077 | 0.0775 | 0.08
| Std Dev. |  | 0.0098319 | 0.0040825 |  |  | 0.0045461 | 0.0054772 | 0.0031716 | 0.0056451 | 0.003219 |  | 0.0016432 | 0.00075277 | 0.0051284 | 0.0017224
| Rel.Std.Dev. |  | 12.55% | 5.212% |  |  | 5.73% | 7.303% | 3.464% | 6.941% | 3.864% |  | 2.176% | 0.9797% | 6.617% | 2.149%
| PDM3 |  | -0.293% | -0.293% |  |  | 0.9799% | -4.536% | 16.53% | 3.526% | 6.031% |  | -3.899% | -2.202% | -1.354% | 2.041%
| Z-Score (Absolute) |  | 0 | 0 | 6.55 | 213.08 | 0.23 | 0.77 | 3.06 | 0.69 | 1.15 | 6.55 | 0.66 | 0.35 | 0.19 | 0.42
| Analytical results for K in OREAS 266 (Certified Value 1.51 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.4982 | 1.6138 | 1.5106 | 1.4263 | 1.49 | 1.31 | 1.49 | 1.54 | 1.46 | 1.55 | 1.5045 | 1.5 | 1.5316 | 1.52 | 1.43 | 1.55 | 1.53 | 1.48 | 1
| 1 | 2 | 1.5076 | 1.6269 | 1.4671 | 1.4019 | 1.5 | 1.25 | 1.52 | 1.54 | 1.46 | 1.33 | 1.5068 | 1.49 | 1.5567 | 1.49 | 1.4 | 1.52 | 1.51 | 1.48 | #N/A
| 1 | 3 | 1.5214 | 1.6089 | 1.5453 | 1.4339 | 1.51 | 1.36 | 1.51 | 1.54 | 1.49 | 1.5 | 1.4991 | 1.5 | 1.5342 | 1.49 | 1.39 | 1.49 | 1.47 | 1.48 | 16
| 1 | 4 | 1.5254 | 1.5912 | 1.5116 | 1.4262 | 1.5 | 1.38 | 1.51 | 1.51 | 1.49 | 1.46 | 1.5197 | 1.52 | 1.5189 | 1.52 | 1.45 | 1.51 | 1.44 | 1.5 | 1.508
| 1 | 5 | 1.5039 | 1.6047 | 1.5746 | 1.373 | 1.48 | 1.41 | 1.53 | 1.55 | 1.47 | 1.49 | 1.5457 | 1.49 | 1.5605 | 1.5 | 1.42 | 1.52 | 1.5 | 1.46 | 37
| 1 | 6 | 1.5187 | 1.6401 | 1.5225 | 1.4074 | 1.5 | 1.45 | 1.5 | 1.55 | 1.44 | 1.51 | 1.5215 | 1.53 | 1.5354 | 1.5 | 1.44 | 1.53 | 1.51 | 1.51
| Mean |  | 1.5125 | 1.6143 | 1.522 | 1.4115 | 1.4967 | 1.36 | 1.51 | 1.5383 | 1.4683 | 1.4733 | 1.5162 | 1.505 | 1.5395 | 1.5033 | 1.4217 | 1.52 | 1.4933 | 1.485
| Median |  | 1.5131 | 1.6114 | 1.5171 | 1.4168 | 1.5 | 1.37 | 1.51 | 1.54 | 1.465 | 1.495 | 1.5133 | 1.5 | 1.5348 | 1.5 | 1.425 | 1.52 | 1.505 | 1.48
| Std Dev. |  | 0.010831 | 0.017215 | 0.036249 | 0.02249 | 0.010328 | 0.071554 | 0.014142 | 0.01472 | 0.019408 | 0.07607 | 0.0169 | 0.016432 | 0.015927 | 0.013663 | 0.023166 | 0.02 | 0.03266 | 0.017607
| Rel.Std.Dev. |  | 0.7161% | 1.066% | 2.382% | 1.593% | 0.6901% | 5.261% | 0.9366% | 0.9569% | 1.322% | 5.163% | 1.115% | 1.092% | 1.035% | 0.9088% | 1.63% | 1.316% | 2.187% | 1.186%
| PDM3 |  | 0.2995% | 7.045% | 0.9243% | -6.404% | -0.7527% | -9.815% | 0.1315% | 2.01% | -2.632% | -2.3% | 0.5442% | -0.2001% | 2.089% | -0.3106% | -5.726% | 0.7946% | -0.9737% | -1.526%
| Z-Score (Absolute) |  | 0.31 | 4.02 | 0.65 | 3.38 | 0.27 | 5.26 | 0.21 | 1.25 | 1.31 | 1.13 | 0.44 | 0.03 | 1.29 | 0.03 | 3.01 | 0.58 | 0.4 | 0.7
| Analytical results for La in OREAS 266 (Certified Value 44.9 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 44.47 | 45.4405 | 46.7841 | 34 | 40 | 45.1558 | 43.9 | 44.6 | 36.0109 | 42.7 | 46.2 | 44.8 | 46.4 | 45.8 | 64.61 |  |  |  | 1
| 1 | 2 | 44.33 | 44.5234 | 45.7259 | 40.2 | 40 | 44.8479 | 44.3 | 44 | 35.6385 | 43.6 | 44.8 | 43.3 | 45.8 | 43.9 | 62.72 |  |  |  | 21
| 1 | 3 | 43.94 | 44.857 | 47.531 | 40.1 | 41 | 44.8946 | 44.5 | 42.2 | 36.4712 | 43.2 | 45.9 | 43.1 | 44.4 | 44.9 | 66.42 |  |  |  | 16
| 1 | 4 | 45.14 | 44.1716 | 46.728 | 39.8 | 41 | 44.5879 | 44.7 | 42.6 | 36.2711 | 44.8 | 46.5 | 44.1 | 45.3 | 44.8 | 64.68 |  |  |  | 44.8686
| 1 | 5 | 44.45 | 44.4313 | 47.4099 | 40 | 41 | 44.0081 | 45.9 | 44 | 37.9502 | 47.5 | 45.8 | 42.9 | 44.1 | 46.7 | 64.1 |  |  |  | 38
| 1 | 6 | 44.48 | 44.4857 | 47.1685 | 41 | 41 | 44.6673 | 44.6 | 44.4 | 35.7578 | 47.6 | 45.1 | 43.8 | 44.7 | 46.4 | 67.26
| Mean |  | 44.4683 | 44.6516 | 46.8912 | 39.1833 | 40.6667 | 44.6936 | 44.65 | 43.6333 | 36.3499 | 44.9 | 45.7167 | 43.6667 | 45.1167 | 45.4167 | 64.965
| Median |  | 44.46 | 44.5045 | 46.9763 | 40.05 | 41 | 44.7576 | 44.55 | 44 | 36.141 | 44.2 | 45.85 | 43.55 | 45 | 45.35 | 64.645
| Std Dev. |  | 0.38727 | 0.44439 | 0.65595 | 2.5725 | 0.5164 | 0.38978 | 0.67454 | 0.9913 | 0.84303 | 2.167 | 0.64936 | 0.71181 | 0.87958 | 1.0685 | 1.6358
| Rel.Std.Dev. |  | 0.8709% | 0.9952% | 1.399% | 6.565% | 1.27% | 0.8721% | 1.511% | 2.272% | 2.319% | 4.826% | 1.42% | 1.63% | 1.95% | 2.353% | 2.518%
| PDM3 |  | -0.8921% | -0.4837% | 4.508% | -12.67% | -9.365% | -0.39% | -0.4872% | -2.753% | -18.99% | 0.06995% | 1.89% | -2.679% | 0.5528% | 1.221% | 44.79%
| Z-Score (Absolute) |  | 0.13 | 0 | 1.53 | 3.74 | 2.73 | 0.03 | 0 | 0.7 | 5.68 | 0.17 | 0.73 | 0.67 | 0.32 | 0.52 | 13.91
| Analytical results for Li in OREAS 266 (Certified Value 29.4 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 28.3 | 30.9045 | 28.8305 | 50.2435 | 28.8 | 30 | 27.2 | 29.6765 | 29.5 | 29.2 | 26.91 | 31 | 31.3606 | 28.5 | 29.4 | 32.3 | 30.1 | 27 | 1
| 1 | 2 | 28.9 | 30.3424 | 28.1741 | 49.5294 | 28.3 | 30 | 26.5 | 29.7637 | 29.5 | 29.3 | 26.832 | 30.7 | 31.7366 | 27.5 | 28.8 | 31.6 | 29.9 | 31 | #N/A
| 1 | 3 | 29 | 29.778 | 29.5914 | 48.5474 | 28.5 | 31 | 26.4 | 28.6898 | 29.7 | 28.9 | 27.066 | 30.8 | 31.7402 | 28.2 | 28.4 | 31.1 | 30.9 | 32 | 16
| 1 | 4 | 29.4 | 29.2968 | 29.1313 | 48.4794 | 28.7 | 30 | 26.4 | 29.6863 | 29.7 | 28.7 | 26.988 | 31.2 | 31.9111 | 28.4 | 29.8 | 31.3 | 30.1 | 27 | 29.3825
| 1 | 5 | 28.9 | 30.5549 | 30.1399 | 49.0911 | 29.3 | 30 | 26.2 | 29.5821 | 29.5 | 28 | 27.534 | 31 | 31.4347 | 27.9 | 29 | 31.4 | 29.7 | 27 | 39
| 1 | 6 | 30.3 | 29.9779 | 29.2471 | 49.9155 | 29.3 | 31 | 26.6 | 30.009 | 29.4 | 28.8 | 25.662 | 31.5 | 31.6664 | 27.7 | 29.5 | 31.7 | 31.7 | 29
| Mean |  | 29.1333 | 30.1424 | 29.1857 | 49.3011 | 28.8167 | 30.3333 | 26.55 | 29.5679 | 29.55 | 28.8167 | 26.832 | 31.0333 | 31.6416 | 28.0333 | 29.15 | 31.5667 | 30.4 | 28.8333
| Median |  | 28.95 | 30.1601 | 29.1892 | 49.3102 | 28.75 | 30 | 26.45 | 29.6814 | 29.5 | 28.85 | 26.949 | 31 | 31.7015 | 28.05 | 29.2 | 31.5 | 30.1 | 28
| Std Dev. |  | 0.67132 | 0.57721 | 0.66802 | 0.72181 | 0.41191 | 0.5164 | 0.34496 | 0.45386 | 0.12247 | 0.46224 | 0.624 | 0.28752 | 0.20677 | 0.39833 | 0.51284 | 0.41793 | 0.75631 | 2.2286
| Rel.Std.Dev. |  | 2.304% | 1.915% | 2.289% | 1.464% | 1.429% | 1.702% | 1.299% | 1.535% | 0.4145% | 1.604% | 2.326% | 0.9265% | 0.6535% | 1.421% | 1.759% | 1.324% | 2.488% | 7.729%
| PDM3 |  | -0.848% | 2.586% | -0.6697% | 67.79% | -1.926% | 3.236% | -9.64% | 0.6309% | 0.5701% | -1.926% | -8.68% | 5.618% | 7.689% | -4.592% | -0.7913% | 7.434% | 3.463% | -1.869%
| Z-Score (Absolute) |  | 0.18 | 0.6 | 0.14 | 15.45 | 0.43 | 0.75 | 2.18 | 0.16 | 0.14 | 0.43 | 1.97 | 1.29 | 1.76 | 1.03 | 0.17 | 1.7 | 0.8 | 0.41
| Analytical results for Lu in OREAS 266 (Certified Value 0.24 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.25 | 0.25984 | 0.26619 | 0.2 | 0.24 | 0.117 | 0.19 | [11 empty] | 1
| 1 | 2 | 0.26 | 0.24282 | 0.25423 | 0.2 | 0.23 | 0.1064 | 0.19 | [11 empty] | 4
| 1 | 3 | 0.24 | 0.22981 | 0.24564 | 0.3 | 0.24 | 0.1173 | 0.19 | [11 empty] | 16
| 1 | 4 | 0.26 | 0.2247 | 0.24721 | 0.3 | 0.24 | 0.1187 | 0.19 | [11 empty] | 0.2362
| 1 | 5 | 0.24 | 0.22367 | 0.22887 | 0.3 | 0.23 | 0.1316 | 0.2 | [11 empty] | 40
| 1 | 6 | 0.26 | 0.23457 | 0.25551 | 0.2 | 0.24 | 0.1044 | 0.2
| Mean |  | 0.25167 | 0.2359 | 0.24961 | 0.25 | 0.23667 | 0.1159 | 0.19333
| Median |  | 0.255 | 0.23219 | 0.25072 | 0.25 | 0.24 | 0.11715 | 0.19
| Std Dev. |  | 0.0098319 | 0.013673 | 0.01251 | 0.054772 | 0.005164 | 0.009798 | 0.005164
| Rel.Std.Dev. |  | 3.907% | 5.796% | 5.012% | 21.91% | 2.182% | 8.454% | 2.671%
| PDM3 |  | 6.55% | -0.1247% | 5.679% | 5.844% | 0.1992% | -50.93% | -18.15%
| Z-Score (Absolute) |  | 0.76 | 0.04 | 0.65 | 0.67 | 0 | 6.11 | 2.19
| Analytical results for Mg in OREAS 266 (Certified Value 1.76 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.7564 | 1.9515 | 1.7417 | 1.9325 | 1.82 | 1.68 | 1.68 | 1.77 | 1.74 | 1.77 | 1.7283 | 1.78 | 1.8174 | 1.81 | 1.67 | 1.81 | 1.79 | 1.676 | 1
| 1 | 2 | 1.7664 | 1.9718 | 1.7325 | 1.8909 | 1.84 | 1.79 | 1.69 | 1.77 | 1.73 | 1.78 | 1.7172 | 1.78 | 1.8549 | 1.8 | 1.64 | 1.78 | 1.76 | 1.667 | #N/A
| 1 | 3 | 1.7822 | 1.9392 | 1.7742 | 1.8784 | 1.81 | 1.84 | 1.7 | 1.75 | 1.75 | 1.79 | 1.7334 | 1.77 | 1.807 | 1.79 | 1.62 | 1.75 | 1.69 | 1.652 | 16
| 1 | 4 | 1.7847 | 1.9148 | 1.7585 | 1.8854 | 1.83 | 1.8 | 1.71 | 1.75 | 1.76 | 1.73 | 1.7396 | 1.79 | 1.8003 | 1.84 | 1.7 | 1.75 | 1.65 | 1.576 | 1.7628
| 1 | 5 | 1.7591 | 1.9306 | 1.8203 | 1.8886 | 1.85 | 1.77 | 1.69 | 1.77 | 1.75 | 1.79 | 1.7894 | 1.78 | 1.8291 | 1.8 | 1.65 | 1.77 | 1.75 | 1.588 | 41
| 1 | 6 | 1.7732 | 1.9464 | 1.759 | 1.9036 | 1.83 | 1.86 | 1.7 | 1.76 | 1.72 | 1.79 | 1.75 | 1.83 | 1.8479 | 1.82 | 1.66 | 1.78 | 1.72 | 1.624
| Mean |  | 1.7703 | 1.9424 | 1.7644 | 1.8966 | 1.83 | 1.79 | 1.695 | 1.7617 | 1.7417 | 1.775 | 1.743 | 1.7883 | 1.8261 | 1.81 | 1.6567 | 1.7733 | 1.7267 | 1.6305
| Median |  | 1.7698 | 1.9428 | 1.7588 | 1.8897 | 1.83 | 1.795 | 1.695 | 1.765 | 1.745 | 1.785 | 1.7365 | 1.78 | 1.8232 | 1.805 | 1.655 | 1.775 | 1.735 | 1.638
| Std Dev. |  | 0.011758 | 0.019329 | 0.031062 | 0.019451 | 0.014142 | 0.063246 | 0.010488 | 0.0098319 | 0.01472 | 0.023452 | 0.025228 | 0.02137 | 0.021984 | 0.017889 | 0.027325 | 0.022509 | 0.050859 | 0.041674
| Rel.Std.Dev. |  | 0.6642% | 0.9951% | 1.76% | 1.026% | 0.7728% | 3.533% | 0.6188% | 0.5581% | 0.8451% | 1.321% | 1.447% | 1.195% | 1.204% | 0.9883% | 1.649% | 1.269% | 2.946% | 2.556%
| PDM3 |  | 0.427% | 10.19% | 0.08892% | 7.588% | 3.812% | 1.543% | -3.846% | -0.06462% | -1.199% | 0.6918% | -1.125% | 1.448% | 3.59% | 2.677% | -6.021% | 0.5972% | -2.05% | -7.505%
| Z-Score (Absolute) |  | 0.03 | 3.37 | 0.15 | 2.46 | 1.15 | 0.36 | 1.52 | 0.2 | 0.6 | 0.06 | 0.57 | 0.33 | 1.07 | 0.75 | 2.27 | 0.03 | 0.89 | 2.79
| Analytical results for Mn in OREAS 266 (Certified Value 0.052 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0526 | 0.050871 | 0.051266 | < 0.1 | 0.0512 | 0.0495 | 0.0518 | 0.0523 | 0.0508 | 0.0504 | 0.044536 | 0.0524 | 0.053615 | 0.0519 | 0.0518 | 0.0546 | 0.052 | 0.0476 | 1
| 1 | 2 | 0.0532 | 0.051032 | 0.050507 | < 0.1 | 0.0507 | 0.0498 | 0.0522 | 0.0523 | 0.0507 | 0.0511 | 0.044431 | 0.0523 | 0.055434 | 0.0516 | 0.051 | 0.0542 | 0.0508 | 0.0472 | 23
| 1 | 3 | 0.0534 | 0.050663 | 0.052236 | < 0.1 | 0.0506 | 0.0507 | 0.0519 | 0.0515 | 0.0526 | 0.0504 | 0.045074 | 0.0513 | 0.053932 | 0.0524 | 0.0504 | 0.0534 | 0.051 | 0.0469 | 16
| 1 | 4 | 0.0538 | 0.050323 | 0.051627 | < 0.1 | 0.0508 | 0.0497 | 0.0516 | 0.0526 | 0.0521 | 0.0498 | 0.04475 | 0.0524 | 0.055177 | 0.0526 | 0.0527 | 0.0532 | 0.0501 | 0.0444 | 0.051686
| 1 | 5 | 0.0529 | 0.050733 | 0.053377 | < 0.1 | 0.0505 | 0.0498 | 0.0522 | 0.0526 | 0.0521 | 0.051 | 0.04604 | 0.052 | 0.054306 | 0.0512 | 0.0511 | 0.0535 | 0.0508 | 0.0446 | 42
| 1 | 6 | 0.0532 | 0.051596 | 0.052292 | < 0.1 | 0.0518 | 0.055 | 0.0519 | 0.0524 | 0.0507 | 0.0507 | 0.044755 | 0.0532 | 0.053764 | 0.0517 | 0.0515 | 0.0544 | 0.0518 | 0.0459
| Mean |  | 0.053183 | 0.05087 | 0.051884 |  | 0.050933 | 0.05075 | 0.051933 | 0.052283 | 0.0515 | 0.050567 | 0.044931 | 0.052267 | 0.054371 | 0.0519 | 0.051417 | 0.053883 | 0.051083 | 0.0461
| Median |  | 0.0532 | 0.050802 | 0.051931 |  | 0.05075 | 0.0498 | 0.0519 | 0.05235 | 0.05145 | 0.05055 | 0.044753 | 0.05235 | 0.054119 | 0.0518 | 0.0513 | 0.05385 | 0.0509 | 0.0464
| Std Dev. |  | 0.00041191 | 0.00042769 | 0.00098603 |  | 0.00048854 | 0.002123 | 0.00023381 | 0.00040702 | 0.00086023 | 0.0004761 | 0.0005866 | 0.0006186 | 0.00076392 | 0.00052154 | 0.00078846 | 0.00058793 | 0.00070545 | 0.0013624
| Rel.Std.Dev. |  | 0.7745% | 0.8408% | 1.9% |  | 0.9592% | 4.183% | 0.4502% | 0.7785% | 1.67% | 0.9415% | 1.306% | 1.184% | 1.405% | 1.005% | 1.533% | 1.091% | 1.381% | 2.955%
| PDM3 |  | 2.897% | -1.579% | 0.3836% |  | -1.456% | -1.811% | 0.4785% | 1.156% | -0.3598% | -2.166% | -13.07% | 1.123% | 5.195% | 0.4141% | -0.5211% | 4.251% | -1.166% | -10.81%
| Z-Score (Absolute) |  | 1.53 | 0.52 | 0.38 | 1.3 | 0.47 | 0.63 | 0.42 | 0.73 | 0.04 | 0.79 | 5.8 | 0.72 | 2.59 | 0.39 | 0.04 | 2.16 | 0.33 | 4.76
| Analytical results for Mo in OREAS 266 (Certified Value 9.31 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 9.3 | 9.4215 | 9.7906 | 9.6 | 6 | 9.85 | 9.0333 | 9.31 | 9.2 | 9.7483 | 8.22 | 8.8894 | 9.07 | 8.6 | 10.3 | 9.4 | 9.48 |  | 1
| 1 | 2 | 9.5 | 9.2669 | 10.2155 | 9.5 | 2 | 9.95 | 9.1447 | 9.43 | 9.1 | 9.7944 | 8.25 | 8.4266 | 8.85 | 8.31 | 10 | 9.16 | 9.79 |  | 24
| 1 | 3 | 9.5 | 9.4971 | 10.111 | 9.5 | 2 | 9.91 | 9.0525 | 9 | 8.8 | 9.9239 | 7.96 | 8.5463 | 8.98 | 8.23 | 10 | 9.63 | 9.31 |  | 16
| 1 | 4 | 9.4 | 9.4156 | 10.2313 | 9.6 | 3 | 9.96 | 8.9337 | 8.88 | 9 | 9.7747 | 8.31 | 9.1008 | 8.91 | 8.43 | 10 | 9.43 | 9.22 |  | 9.3148
| 1 | 5 | 9 | 9.7205 | 10.4938 | 9.5 | 4 | 9.89 | 8.96 | 9.23 | 8.9 | 9.8638 | 8.93 | 8.8679 | 8.8 | 8.35 | 9.89 | 8.75 | 9.35 |  | 43
| 1 | 6 | 9.7 | 9.5334 | 9.9241 | 9.4 | 2 | 9.86 | 9.2675 | 9.28 | 9.2 | 9.7717 | 8.86 | 9.0016 | 8.94 | 8.66 | 10.25 | 10.2 | 9.43
| Mean |  | 9.4 | 9.4758 | 10.1277 | 9.5167 | 3.1667 | 9.9033 | 9.0653 | 9.1883 | 9.0333 | 9.8128 | 8.4217 | 8.8055 | 8.925 | 8.43 | 10.0733 | 9.4283 | 9.43
| Median |  | 9.45 | 9.4593 | 10.1633 | 9.5 | 2.5 | 9.9 | 9.0429 | 9.255 | 9.05 | 9.7845 | 8.28 | 8.8787 | 8.925 | 8.39 | 10 | 9.415 | 9.39
| Std Dev. |  | 0.23664 | 0.15089 | 0.24813 | 0.075277 | 1.6021 | 0.045461 | 0.12392 | 0.20683 | 0.1633 | 0.067207 | 0.38634 | 0.26358 | 0.095656 | 0.16888 | 0.16269 | 0.48355 | 0.19849
| Rel.Std.Dev. |  | 2.517% | 1.592% | 2.45% | 0.791% | 50.59% | 0.459% | 1.367% | 2.251% | 1.808% | 0.6849% | 4.587% | 2.993% | 1.072% | 2.003% | 1.615% | 5.129% | 2.105%
| PDM3 |  | 0.9145% | 1.729% | 8.727% | 2.167% | -66% | 6.318% | -2.679% | -1.358% | -3.022% | 5.346% | -9.589% | -5.468% | -4.185% | -9.499% | 8.143% | 1.219% | 1.237%
| Z-Score (Absolute) |  | 0.02 | 0.11 | 1.23 | 0.18 | NA | 0.85 | 0.6 | 0.39 | 0.66 | 0.69 | 1.72 | 1.05 | 0.85 | 1.7 | 1.14 | 0.02 | 0.03
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 266 (Certified Value 1.15 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.1243 | 1.0217 | 1.1341 | 2.9325 | 1.13 | 1.1 | 1.11 | 1.2 | 1.12 | 1.184 | 1.1683 | 1.15 | 1.1995 | 1.16 | 1.1 | 1.18 | 1.16 | 1.054 | 1
| 1 | 2 | 1.136 | 1.0286 | 1.1064 | 2.8957 | 1.14 | 1.12 | 1.14 | 1.2 | 1.12 | 1.183 | 1.1633 | 1.14 | 1.2209 | 1.13 | 1.09 | 1.16 | 1.15 | 1.074 | #N/A
| 1 | 3 | 1.144 | 1.0178 | 1.16 | 2.7096 | 1.14 | 1.14 | 1.13 | 1.21 | 1.13 | 1.151 | 1.1715 | 1.13 | 1.1845 | 1.15 | 1.07 | 1.15 | 1.13 | 1.086 | 16
| 1 | 4 | 1.1491 | 1.0071 | 1.1497 | 2.7789 | 1.13 | 1.14 | 1.14 | 1.18 | 1.13 | 1.169 | 1.1851 | 1.15 | 1.1597 | 1.18 | 1.12 | 1.15 | 1.11 | 1.095 | 1.1464
| 1 | 5 | 1.1329 | 1.0177 | 1.1821 | 2.8682 | 1.13 | 1.12 | 1.14 | 1.22 | 1.12 | 1.181 | 1.2079 | 1.14 | 1.2101 | 1.15 | 1.09 | 1.16 | 1.14 | 1.144 | 44
| 1 | 6 | 1.1366 | 1.0359 | 1.1473 | 2.977 | 1.16 | 1.16 | 1.14 | 1.2 | 1.12 | 1.184 | 1.1829 | 1.17 | 1.1659 | 1.16 | 1.1 | 1.16 | 1.15 | 1.04
| Mean |  | 1.1371 | 1.0215 | 1.1466 | 2.8603 | 1.1383 | 1.13 | 1.1333 | 1.2017 | 1.1233 | 1.1753 | 1.1798 | 1.1467 | 1.1901 | 1.155 | 1.095 | 1.16 | 1.14 | 1.0822
| Median |  | 1.1363 | 1.0198 | 1.1485 | 2.8819 | 1.135 | 1.13 | 1.14 | 1.2 | 1.12 | 1.182 | 1.1772 | 1.145 | 1.192 | 1.155 | 1.095 | 1.16 | 1.145 | 1.08
| Std Dev. |  | 0.0086549 | 0.0099333 | 0.025373 | 0.099504 | 0.01169 | 0.020976 | 0.012111 | 0.013292 | 0.005164 | 0.013216 | 0.016129 | 0.013663 | 0.024397 | 0.016432 | 0.016432 | 0.010954 | 0.017889 | 0.036444
| Rel.Std.Dev. |  | 0.7611% | 0.9725% | 2.213% | 3.479% | 1.027% | 1.856% | 1.069% | 1.106% | 0.4597% | 1.124% | 1.367% | 1.192% | 2.05% | 1.423% | 1.501% | 0.9443% | 1.569% | 3.368%
| PDM3 |  | -0.8098% | -10.9% | 0.01462% | 149.5% | -0.7066% | -1.434% | -1.143% | 4.818% | -2.015% | 2.521% | 2.913% | 0.02027% | 3.809% | 0.7472% | -4.486% | 1.183% | -0.5612% | -5.606%
| Z-Score (Absolute) |  | 0.23 | 4.48 | 0.12 | 63.15 | 0.18 | 0.49 | 0.37 | 2.15 | 0.73 | 1.18 | 1.34 | 0.12 | 1.72 | 0.43 | 1.78 | 0.61 | 0.12 | 2.25
| Analytical results for Nb in OREAS 266 (Certified Value 32.4 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 31.64 | 38.1892 | 115.0191 | 31.4 | 9.9 | 28 | 29.4577 | 30.9 | 33.1 | 43.3637 | 29.5 | 36.2306 | 31.8 | 33.4 | 34.7 | 34.2 | 32.9 |  | 1
| 1 | 2 | 31.83 | 39.6273 | 73.389 | 31 | 0.4 | 27.8 | 29.7123 | 30.4 | 34 | 43.0149 | 30.1 | 36.9815 | 31.2 | 32.6 | 33.7 | 31.9 | 33.1 |  | 25
| 1 | 3 | 31.61 | 37.5973 | 185.1046 | 30.3 | 0.3 | 28.3 | 29.0369 | 30.9 | 33.8 | 44.2909 | 29.1 | 37.1508 | 31.8 | 32.3 | 35.2 | 31.2 | 33.1 |  | 16
| 1 | 4 | 32.41 | 37.6316 | 68.8008 | 31 | 0.7 | 26.9 | 29.509 | 31.6 | 33.3 | 44.0467 | 30 | 37.0296 | 32.2 | 33.3 | 33.3 | 32 | 32.6 |  | 32.3713
| 1 | 5 | 31.28 | 35.8332 | 64.6085 | 31.3 | 0.5 | 26.4 | 30.6317 | 31.9 | 32.1 | 45.4784 | 31.5 | 36.8293 | 31.8 | 32.3 | 35.1 | 30.8 | 32.3 |  | 45
| 1 | 6 | 31.72 | 38.0506 | 70.72 | 31.4 | 0.4 | 26.5 | 29.6438 | 30.8 | 33.7 | 42.3276 | 31.9 | 36.2155 | 31.8 | 33.1 | 35.8 | 33.3 | 32.8
| Mean |  | 31.7483 | 37.8215 | 96.2737 | 31.0667 | 2.0333 | 27.3167 | 29.6652 | 31.0833 | 33.3333 | 43.7537 | 30.35 | 36.7395 | 31.7667 | 32.8333 | 34.6333 | 32.2333 | 32.8
| Median |  | 31.68 | 37.8411 | 72.0545 | 31.15 | 0.45 | 27.35 | 29.5764 | 30.9 | 33.5 | 43.7052 | 30.05 | 36.9054 | 31.8 | 32.85 | 34.9 | 31.95 | 32.85
| Std Dev. |  | 0.37296 | 1.2238 | 47.2789 | 0.41793 | 3.8563 | 0.81833 | 0.52897 | 0.55648 | 0.68896 | 1.1028 | 1.1131 | 0.41316 | 0.32042 | 0.49666 | 0.95429 | 1.2879 | 0.30984
| Rel.Std.Dev. |  | 1.175% | 3.236% | 49.11% | 1.345% | 189.7% | 2.996% | 1.783% | 1.79% | 2.067% | 2.52% | 3.668% | 1.125% | 1.009% | 1.513% | 2.755% | 3.996% | 0.9446%
| PDM3 |  | -1.924% | 16.84% | 197.4% | -4.03% | -93.72% | -15.61% | -8.359% | -3.979% | 2.972% | 35.16% | -6.244% | 13.49% | -1.868% | 1.427% | 6.988% | -0.4262% | 1.324%
| Z-Score (Absolute) |  | 0.17 | 2 | 22.93 | 0.42 | 10.81 | 1.76 | 0.92 | 0.41 | 0.39 | 4.12 | 0.67 | 1.61 | 0.17 | 0.21 | 0.86 | 0 | 0.2
| Analytical results for Nd in OREAS 266 (Certified Value 37.4 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.14 | 1.22 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 1
1
| 1 | 1 | 38.48 | 38.3676 | 32.3 | 36.9 | 29.7786 | 44 | [12 empty] | 1
| 1 | 2 | 37.29 | 37.4613 | 36.5 | 36.6 | 29.4233 | 44.4 | [12 empty] | 5
| 1 | 3 | 37.44 | 37.6355 | 36.3 | 37.5 | 30.0773 | 46.5 | [12 empty] | 16
| 1 | 4 | 37.58 | 37.8094 | 36.2 | 37.2 | 29.9738 | 45.5 | [12 empty] | 37.3789
| 1 | 5 | 37.94 | 38.5108 | 36.8 | 36.6 | 31.1782 | 43.3 | [12 empty] | 46
| 1 | 6 | 37.69 | 37.7336 | 37.6 | 37.3 | 29.5912 | 45.8
| Mean |  | 37.7367 | 37.9197 | 35.95 | 37.0167 | 30.0037 | 44.9167
| Median |  | 37.635 | 37.7715 | 36.4 | 37.05 | 29.8762 | 44.95
| Std Dev. |  | 0.42618 | 0.42134 | 1.8577 | 0.37639 | 0.62345 | 1.2123
| Rel.Std.Dev. |  | 1.129% | 1.111% | 5.167% | 1.017% | 2.078% | 2.699%
| PDM3 |  | 0.9571% | 1.447% | -3.823% | -0.9691% | -19.73% | 20.17%
| Z-Score (Absolute) |  | 0.25 | 0.37 | 0.98 | 0.25 | 5.05 | 5.16
| Analytical results for Ni in OREAS 266 (Certified Value 95 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 96.7 | 96.4332 | 80.5826 | 93 | 95 | 95 | 89.5533 | 93.6 | 90.3 | 75.4469 | 87.8 | 98.4755 | 96.7 | 98.5 | 102.5 | 98.4 | 100 |  | 1
| 1 | 2 | 98.2 | 95.3317 | 79.4578 | 92 | 97 | 94 | 89.76 | 92.9 | 91.8 | 75.7722 | 89.9 | 97.3032 | 94.7 | 95.1 | 99.4 | 94.7 | 101 |  | 27
| 1 | 3 | 98.1 | 98.6965 | 80.3647 | 92 | 98 | 95 | 87.7 | 93.8 | 89.1 | 75.678 | 87.9 | 94.7226 | 97.4 | 96.4 | 96.8 | 95.8 | 100 |  | 16
| 1 | 4 | 99.6 | 96.9076 | 79.9762 | 92 | 97 | 94 | 87.7533 | 94.7 | 89.7 | 75.909 | 91.4 | 94.1113 | 99.2 | 99.5 | 97.4 | 95 | 101 |  | 95.4208
| 1 | 5 | 96.7 | 100.8478 | 78.4222 | 92 | 98 | 95 | 88.6333 | 95.8 | 92.1 | 75.5745 | 94.5 | 98.8003 | 96.9 | 96 | 96.8 | 93.3 | 100 |  | 47
| 1 | 6 | 100.5 | 98.1275 | 77.4287 | 92 | 97 | 94 | 90.4067 | 93.3 | 93.5 | 75.755 | 95.9 | 98.0123 | 97.3 | 99.5 | 98.7 | 99.5 | 99
| Mean |  | 98.3 | 97.724 | 79.372 | 92.1667 | 97 | 94.5 | 88.9678 | 94.0167 | 91.0833 | 75.6892 | 91.2333 | 96.9042 | 97.0333 | 97.5 | 98.6 | 96.1167 | 100.1667
| Median |  | 98.15 | 97.5176 | 79.717 | 92 | 97 | 94.5 | 89.0933 | 93.7 | 91.05 | 75.7165 | 90.65 | 97.6577 | 97.1 | 97.45 | 98.05 | 95.4 | 100
| Std Dev. |  | 1.5297 | 1.9454 | 1.2249 | 0.40825 | 1.0954 | 0.54772 | 1.1167 | 1.061 | 1.6642 | 0.16219 | 3.3809 | 2.0005 | 1.4473 | 1.9089 | 2.1808 | 2.3642 | 0.75277
| Rel.Std.Dev. |  | 1.556% | 1.991% | 1.543% | 0.4429% | 1.129% | 0.5796% | 1.255% | 1.128% | 1.827% | 0.2143% | 3.706% | 2.064% | 1.492% | 1.958% | 2.212% | 2.46% | 0.7515%
| PDM3 |  | 3.017% | 2.414% | -16.82% | -3.41% | 1.655% | -0.965% | -6.763% | -1.472% | -4.546% | -20.68% | -4.388% | 1.555% | 1.69% | 2.179% | 3.332% | 0.7292% | 4.974%
| Z-Score (Absolute) |  | 0.67 | 0.5 | 5.17 | 1.22 | 0.27 | 0.5 | 2.21 | 0.65 | 1.55 | 6.31 | 1.51 | 0.24 | 0.28 | 0.43 | 0.77 | 0 | 1.25
| Analytical results for P in OREAS 266 (Certified Value 0.103 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.1007 | 0.10173 | 0.10163 | 0.106 | 0.099 | 0.1088 | 0.103 | 0.103 | 0.105 | 0.095261 | 0.105 | 0.10855 | 0.11 | 0.099 | 0.109 | 0.102 | 0.104 |  | 1
| 1 | 2 | 0.1016 | 0.099647 | < 0.1 | 0.106 | 0.094 | 0.1089 | 0.1038 | 0.102 | 0.105 | 0.094862 | 0.104 | 0.11097 | 0.1 | 0.098 | 0.106 | 0.102 | 0.108 |  | #N/A
| 1 | 3 | 0.1026 | 0.10341 | 0.10131 | 0.106 | 0.085 | 0.1094 | 0.1029 | 0.101 | 0.105 | 0.095961 | 0.105 | 0.11101 | 0.11 | 0.098 | 0.104 | 0.102 | 0.101 |  | 16
| 1 | 4 | 0.1025 | 0.10208 | 0.10025 | 0.106 | 0.094 | 0.1086 | 0.0998 | 0.101 | 0.104 | 0.09626 | 0.106 | 0.10744 | 0.11 | 0.1 | 0.105 | 0.097 | 0.098 |  | 0.10349
| 1 | 5 | 0.1025 | 0.10655 | < 0.1 | 0.106 | 0.091 | 0.1085 | 0.1015 | 0.098 | 0.105 | 0.097856 | 0.105 | 0.10956 | 0.11 | 0.095 | 0.104 | 0.102 | 0.097 |  | 48
| 1 | 6 | 0.1017 | 0.10332 | < 0.1 | 0.108 | 0.093 | 0.1083 | 0.1036 | 0.101 | 0.105 | 0.095761 | 0.107 | 0.10792 | 0.11 | 0.097 | 0.106 | 0.103 | 0.101
| Mean |  | 0.10193 | 0.10279 | 0.10106 | 0.10633 | 0.092667 | 0.10875 | 0.10243 | 0.101 | 0.10483 | 0.095994 | 0.10533 | 0.10924 | 0.10833 | 0.097833 | 0.10567 | 0.10133 | 0.1015
| Median |  | 0.1021 | 0.1027 | 0.10131 | 0.106 | 0.0935 | 0.1087 | 0.10295 | 0.101 | 0.105 | 0.095861 | 0.105 | 0.10905 | 0.11 | 0.098 | 0.1055 | 0.102 | 0.101
| Std Dev. |  | 0.00074476 | 0.0022933 | 0.00072039 | 0.0008165 | 0.0045898 | 0.00038341 | 0.0015214 | 0.0016733 | 0.00040825 | 0.0010404 | 0.0010328 | 0.0015304 | 0.0040825 | 0.0017224 | 0.0018619 | 0.0021602 | 0.0040373
| Rel.Std.Dev. |  | 0.7306% | 2.231% | 0.7128% | 0.7679% | 4.953% | 0.3526% | 1.485% | 1.657% | 0.3894% | 1.084% | 0.9805% | 1.401% | 3.768% | 1.761% | 1.762% | 2.132% | 3.978%
| PDM3 |  | -1.5% | -0.6718% | -2.34% | 2.752% | -10.45% | 5.087% | -1.017% | -2.402% | 1.303% | -7.239% | 1.786% | 5.561% | 4.685% | -5.462% | 2.108% | -2.079% | -1.918%
| Z-Score (Absolute) |  | 0.1 | 0.07 | 5.61 | 0.81 | 2.04 | 1.32 | 0 | 0.3 | 0.5 | 1.34 | 0.6 | 1.42 | 1.23 | 0.96 | 0.67 | 0.23 | 0.19
| Analytical results for Pb in OREAS 266 (Certified Value 15.2 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 15.4 | 16.6411 | 6.8406 | 17 | 11 | 11.1 | 13.9961 | 14.6 | 14.8 | 11.2532 | 14.1 | 14.7585 | 15.2 | 13.7 | 15 | 16 | 16.2 |  | 1
| 1 | 2 | 15.2 | 15.8609 | 8.6562 | 17 | 14 | 11.1 | 12.4378 | 14.8 | 17 | 11.1441 | 14.5 | 14.811 | 15 | 13.3 | 15.2 | 15.4 | 17.1 |  | 28
| 1 | 3 | 15.8 | 17.012 | 7.5525 | 17 | 15 | 11 | 12.9771 | 15.2 | 16.7 | 11.3308 | 13.6 | 15.3425 | 14.8 | 13.6 | 15.8 | 15.6 | 17.9 |  | 16
| 1 | 4 | 16 | 16.4837 | 8.9889 | 17 | 13 | 11.1 | 13.5196 | 14.9 | 15.8 | 11.7196 | 14.5 | 15.5345 | 15.2 | 13.9 | 15.1 | 14.6 | 17.7 |  | 15.1841
| 1 | 5 | 15.8 | 16.373 | 9.3453 | 17 | 14 | 11.4 | 12.2173 | 15.2 | 14.5 | 11.0546 | 15.4 | 15.4804 | 15.4 | 13.5 | 15.5 | 15.2 | 16.1 |  | 49
| 1 | 6 | 16.2 | 16.4739 | 8.9296 | 17 | 13 | 10.9 | 12.7392 | 14.5 | 16.8 | 11.1931 | 15.1 | 14.993 | 15 | 13.5 | 15.9 | 15.8 | 17
| Mean |  | 15.7333 | 16.4741 | 8.3855 | 17 | 13.3333 | 11.1 | 12.9812 | 14.8667 | 15.9333 | 11.2826 | 14.5333 | 15.1533 | 15.1 | 13.5833 | 15.4167 | 15.4333 | 17
| Median |  | 15.8 | 16.4788 | 8.7929 | 17 | 13.5 | 11.1 | 12.8581 | 14.85 | 16.25 | 11.2231 | 14.5 | 15.1677 | 15.1 | 13.55 | 15.35 | 15.5 | 17.05
| Std Dev. |  | 0.37238 | 0.37511 | 0.97316 | 0 | 1.3663 | 0.16733 | 0.67192 | 0.29439 | 1.0801 | 0.2338 | 0.6532 | 0.34259 | 0.20976 | 0.20412 | 0.37639 | 0.49666 | 0.74297
| Rel.Std.Dev. |  | 2.367% | 2.277% | 11.61% | 0% | 10.25% | 1.507% | 5.176% | 1.98% | 6.779% | 2.072% | 4.494% | 2.261% | 1.389% | 1.503% | 2.441% | 3.218% | 4.37%
| PDM3 |  | 3.618% | 8.496% | -44.77% | 11.96% | -12.19% | -26.9% | -14.51% | -2.09% | 4.935% | -25.69% | -4.286% | -0.2025% | -0.5535% | -10.54% | 1.532% | 1.642% | 11.96%
| Z-Score (Absolute) |  | 0.51 | 1.11 | 5.43 | NA | NA | 3.24 | 1.71 | 0.19 | 0.67 | 3.09 | 0.46 | 0.04 | 0 | 1.23 | 0.26 | 0.27 | 1.54
Results from laboratories 1.07 and 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 266 (Certified Value 9.92 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 10.16 | 10.3073 | 10.1526 | 8.2 | 9.52 | 4.995 | 12.07 | [11 empty] | 1
| 1 | 2 | 10.1 | 9.8572 | 11.0151 | 9.4 | 9.37 | 4.9753 | 11.55 | [11 empty] | 7
| 1 | 3 | 9.99 | 9.8507 | 10.2498 | 9.4 | 9.54 | 5.0858 | 11.87 | [11 empty] | 16
| 1 | 4 | 10.32 | 9.8509 | 10.5463 | 9.4 | 9.59 | 5.0847 | 12.24 | [11 empty] | 9.9176
| 1 | 5 | 10.52 | 10.0341 | 9.6524 | 9.6 | 9.44 | 5.0609 | 11.76 | [11 empty] | 50
| 1 | 6 | 10.32 | 10.149 | 10.3715 | 9.8 | 9.5 | 5.0555 | 12.29
| Mean |  | 10.235 | 10.0082 | 10.3313 | 9.3 | 9.4933 | 5.0429 | 11.9633
| Median |  | 10.24 | 9.9456 | 10.3106 | 9.4 | 9.51 | 5.0582 | 11.97
| Std Dev. |  | 0.1895 | 0.19096 | 0.45053 | 0.56214 | 0.077889 | 0.046743 | 0.28828
| Rel.Std.Dev. |  | 1.851% | 1.908% | 4.361% | 6.045% | 0.8205% | 0.9269% | 2.41%
| PDM3 |  | 3.201% | 0.9137% | 4.172% | -6.227% | -4.278% | -49.15% | 20.63%
| Z-Score (Absolute) |  | 0.3 | 0 | 0.42 | 0.93 | 0.67 | 6.5 | 2.56
| Analytical results for Rb in OREAS 266 (Certified Value 78 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 79.59 | 81.2746 | 78.6 | 49.5 | 77.2 | 76.3 | 74.4 | 49.1578 | 69.1 | 77.7972 | 73.3 | 77.4 | 83.2 | 78.2 | 91 |  |  |  | 1
| 1 | 2 | 79.68 | 79.3122 | 77.6 | 59.1 | 78 | 76.4 | 63 | 49.8877 | 70.9 | 78.9933 | 70.7 | 77.5 | 80.3 | 75.8 | 90.9 |  |  |  | 30
| 1 | 3 | 77.77 | 83.604 | 77.7 | 63.1 | 77.9 | 75.6 | 73.4 | 49.1151 | 68.6 | 79.4241 | 72.3 | 74.9 | 82.3 | 78.2 | 93.3 |  |  |  | 16
| 1 | 4 | 81.68 | 81.6221 | 77.9 | 61.7 | 79.4 | 76.1 | 67 | 50.5512 | 71.5 | 78.9881 | 74.2 | 78.8 | 78.7 | 74.8 | 91.3 |  |  |  | 78.0318
| 1 | 5 | 78.53 | 84.379 | 77.5 | 63.9 | 77.8 | 77.3 | 67.9 | 55.9996 | 74.9 | 79.1715 | 72 | 77.8 | 82 | 75.1 | 90.1 |  |  |  | 51
| 1 | 6 | 81.91 | 82.4285 | 78.1 | 65.8 | 77.9 | 76.5 | 70.7 | 50.641 | 74.6 | 77.709 | 72.4 | 77.6 | 82.6 | 81.3 | 91.9
| Mean |  | 79.86 | 82.1034 | 77.9 | 60.5167 | 78.0333 | 76.3667 | 69.4 | 50.8921 | 71.6 | 78.6805 | 72.4833 | 77.3333 | 81.5167 | 77.2333 | 91.4167
| Median |  | 79.635 | 82.0253 | 77.8 | 62.4 | 77.9 | 76.35 | 69.3 | 50.2195 | 71.2 | 78.9907 | 72.35 | 77.55 | 82.15 | 77 | 91.15
| Std Dev. |  | 1.6587 | 1.8025 | 0.40497 | 5.8448 | 0.72847 | 0.55737 | 4.2816 | 2.5861 | 2.6698 | 0.73629 | 1.189 | 1.2956 | 1.6893 | 2.4873 | 1.0926
| Rel.Std.Dev. |  | 2.077% | 2.195% | 0.5199% | 9.658% | 0.9335% | 0.7299% | 6.169% | 5.082% | 3.729% | 0.9358% | 1.64% | 1.675% | 2.072% | 3.221% | 1.195%
| PDM3 |  | 2.343% | 5.218% | -0.169% | -22.45% | 0.001914% | -2.134% | -11.06% | -34.78% | -8.243% | 0.8313% | -7.111% | -0.8952% | 4.466% | -1.023% | 17.15%
| Z-Score (Absolute) |  | 0.41 | 0.77 | 0.09 | 2.71 | 0.11 | 0.16 | 1.28 | 4.26 | 0.92 | 0.22 | 0.78 | 0 | 0.67 | 0.02 | 2.27
| Analytical results for Re in OREAS 266 (Certified Value < 0.002 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0124 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.013 | 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | < 0.002 | [5 empty] | 31
| 1 | 3 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0131 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | 16
| 1 | 4 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.0136 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0145 | 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | 52
| 1 | 6 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.0131 | 0.002 | < 0.05 | < 0.002 | 0.002 | 0.002 | < 0.002
| Mean | [5 empty] | 0.002 |  | 0.013283 | 0.002 |  |  | 0.002 | 0.002 | 0.002
| Median | [5 empty] | 0.002 |  | 0.0131 | 0.002 |  |  | 0.002 | 0.002 | 0.002
| Std Dev. | [5 empty] | 0 |  | 0.00070828 | 0 |  |  | 0 |  | 0
| Rel.Std.Dev. | [5 empty] | 0% |  | 5.332% | 0% |  |  | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 31.69 | 1.35 | 0.67 | 0.22 | 1.35 | 15.89 | 0 | 31.69 | 0.67 | 0.67 | 0.45 | 0.22
| Analytical results for S in OREAS 266 (Certified Value 0.032 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0312 | 0.036018 | 0.4739 | 0.029 | 0.04 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.028126 | 0.03 | 0.031953 | 0.03 | 0.05 | 0.04 | 0.03 | 0.031 |  | 1
| 1 | 2 | 0.0306 | 0.034187 | 0.47135 | 0.029 | 0.04 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.027376 | 0.03 | 0.033144 | 0.03 | 0.04 | 0.04 | 0.03 | 0.03 |  | 16
| 1 | 3 | 0.0308 | 0.03739 | 0.44248 | 0.029 | 0.03 | 0.03 | 0.04 | 0.03 | < 0.1 | 0.028541 | 0.03 | 0.03135 | 0.03 | 0.04 | 0.04 | 0.03 | 0.031 |  | 16
| 1 | 4 | 0.031 | 0.036062 | 0.47252 | 0.029 | 0.03 | 0.03 | 0.03 | 0.03 | < 0.1 | 0.027904 | 0.03 | 0.034069 | 0.033 | 0.05 | 0.04 | 0.03 | 0.029 |  | 0.031599
| 1 | 5 | 0.0304 | 0.040102 | 0.48397 | 0.029 | 0.03 | 0.03 | 0.03 | 0.03 | < 0.1 | 0.028021 | 0.03 | 0.034541 | 0.033 | 0.04 | 0.04 | 0.03 | 0.031 |  | 53
| 1 | 6 | 0.0309 | 0.03778 | 0.48391 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.029234 | 0.03 | 0.032047 | 0.03 | 0.04 | 0.04 | 0.03 | 0.029
| Mean |  | 0.030817 | 0.036923 | 0.47135 | 0.029167 | 0.035 | 0.03 | 0.031667 | 0.035 |  | 0.0282 | 0.03 | 0.032851 | 0.031 | 0.043333 | 0.04 | 0.03 | 0.030167
| Median |  | 0.03085 | 0.036726 | 0.47321 | 0.029 | 0.035 | 0.03 | 0.03 | 0.035 |  | 0.028073 | 0.03 | 0.032596 | 0.03 | 0.04 | 0.04 | 0.03 | 0.0305
| Std Dev. |  | 0.00028577 | 0.0020053 | 0.015221 | 0.00040825 | 0.0054772 | 0 | 0.0040825 | 0.0054772 |  | 0.00063056 | 0 | 0.0012753 | 0.0015492 | 0.005164 | 0 | 0 | 0.00098319
| Rel.Std.Dev. |  | 0.9273% | 5.431% | 3.229% | 1.4% | 15.65% | 0% | 12.89% | 15.65% |  | 2.236% | 0% | 3.882% | 4.997% | 11.92% | 0% | 0% | 3.259%
| PDM3 |  | -2.477% | 16.85% | 1392% | -7.698% | 10.76% | -5.061% | 0.2132% | 10.76% |  | -10.76% | -5.061% | 3.96% | -1.897% | 37.13% | 26.59% | -5.061% | -4.534%
| Z-Score (Absolute) |  | 0.23 | 1.42 | 118.59 | 0.67 | 0.9 | 0.45 | 0 | 0.9 | 4.94 | 0.93 | 0.45 | 0.32 | 0.18 | 3.15 | 2.25 | 0.45 | 0.4
| Analytical results for Sb in OREAS 266 (Certified Value 10.3 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 10.63 | 9.9644 | 8.6504 | 9.9 | 2.4 | 9.44 | 11.3067 | 10.7 | 10.3 | 10.9213 | 9.26 | 10.3094 | 9.9 | 10.7 | 11.1 | 10.15 | 10.22 |  | 1
| 1 | 2 | 10.9 | 9.8342 | 11.407 | 10 | 0.2 | 9.5 | 11.46 | 10.6 | 10.5 | 11.1522 | 9.49 | 11.0423 | 9.77 | 10.35 | 10.75 | 10.05 | 10.34 |  | 32
| 1 | 3 | 10.74 | 10.1639 | 9.0608 | 9.9 | 0.1 | 9.53 | 11.3733 | 9.62 | 10.3 | 11.0843 | 9.01 | 10.6055 | 9.89 | 10.3 | 11.2 | 9.71 | 10.35 |  | 16
| 1 | 4 | 11.17 | 9.821 | 8.8422 | 9.9 | 0.2 | 9.55 | 10.4 | 9.89 | 10.2 | 11.1141 | 9.54 | 11.2849 | 10 | 10.8 | 11.1 | 10 | 10.8 |  | 10.3051
| 1 | 5 | 10.7 | 10.059 | 5.539 | 10.2 | 0.2 | 9.33 | 9.73 | 10.05 | 9.9 | 11.3123 | 10.1 | 10.478 | 10.05 | 10.6 | 11.25 | 9.64 | 10.39 |  | 54
| 1 | 6 | 11.18 | 9.9084 | 10.9283 | 10 | 0.2 | 9.38 | 10.9433 | 10.85 | 10.3 | 10.9065 | 10.05 | 9.9672 | 9.99 | 10.9 | 11.55 | 10.55 | 10.49
| Mean |  | 10.8867 | 9.9585 | 9.0713 | 9.9833 | 0.55 | 9.455 | 10.8689 | 10.285 | 10.25 | 11.0818 | 9.575 | 10.6146 | 9.9333 | 10.6083 | 11.1583 | 10.0167 | 10.4317
| Median |  | 10.82 | 9.9364 | 8.9515 | 9.95 | 0.2 | 9.47 | 11.125 | 10.325 | 10.3 | 11.0992 | 9.515 | 10.5417 | 9.945 | 10.65 | 11.15 | 10.025 | 10.37
| Std Dev. |  | 0.24031 | 0.13357 | 2.079 | 0.1169 | 0.90719 | 0.08735 | 0.68088 | 0.49883 | 0.19748 | 0.15207 | 0.43085 | 0.48235 | 0.10093 | 0.2417 | 0.25965 | 0.32849 | 0.20034
| Rel.Std.Dev. |  | 2.207% | 1.341% | 22.92% | 1.171% | 164.9% | 0.9238% | 6.264% | 4.85% | 1.927% | 1.372% | 4.5% | 4.544% | 1.016% | 2.278% | 2.327% | 3.279% | 1.921%
| PDM3 |  | 5.644% | -3.363% | -11.97% | -3.122% | -94.66% | -8.249% | 5.471% | -0.1948% | -0.5344% | 7.537% | -7.085% | 3.003% | -3.607% | 2.943% | 8.28% | -2.799% | 1.228%
| Z-Score (Absolute) |  | 1.18 | 0.54 | 2.18 | 0.49 | 17.94 | 1.47 | 1.14 | 0.06 | 0 | 1.54 | 1.25 | 0.67 | 0.59 | 0.66 | 1.68 | 0.43 | 0.34
| Analytical results for Sc in OREAS 266 (Certified Value 12.8 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 13.4 | 12.8084 | 12 | 13 | 12.5 | 12.5097 | 12.8 | 13 | 11.8726 | 10.6 | 13.6443 | 12.7 | 13.2 | 13.8 | 12.8 | 14.5 |  |  | 1
| 1 | 2 | 13.4 | 12.5776 | 12 | 14 | 12.2 | 12.5208 | 12.8 | 13 | 11.9229 | 10.8 | 13.7315 | 12.4 | 12.1 | 13.4 | 12.2 | 14.4 |  |  | 33
| 1 | 3 | 13.3 | 13.2686 | 12 | 14 | 12.5 | 11.9954 | 12.5 | 13 | 11.9024 | 10.4 | 13.6878 | 12.5 | 12 | 14.4 | 12.7 | 14.2 |  |  | 16
| 1 | 4 | 13.7 | 12.9111 | 12 | 13 | 12.4 | 12.0953 | 12.6 | 12 | 11.9422 | 11 | 13.2078 | 12.8 | 12.5 | 13.2 | 12.6 | 14.6 |  |  | 12.8293
| 1 | 5 | 13.3 | 13.3686 | 12 | 13 | 12.2 | 12.1878 | 12.6 | 13 | 12.2001 | 11.6 | 13.817 | 12.6 | 12.4 | 14.7 | 12.2 | 14.5 |  |  | 55
| 1 | 6 | 13.9 | 13.1375 | 12 | 14 | 12.1 | 12.4468 | 13 | 13 | 11.813 | 11.6 | 13.1622 | 12.5 | 12.8 | 14.7 | 13.6 | 14.8
| Mean |  | 13.5 | 13.012 | 12 | 13.5 | 12.3167 | 12.2926 | 12.7167 | 12.8333 | 11.9422 | 11 | 13.5418 | 12.5833 | 12.5 | 14.0333 | 12.6833 | 14.5
| Median |  | 13.4 | 13.0243 | 12 | 13.5 | 12.3 | 12.3173 | 12.7 | 13 | 11.9126 | 10.9 | 13.666 | 12.55 | 12.45 | 14.1 | 12.65 | 14.5
| Std Dev. |  | 0.24495 | 0.29971 | 0 | 0.54772 | 0.17224 | 0.22857 | 0.18348 | 0.40825 | 0.13419 | 0.50596 | 0.28257 | 0.1472 | 0.44721 | 0.65929 | 0.51543 | 0.2
| Rel.Std.Dev. |  | 1.814% | 2.303% | 0% | 4.057% | 1.398% | 1.859% | 1.443% | 3.181% | 1.124% | 4.6% | 2.087% | 1.17% | 3.578% | 4.698% | 4.064% | 1.379%
| PDM3 |  | 5.228% | 1.424% | -6.464% | 5.228% | -3.996% | -4.183% | -0.8777% | 0.03172% | -6.914% | -14.26% | 5.554% | -1.917% | -2.567% | 9.385% | -1.137% | 13.02%
| Z-Score (Absolute) |  | 1.41 | 0.57 | NA | NA | 0.63 | 0.67 | 0.06 | NA | 1.28 | 2.91 | 1.48 | 0.17 | 0.32 | 2.33 | 0 | 3.14
Results from laboratories 1.07, 1.09 and 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 266 (Indicative Value 0.97 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.4 | < 1.5 | 14.2599 | < 1 | 0.9 | 1 | < 5 | 1 | < 1 | 0.4473 | < 1 | < 2 | 2.2 | 1 | < 1 | < 1 |  |  | 1
| 1 | 2 | 1.3 | < 1.5 | 10.8973 | < 1 | 0.9 | 1 | < 5 | 3 | 1 | 0.4466 | < 1 | < 2 | 1.9 | 1 | < 1 | < 1 |  |  | 34
| 1 | 3 | 1.1 | < 1.5 | 12.4663 | < 1 | 1.1 | 1 | < 5 | 1 | 1 | 0.4141 | 1 | < 2 | 2 | 1 | < 1 | 1 |  |  | 16
| 1 | 4 | 1.6 | < 1.5 | 12.9206 | < 1 | 0.9 | 1 | < 5 | 1 | < 1 | 0.4626 | < 1 | < 2 | 2 | 1 | < 1 | 1 |  |  | 0.96742
| 1 | 5 | 1.4 | < 1.5 | 12.3886 | < 1 | 0.9 | 1 | < 5 | 1 | 1 | 0.4836 | < 1 | < 2 | 2.1 | 1 | < 1 | 1 |  |  | 7
| 1 | 6 | 1.1 | 1.9037 | 14.8214 | < 1 | 0.9 | 1 | < 5 | 1 | < 1 | 0.4864 | 1 | < 2 | 1.9 | 1 | < 1 | 1
| Mean |  | 1.3167 | 1.9037 | 12.959 |  | 0.93333 | 1 |  | 1.3333 | 1 | 0.45677 | 1 |  | 2.0167 | 1 |  | 1
| Median |  | 1.35 | 1.9037 | 12.6935 |  | 0.9 | 1 |  | 1 | 1 | 0.45495 | 1 |  | 2 | 1 |  | 1
| Std Dev. |  | 0.19408 |  | 1.4129 |  | 0.08165 | 0 |  | 0.8165 | 0 | 0.027002 | 0 |  | 0.1169 | 0 |  | 0
| Rel.Std.Dev. |  | 14.74% |  | 10.9% |  | 8.748% | 0% |  | 61.24% | 0% | 5.912% | 0% |  | 5.797% | 0% |  | 0%
| PDM3 |  | 36.1% | 96.78% | 1240% |  | -3.523% | 3.368% |  | 37.82% | 3.368% | -52.78% | 3.368% |  | 108.5% | 3.368% |  | 3.368%
| Z-Score (Absolute) |  | 0.72 | 0.06 | 24.87 | 0.98 | 0.08 | 0.06 | 3.17 | 0.75 | 0.46 | 1.07 | 0.63 | 0.06 | 2.17 | 0.06 | 0.98 | 0.29
| Analytical results for Sm in OREAS 266 (Certified Value 7.52 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 7.5 | 7.3614 | 7.8253 | 6.4 | 7.18 | 7.5196 | 8.57 | [11 empty] | 1
| 1 | 2 | 7.29 | 6.9435 | 8.0502 | 6.9 | 7.23 | 7.3079 | 8.46 | [11 empty] | 5
| 1 | 3 | 7.31 | 6.9804 | 8.2315 | 6.8 | 7.22 | 7.5649 | 8.77 | [11 empty] | 16
| 1 | 4 | 7.56 | 6.9407 | 7.8269 | 6.5 | 7.22 | 7.3904 | 8.51 | [11 empty] | 7.5228
| 1 | 5 | 7.4 | 7.1579 | 7.8447 | 6.7 | 7.19 | 7.8239 | 8.56 | [11 empty] | 56
| 1 | 6 | 7.69 | 7.2835 | 8.0183 | 7.4 | 7.27 | 7.497 | 8.76
| Mean |  | 7.4583 | 7.1112 | 7.9662 | 6.7833 | 7.2183 | 7.5173 | 8.605
| Median |  | 7.45 | 7.0692 | 7.9315 | 6.75 | 7.22 | 7.5083 | 8.565
| Std Dev. |  | 0.15459 | 0.18373 | 0.16383 | 0.35449 | 0.031885 | 0.17699 | 0.13004
| Rel.Std.Dev. |  | 2.073% | 2.584% | 2.057% | 5.226% | 0.4417% | 2.354% | 1.511%
| PDM3 |  | -0.8571% | -5.471% | 5.893% | -9.83% | -4.047% | -0.07363% | 14.39%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.99 | 1.31 | 0.47 | 0.11 | 2.23
| Analytical results for Sn in OREAS 266 (Certified Value 6.04 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 6.5 | 6.0085 | 37.3359 | 6 | 4 | 5.6 | < 10 | 6.1 | 6.2 | 5.6468 | 5.4 | 6.1925 | 6.3 | 5.6 | 6.5 | 6.1 | 5.9 |  | 1
| 1 | 2 | 6.8 | 5.7044 | 39.9737 | 6 | 2 | 5.6 | < 10 | 6.3 | 6.5 | 5.64 | 5.6 | 6.5701 | 6.1 | 5.5 | 6.1 | 5.8 | 6 |  | 19
| 1 | 3 | 6.5 | 6.1226 | 38.7828 | 6 | 2 | 5.8 | < 10 | 6.1 | 6.1 | 5.5985 | 5.4 | 6.3044 | 6.2 | 5.5 | 6.2 | 6.1 | 6.1 |  | 16
| 1 | 4 | 6.8 | 5.8937 | 42.6972 | 6.2 | 2 | 5.7 | < 10 | 6.1 | 6.2 | 5.6833 | 5.7 | 6.4912 | 6.4 | 5.4 | 6.1 | 5.9 | 5.8 |  | 6.0402
| 1 | 5 | 6.6 | 6.1141 | 40.8906 | 6.2 | 2 | 5.7 | < 10 | 6.7 | 6.3 | 5.7828 | 5.8 | 6.2643 | 6.2 | 5.4 | 6.4 | 5.8 | 5.9 |  | 57
| 1 | 6 | 6.9 | 6.0342 | 39.8506 | 6.3 | 3 | 5.6 | < 10 | 6.3 | 6.2 | 5.6337 | 6.1 | 6.1508 | 6.3 | 5.6 | 6.8 | 7.5 | 6
| Mean |  | 6.6833 | 5.9796 | 39.9218 | 6.1167 | 2.5 | 5.6667 |  | 6.2667 | 6.25 | 5.6642 | 5.6667 | 6.3289 | 6.25 | 5.5 | 6.35 | 6.2 | 5.95
| Median |  | 6.7 | 6.0214 | 39.9121 | 6.1 | 2 | 5.65 |  | 6.2 | 6.2 | 5.6434 | 5.65 | 6.2843 | 6.25 | 5.5 | 6.3 | 6 | 5.95
| Std Dev. |  | 0.17224 | 0.15844 | 1.8239 | 0.13292 | 0.83666 | 0.08165 |  | 0.23381 | 0.13784 | 0.064134 | 0.26583 | 0.16713 | 0.10488 | 0.089443 | 0.27386 | 0.65115 | 0.10488
| Rel.Std.Dev. |  | 2.577% | 2.65% | 4.569% | 2.173% | 33.47% | 1.441% |  | 3.731% | 2.205% | 1.132% | 4.691% | 2.641% | 1.678% | 1.626% | 4.313% | 10.5% | 1.763%
| PDM3 |  | 10.65% | -1.003% | 560.9% | 1.266% | -58.61% | -6.184% |  | 3.75% | 3.474% | -6.225% | -6.184% | 4.779% | 3.474% | -8.943% | 5.129% | 2.646% | -1.493%
| Z-Score (Absolute) |  | 1.77 | 0.6 | 113.84 | 0.14 | NA | 1.66 | 3.91 | 0.37 | 0.31 | 1.67 | 1.66 | 0.57 | 0.31 | 2.22 | 0.65 | 0.14 | 0.7
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 266 (Certified Value 335 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 329.01 | 359.7804 | 332.6028 | 453.3506 | 323 | 317 | 331 | 354 | 337 | 347 | 301.1807 | 341 | 325.223 | 342 | 326 | 360 | 345 | 304 | 1
| 1 | 2 | 328.8 | 362.1672 | 325.214 | 442.4948 | 326 | 322 | 331 | 355 | 335 | 348 | 301.2938 | 338 | 342.269 | 328 | 318 | 353 | 340 | 322 | 20
| 1 | 3 | 326.65 | 357.6852 | 343.2528 | 433.9195 | 321 | 331 | 333 | 341 | 340 | 351 | 309.6039 | 334 | 348.383 | 337 | 315 | 345 | 337 | 310 | 16
| 1 | 4 | 338.48 | 351.9288 | 335.7713 | 437.9918 | 322 | 323 | 332 | 349 | 341 | 339 | 314.478 | 340 | 328.4984 | 343 | 331 | 347 | 329 | 310 | 334.5849
| 1 | 5 | 325.2 | 357.7068 | 343.804 | 448.4635 | 324 | 327 | 330 | 351 | 337 | 344 | 310.6884 | 336 | 346.7468 | 335 | 322 | 352 | 340 | 315 | 58
| 1 | 6 | 340.7 | 364.8132 | 338.1977 | 451.419 | 330 | 331 | 332 | 355 | 335 | 343 | 318.9592 | 345 | 332.5689 | 336 | 325 | 354 | 343 | 300
| Mean |  | 331.4733 | 359.0136 | 336.4738 | 444.6065 | 324.3333 | 325.1667 | 331.5 | 350.8333 | 337.5 | 345.3333 | 309.3673 | 339 | 337.2815 | 336.8333 | 322.8333 | 351.8333 | 339 | 310.1667
| Median |  | 328.905 | 358.7436 | 336.9845 | 445.4792 | 323.5 | 325 | 331.5 | 352.5 | 337 | 345.5 | 310.1462 | 339 | 337.419 | 336.5 | 323.5 | 352.5 | 340 | 310
| Std Dev. |  | 6.4816 | 4.4235 | 6.9988 | 7.7482 | 3.266 | 5.5287 | 1.0488 | 5.3821 | 2.51 | 4.2269 | 7.1011 | 3.8987 | 9.8231 | 5.4191 | 5.7764 | 5.3448 | 5.6214 | 7.8081
| Rel.Std.Dev. |  | 1.955% | 1.232% | 2.08% | 1.743% | 1.007% | 1.7% | 0.3164% | 1.534% | 0.7437% | 1.224% | 2.295% | 1.15% | 2.912% | 1.609% | 1.789% | 1.519% | 1.658% | 2.517%
| PDM3 |  | -0.93% | 7.301% | 0.5645% | 32.88% | -3.064% | -2.815% | -0.922% | 4.856% | 0.8713% | 3.212% | -7.537% | 1.32% | 0.806% | 0.672% | -3.512% | 5.155% | 1.32% | -7.298%
| Z-Score (Absolute) |  | 0.37 | 1.47 | 0.04 | 7.19 | 0.85 | 0.8 | 0.37 | 0.92 | 0.03 | 0.55 | 1.85 | 0.13 | 0.01 | 0.01 | 0.95 | 0.99 | 0.13 | 1.8
| Analytical results for Ta in OREAS 266 (Certified Value 2.12 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 2.18 | 2.5273 | 1.95 | 0.5 | 1.21 | 8.9267 | 2.14 | 2.1 | 2.6783 | 1.75 | 2.0572 | 1.9 | 2.04 | 2.29 | 2.13 | 1.84 |  |  | 1
| 1 | 2 | 2.19 | 2.449 | 1.9 | < 0.1 | 1.18 | 8.3433 | 2.1 | 2.3 | 2.5866 | 1.86 | 2.0931 | 2 | 2.03 | 2.27 | 2.03 | 1.85 |  |  | 21
| 1 | 3 | 2.24 | 2.5799 | 1.88 | < 0.1 | 1.23 | 7.9967 | 2.04 | 2.2 | 2.6979 | 1.72 | 2.1071 | 1.8 | 1.96 | 2.29 | 2.06 | 1.83 |  |  | 16
| 1 | 4 | 2.28 | 2.4237 | 1.89 | < 0.1 | 1.21 | 9.1367 | 2.16 | 2.1 | 2.6588 | 1.78 | 2.0827 | 1.9 | 2.05 | 2.2 | 1.99 | 1.8 |  |  | 2.1211
| 1 | 5 | 2.19 | 2.6168 | 1.93 | < 0.1 | 1.24 | 8.3467 | 2.07 | 2 | 2.7662 | 1.89 | 2.0965 | 2 | 2.03 | 2.34 | 2.02 | 1.81 |  |  | 59
| 1 | 6 | 2.29 | 2.4628 | 2.07 | < 0.1 | 1.19 | 8.8367 | 2.15 | 2.2 | 2.5905 | 1.9 | 2.0618 | 1.8 | 1.97 | 2.37 | 2.18 | 1.84
| Mean |  | 2.2283 | 2.5099 | 1.9367 | 0.5 | 1.21 | 8.5978 | 2.11 | 2.15 | 2.663 | 1.8167 | 2.0831 | 1.9 | 2.0133 | 2.2933 | 2.0683 | 1.8283
| Median |  | 2.215 | 2.4951 | 1.915 | 0.5 | 1.21 | 8.5917 | 2.12 | 2.15 | 2.6685 | 1.82 | 2.0879 | 1.9 | 2.03 | 2.29 | 2.045 | 1.835
| Std Dev. |  | 0.048751 | 0.077433 | 0.070333 |  | 0.022804 | 0.43469 | 0.048166 | 0.10488 | 0.068137 | 0.076594 | 0.019908 | 0.089443 | 0.038297 | 0.058878 | 0.072503 | 0.019408
| Rel.Std.Dev. |  | 2.188% | 3.085% | 3.632% |  | 1.885% | 5.056% | 2.283% | 4.878% | 2.559% | 4.216% | 0.9557% | 4.708% | 1.902% | 2.567% | 3.505% | 1.062%
| PDM3 |  | 5.055% | 18.33% | -8.695% | -76.43% | -42.95% | 305.3% | -0.5236% | 1.362% | 25.55% | -14.35% | -1.794% | -10.42% | -5.081% | 8.12% | -2.488% | -13.8%
| Z-Score (Absolute) |  | 0.54 | 1.57 | 0.54 | NA | 3.22 | 24 | 0.1 | 0.25 | 2.14 | 0.98 | 0 | 0.67 | 0.26 | 0.77 | 0.05 | 0.94
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 266 (Certified Value 0.83 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.84 | 0.93636 | 0.80917 | 0.7 | 0.82 | 0.6473 | 0.9 | 0.79 | [10 empty] | 1
| 1 | 2 | 0.81 | 0.84437 | 0.86893 | 0.8 | 0.84 | 0.64929 | 0.9 | 0.77 | [10 empty] | 22
| 1 | 3 | 0.82 | 0.86945 | 0.82776 | 0.8 | 0.85 | 0.69433 | 0.9 | 0.77 | [10 empty] | 16
| 1 | 4 | 0.85 | 0.82513 | 0.79472 | 0.8 | 0.85 | 0.64708 | 0.9 | 0.75 | [10 empty] | 0.82788
| 1 | 5 | 0.84 | 0.86945 | 0.75955 | 0.8 | 0.86 | 0.6766 | 0.9 | 0.78 | [10 empty] | 60
| 1 | 6 | 0.86 | 0.90023 | 0.81117 | 0.8 | 0.83 | 0.6704 | 0.9 | 0.79
| Mean |  | 0.83667 | 0.87416 | 0.81188 | 0.78333 | 0.84167 | 0.66417 | 0.9 | 0.775
| Median |  | 0.84 | 0.86945 | 0.81017 | 0.8 | 0.845 | 0.65985 | 0.9 | 0.775
| Std Dev. |  | 0.018619 | 0.039709 | 0.03619 | 0.040825 | 0.01472 | 0.019496 | 0 | 0.015166
| Rel.Std.Dev. |  | 2.225% | 4.543% | 4.458% | 5.212% | 1.749% | 2.935% | 0% | 1.957%
| PDM3 |  | 1.062% | 5.591% | -1.932% | -5.38% | 1.666% | -19.77% | 8.712% | -6.387%
| Z-Score (Absolute) |  | 0.25 | 1.01 | 0.25 | NA | 0.35 | 3.24 | NA | 1
Results from laboratories 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 266 (Certified Value 0.085 ppm) | [19 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | < 0.2 | 0.10661 | < 0.1 | < 0.1 | 0.09 | 9.1573 | 0.08 | < 0.5 | 0.10301 | 0.07 | 0.17418 | 0.09 | 0.1 | 0.08 | 0.07 |  |  |  | 1
| 1 | 2 | < 0.2 | 0.11428 | < 0.1 | < 0.1 | 0.08 | 9.24 | 0.08 | < 0.5 | 0.1012 | 0.1 | 0.13875 | 0.08 | 0.09 | 0.09 | 0.07 |  |  |  | 23
| 1 | 3 | < 0.2 | 0.12264 | < 0.1 | < 0.1 | 0.09 | 8.424 | 0.07 | < 0.5 | 0.10108 | 0.06 | 0.27231 | 0.08 | 0.09 | 0.09 | < 0.05 |  |  |  | 16
| 1 | 4 | < 0.2 | 0.16141 | < 0.1 | < 0.1 | 0.09 | 9.1493 | 0.08 | < 0.5 | 0.0919 | 0.11 | 0.25277 | 0.07 | 0.11 | 0.08 | < 0.05 |  |  |  | 0.084937
| 1 | 5 | < 0.2 | < 0.1 | < 0.1 | < 0.1 | 0.1 | 8.4667 | 0.08 | < 0.5 | 0.0964 | 0.11 | 0.2034 | 0.09 | 0.09 | 0.09 | 0.05 |  |  |  | 61
| 1 | 6 | < 0.2 | < 0.1 | < 0.1 | < 0.1 | 0.09 | 9.0187 | 0.08 | < 0.5 | 0.0934 | 0.07 | 0.14535 | 0.08 | 0.09 | 0.09 | < 0.05
| Mean |  |  | 0.12624 |  |  | 0.09 | 8.9093 | 0.078333 |  | 0.097832 | 0.086667 | 0.19779 | 0.081667 | 0.095 | 0.086667 | 0.063333
| Median |  |  | 0.11846 |  |  | 0.09 | 9.084 | 0.08 |  | 0.09874 | 0.085 | 0.18879 | 0.08 | 0.09 | 0.09 | 0.07
| Std Dev. |  |  | 0.024348 |  |  | 0.0063246 | 0.36658 | 0.0040825 |  | 0.0045954 | 0.022509 | 0.055498 | 0.0075277 | 0.0083666 | 0.005164 | 0.011547
| Rel.Std.Dev. |  |  | 19.29% |  |  | 7.027% | 4.115% | 5.212% |  | 4.697% | 25.97% | 28.06% | 9.218% | 8.807% | 5.958% | 18.23%
| PDM3 |  |  | 48.62% |  |  | 5.961% | 10389% | -7.775% |  | 15.18% | 2.036% | 132.9% | -3.851% | 11.85% | 2.036% | -25.44%
| Z-Score (Absolute) |  | 0.62 | 0.67 | 2.49 | 2.49 | 0 | 549.37 | 0.73 | 9.97 | 0.49 | 0.21 | 6.71 | 0.52 | 0.31 | 0.21 | 2.86
| Analytical results for Th in OREAS 266 (Certified Value 10.9 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 11.28 | 10.162 | 11.6288 | 10.7 | 9.2 | 8.71 | 10.6 | 11.4 | 13.9622 | 8.98 | 10.9243 | 9.97 | 11.25 | 11.05 | 11 | 14.81 |  |  | 1
| 1 | 2 | 11 | 9.8592 | 11.6689 | 10.8 | 11.5 | 8.56 | 11.1 | 12.4 | 13.8002 | 9.37 | 11.221 | 10.23 | 10.5 | 11.25 | 10.5 | 14.64 |  |  | 24
| 1 | 3 | 11.11 | 9.8223 | 11.753 | 10.5 | 11.7 | 8.83 | 10.9 | 13.6 | 14.1682 | 8.98 | 11.2171 | 9.76 | 11.3 | 11.35 | 10.75 | 14.61 |  |  | 16
| 1 | 4 | 11.55 | 9.9697 | 11.8206 | 10.5 | 11.5 | 8.97 | 11.3 | 11.5 | 13.9911 | 9.46 | 11.349 | 9.98 | 11.2 | 11.15 | 10.4 | 14.32 |  |  | 10.9234
| 1 | 5 | 11.2 | 10.0992 | 11.8555 | 11 | 11.5 | 8.76 | 11.35 | 10.8 | 14.5979 | 10.2 | 11.1945 | 10.23 | 11.15 | 11.55 | 10.15 | 14.56 |  |  | 62
| 1 | 6 | 11.56 | 10.0629 | 12.4504 | 10.9 | 12 | 9.02 | 10.75 | 12.8 | 13.9174 | 9.99 | 11.005 | 9.76 | 11.05 | 11.65 | 11.2 | 14.99
| Mean |  | 11.2833 | 9.9959 | 11.8629 | 10.7333 | 11.2333 | 8.8083 | 11 | 12.0833 | 14.0728 | 9.4967 | 11.1518 | 9.9883 | 11.075 | 11.3333 | 10.6667 | 14.655
| Median |  | 11.24 | 10.0163 | 11.7868 | 10.75 | 11.5 | 8.795 | 11 | 11.95 | 13.9766 | 9.415 | 11.2058 | 9.975 | 11.175 | 11.3 | 10.625 | 14.625
| Std Dev. |  | 0.23019 | 0.13583 | 0.30048 | 0.20656 | 1.0152 | 0.17034 | 0.30166 | 1.0362 | 0.28362 | 0.50781 | 0.15688 | 0.21047 | 0.29453 | 0.23166 | 0.392 | 0.22793
| Rel.Std.Dev. |  | 2.04% | 1.359% | 2.533% | 1.924% | 9.038% | 1.934% | 2.742% | 8.575% | 2.015% | 5.347% | 1.407% | 2.107% | 2.659% | 2.044% | 3.675% | 1.555%
| PDM3 |  | 3.295% | -8.492% | 8.6% | -1.74% | 2.837% | -19.36% | 0.7009% | 10.62% | 28.83% | -13.06% | 2.091% | -8.561% | 1.387% | 3.752% | -2.351% | 34.16%
| Z-Score (Absolute) |  | 0.19 | 1.26 | 0.84 | 0.43 | 0.14 | 2.6 | 0.13 | 1.09 | 3.34 | 1.82 | 0.04 | 1.27 | 0.04 | 0.25 | 0.5 | 3.99
| Analytical results for Ti in OREAS 266 (Certified Value 0.65 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.6708 | 0.70289 | 0.64911 | 0.57844 | 0.652 | 0.506 | 0.62 | 0.62 | 0.654 | 0.656 | 0.47207 | 0.668 | 0.702 | 0.66 | 0.619 | 0.682 | 0.676 | 0.607 | 1
| 1 | 2 | 0.6812 | 0.70057 | 0.63616 | 0.57523 | 0.656 | 0.389 | 0.61 | 0.63 | 0.663 | 0.656 | 0.46319 | 0.661 | 0.71717 | 0.64 | 0.624 | 0.664 | 0.665 | 0.631 | 25
| 1 | 3 | 0.6759 | 0.70602 | 0.65359 | 0.57675 | 0.649 | 0.411 | 0.63 | 0.62 | 0.668 | 0.645 | 0.47611 | 0.648 | 0.70187 | 0.65 | 0.611 | 0.66 | 0.659 | 0.609 | 16
| 1 | 4 | 0.6837 | 0.70082 | 0.65558 | < 0.1 | 0.658 | 0.392 | 0.62 | 0.61 | 0.667 | 0.653 | 0.46872 | 0.658 | 0.69001 | 0.67 | 0.634 | 0.663 | 0.641 | 0.61 | 0.64996
| 1 | 5 | 0.6804 | 0.7076 | 0.66305 | 0.55977 | 0.665 | 0.372 | 0.62 | 0.62 | 0.659 | 0.666 | 0.4808 | 0.656 | 0.71515 | 0.66 | 0.617 | 0.67 | 0.672 | 0.614 | 63
| 1 | 6 | 0.6816 | 0.7063 | 0.65518 | 0.56929 | 0.651 | 0.454 | 0.61 | 0.63 | 0.648 | 0.658 | 0.4551 | 0.673 | 0.69666 | 0.66 | 0.617 | 0.671 | 0.661 | 0.586
| Mean |  | 0.67893 | 0.70403 | 0.65211 | 0.5719 | 0.65517 | 0.42067 | 0.61833 | 0.62167 | 0.65983 | 0.65567 | 0.46933 | 0.66067 | 0.70381 | 0.65667 | 0.62033 | 0.66833 | 0.66233 | 0.6095
| Median |  | 0.6808 | 0.70445 | 0.65439 | 0.57523 | 0.654 | 0.4015 | 0.62 | 0.62 | 0.661 | 0.656 | 0.47039 | 0.6595 | 0.70193 | 0.66 | 0.618 | 0.667 | 0.663 | 0.6095
| Std Dev. |  | 0.0047429 | 0.0030152 | 0.0090191 | 0.007604 | 0.0058452 | 0.050357 | 0.0075277 | 0.0075277 | 0.0077825 | 0.0068313 | 0.0092259 | 0.0088919 | 0.010542 | 0.010328 | 0.0078909 | 0.0079162 | 0.012291 | 0.014433
| Rel.Std.Dev. |  | 0.6986% | 0.4283% | 1.383% | 1.33% | 0.8922% | 11.97% | 1.217% | 1.211% | 1.179% | 1.042% | 1.966% | 1.346% | 1.498% | 1.573% | 1.272% | 1.184% | 1.856% | 2.368%
| PDM3 |  | 4.458% | 8.32% | 0.3317% | -12.01% | 0.8018% | -35.28% | -4.865% | -4.352% | 1.52% | 0.8787% | -27.79% | 1.648% | 8.286% | 1.033% | -4.558% | 2.828% | 1.904% | -6.224%
| Z-Score (Absolute) |  | 0.55 | 1.14 | 0.08 | 4.02 | 0.01 | 5.53 | 0.87 | 0.79 | 0.1 | 0.01 | 4.38 | 0.12 | 1.14 | 0.03 | 0.83 | 0.3 | 0.16 | 1.08
| Analytical results for Tl in OREAS 266 (Certified Value 0.42 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.43 | < 0.1 | 0.45123 | 0.44 | 0.43 | 0.28 | 0.38 | < 0.5 | < 0.6453 | 0.38 | 0.41558 | 0.42 | 0.39 | 0.41 | 0.41 | 0.41 |  |  | 1
| 1 | 2 | 0.43 | 0.10796 | 0.43094 | 0.44 | 0.42 | 0.27 | 0.39 | < 0.5 | < 0.6326 | 0.37 | 0.43254 | 0.43 | 0.38 | 0.4 | 0.41 | 0.42 |  |  | 26
| 1 | 3 | 0.45 | 0.10184 | 0.45929 | 0.44 | 0.43 | 0.25 | 0.38 | < 0.5 | < 0.6453 | 0.37 | 0.41693 | 0.42 | 0.38 | 0.41 | 0.41 | 0.43 |  |  | 16
| 1 | 4 | 0.46 | 0.11052 | 0.46359 | 0.44 | 0.43 | 0.3 | 0.38 | < 0.5 | < 0.6412 | 0.38 | 0.4454 | 0.43 | 0.39 | 0.4 | 0.41 | 0.42 |  |  | 0.41794
| 1 | 5 | 0.43 | 0.11438 | 0.45865 | 0.43 | 0.43 | 0.28 | 0.38 | < 0.5 | < 0.6302 | 0.41 | 0.44549 | 0.44 | 0.39 | 0.41 | 0.38 | 0.42 |  |  | 64
| 1 | 6 | 0.44 | 0.11162 | 0.45122 | 0.45 | 0.44 | 0.26 | 0.41 | < 0.5 | < 0.6406 | 0.41 | 0.40663 | 0.43 | 0.38 | 0.42 | 0.42 | 0.44
| Mean |  | 0.44 | 0.10926 | 0.45249 | 0.44 | 0.43 | 0.27333 | 0.38667 |  |  | 0.38667 | 0.42709 | 0.42833 | 0.385 | 0.40833 | 0.40667 | 0.42333
| Median |  | 0.435 | 0.11052 | 0.45494 | 0.44 | 0.43 | 0.275 | 0.38 |  |  | 0.38 | 0.42474 | 0.43 | 0.385 | 0.41 | 0.41 | 0.42
| Std Dev. |  | 0.012649 | 0.0047456 | 0.011618 | 0.0063246 | 0.0063246 | 0.017512 | 0.012111 |  |  | 0.018619 | 0.016476 | 0.0075277 | 0.0054772 | 0.0075277 | 0.013663 | 0.010328
| Rel.Std.Dev. |  | 2.875% | 4.343% | 2.568% | 1.437% | 1.471% | 6.407% | 3.132% |  |  | 4.815% | 3.858% | 1.757% | 1.423% | 1.844% | 3.36% | 2.44%
| PDM3 |  | 5.279% | -73.86% | 8.267% | 5.279% | 2.886% | -34.6% | -7.482% |  |  | -7.482% | 2.191% | 2.487% | -7.881% | -2.298% | -2.697% | 1.291%
| Z-Score (Absolute) |  | 0.97 | 9.23 | 1.35 | 0.97 | 0.67 | 4.02 | 0.62 | 4.72 | 2.63 | 0.62 | 0.59 | 0.62 | 0.67 | 0.02 | 0.02 | 0.47
| Analytical results for Tm in OREAS 266 (Certified Value 0.26 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.26 | 0.26576 | 0.25629 | 0.3 | 0.26 | 0.2573 | 0.24 | [11 empty] | 1
| 1 | 2 | 0.25 | 0.25352 | 0.26664 | 0.3 | 0.25 | 0.2542 | 0.24 | [11 empty] | 6
| 1 | 3 | 0.31 | 0.2518 | 0.30353 | 0.3 | 0.26 | 0.2612 | 0.23 | [11 empty] | 16
| 1 | 4 | 0.26 | 0.25835 | 0.2472 | 0.3 | 0.26 | 0.2619 | 0.23 | [11 empty] | 0.25595
| 1 | 5 | 0.26 | 0.24766 | 0.2315 | 0.3 | 0.26 | 0.2619 | 0.23 | [11 empty] | 65
| 1 | 6 | 0.27 | 0.25162 | 0.2418 | 0.3 | 0.26 | 0.2513 | 0.24
| Mean |  | 0.26833 | 0.25478 | 0.25783 | 0.3 | 0.25833 | 0.25797 | 0.235
| Median |  | 0.26 | 0.25266 | 0.25175 | 0.3 | 0.26 | 0.25925 | 0.235
| Std Dev. |  | 0.02137 | 0.0063943 | 0.025423 | 0 | 0.0040825 | 0.0044827 | 0.0054772
| Rel.Std.Dev. |  | 7.964% | 2.51% | 9.86% | 0% | 1.58% | 1.738% | 2.331%
| PDM3 |  | 4.836% | -0.457% | 0.7315% | 17.21% | 0.9295% | 0.7863% | -8.187%
| Z-Score (Absolute) |  | 3.97 | 1.18 | 0.03 | NA | 0.17 | 0.03 | 8.7
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 266 (Certified Value 2.08 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.18 | 1.8621 | 2.249 | 2.07 | 2 | 2 | 2 | 2 | 2.2191 | 1.8 | 2.1283 | 2.08 | 1.9 | 2.2 | 2 | 2.28 |  |  | 1
| 1 | 2 | 2.12 | 1.8927 | 2.0091 | 2.06 | 2.2 | 1.9 | 2.1 | 2.2 | 2.1842 | 1.8 | 2.1511 | 2.03 | 1.8 | 2.2 | 1.9 | 2.26 |  |  | 28
| 1 | 3 | 2.11 | 1.8565 | 2.151 | 2.05 | 2.3 | 2 | 2 | 2.2 | 2.2405 | 1.8 | 2.1275 | 2.04 | 1.9 | 2.1 | 2.1 | 2.29 |  |  | 16
| 1 | 4 | 2.14 | 1.8562 | 2.0955 | 2.06 | 2.3 | 2 | 2.1 | 2 | 2.1977 | 1.9 | 2.157 | 2.12 | 2 | 2.1 | 2 | 2.29 |  |  | 2.078
| 1 | 5 | 2.09 | 1.8826 | 2.1462 | 2.07 | 2.3 | 1.9 | 2.1 | 1.9 | 2.2963 | 2 | 2.1549 | 2.12 | 1.9 | 2.1 | 1.9 | 2.21 |  |  | 66
| 1 | 6 | 2.18 | 1.9068 | 2.1159 | 2.12 | 2.3 | 2 | 2 | 2.2 | 2.1442 | 2 | 2.1601 | 2.08 | 1.9 | 2.1 | 2.2 | 2.3
| Mean |  | 2.1367 | 1.8761 | 2.1278 | 2.0717 | 2.2333 | 1.9667 | 2.05 | 2.0833 | 2.2137 | 1.8833 | 2.1465 | 2.0783 | 1.9 | 2.1333 | 2.0167 | 2.2717
| Median |  | 2.13 | 1.8723 | 2.1311 | 2.065 | 2.3 | 2 | 2.05 | 2.1 | 2.2084 | 1.85 | 2.153 | 2.08 | 1.9 | 2.1 | 2 | 2.285
| Std Dev. |  | 0.037238 | 0.021144 | 0.078517 | 0.024833 | 0.12111 | 0.05164 | 0.054772 | 0.13292 | 0.052013 | 0.098319 | 0.014728 | 0.038166 | 0.063246 | 0.05164 | 0.1169 | 0.033116
| Rel.Std.Dev. |  | 1.743% | 1.127% | 3.69% | 1.199% | 5.423% | 2.626% | 2.672% | 6.38% | 2.35% | 5.22% | 0.6862% | 1.836% | 3.329% | 2.421% | 5.797% | 1.458%
| PDM3 |  | 2.823% | -9.714% | 2.396% | -0.3051% | 7.475% | -5.358% | -1.348% | 0.2564% | 6.528% | -9.368% | 3.296% | 0.01574% | -8.566% | 2.663% | -2.952% | 9.32%
| Z-Score (Absolute) |  | 0.58 | 2.13 | 0.49 | 0.1 | 1.58 | 1.19 | 0.32 | 0.03 | 1.38 | 2.05 | 0.68 | 0.03 | 1.88 | 0.55 | 0.67 | 1.98
| Analytical results for V in OREAS 266 (Certified Value 92 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 92 | 88.0922 | 90.6922 | 75.632 | 92 | 73 | 85 | 94.4267 | 92 | 99 | 79.3805 | 95 | 98.5992 | 92 | 88 | 96 | 93 | 90 | 1
| 1 | 2 | 93 | 89.8268 | 89.7597 | 74.7048 | 93 | 63 | 86 | 95.3333 | 94 | 99 | 79.3658 | 94 | 99.918 | 89 | 88 | 94 | 92 | 87 | 29
| 1 | 3 | 96 | 89.8268 | 92.5521 | 75.535 | 93 | 68 | 86 | 97.02 | 94 | 98 | 79.4409 | 93 | 97.8746 | 90 | 86 | 93 | 94 | 87 | 16
| 1 | 4 | 91 | 87.612 | 92.4728 | 75.4381 | 93 | 66 | 87 | 92.1933 | 93 | 98 | 80.1457 | 94 | 98.4692 | 93 | 86 | 93 | 91 | 87 | 92.315
| 1 | 5 | 89 | 87.2396 | 94.5068 | 74.1785 | 92 | 65 | 87 | 93.1433 | 93 | 100 | 81.1597 | 93 | 101.4626 | 91 | 88 | 93 | 92 | 88 | 67
| 1 | 6 | 91 | 89.5132 | 92.0431 | 74.249 | 92 | 72 | 87 | 95.07 | 93 | 99 | 79.8436 | 95 | 100.6989 | 91 | 88 | 94 | 94 | 91
| Mean |  | 92 | 88.6851 | 92.0045 | 74.9562 | 92.5 | 67.8333 | 86.3333 | 94.5311 | 93.1667 | 98.8333 | 79.8894 | 94 | 99.5038 | 91 | 87.3333 | 93.8333 | 92.6667 | 88.3333
| Median |  | 91.5 | 88.8027 | 92.258 | 75.0714 | 92.5 | 67 | 86.5 | 94.7483 | 93 | 99 | 79.6422 | 94 | 99.2586 | 91 | 88 | 93.5 | 92.5 | 87.5
| Std Dev. |  | 2.3664 | 1.1735 | 1.6458 | 0.66213 | 0.54772 | 3.9707 | 0.8165 | 1.7047 | 0.75277 | 0.75277 | 0.6951 | 0.89443 | 1.4128 | 1.4142 | 1.0328 | 1.169 | 1.2111 | 1.7512
| Rel.Std.Dev. |  | 2.572% | 1.323% | 1.789% | 0.8834% | 0.5921% | 5.854% | 0.9457% | 1.803% | 0.808% | 0.7617% | 0.8701% | 0.9515% | 1.42% | 1.554% | 1.183% | 1.246% | 1.307% | 1.982%
| PDM3 |  | -0.3412% | -3.932% | -0.3364% | -18.8% | 0.2004% | -26.52% | -6.48% | 2.401% | 0.9226% | 7.061% | -13.46% | 1.825% | 7.787% | -1.424% | -5.396% | 1.645% | 0.381% | -4.313%
| Z-Score (Absolute) |  | 0 | 0.77 | 0 | 3.93 | 0.11 | 5.58 | 1.31 | 0.58 | 0.27 | 1.58 | 2.79 | 0.46 | 1.73 | 0.23 | 1.08 | 0.42 | 0.15 | 0.85
| Analytical results for W in OREAS 266 (Certified Value 32.1 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 34.1 | 30.9343 | 30.9061 | 35.3958 | 30.5 | 4.2 | 22.4 | 35.7315 | 31.6 | 29.8 | 35.3459 | 28.4 | 30.6349 | 25.2 | 37 | 35.2 | 33.3 | 30.7 | 1
| 1 | 2 | 32.9 | 30.1991 | 31.3528 | 34.5641 | 30.1 | 0.4 | 23.4 | 39.4327 | 30.7 | 34 | 36.6124 | 29.5 | 32.9617 | 24 | 33.2 | 34.6 | 31.5 | 30.6 | 30
| 1 | 3 | 33.8 | 30.0048 | 29.7801 | 34.8698 | 29.9 | 0.3 | 22.7 | 34.5928 | 29.5 | 33.7 | 36.1033 | 28.5 | 32.0807 | 25.1 | 34.5 | 34.2 | 33 | 32 | 16
| 1 | 4 | 34.4 | 30.0288 | 29.778 | 36.7003 | 30.8 | 0.5 | 22.3 | 35.0355 | 30.1 | 32.5 | 34.5045 | 29.4 | 32.0407 | 25.8 | 36.4 | 33.7 | 31.8 | 29.4 | 32.083
| 1 | 5 | 32.7 | 30.9093 | 30.315 | 35.4561 | 29.9 | 0.5 | 22 | 36.1923 | 30.3 | 29.1 | 36.8223 | 31 | 31.4075 | 25.6 | 35.9 | 34.4 | 31.4 | 30.1 | 68
| 1 | 6 | 34.5 | 30.0174 | 28.9071 | 33.7931 | 31 | 0.4 | 22.2 | 33.2828 | 31.4 | 34.1 | 33.8233 | 31.2 | 30.5529 | 24.7 | 36.2 | 34.8 | 35.4 | 29.8
| Mean |  | 33.7333 | 30.349 | 30.1732 | 35.1299 | 30.3667 | 1.05 | 22.5 | 35.7112 | 30.6 | 32.2 | 35.5353 | 29.6667 | 31.6131 | 25.0667 | 35.5333 | 34.4833 | 32.7333 | 30.4333
| Median |  | 33.95 | 30.114 | 30.0475 | 35.1328 | 30.3 | 0.45 | 22.35 | 35.3835 | 30.5 | 33.1 | 35.7246 | 29.45 | 31.7241 | 25.15 | 36.05 | 34.5 | 32.4 | 30.35
| Std Dev. |  | 0.76594 | 0.44941 | 0.87893 | 0.98164 | 0.47188 | 1.545 | 0.498 | 2.0826 | 0.8 | 2.2163 | 1.1974 | 1.1994 | 0.93202 | 0.65013 | 1.4137 | 0.51543 | 1.528 | 0.90921
| Rel.Std.Dev. |  | 2.271% | 1.481% | 2.913% | 2.794% | 1.554% | 147.1% | 2.213% | 5.832% | 2.614% | 6.883% | 3.37% | 4.043% | 2.948% | 2.594% | 3.979% | 1.495% | 4.668% | 2.988%
| PDM3 |  | 5.144% | -5.405% | -5.953% | 9.497% | -5.35% | -96.73% | -29.87% | 11.31% | -4.622% | 0.3646% | 10.76% | -7.532% | -1.465% | -21.87% | 10.75% | 7.482% | 2.027% | -5.142%
| Z-Score (Absolute) |  | 0.83 | 0.24 | 0.3 | 1.28 | 0.23 | 9.53 | 2.73 | 1.46 | 0.16 | 0.35 | 1.4 | 0.46 | 0.16 | 1.91 | 1.4 | 1.07 | 0.52 | 0.21
| Analytical results for Y in OREAS 266 (Certified Value 19.5 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 20.05 | 20.4732 | 18.8121 | 20.2 | 16.5 | 18.1 | 18.2367 | 19.8 | 18.9 | 16.547 | 18.4 | 20.1513 | 20.1 | 20.5 | 21.3 | 19.8 | 25.3 |  | 1
| 1 | 2 | 20.03 | 19.9907 | 18.7946 | 19.9 | 18.7 | 18.2 | 18.4067 | 19.7 | 18.9 | 16.2872 | 19 | 20.0609 | 19.5 | 19.4 | 20.2 | 19.2 | 25.5 |  | 31
| 1 | 3 | 19.86 | 20.9104 | 18.7503 | 19.6 | 19.2 | 18.2 | 18.1933 | 19 | 19.4 | 16.8769 | 18.4 | 20.1037 | 19.8 | 19.4 | 20.8 | 19.4 | 25.2 |  | 16
| 1 | 4 | 20.3 | 20.591 | 18.7975 | 19.9 | 18.7 | 18.7 | 17.81 | 19.2 | 17.7 | 16.567 | 19.2 | 20.4778 | 20.2 | 20.2 | 20.2 | 19.4 | 25.2 |  | 19.5221
| 1 | 5 | 20.47 | 21.1901 | 18.4905 | 20 | 19.2 | 18.4 | 17.8867 | 21.9 | 17.9 | 17.0665 | 20.9 | 20.5059 | 19.8 | 19.7 | 21.1 | 18.7 | 25.8 |  | 69
| 1 | 6 | 20.69 | 20.887 | 18.4594 | 20.2 | 19.6 | 18.4 | 18.46 | 20.1 | 18.5 | 16.3471 | 20.9 | 19.9717 | 20.2 | 25.4 | 21.4 | 20.9 | 25.8
| Mean |  | 20.2333 | 20.6737 | 18.6841 | 19.9667 | 18.65 | 18.3333 | 18.1656 | 19.95 | 18.55 | 16.6153 | 19.4667 | 20.2119 | 19.9333 | 20.7667 | 20.8333 | 19.5667 | 25.4667
| Median |  | 20.175 | 20.739 | 18.7724 | 19.95 | 18.95 | 18.3 | 18.215 | 19.75 | 18.7 | 16.557 | 19.1 | 20.1275 | 19.95 | 19.95 | 20.95 | 19.4 | 25.4
| Std Dev. |  | 0.31091 | 0.41994 | 0.16359 | 0.22509 | 1.1077 | 0.21602 | 0.26641 | 1.0368 | 0.65038 | 0.30285 | 1.1553 | 0.22494 | 0.28048 | 2.3123 | 0.53166 | 0.74476 | 0.28048
| Rel.Std.Dev. |  | 1.537% | 2.031% | 0.8756% | 1.127% | 5.939% | 1.178% | 1.467% | 5.197% | 3.506% | 1.823% | 5.935% | 1.113% | 1.407% | 11.13% | 2.552% | 3.806% | 1.101%
| PDM3 |  | 3.643% | 5.899% | -4.293% | 2.277% | -4.467% | -6.089% | -6.949% | 2.192% | -4.98% | -14.89% | -0.2841% | 3.533% | 2.106% | 6.375% | 6.717% | 0.2282% | 30.45%
| Z-Score (Absolute) |  | 0.24 | 0.6 | 1.01 | 0.03 | 1.04 | 1.29 | 1.43 | 0.01 | 1.12 | 2.68 | 0.38 | 0.23 | 0 | 0.67 | 0.73 | 0.3 | 4.48
| Analytical results for Yb in OREAS 266 (Certified Value 1.67 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 1.77 | 1.7113 | 1.6475 | 1.6 | 1.64 | 1.5565 | 1.6 | 1.4 | [10 empty] | 1
| 1 | 2 | 1.75 | 1.7231 | 1.7121 | 1.7 | 1.65 | 1.5065 | 1.6 | 1.5 | [10 empty] | 32
| 1 | 3 | 1.8 | 1.579 | 1.6459 | 1.8 | 1.67 | 1.6364 | 1.6 | 1.5 | [10 empty] | 16
| 1 | 4 | 1.8 | 1.6257 | 1.6273 | 1.8 | 1.66 | 1.6165 | 1.7 | 1.4 | [10 empty] | 1.6734
| 1 | 5 | 1.7 | 1.6397 | 1.5192 | 1.8 | 1.7 | 1.6959 | 1.6 | 1.5 | [10 empty] | 70
| 1 | 6 | 1.77 | 1.71 | 1.6407 | 1.8 | 1.64 | 1.2063 | 1.6 | 1.5
| Mean |  | 1.765 | 1.6648 | 1.6321 | 1.75 | 1.66 | 1.5363 | 1.6167 | 1.4667
| Median |  | 1.77 | 1.6749 | 1.6433 | 1.8 | 1.655 | 1.5865 | 1.6 | 1.5
| Std Dev. |  | 0.037283 | 0.058509 | 0.06272 | 0.083666 | 0.022804 | 0.1744 | 0.040825 | 0.05164
| Rel.Std.Dev. |  | 2.112% | 3.514% | 3.843% | 4.781% | 1.374% | 11.35% | 2.525% | 3.521%
| PDM3 |  | 5.477% | -0.5111% | -2.466% | 4.58% | -0.7981% | -8.188% | -3.388% | -12.35%
| Z-Score (Absolute) |  | 1.2 | 0.19 | 0.14 | 1.05 | 0.14 | 1.11 | 0.3 | 1.81
| Analytical results for Zn in OREAS 266 (Certified Value 90 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 92 | 99.312 | 89.7092 | 74.5678 | 91 | 89 | 89 | 84.6546 | 90 | 88 | 78.2321 | 94 | 92.228 | 93.1 | 89 | 99 | 90 | 85 | 1
| 1 | 2 | 93 | 97.013 | 87.5554 | 73.6829 | 91 | 90 | 87 | 83.1775 | 91 | 88 | 78.4265 | 93 | 95.6836 | 91.2 | 87 | 97 | 89 | 82 | 33
| 1 | 3 | 94 | 93.944 | 91.5613 | 74.1485 | 91 | 90 | 87 | 81.3284 | 95 | 87 | 79.244 | 93 | 92.0671 | 92.5 | 86 | 96 | 90 | 81 | 16
| 1 | 4 | 97 | 95.418 | 90.4337 | 76.833 | 91 | 91 | 88 | 81.3231 | 95 | 87 | 79.8381 | 94 | 94.2721 | 94.5 | 90 | 98 | 88 | 78 | 90.4702
| 1 | 5 | 92 | 94.483 | 92.8136 | 72.9315 | 91 | 90 | 86 | 80.8742 | 95 | 89 | 82.4404 | 93 | 95.37 | 91.2 | 87 | 97 | 88 | 80 | 71
| 1 | 6 | 97 | 97.299 | 91.1769 | 73.3796 | 90 | 90 | 88 | 85.0232 | 89 | 88 | 78.2417 | 95 | 90.523 | 91.4 | 88 | 98 | 92 | 80
| Mean |  | 94.1667 | 96.2448 | 90.5417 | 74.2572 | 90.8333 | 90 | 87.5 | 82.7302 | 92.5 | 87.8333 | 79.4038 | 93.6667 | 93.3573 | 92.3167 | 87.8333 | 97.5 | 89.5 | 81
| Median |  | 93.5 | 96.2155 | 90.8053 | 73.9157 | 91 | 90 | 87.5 | 82.2529 | 93 | 88 | 78.8352 | 93.5 | 93.25 | 91.95 | 87.5 | 97.5 | 89.5 | 80.5
| Std Dev. |  | 2.3166 | 2.0095 | 1.8015 | 1.3857 | 0.40825 | 0.63246 | 1.0488 | 1.8194 | 2.8107 | 0.75277 | 1.6195 | 0.8165 | 2.0627 | 1.3227 | 1.472 | 1.0488 | 1.5166 | 2.3664
| Rel.Std.Dev. |  | 2.46% | 2.088% | 1.99% | 1.866% | 0.4494% | 0.7027% | 1.199% | 2.199% | 3.039% | 0.857% | 2.04% | 0.8717% | 2.209% | 1.433% | 1.676% | 1.076% | 1.694% | 2.922%
| PDM3 |  | 4.086% | 6.383% | 0.07893% | -17.92% | 0.4013% | -0.5198% | -3.283% | -8.555% | 2.244% | -2.915% | -12.23% | 3.533% | 3.191% | 2.041% | -2.915% | 7.77% | -1.072% | -10.47%
| Z-Score (Absolute) |  | 0.9 | 1.38 | 0.06 | 3.69 | 0.13 | 0.06 | 0.64 | 1.74 | 0.51 | 0.56 | 2.5 | 0.78 | 0.71 | 0.47 | 0.56 | 1.66 | 0.18 | 2.13
| Analytical results for Zr in OREAS 266 (Certified Value 205 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 210.8 | 196.5276 | 206.0927 | 241.332 | 204 | 183 | 196.4 | 186 | 206 | 209.3 | 167.4604 | 194.5 | 204.833 | 213 | 206 | 214 | 205 | 171 | 1
| 1 | 2 | 210.9 | 198.072 | 199.1061 | 222.4219 | 207 | 163 | 193.8 | 185 | 206 | 208.9 | 166.6276 | 198.5 | 210.8206 | 213 | 199.5 | 212 | 197 | 178 | 7
| 1 | 3 | 206.3 | 196.074 | 207.3449 | 216.3417 | 204 | 60 | 201.8 | 187 | 206 | 209.9 | 168.3253 | 192 | 209.5474 | 214 | 199.5 | 211 | 200 | 178 | 16
| 1 | 4 | 214.2 | 194.3136 | 206.5005 | 219.6116 | 203 | 175 | 203.7 | 181 | 206 | 207.4 | 168.484 | 202 | 208.6831 | 211 | 216 | 208 | 202 | 171 | 205.1445
| 1 | 5 | 206.6 | 197.6076 | 211.4573 | 205.28 | 209 | 174 | 202.3 | 189 | 208 | 210.6 | 174.7886 | 205 | 210.407 | 211 | 201 | 213 | 194.5 | 175 | 72
| 1 | 6 | 217.8 | 199.7136 | 207.4997 | 214.2688 | 211 | 168 | 198.8 | 185 | 207 | 207.9 | 164.1365 | 209 | 204.3185 | 210 | 206 | 214 | 213 | 176
| Mean |  | 211.1 | 197.0514 | 206.3335 | 219.876 | 206.3333 | 153.8333 | 199.4667 | 185.5 | 206.5 | 209 | 168.3037 | 200.1667 | 208.1016 | 212 | 204.6667 | 212 | 201.9167 | 174.8333
| Median |  | 210.85 | 197.0676 | 206.9227 | 217.9767 | 205.5 | 171 | 200.3 | 185.5 | 206 | 209.1 | 167.8928 | 200.25 | 209.1152 | 212 | 203.5 | 212.5 | 201 | 175.5
| Std Dev. |  | 4.4254 | 1.8528 | 4.0239 | 12.0323 | 3.2042 | 46.4647 | 3.8302 | 2.6646 | 0.83666 | 1.2033 | 3.5489 | 6.4239 | 2.8331 | 1.5492 | 6.3061 | 2.2804 | 6.5606 | 3.1885
| Rel.Std.Dev. |  | 2.096% | 0.9402% | 1.95% | 5.472% | 1.553% | 30.2% | 1.92% | 1.436% | 0.4052% | 0.5758% | 2.109% | 3.209% | 1.361% | 0.7308% | 3.081% | 1.076% | 3.249% | 1.824%
| PDM3 |  | 2.903% | -3.945% | 0.5796% | 7.181% | 0.5795% | -25.01% | -2.768% | -9.576% | 0.6608% | 1.879% | -17.96% | -2.426% | 1.441% | 3.342% | -0.2329% | 3.342% | -1.573% | -14.78%
| Z-Score (Absolute) |  | 0.65 | 0.98 | 0.1 | 1.67 | 0.1 | 5.99 | 0.7 | 2.32 | 0.12 | 0.41 | 4.31 | 0.62 | 0.3 | 0.75 | 0.1 | 0.75 | 0.42 | 3.56
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C833: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 266 (Certified Value 4.3 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 4.34 | 4.2286 | < 10 | 4.14 | 4.29 | 4.55 | 4.355 | 4.33 | 3.8 | 4.774 | 4.2 | 4.3198 | 4.2 | 4.59 | 4.28 | 4.26 |  |  |  | 1
| 1 | 2 | 4.37 | 4.2535 | < 10 | 4.07 | 3.94 | 4.48 | 4.29 | 4.32 | 3.8 | 4.591 | 4.28 | 4.4605 | 4.3 | 4.25 | 4.35 | 4.21 |  |  |  | 18
| 1 | 3 | 4.32 | 4.3295 | < 10 | 4.1 | 4.18 | 4.42 | 4.31 | 4.35 | 4.2 | 4.706 | 4.22 | 4.3385 | 4.3 | 4.31 | 4.37 | 4.32 |  |  |  | 16
| 1 | 4 | 4.37 | 4.2423 | < 10 | 4.1 | 4.13 | 4.52 | 4.305 | 4.2 | 4 | 4.535 | 4.25 | 4.2831 | 4.3 | 4.58 | 4.38 | 4.32 |  |  |  | 4.3007
| 1 | 5 | 4.42 | 4.2943 | < 10 | 4.1 | 4.25 | 4.53 | 4.3 | 4.36 | 4 | 4.658 | 4.27 | 4.445 | 4.2 | 4.9 | 4.26 | 4.13 |  |  |  | 74
| 1 | 6 | 4.32 | 4.2922 | < 10 | 4.08 | 4.01 | 4.52 | 4.295 | 4.3 | 4.1 | 4.607 | 4.25 | 4.294 | 4.5 | 4.27 | 4.41 | 4.45
| Mean |  | 4.3567 | 4.2734 |  | 4.0983 | 4.1333 | 4.5033 | 4.3092 | 4.31 | 3.9833 | 4.6452 | 4.245 | 4.3568 | 4.3 | 4.4833 | 4.3417 | 4.2817
| Median |  | 4.355 | 4.2729 |  | 4.1 | 4.155 | 4.52 | 4.3025 | 4.325 | 4 | 4.6325 | 4.25 | 4.3292 | 4.3 | 4.445 | 4.36 | 4.29
| Std Dev. |  | 0.038297 | 0.038232 |  | 0.024014 | 0.13633 | 0.046762 | 0.023541 | 0.057966 | 0.16021 | 0.086013 | 0.030166 | 0.076948 | 0.10954 | 0.25469 | 0.059133 | 0.10944
| Rel.Std.Dev. |  | 0.879% | 0.8947% |  | 0.5859% | 3.298% | 1.038% | 0.5463% | 1.345% | 4.022% | 1.852% | 0.7106% | 1.766% | 2.548% | 5.681% | 1.362% | 2.556%
| PDM3 |  | 1.301% | -0.6352% |  | -4.705% | -3.892% | 4.712% | 0.1969% | 0.2162% | -7.379% | 8.01% | -1.295% | 1.305% | -0.01628% | 4.247% | 0.9526% | -0.4426%
| Z-Score (Absolute) |  | 0.57 | 0.44 | 8.33 | 2.55 | 2.13 | 2.34 | 0.01 | 0.01 | 3.93 | 4.05 | 0.78 | 0.57 | 0.12 | 2.1 | 0.39 | 0.34
| Analytical results for Al in OREAS 266 (Certified Value 1.55 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 1.6124 | 1.8887 | 1.51 | 1.19 | 1.76 | 1.8 | 1.51 | 1.64 | 1.44 | 1.39 | 1.45 | 1.6942 | 1.83 | 1.63 | 1.45 | 1.44 | 1.242 |  |  | 1
| 1 | 2 | 1.6132 | 1.8212 | 1.5006 | 1.29 | 1.78 | 1.59 | 1.43 | 1.61 | 1.43 | 1.36 | 1.49 | 1.7427 | 1.81 | 1.5 | 1.49 | 1.44 | 1.283 |  |  | #N/A
| 1 | 3 | 1.607 | 1.8038 | 1.5611 | 1.29 | 1.81 | 1.67 | 1.47 | 1.65 | 1.46 | 1.45 | 1.36 | 1.7237 | 1.86 | 1.52 | 1.4 | 1.4 | 1.346 |  |  | 16
| 1 | 4 | 1.6123 | 1.8735 | 1.5174 | 1.24 | 1.83 | 1.71 | 1.47 | 1.65 | 1.46 | 1.26 | 1.41 | 1.6801 | 1.84 | 1.6 | 1.46 | 1.41 | 1.2 |  |  | 1.5465
| 1 | 5 | 1.5945 | 1.8564 | 1.4998 | 1.38 | 1.8 | 1.7 | 1.5 | 1.64 | 1.46 | 1.4 | 1.39 | 1.74 | 1.83 | 1.6 | 1.37 | 1.36 | 1.275 |  |  | 75
| 1 | 6 | 1.6202 | 1.8668 | 1.5146 | 1.15 | 1.83 | 1.61 | 1.45 | 1.64 | 1.45 | 1.43 | 1.36 | 1.6955 | 1.86 | 1.51 | 1.4 | 1.44 | 1.313
| Mean |  | 1.6099 | 1.8517 | 1.5172 | 1.2567 | 1.8017 | 1.68 | 1.4717 | 1.6383 | 1.45 | 1.3817 | 1.41 | 1.7127 | 1.8383 | 1.56 | 1.4283 | 1.415 | 1.2765
| Median |  | 1.6123 | 1.8616 | 1.5123 | 1.265 | 1.805 | 1.685 | 1.47 | 1.64 | 1.455 | 1.395 | 1.4 | 1.7096 | 1.835 | 1.56 | 1.425 | 1.425 | 1.279
| Std Dev. |  | 0.0086551 | 0.032611 | 0.022669 | 0.081894 | 0.027869 | 0.075895 | 0.029944 | 0.01472 | 0.012649 | 0.067355 | 0.051769 | 0.026314 | 0.019408 | 0.056214 | 0.04535 | 0.032094 | 0.051477
| Rel.Std.Dev. |  | 0.5376% | 1.761% | 1.494% | 6.517% | 1.547% | 4.518% | 2.035% | 0.8984% | 0.8724% | 4.875% | 3.672% | 1.536% | 1.056% | 3.603% | 3.175% | 2.268% | 4.033%
| PDM3 |  | 4.1% | 19.73% | -1.893% | -18.74% | 16.5% | 8.63% | -4.841% | 5.936% | -6.242% | -10.66% | -8.828% | 10.75% | 18.87% | 0.871% | -7.643% | -8.505% | -17.46%
| Z-Score (Absolute) |  | 0.52 | 1.86 | 0 | 1.45 | 1.58 | 0.91 | 0.25 | 0.67 | 0.37 | 0.76 | 0.6 | 1.09 | 1.79 | 0.24 | 0.5 | 0.57 | 1.34
| Analytical results for As in OREAS 266 (Certified Value 168 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 0
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 0
| 1 | 1 | 172 | 177.5236 | 148.254 | 151 | 165 | 167 | 175.5 | 167 | 177 | 159.6 | 165.961 | 176 | 164.2288 | 173 | 179.5 | 174 | 169 | 161 |  | 1
| 1 | 2 | 172 | 178.7726 | 142.321 | 160 | 161 | 153 | 169.4 | 165 | 174 | 157.3 | 166.303 | 182 | 168.9764 | 170 | 174 | 179.5 | 168.5 | 160 |  | 19
| 1 | 3 | 170 | 176.3898 | 136.891 | 164 | 164 | 158 | 173.8 | 166 | 174 | 159.1 | 165.558 | 175 | 168.376 | 171 | 177.5 | 173 | 165 | 159 |  | 16
| 1 | 4 | 173 | 177.1857 | 142.026 | 158 | 161 | 161 | 174.1 | 164 | 168.5 | 149.5 | 166.278 | 176 | 160.3306 | 175 | 176.5 | 172 | 165.5 | 159 |  | 168.0743
| 1 | 5 | 171 | 158.4103 | 132.395 | 158 | 162 | 167 | 176.8 | 164 | 171.5 | 161.7 | 165.711 | 176 | 168.565 | 171 | 183 | 167.5 | 164 | 158 |  | 76
| 1 | 6 | 171 | 176.1634 | 105.172 | 152 | 163 | 156 | 169.2 | 167 | 172.5 | 161.5 | 166.378 | 174.5 | 160.8283 | 176 | 176.5 | 171.5 | 171 | 161
| Mean |  | 171.5 | 174.0742 | 134.5098 | 157.1667 | 162.6667 | 160.3333 | 173.1333 | 165.5 | 172.9167 | 158.1167 | 166.0315 | 176.5833 | 165.2175 | 172.6667 | 177.8333 | 172.9167 | 167.1667 | 159.6667
| Median |  | 171.5 | 176.7877 | 139.4585 | 158 | 162.5 | 159.5 | 173.95 | 165.5 | 173.25 | 159.35 | 166.1195 | 176 | 166.3024 | 172 | 177 | 172.5 | 167 | 159.5
| Std Dev. |  | 1.0488 | 7.7297 | 15.3444 | 4.916 | 1.633 | 5.785 | 3.157 | 1.3784 | 2.8534 | 4.5248 | 0.34247 | 2.7279 | 3.9859 | 2.4221 | 3.093 | 3.9168 | 2.7325 | 1.2111
| Rel.Std.Dev. |  | 0.6116% | 4.44% | 11.41% | 3.128% | 1.004% | 3.608% | 1.823% | 0.8329% | 1.65% | 2.862% | 0.2063% | 1.545% | 2.412% | 1.403% | 1.739% | 2.265% | 1.635% | 0.7585%
| PDM3 |  | 2.038% | 3.57% | -19.97% | -6.49% | -3.217% | -4.606% | 3.01% | -1.532% | 2.881% | -5.925% | -1.215% | 5.063% | -1.7% | 2.732% | 5.806% | 2.881% | -0.54% | -5.002%
| Z-Score (Absolute) |  | 0.52 | 0.8 | 3.43 | 1.01 | 0.42 | 0.67 | 0.7 | 0.12 | 0.67 | 0.91 | 0.06 | 1.07 | 0.15 | 0.65 | 1.2 | 0.67 | 0.06 | 0.74
| Analytical results for B in OREAS 266 (Certified Value < 10 ppm) | [20 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 1
| 1 | 1 | < 10 | 4.2 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 1
| 1 | 2 | < 10 | 3.9 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | #N/A
| 1 | 3 | < 10 | 3.1 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 16
| 1 | 4 | < 10 | 3.2 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | < 10
| 1 | 5 | < 10 | 23.4 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 77
| 1 | 6 | < 10 | 2.4 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10
| Mean |  |  | 6.7 | 5
| Median |  |  | 3.55 | 5
| Std Dev. |  |  | 8.2059 | 0
| Rel.Std.Dev. |  |  | 122.5% | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 266 (Certified Value 100 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 104 | 116.046 | 101.5856 | 108.369 | 102 | 100 | 99.2 | 98 | 100.5833 | 100 | 88 | 100 | 101.1457 | 110 | 100 | 100 | 93 |  |  | 1
| 1 | 2 | 103 | 110.6352 | 103.9588 | 107.677 | 107 | 99 | 92.6 | 96 | 97.95 | 100 | 87 | 100 | 107.9812 | 100 | 100 | 100 | 92 |  |  | 20
| 1 | 3 | 104 | 109.2096 | 104.9332 | 107.442 | 106 | 102 | 91.3 | 99 | 98.85 | 100 | 98 | 100 | 108.0134 | 100 | 100 | 90 | 94 |  |  | 16
| 1 | 4 | 103 | 115.2684 | 100.9527 | 104.637 | 104 | 102 | 98.6 | 100 | 100.2433 | 100 | 89 | 100 | 100.0906 | 110 | 100 | 90 | 88 |  |  | 99.9199
| 1 | 5 | 104 | 115.8624 | 93.6092 | 105.524 | 111 | 103 | 99 | 100 | 99.92 | 100 | 88 | 100 | 105.6527 | 110 | 100 | 100 | 86 |  |  | 78
| 1 | 6 | 102 | 117.072 | 102.0795 | 104.698 | 101 | 101 | 94.1 | 98 | 100.47 | 100 | 93 | 100 | 104.056 | 100 | 100 | 100 | 92
| Mean |  | 103.3333 | 114.0156 | 101.1865 | 106.3912 | 105.1667 | 101.1667 | 95.8 | 98.5 | 99.6694 | 100 | 90.5 | 100 | 104.4899 | 105 | 100 | 96.6667 | 90.8333
| Median |  | 103.5 | 115.5654 | 101.8325 | 106.483 | 105 | 101.5 | 96.35 | 98.5 | 100.0817 | 100 | 88.5 | 100 | 104.8544 | 105 | 100 | 100 | 92
| Std Dev. |  | 0.8165 | 3.2548 | 4.0038 | 1.635 | 3.656 | 1.472 | 3.5502 | 1.5166 | 1.0484 | 0 | 4.2308 | 0 | 3.367 | 5.4772 | 0 | 5.164 | 3.1252
| Rel.Std.Dev. |  | 0.7902% | 2.855% | 3.957% | 1.537% | 3.476% | 1.455% | 3.706% | 1.54% | 1.052% | 0% | 4.675% | 0% | 3.222% | 5.216% | 0% | 5.342% | 3.441%
| PDM3 |  | 3.416% | 14.11% | 1.268% | 6.476% | 5.251% | 1.248% | -4.123% | -1.421% | -0.2507% | 0.08011% | -9.427% | 0.08011% | 4.574% | 5.084% | 0.08011% | -3.256% | -9.094%
| Z-Score (Absolute) |  | 0.67 | 2.84 | 0.24 | 1.29 | 1.05 | 0.24 | 0.85 | 0.3 | 0.07 | 0 | 1.92 | 0 | 0.91 | 1.01 | 0 | 0.67 | 1.85
| Analytical results for Be in OREAS 266 (Certified Value 1.07 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 1.09 | 1.0929 | < 10 | 1.2 | 1 | 1.1 | 1.09 | < 5 | 1.04 | 1.03 | 1.0882 | 1.1 | 1.12 | 1.07 | 1.01 | 1.17 |  |  |  | 1
| 1 | 2 | 1.11 | 1.102 | < 10 | 1.2 | 1 | 1 | 1.05 | < 5 | 1.03 | 1.05 | 1.0986 | 1.2 | 1.08 | 1.1 | 1.03 | 1.12 |  |  |  | 21
| 1 | 3 | 1.11 | 1.1383 | < 10 | 1.3 | 1 | 1 | 1.07 | < 5 | 1.1 | 1 | 1.0705 | 1.2 | 1.09 | 1.05 | 1 | 1.15 |  |  |  | 16
| 1 | 4 | 1.07 | 1.1195 | < 10 | 1.2 | 1 | 1 | 1.07 | < 5 | 1.1 | 1.01 | 1.0657 | 1.2 | 1.1 | 1.07 | 0.99 | 1.06 |  |  |  | 1.0736
| 1 | 5 | 1.11 | 1.0111 | < 10 | 1.3 | 1 | 1 | 1.07 | < 5 | 1.1 | 1.02 | 1.0788 | 1.1 | 1.1 | 1.01 | 1 | 1.12 |  |  |  | 79
| 1 | 6 | 1.03 | 1.1092 | < 10 | 1.2 | 1 | 1 | 1.07 | < 5 | 1.04 | 1 | 1.0942 | 1.1 | 1.08 | 1.03 | 1.01 | 1.11
| Mean |  | 1.0867 | 1.0955 |  | 1.2333 | 1 | 1.0167 | 1.07 |  | 1.0683 | 1.0183 | 1.0827 | 1.15 | 1.095 | 1.055 | 1.0067 | 1.1217
| Median |  | 1.1 | 1.1056 |  | 1.2 | 1 | 1 | 1.07 |  | 1.07 | 1.015 | 1.0835 | 1.15 | 1.095 | 1.06 | 1.005 | 1.12
| Std Dev. |  | 0.032042 | 0.04421 |  | 0.05164 | 0 | 0.040825 | 0.012649 |  | 0.034881 | 0.019408 | 0.013187 | 0.054772 | 0.015166 | 0.032094 | 0.013663 | 0.037639
| Rel.Std.Dev. |  | 2.949% | 4.036% |  | 4.187% | 0% | 4.016% | 1.182% |  | 3.265% | 1.906% | 1.218% | 4.763% | 1.385% | 3.042% | 1.357% | 3.356%
| PDM3 |  | 1.216% | 2.039% |  | 14.88% | -6.857% | -5.304% | -0.3368% |  | -0.492% | -5.149% | 0.8422% | 7.115% | 1.992% | -1.734% | -6.236% | 4.476%
| Z-Score (Absolute) |  | 0 | 0.17 | 75.39 | 2.83 | NA | 1.35 | 0.32 | 27.23 | 0.35 | 1.32 | 0.08 | 1.22 | 0.16 | 0.61 | 1.54 | 0.67
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 266 (Certified Value 2.65 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.79 | 2.6879 | < 10 | < 0.3 | 2.91 | 2.79 | < 5 | < 2 | 2.7 | 2.7 | 2.718 | 2.31 | 2.8414 | 3.39 | 2.52 | 2.79 | 2.62 | 2.67 |  | 1
| 1 | 2 | 2.8 | 2.6723 | < 10 | < 0.3 | 2.89 | 2.57 | < 5 | < 2 | 2.43 | 2.2 | 2.619 | 2.31 | 2.9404 | 3.19 | 2.96 | 2.77 | 2.59 | 2.4 |  | 22
| 1 | 3 | 2.78 | 2.6091 | < 10 | < 0.3 | 2.89 | 2.73 | < 5 | < 2 | 2.62 | 2.5 | 2.645 | 2.5 | 2.9495 | 3.12 | 2.76 | 2.71 | 2.2 | 2.51 |  | 16
| 1 | 4 | 2.78 | 2.7174 | < 10 | < 0.3 | 2.97 | 2.69 | < 5 | < 2 | 2.45 | 2.4 | 2.621 | 2.36 | 2.8571 | 3.47 | 2.46 | 2.84 | 2.29 | 2.48 |  | 2.6475
| 1 | 5 | 2.76 | 2.126 | < 10 | < 0.3 | 2.98 | 2.73 | < 5 | < 2 | 2.76 | 2.4 | 2.657 | 2.61 | 2.8248 | 3.16 | 2.65 | 2.59 | 2.57 | 2.4 |  | 80
| 1 | 6 | 2.76 | 2.5908 | < 10 | 30.1 | 2.86 | 2.52 | < 5 | < 2 | 2.5 | 2.4 | 2.682 | 2.31 | 2.8485 | 3.33 | 2.83 | 2.88 | 2.79 | 2.68
| Mean |  | 2.7783 | 2.5672 |  | 30.1 | 2.9167 | 2.6717 |  |  | 2.5767 | 2.4333 | 2.657 | 2.4 | 2.8769 | 3.2767 | 2.6967 | 2.7633 | 2.51 | 2.5233
| Median |  | 2.78 | 2.6407 |  | 30.1 | 2.9 | 2.71 |  |  | 2.56 | 2.4 | 2.651 | 2.335 | 2.8528 | 3.26 | 2.705 | 2.78 | 2.58 | 2.495
| Std Dev. |  | 0.016021 | 0.22143 |  |  | 0.048028 | 0.10439 |  |  | 0.13721 | 0.1633 | 0.038026 | 0.12649 | 0.053815 | 0.14052 | 0.19002 | 0.10309 | 0.22136 | 0.12533
| Rel.Std.Dev. |  | 0.5766% | 8.625% |  |  | 1.647% | 3.907% |  |  | 5.325% | 6.711% | 1.431% | 5.27% | 1.871% | 4.289% | 7.046% | 3.73% | 8.819% | 4.967%
| PDM3 |  | 4.942% | -3.032% |  | 1037% | 10.17% | 0.913% |  |  | -2.675% | -8.089% | 0.359% | -9.348% | 8.667% | 23.76% | 1.857% | 4.375% | -5.193% | -4.69%
| Z-Score (Absolute) |  | 0.48 | 0.41 | 9.88 | 10.48 | 1.07 | 0.03 | 0.7 | 7.04 | 0.37 | 0.98 | 0.03 | 1.12 | 0.9 | 2.59 | 0.14 | 0.42 | 0.65 | 0.6
| Analytical results for Ca in OREAS 266 (Certified Value 0.632 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 0.66 | 0.74957 | 0.63126 | 0.52122 | 1.15 | 0.68 | 0.83 | 0.67 | 0.67 | 0.62 | 0.58 | 0.49669 | 0.6 | 0.66579 | 0.75 | 0.66 | 0.61 | 0.59 | 0.507 | 1
| 1 | 2 | 0.67 | 0.74819 | 0.63765 | 0.51673 | 1.19 | 0.68 | 0.74 | 0.63 | 0.65 | 0.61 | 0.56 | 0.49879 | 0.61 | 0.69594 | 0.74 | 0.65 | 0.63 | 0.59 | 0.518 | #N/A
| 1 | 3 | 0.66 | 0.74787 | 0.64387 | 0.50571 | 1.18 | 0.68 | 0.78 | 0.63 | 0.65 | 0.62 | 0.59 | 0.49439 | 0.57 | 0.70219 | 0.75 | 0.65 | 0.6 | 0.57 | 0.525 | 16
| 1 | 4 | 0.67 | 0.74472 | 0.62874 | 0.51329 | 1.18 | 0.69 | 0.8 | 0.64 | 0.67 | 0.61 | 0.54 | 0.49989 | 0.59 | 0.6461 | 0.75 | 0.65 | 0.62 | 0.57 | 0.522 | 0.63202
| 1 | 5 | 0.64 | 0.75341 | 0.62596 | 0.5023 | 1.24 | 0.68 | 0.8 | 0.59 | 0.67 | 0.62 | 0.59 | 0.51579 | 0.58 | 0.69614 | 0.76 | 0.65 | 0.58 | 0.61 | 0.507 | 81
| 1 | 6 | 0.67 | 0.74325 | 0.62747 | 0.50804 | 1.15 | 0.69 | 0.77 | 0.62 | 0.67 | 0.62 | 0.58 | 0.50199 | 0.57 | 0.6538 | 0.75 | 0.65 | 0.59 | 0.58 | 0.52
| Mean |  | 0.66167 | 0.74783 | 0.63249 | 0.51121 | 1.1817 | 0.68333 | 0.78667 | 0.63 | 0.66333 | 0.61667 | 0.57333 | 0.50126 | 0.58667 | 0.67666 | 0.75 | 0.65167 | 0.605 | 0.585 | 0.5165
| Median |  | 0.665 | 0.74803 | 0.63 | 0.51067 | 1.18 | 0.68 | 0.79 | 0.63 | 0.67 | 0.62 | 0.58 | 0.49934 | 0.585 | 0.68086 | 0.75 | 0.65 | 0.605 | 0.585 | 0.519
| Std Dev. |  | 0.01169 | 0.0036053 | 0.0069198 | 0.0071373 | 0.033116 | 0.005164 | 0.030768 | 0.026077 | 0.010328 | 0.005164 | 0.019664 | 0.0075839 | 0.01633 | 0.024404 | 0.0063246 | 0.0040825 | 0.018708 | 0.015166 | 0.0077136
| Rel.Std.Dev. |  | 1.767% | 0.4821% | 1.094% | 1.396% | 2.802% | 0.7557% | 3.911% | 4.139% | 1.557% | 0.8374% | 3.43% | 1.513% | 2.784% | 3.607% | 0.8433% | 0.6265% | 3.092% | 2.592% | 1.493%
| PDM3 |  | 4.691% | 18.32% | 0.07441% | -19.11% | 86.97% | 8.119% | 24.47% | -0.3198% | 4.954% | -2.429% | -9.286% | -20.69% | -7.176% | 7.063% | 18.67% | 3.108% | -4.275% | -7.44% | -18.28%
| Z-Score (Absolute) |  | 0.41 | 1.64 | 0 | 1.72 | 7.8 | 0.72 | 2.19 | 0.04 | 0.44 | 0.22 | 0.84 | 1.86 | 0.65 | 0.63 | 1.67 | 0.27 | 0.39 | 0.67 | 1.65
| Analytical results for Cd in OREAS 266 (Certified Value 0.063 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.06 | 0.13 | 1.1163 | 0.07 | < 0.1 | 0.06 | 0.070144 | 0.09 | 0.06 | 0.06 | 0.05 |  |  |  | 1
| 1 | 2 | 0.05 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.04 | 0.12 | 1.072 | 0.06 | < 0.1 | 0.07 | 0.079679 | 0.09 | 0.06 | 0.06 | 0.05 |  |  |  | 23
| 1 | 3 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.07 | 0.13 | 1.1303 | 0.04 | < 0.1 | 0.06 | 0.074063 | 0.1 | 0.06 | 0.06 | 0.05 |  |  |  | 16
| 1 | 4 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.07 | 0.13 | 1.1013 | 0.07 | 0.1 | 0.07 | 0.083376 | 0.11 | 0.07 | 0.06 | 0.05 |  |  |  | 0.063413
| 1 | 5 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.03 | 0.14 | 1.1357 | 0.07 | < 0.1 | 0.06 | 0.071589 | 0.09 | 0.06 | 0.06 | 0.07 |  |  |  | 82
| 1 | 6 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.06 | 0.13 | 1.1127 | 0.06 | < 0.1 | 0.07 | 0.065458 | 0.1 | 0.06 | 0.04 | 0.06
| Mean |  | 0.065 |  |  |  | 0.071667 | 0.055 | 0.13 | 1.1114 | 0.061667 | 0.1 | 0.065 | 0.074052 | 0.096667 | 0.061667 | 0.056667 | 0.055
| Median |  | 0.07 |  |  |  | 0.07 | 0.06 | 0.13 | 1.1145 | 0.065 | 0.1 | 0.065 | 0.072826 | 0.095 | 0.06 | 0.06 | 0.05
| Std Dev. |  | 0.0083666 |  |  |  | 0.0040825 | 0.016432 | 0.0063246 | 0.022922 | 0.01169 |  | 0.0054772 | 0.006539 | 0.008165 | 0.0040825 | 0.008165 | 0.0083666
| Rel.Std.Dev. |  | 12.87% |  |  |  | 5.696% | 29.88% | 4.865% | 2.062% | 18.96% |  | 8.427% | 8.83% | 8.447% | 6.62% | 14.41% | 15.21%
| PDM3 |  | 2.502% |  |  |  | 13.02% | -13.27% | 105% | 1653% | -2.754% | 57.7% | 2.502% | 16.78% | 52.44% | -2.754% | -10.64% | -13.27%
| Z-Score (Absolute) |  | 0.27 | 1.15 | 199.39 | 3.17 | 0 | 0.67 | 2.36 | 42.07 | 0.4 | NA | 0.27 | 0.1 | 1.01 | 0.4 | 0.61 | 0.67
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 266 (Certified Value 47.7 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 48.518 | 45.424 | 50.4 | 52.4 | 43.7 | 47 | 47.4 | 53.8324 | 44.6 | 52.2 | 49.6 | 45.6 | 46.43 | [6 empty] | 1
| 1 | 2 | 48.701 | 45.96 | 50 | 48 | 42.38 | 47.8 | 49.3 | 53.4784 | 44.8 | 48.1 | 51.4 | 44.6 | 47.38 | [6 empty] | 24
| 1 | 3 | 48.267 | 46.1869 | 50.9 | 48.1 | 41.75 | 46.6 | 46 | 54.1552 | 45.5 | 49.6 | 47 | 42.9 | 47.36 | [6 empty] | 16
| 1 | 4 | 48.269 | 45.2857 | 51.3 | 50.2 | 42.21 | 43.8 | 47 | 52.2041 | 46 | 52.3 | 52 | 44.7 | 46.29 | [6 empty] | 47.7415
| 1 | 5 | 48.475 | 45.145 | 51.6 | 51.1 | 42.89 | 46.2 | 45.3 | 51.9021 | 44.2 | 55.6 | 45.9 | 42.9 | 46.31 | [6 empty] | 83
| 1 | 6 | 47.783 | 46.0172 | 50.8 | 48.2 | 43.17 | 47.4 | 46.4 | 52.2141 | 47.3 | 48 | 47.9 | 45.7 | 47.35
| Mean |  | 48.3355 | 45.6698 | 50.8333 | 49.6667 | 42.6833 | 46.4667 | 46.9 | 52.9644 | 45.4 | 50.9667 | 48.9667 | 44.4 | 46.8533
| Median |  | 48.372 | 45.692 | 50.85 | 49.2 | 42.635 | 46.8 | 46.7 | 52.8462 | 45.15 | 50.9 | 48.75 | 44.65 | 46.89
| Std Dev. |  | 0.31646 | 0.43717 | 0.58195 | 1.8544 | 0.70715 | 1.4236 | 1.3885 | 0.97007 | 1.1331 | 2.9548 | 2.4468 | 1.2458 | 0.56081
| Rel.Std.Dev. |  | 0.6547% | 0.9573% | 1.145% | 3.734% | 1.657% | 3.064% | 2.961% | 1.832% | 2.496% | 5.797% | 4.997% | 2.806% | 1.197%
| PDM3 |  | 1.244% | -4.339% | 6.476% | 4.032% | -10.59% | -2.67% | -1.763% | 10.94% | -4.905% | 6.755% | 2.566% | -6.999% | -1.86%
| Z-Score (Absolute) |  | 0.47 | 0.4 | 1.28 | 0.9 | 1.38 | 0.14 | 0 | 1.98 | 0.49 | 1.33 | 0.67 | 0.82 | 0.02
| Analytical results for Co in OREAS 266 (Certified Value 16.2 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 16.1 | 19.0836 | 16.5904 | 21.758 | 8 | 16.5 | 17.8 | 15.6 | 17.0633 | 15.2 | 14.4 | 16 | 17.2821 | 17.65 | 16.2 | 15.4 | 15.5 | 14.7 |  | 1
| 1 | 2 | 16.1 | 18.1116 | 16.3943 | 23.404 | 23 | 16.5 | 16.5 | 14.7 | 16.7533 | 15.4 | 14.1 | 16.4 | 17.6218 | 18.11 | 16 | 16.3 | 15.1 | 15.1 |  | 25
| 1 | 3 | 15.9 | 18.1656 | 16.5805 | 22.278 | 13 | 16.9 | 16.6 | 14.8 | 16.74 | 15.4 | 16 | 15.8 | 17.6328 | 18.18 | 16.2 | 14.7 | 14.5 | 14.4 |  | 16
| 1 | 4 | 15.9 | 18.8352 | 16.2157 | 19.844 | 10 | 16.8 | 16.9 | 14.9 | 16.9433 | 14.7 | 14.5 | 16 | 17.2318 | 18.05 | 16.4 | 16.1 | 14.8 | 14.5 |  | 16.1926
| 1 | 5 | 16.1 | 18.7596 | 16.1539 | 20.732 | 14 | 16.8 | 17.8 | 15 | 17.0033 | 15 | 14.1 | 15.5 | 17.2962 | 17.99 | 17.3 | 14.8 | 14.6 | 14.4 |  | 84
| 1 | 6 | 15.6 | 19.0296 | 16.5803 | 18.903 | 9 | 16.7 | 16.3 | 15 | 16.99 | 15.4 | 15.2 | 15.8 | 17.1947 | 18.68 | 16.2 | 15.3 | 15.1 | 15.1
| Mean |  | 15.95 | 18.6642 | 16.4192 | 21.1532 | 12.8333 | 16.7 | 16.9833 | 15 | 16.9156 | 15.1833 | 14.7167 | 15.9167 | 17.3766 | 18.11 | 16.3833 | 15.4333 | 14.9333 | 14.7
| Median |  | 16 | 18.7974 | 16.4873 | 21.245 | 11.5 | 16.75 | 16.75 | 14.95 | 16.9667 | 15.3 | 14.45 | 15.9 | 17.2892 | 18.08 | 16.2 | 15.35 | 14.95 | 14.6
| Std Dev. |  | 0.19748 | 0.42467 | 0.19681 | 1.6524 | 5.4924 | 0.16733 | 0.66156 | 0.31623 | 0.13638 | 0.28577 | 0.74677 | 0.29944 | 0.19761 | 0.33448 | 0.46655 | 0.65625 | 0.37238 | 0.32863
| Rel.Std.Dev. |  | 1.238% | 2.275% | 1.199% | 7.812% | 42.8% | 1.002% | 3.895% | 2.108% | 0.8062% | 1.882% | 5.074% | 1.881% | 1.137% | 1.847% | 2.848% | 4.252% | 2.494% | 2.236%
| PDM3 |  | -1.498% | 15.26% | 1.399% | 30.63% | -20.75% | 3.133% | 4.883% | -7.365% | 4.464% | -6.233% | -9.115% | -1.704% | 7.312% | 11.84% | 1.178% | -4.689% | -7.777% | -9.218%
| Z-Score (Absolute) |  | 0.29 | 1.55 | 0.02 | 3.24 | NA | 0.21 | 0.41 | 0.94 | 0.36 | 0.81 | 1.13 | 0.32 | 0.67 | 1.17 | 0 | 0.64 | 0.98 | 1.14
Results from laboratory 1.06 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cr in OREAS 266 (Certified Value 49.1 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 1
| 1 | 1 | 53 | 57.7476 | 49.9653 | 37.844 | 37 | 61 | 56 | 45 | 50.0562 | 47 | 42 | 46 | 57.0053 | 61 | 51 | 49 | 45 | 42 |  | 1
| 1 | 2 | 53 | 56.9676 | 49.8675 | 37.009 | 40 | 61 | 51 | 42 | 49.1243 | 48 | 41 | 47 | 58.5893 | 63 | 48 | 50 | 45 | 43 |  | 26
| 1 | 3 | 53 | 56.8728 | 50.9085 | 38.273 | 40 | 61 | 52 | 43 | 49.8459 | 46 | 43 | 43 | 59.1309 | 63 | 49 | 47 | 44 | 42 |  | 16
| 1 | 4 | 53 | 57.5316 | 49.5773 | 36.33 | 41 | 62 | 54 | 42 | 49.6325 | 45 | 39 | 45 | 56.3922 | 63 | 50 | 49 | 44 | 40 |  | 49.0785
| 1 | 5 | 51 | 56.5704 | 45.7204 | 36.61 | 47 | 60 | 54 | 43 | 49.8004 | 46 | 42 | 45 | 57.1912 | 62 | 50 | 47 | 45 | 41 |  | 85
| 1 | 6 | 53 | 57.888 | 50.2071 | 36.63 | 37 | 62 | 53 | 43 | 50.3317 | 48 | 42 | 43 | 58.4782 | 65 | 48 | 50 | 46 | 42
| Mean |  | 52.6667 | 57.263 | 49.3744 | 37.116 | 40.3333 | 61.1667 | 53.3333 | 43 | 49.7985 | 46.6667 | 41.5 | 44.8333 | 57.7978 | 62.8333 | 49.3333 | 48.6667 | 44.8333 | 41.6667
| Median |  | 53 | 57.2496 | 49.9164 | 36.8195 | 40 | 61 | 53.5 | 43 | 49.8232 | 46.5 | 42 | 45 | 57.8347 | 63 | 49.5 | 49 | 45 | 42
| Std Dev. |  | 0.8165 | 0.53232 | 1.8456 | 0.77332 | 3.6697 | 0.75277 | 1.7512 | 1.0954 | 0.40867 | 1.2111 | 1.3784 | 1.6021 | 1.0806 | 1.3292 | 1.2111 | 1.3663 | 0.75277 | 1.0328
| Rel.Std.Dev. |  | 1.55% | 0.9296% | 3.738% | 2.084% | 9.098% | 1.231% | 3.283% | 2.548% | 0.8206% | 2.595% | 3.321% | 3.573% | 1.87% | 2.115% | 2.455% | 2.807% | 1.679% | 2.479%
| PDM3 |  | 7.311% | 16.68% | 0.6027% | -24.37% | -17.82% | 24.63% | 8.669% | -12.39% | 1.467% | -4.914% | -15.44% | -8.65% | 17.77% | 28.03% | 0.5191% | -0.8392% | -8.65% | -15.1%
| Z-Score (Absolute) |  | 0.48 | 1.08 | 0.05 | 1.55 | 1.13 | 1.59 | 0.57 | 0.78 | 0.1 | 0.3 | 0.98 | 0.54 | 1.15 | 1.81 | 0.04 | 0.04 | 0.54 | 0.96
| Analytical results for Cs in OREAS 266 (Certified Value 1.21 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 1.28 | 1.2804 | 1.52 | 1.07 | 1.27 | 1.04 | 0.99 | 1.2444 | 1.56 | 1.15 | 1.09 | 1 | 1.26 | [6 empty] | 1
| 1 | 2 | 1.28 | 1.2125 | 1.47 | 1.04 | 1.23 | 1.06 | 1 | 1.3315 | 1.62 | 1.05 | 1.19 | 0.97 | 1.3 | [6 empty] | 27
| 1 | 3 | 1.28 | 1.269 | 1.51 | 1.02 | 1.23 | 1.13 | 0.97 | 1.2905 | 1.61 | 1.07 | 1.1 | 0.96 | 1.31 | [6 empty] | 16
| 1 | 4 | 1.25 | 1.2614 | 1.51 | 1.12 | 1.22 | 1.06 | 0.99 | 1.2359 | 1.6 | 1.16 | 1.1 | 0.97 | 1.25 | [6 empty] | 1.215
| 1 | 5 | 1.3 | 0.80413 | 1.49 | 1.19 | 1.26 | 1.15 | 0.98 | 1.303 | 1.57 | 1.24 | 1.05 | 0.96 | 1.28 | [6 empty] | 86
| 1 | 6 | 1.23 | 1.3196 | 1.51 | 1.09 | 1.25 | 1.06 | 1 | 1.2713 | 1.66 | 1.08 | 1.1 | 1 | 1.27
| Mean |  | 1.27 | 1.1912 | 1.5017 | 1.0883 | 1.2433 | 1.0833 | 0.98833 | 1.2795 | 1.6033 | 1.125 | 1.105 | 0.97667 | 1.2783
| Median |  | 1.28 | 1.2652 | 1.51 | 1.08 | 1.24 | 1.06 | 0.99 | 1.2809 | 1.605 | 1.115 | 1.1 | 0.97 | 1.275
| Std Dev. |  | 0.025298 | 0.19271 | 0.018348 | 0.061128 | 0.019664 | 0.045019 | 0.01169 | 0.036252 | 0.036148 | 0.071764 | 0.045935 | 0.018619 | 0.023166
| Rel.Std.Dev. |  | 1.992% | 16.18% | 1.222% | 5.617% | 1.582% | 4.156% | 1.183% | 2.833% | 2.255% | 6.379% | 4.157% | 1.906% | 1.812%
| PDM3 |  | 4.531% | -1.958% | 23.6% | -10.42% | 2.336% | -10.83% | -18.65% | 5.309% | 31.97% | -7.404% | -9.05% | -19.61% | 5.217%
| Z-Score (Absolute) |  | 0.6 | 0 | 2.37 | 0.79 | 0.4 | 0.82 | 1.55 | 0.67 | 3.15 | 0.51 | 0.66 | 1.64 | 0.67
| Analytical results for Cu in OREAS 266 (Certified Value 26.7 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 27.3 | 23.22 | 27.3966 | 23.213 | 27 | 30 | 29 | 30.1 | 28.95 | 24.8 | 29 | 22.8328 | 25.1 | 24.7937 | 29.6 | 28.9 | 28.5 | 26.9 | 21.1 | 1
| 1 | 2 | 27.3 | 23.4252 | 27.3373 | 22.761 | 31 | 30 | 29.4 | 28.4 | 29.3067 | 24.8 | 28.6 | 23.023 | 25.3 | 24.9209 | 30.9 | 28.1 | 29.7 | 26.6 | 21.7 | 28
| 1 | 3 | 27.3 | 23.976 | 27.3882 | 23.559 | 27 | 30 | 28.9 | 28.2 | 28.5333 | 25.8 | 29.4 | 22.8235 | 24 | 25.044 | 30.7 | 28.5 | 28.7 | 24.7 | 21.5 | 16
| 1 | 4 | 27.6 | 21.0276 | 27.0893 | 22.023 | 31 | 30 | 28.7 | 28.7 | 28.6967 | 24.3 | 27 | 22.43 | 25 | 24.6521 | 30.4 | 28.3 | 29.2 | 25 | 20.1 | 26.7025
| 1 | 5 | 27.4 | 21.2544 | 24.7769 | 22.475 | 30 | 30 | 29.5 | 29.9 | 28.7367 | 25.4 | 29.8 | 23.2235 | 23.6 | 25.1111 | 30 | 30.3 | 26.9 | 26.1 | 20.8 | 87
| 1 | 6 | 27 | 22.0752 | 27.0628 | 22.295 | 31 | 31 | 29.6 | 29.1 | 28.7667 | 24.6 | 29.7 | 22.7057 | 23.7 | 24.763 | 31.5 | 27.9 | 27.3 | 26.9 | 21.5
| Mean |  | 27.3167 | 22.4964 | 26.8419 | 22.721 | 29.5 | 30.1667 | 29.1833 | 29.0667 | 28.8317 | 24.95 | 28.9167 | 22.8398 | 24.45 | 24.8808 | 30.5167 | 28.6667 | 28.3833 | 26.0333 | 21.1167
| Median |  | 27.3 | 22.6476 | 27.2133 | 22.618 | 30.5 | 30 | 29.2 | 28.9 | 28.7517 | 24.8 | 29.2 | 22.8281 | 24.5 | 24.8573 | 30.55 | 28.4 | 28.6 | 26.35 | 21.3
| Std Dev. |  | 0.19408 | 1.2209 | 1.0223 | 0.57881 | 1.9748 | 0.40825 | 0.3656 | 0.78655 | 0.26837 | 0.55045 | 1.04 | 0.27096 | 0.76616 | 0.17609 | 0.67355 | 0.87101 | 1.0852 | 0.96678 | 0.5947
| Rel.Std.Dev. |  | 0.7105% | 5.427% | 3.809% | 2.547% | 6.694% | 1.353% | 1.253% | 2.706% | 0.9308% | 2.206% | 3.597% | 1.186% | 3.134% | 0.7077% | 2.207% | 3.038% | 3.823% | 3.714% | 2.816%
| PDM3 |  | 2.3% | -15.75% | 0.5219% | -14.91% | 10.48% | 12.97% | 9.291% | 8.854% | 7.974% | -6.563% | 8.292% | -14.47% | -8.436% | -6.822% | 14.28% | 7.356% | 6.295% | -2.506% | -20.92%
| Z-Score (Absolute) |  | 0 | 1.49 | 0.15 | 1.42 | 0.67 | 0.88 | 0.58 | 0.54 | 0.47 | 0.73 | 0.49 | 1.38 | 0.89 | 0.75 | 0.99 | 0.42 | 0.33 | 0.4 | 1.91
| Analytical results for Dy in OREAS 266 (Indicative Value 2.45 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 2.479 | 2.3417 | 2.6 | [16 empty] | 1
| 1 | 2 | 2.51 | 2.2638 | 2.3 | [16 empty] | 3
| 1 | 3 | 2.501 | 2.4761 | 2.5 | [16 empty] | 16
| 1 | 4 | 2.539 | 2.4538 | 2.4 | [16 empty] | 2.4469
| 1 | 5 | 2.486 | 1.4614 | 2.6 | [16 empty] | 9
| 1 | 6 | 2.418 | 2.4743 | 2.3
| Mean |  | 2.4888 | 2.2452 | 2.45
| Median |  | 2.4935 | 2.3978 | 2.45
| Std Dev. |  | 0.040583 | 0.39327 | 0.13784
| Rel.Std.Dev. |  | 1.631% | 17.52% | 5.626%
| PDM3 |  | 1.713% | -8.245% | 0.1257%
| Z-Score (Absolute) |  | 0.67 | 3.56 | 0
| Analytical results for Er in OREAS 266 (Indicative Value 0.95 ppm) | [20 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 1.027 | 0.90048 | 1 | [16 empty] | 1
| 1 | 2 | 1.012 | 0.86347 | 0.8 | [16 empty] | 4
| 1 | 3 | 1.034 | 0.96779 | 0.9 | [16 empty] | 16
| 1 | 4 | 1.025 | 0.94604 | 1 | [16 empty] | 0.95058
| 1 | 5 | 1.024 | 0.56612 | 0.9 | [16 empty] | 10
| 1 | 6 | 1.011 | 0.97001 | 0.8
| Mean |  | 1.0222 | 0.86899 | 0.9
| Median |  | 1.0245 | 0.92326 | 0.9
| Std Dev. |  | 0.0089759 | 0.15405 | 0.089443
| Rel.Std.Dev. |  | 0.8781% | 17.73% | 9.938%
| PDM3 |  | 7.531% | -8.583% | -5.32%
| Z-Score (Absolute) |  | 2.66 | 0.67 | 0
| Analytical results for Eu in OREAS 266 (Indicative Value 1.02 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 1.105 | 0.96968 | 1 | [16 empty] | 1
| 1 | 2 | 1.124 | 0.92016 | 0.9 | [16 empty] | 5
| 1 | 3 | 1.115 | 0.95331 | 1 | [16 empty] | 16
| 1 | 4 | 1.128 | 0.97597 | 1 | [16 empty] | 1.0157
| 1 | 5 | 1.096 | 0.58067 | 1 | [16 empty] | 11
| 1 | 6 | 1.076 | 0.96334 | 1
| Mean |  | 1.1073 | 0.89385 | 0.98333
| Median |  | 1.11 | 0.95832 | 1
| Std Dev. |  | 0.019387 | 0.15468 | 0.040825
| Rel.Std.Dev. |  | 1.751% | 17.31% | 4.152%
| PDM3 |  | 9.02% | -12% | -3.188%
| Z-Score (Absolute) |  | 0.93 | 0.67 | 0
| Analytical results for Fe in OREAS 266 (Certified Value 3.62 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 3.71 | 3.9604 | 3.6614 | 3.2273 | 3.42 | 4.08 | 4.17 | 3.01 | 3.64 | 3.57 | 3.53 | 3.2134 | 3.48 | 4.1496 | 4.11 | 3.77 | 3.45 | 3.45 | 3.1 | 1
| 1 | 2 | 3.73 | 3.9619 | 3.668 | 3.2061 | 3.64 | 4.08 | 4.22 | 2.96 | 3.53 | 3.54 | 3.46 | 3.2246 | 3.55 | 4.1583 | 4.07 | 3.63 | 3.54 | 3.46 | 3.17 | #N/A
| 1 | 3 | 3.71 | 3.9628 | 3.7168 | 3.1774 | 3.66 | 4.07 | 4.12 | 2.97 | 3.58 | 3.54 | 3.55 | 3.2845 | 3.36 | 4.174 | 4.12 | 3.66 | 3.35 | 3.38 | 3.2 | 16
| 1 | 4 | 3.74 | 4.0707 | 3.6561 | 3.1751 | 3.55 | 4.1 | 4.13 | 3.01 | 3.61 | 3.55 | 3.29 | 3.2741 | 3.44 | 4.1157 | 4.11 | 3.7 | 3.47 | 3.37 | 2.94 | 3.6153
| 1 | 5 | 3.59 | 4.0578 | 3.6554 | 3.1112 | 3.85 | 4.07 | 4.2 | 3.04 | 3.63 | 3.57 | 3.58 | 3.2764 | 3.43 | 4.1873 | 4.16 | 3.72 | 3.27 | 3.48 | 2.9 | 88
| 1 | 6 | 3.76 | 4.042 | 3.6658 | 3.1519 | 3.38 | 4.09 | 4.22 | 2.99 | 3.64 | 3.57 | 3.56 | 3.2163 | 3.36 | 4.1056 | 4.13 | 3.65 | 3.32 | 3.47 | 3.27
| Mean |  | 3.7067 | 4.0093 | 3.6706 | 3.1748 | 3.5833 | 4.0817 | 4.1767 | 2.9967 | 3.605 | 3.5567 | 3.495 | 3.2482 | 3.4367 | 4.1484 | 4.1167 | 3.6883 | 3.4 | 3.435 | 3.0967
| Median |  | 3.72 | 4.0024 | 3.6636 | 3.1762 | 3.595 | 4.08 | 4.185 | 3 | 3.62 | 3.56 | 3.54 | 3.2494 | 3.435 | 4.154 | 4.115 | 3.68 | 3.4 | 3.455 | 3.135
| Std Dev. |  | 0.060222 | 0.052907 | 0.023193 | 0.040749 | 0.17282 | 0.01169 | 0.044121 | 0.029439 | 0.043243 | 0.015055 | 0.10858 | 0.033374 | 0.072847 | 0.032155 | 0.029439 | 0.051929 | 0.10276 | 0.047645 | 0.14787
| Rel.Std.Dev. |  | 1.625% | 1.32% | 0.6319% | 1.284% | 4.823% | 0.2864% | 1.056% | 0.9824% | 1.2% | 0.4233% | 3.107% | 1.027% | 2.12% | 0.7751% | 0.7151% | 1.408% | 3.022% | 1.387% | 4.775%
| PDM3 |  | 2.527% | 10.9% | 1.529% | -12.18% | -0.8842% | 12.9% | 15.53% | -17.11% | -0.2848% | -1.622% | -3.327% | -10.15% | -4.941% | 14.75% | 13.87% | 2.02% | -5.955% | -4.987% | -14.35%
| Z-Score (Absolute) |  | 0.45 | 1.57 | 0.32 | 1.5 | 0 | 1.83 | 2.18 | 2.16 | 0.08 | 0.1 | 0.32 | 1.23 | 0.54 | 2.08 | 1.96 | 0.39 | 0.67 | 0.55 | 1.79
| Analytical results for Ga in OREAS 266 (Certified Value 5.82 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 5.85 | 5.6414 | 6.8 | 6.76 | 5.82 | 6.3351 | 5.52 | 5 | 5.12 | 5.9681 | 7.21 | 5.87 | 6.19 | 4.93 | [5 empty] | 1
| 1 | 2 | 5.98 | 5.4482 | 6.8 | 6.24 | 5.41 | 6.1205 | 5.59 | 5 | 5.4 | 6.3494 | 7.27 | 5.56 | 6.29 | 4.86 | [5 empty] | 32
| 1 | 3 | 5.85 | 5.6998 | 6.7 | 6.32 | 5.56 | 6.1525 | 5.79 | 6 | 5.03 | 6.2596 | 7.17 | 5.65 | 5.36 | 4.81 | [5 empty] | 16
| 1 | 4 | 6.07 | 5.7002 | 7 | 6.31 | 5.62 | 6.2407 | 5.36 | 5 | 5.13 | 5.9692 | 7.57 | 5.76 | 6.38 | 4.88 | [5 empty] | 5.8248
| 1 | 5 | 5.82 | 5.1375 | 7.2 | 6.51 | 5.68 | 6.3068 | 5.58 | 5 | 5.02 | 6.1768 | 6.84 | 6.13 | 5.24 | 4.98 | [5 empty] | 89
| 1 | 6 | 5.71 | 5.7269 | 6.8 | 6.1 | 5.62 | 6.2785 | 5.6 | 5 | 5.01 | 6.396 | 7.76 | 5.57 | 5.49 | 5.06
| Mean |  | 5.88 | 5.559 | 6.8833 | 6.3733 | 5.6183 | 6.239 | 5.5733 | 5.1667 | 5.1183 | 6.1865 | 7.3033 | 5.7567 | 5.825 | 4.92
| Median |  | 5.85 | 5.6706 | 6.8 | 6.315 | 5.62 | 6.2596 | 5.585 | 5 | 5.075 | 6.2182 | 7.24 | 5.705 | 5.84 | 4.905
| Std Dev. |  | 0.12681 | 0.23006 | 0.18348 | 0.23132 | 0.13512 | 0.085913 | 0.1388 | 0.40825 | 0.14743 | 0.18482 | 0.32309 | 0.21778 | 0.51539 | 0.090111
| Rel.Std.Dev. |  | 2.157% | 4.139% | 2.666% | 3.629% | 2.405% | 1.377% | 2.491% | 7.902% | 2.881% | 2.987% | 4.424% | 3.783% | 8.848% | 1.832%
| PDM3 |  | 0.9479% | -4.563% | 18.17% | 9.417% | -3.544% | 7.112% | -4.317% | -11.3% | -12.13% | 6.21% | 25.38% | -1.17% | 0.003612% | -15.53%
| Z-Score (Absolute) |  | 0.1 | 0.5 | 1.97 | 1.02 | 0.39 | 0.77 | 0.47 | NA | 1.32 | 0.67 | 2.76 | 0.13 | 0 | 1.69
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 266 (Indicative Value 3.76 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.839 | 3.7519 | 4.1 | [16 empty] | 1
| 1 | 2 | 3.746 | 3.6126 | 3.5 | [16 empty] | 6
| 1 | 3 | 3.726 | 3.91 | 3.6 | [16 empty] | 16
| 1 | 4 | 3.799 | 3.9131 | 3.8 | [16 empty] | 3.7649
| 1 | 5 | 3.664 | 2.3445 | 3.9 | [16 empty] | 12
| 1 | 6 | 3.73 | 3.8655 | 3.5
| Mean |  | 3.7507 | 3.5663 | 3.7333
| Median |  | 3.738 | 3.8087 | 3.7
| Std Dev. |  | 0.061168 | 0.60948 | 0.24221
| Rel.Std.Dev. |  | 1.631% | 17.09% | 6.488%
| PDM3 |  | -0.3775% | -5.275% | -0.8379%
| Z-Score (Absolute) |  | 0.67 | 6.5 | 0
| Analytical results for Ge in OREAS 266 (Certified Value 0.11 ppm) | [20 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.09 | 1.10 | 1.12 | 1.15 | 1.19 | 1.20 | 1.21 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [11 empty] | 2
|  |  | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [11 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.16 | 0.15 | 0.1 | 0.11 | 0.1 | 0.1 | [11 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.16 | 0.15 | 0.1 | 0.1 | 0.15 | 0.09 | [11 empty] | 1
| 1 | 3 | < 0.2 | < 0.1 | 0.16 | 0.14 | 0.11 | 0.11 | 0.1 | 0.09 | [11 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.15 | 0.11 | 0.1 | 0.11 | 0.11 | 0.09 | [11 empty] | 0.10867
| 1 | 5 | < 0.2 | < 0.1 | 0.17 | 0.12 | 0.1 | 0.12 | 0.11 | 0.07 | [11 empty] | 90
| 1 | 6 | < 0.2 | < 0.1 | 0.15 | 0.14 | 0.1 | 0.11 | 0.08 | 0.09
| Mean |  |  |  | 0.15833 | 0.135 | 0.10167 | 0.11 | 0.10833 | 0.088333
| Median |  |  |  | 0.16 | 0.14 | 0.1 | 0.11 | 0.105 | 0.09
| Std Dev. |  |  |  | 0.0075277 | 0.016432 | 0.0040825 | 0.0063246 | 0.023166 | 0.0098319
| Rel.Std.Dev. |  |  |  | 4.754% | 12.17% | 4.016% | 5.75% | 21.38% | 11.13%
| PDM3 |  |  |  | 45.71% | 24.23% | -6.442% | 1.227% | -0.3067% | -18.71%
| Z-Score (Absolute) |  | 0.31 | 3.42 | 3.32 | 1.87 | 0.21 | 0.31 | 0.21 | 1.04
| Analytical results for Hf in OREAS 266 (Indicative Value 0.68 ppm) | [20 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.12 | 0.82545 | 1.21 | < 0.1 | 0.75 | 0.45 | 0.5 | 1.0273 | 0.89 | 0.37 | 0.37 | 0.31 | 1.03 | [6 empty] | 1
| 1 | 2 | 0.15 | 0.7627 | 1.17 | < 0.1 | 0.76 | 0.49 | 0.3 | 1.0252 | 0.85 | 0.34 | 0.58 | 0.31 | 1.07 | [6 empty] | 7
| 1 | 3 | 0.13 | 0.76885 | 1.18 | < 0.1 | 0.76 | 0.7 | 0.35 | 1.0548 | 0.87 | 0.3 | 0.7 | 0.29 | 1.07 | [6 empty] | 16
| 1 | 4 | 0.19 | 0.82844 | 1.21 | < 0.1 | 0.78 | 0.75 | 0.54 | 0.98012 | 0.94 | 0.39 | 0.34 | 0.29 | 1.01 | [6 empty] | 0.67948
| 1 | 5 | 0.15 | 0.48586 | 1.14 | < 0.1 | 0.73 | 0.94 | 0.54 | 1.0696 | 0.89 | 0.43 | 0.61 | 0.4 | 1.02 | [6 empty] | 13
| 1 | 6 | 0.13 | 0.87477 | 1.1 | < 0.1 | 0.72 | 0.52 | 0.49 | 0.98305 | 0.94 | 0.44 | 0.62 | 0.21 | 1.08
| Mean |  | 0.145 | 0.75768 | 1.1683 |  | 0.75 | 0.64167 | 0.45333 | 1.0234 | 0.89667 | 0.37833 | 0.53667 | 0.30167 | 1.0467
| Median |  | 0.14 | 0.79715 | 1.175 |  | 0.755 | 0.61 | 0.495 | 1.0263 | 0.89 | 0.38 | 0.595 | 0.3 | 1.05
| Std Dev. |  | 0.0251 | 0.13953 | 0.042622 |  | 0.021909 | 0.18904 | 0.1027 | 0.036442 | 0.036697 | 0.053448 | 0.14652 | 0.0608 | 0.030111
| Rel.Std.Dev. |  | 17.31% | 18.42% | 3.648% |  | 2.921% | 29.46% | 22.65% | 3.561% | 4.093% | 14.13% | 27.3% | 20.15% | 2.877%
| PDM3 |  | -78.66% | 11.51% | 71.95% |  | 10.38% | -5.565% | -33.28% | 50.61% | 31.96% | -44.32% | -21.02% | -55.6% | 54.04%
| Z-Score (Absolute) |  | 1.27 | 0.3 | 1.35 | 1.52 | 0.28 | 0 | 0.48 | 0.98 | 0.65 | 0.67 | 0.27 | 0.87 | 1.04
| Analytical results for Hg in OREAS 266 (Certified Value 0.039 ppm) | [20 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [5 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | < 0.1 | < 1 | 0.04 | < 2 | 0.04 | 0.03 | 0.16 | 0.05 | < 0.05 | 1 | 0.05 | 0.02 | 0.04 | 0.38 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 1 | 0.05 | < 2 | 0.06 | 0.05 | 0.16 | 0.03 | < 0.05 | 1 | 0.04 | 0.03 | 0.04 | 0.38 | [5 empty] | 1
| 1 | 3 | < 0.1 | < 1 | 0.02 | < 2 | 0.05 | 0.03 | 0.159 | 0.04 | < 0.05 | 1 | 0.05 | 0.03 | 0.04 | 0.38 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 1 | 0.04 | < 2 | 0.04 | 0.03 | 0.153 | 0.03 | < 0.05 | 1 | 0.06 | 0.05 | 0.05 | 0.36 | [5 empty] | 0.039143
| 1 | 5 | < 0.1 | < 1 | 0.04 | < 2 | 0.05 | 0.07 | 0.158 | 0.03 | < 0.05 | 1 | 0.05 | 0.02 | 0.07 | 0.4 | [5 empty] | 91
| 1 | 6 | < 0.1 | < 1 | 0.08 | < 2 | 0.05 | 0.03 | 0.157 | 0.03 | < 0.05 | 1 | 0.05 | 0.01 | 0.04 | 0.38
| Mean |  |  |  | 0.045 |  | 0.048333 | 0.04 | 0.15783 | 0.035 |  | 1 | 0.05 | 0.026667 | 0.046667 | 0.38
| Median |  |  |  | 0.04 |  | 0.05 | 0.03 | 0.1585 | 0.03 |  | 1 | 0.05 | 0.025 | 0.04 | 0.38
| Std Dev. |  |  |  | 0.019748 |  | 0.0075277 | 0.016733 | 0.0026394 | 0.0083666 |  | 0 | 0.0063246 | 0.013663 | 0.012111 | 0.012649
| Rel.Std.Dev. |  |  |  | 43.89% |  | 15.57% | 41.83% | 1.672% | 23.9% |  | 0% | 12.65% | 51.23% | 25.95% | 3.329%
| PDM3 |  |  |  | 14.96% |  | 23.48% | 2.19% | 303.2% | -10.58% |  | 2455% | 27.74% | -31.87% | 19.22% | 870.8%
| Z-Score (Absolute) |  | 0.08 | 22.84 | 0.17 | 48.13 | 0 | 0.42 | 5.54 | 0.67 | 1.18 | NA | 0.08 | 1.1 | 0.08 | 16.77
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 266 (Indicative Value 0.4 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.395 | 0.37796 | 0.4 | [16 empty] | 1
| 1 | 2 | 0.407 | 0.37002 | 0.4 | [16 empty] | 3
| 1 | 3 | 0.394 | 0.37317 | 0.4 | [16 empty] | 16
| 1 | 4 | 0.404 | 0.37649 | 0.4 | [16 empty] | 0.3995
| 1 | 5 | 0.395 | 0.23541 | 0.4 | [16 empty] | 9
| 1 | 6 | 0.399 | 0.38104 | 0.4
| Mean |  | 0.399 | 0.35235 | 0.4
| Median |  | 0.397 | 0.37483 | 0.4
| Std Dev. |  | 0.0054037 | 0.057416 | 6.0809e-17
| Rel.Std.Dev. |  | 1.354% | 16.3% | 1.52e-14%
| PDM3 |  | -0.1252% | -11.8% | 0.1252%
| Z-Score (Absolute) |  | 0 | 31.46 | 0.67
| Analytical results for In in OREAS 266 (Certified Value 0.031 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.03 | < 0.05 | 0.04 | < 5 | 0.027 | 0.029 | < 0.05 | 0.1 | 0.032 | 0.027 | 0.028 | 0.018 | [7 empty] | 1
| 1 | 2 | 0.04 | < 0.05 | 0.04 | < 5 | 0.028 | 0.029 | < 0.05 | 0.1 | 0.025 | 0.027 | 0.027 | 0.021 | [7 empty] | 20
| 1 | 3 | 0.03 | < 0.05 | 0.04 | < 5 | 0.03 | 0.028 | < 0.05 | 0.1 | 0.028 | 0.028 | 0.025 | 0.023 | [7 empty] | 16
| 1 | 4 | 0.04 | < 0.05 | 0.04 | < 5 | 0.029 | 0.032 | < 0.05 | 0.1 | 0.034 | 0.032 | 0.025 | 0.018 | [7 empty] | 0.030667
| 1 | 5 | 0.03 | < 0.05 | 0.04 | < 5 | 0.026 | 0.027 | < 0.05 | 0.1 | 0.036 | 0.028 | 0.027 | 0.017 | [7 empty] | 92
| 1 | 6 | 0.03 | < 0.05 | 0.03 | < 5 | 0.027 | 0.028 | < 0.05 | 0.1 | 0.027 | 0.034 | 0.028 | 0.02
| Mean |  | 0.033333 |  | 0.038333 |  | 0.027833 | 0.028833 |  | 0.1 | 0.030333 | 0.029333 | 0.026667 | 0.0195
| Median |  | 0.03 |  | 0.04 |  | 0.0275 | 0.0285 |  | 0.1 | 0.03 | 0.028 | 0.027 | 0.019
| Std Dev. |  | 0.005164 |  | 0.0040825 |  | 0.001472 | 0.0017224 |  | 1.5202e-17 | 0.0043205 | 0.0029439 | 0.0013663 | 0.0022583
| Rel.Std.Dev. |  | 15.49% |  | 10.65% |  | 5.288% | 5.974% |  | 1.52e-14% | 14.24% | 10.04% | 5.123% | 11.58%
| PDM3 |  | 8.696% |  | 25% |  | -9.239% | -5.978% |  | 226.1% | -1.087% | -4.348% | -13.04% | -36.41%
| Z-Score (Absolute) |  | 0.7 | 0.67 | 1.53 | 408.04 | 0.21 | 0.04 | 0.67 | 11.71 | 0.21 | 0.04 | 0.4 | 1.58
| Analytical results for K in OREAS 266 (Certified Value 0.275 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.2712 | 0.26821 | 0.23265 | 0.33 | 0.31 | 0.27 | 0.24 | 0.29 | 0.27 | 0.26 | 0.26 | 0.29738 | 0.33 | 0.29 | 0.26 | 0.25 | 0.212 |  |  | 1
| 1 | 2 | 0.2723 | 0.27128 | 0.23131 | 0.33 | 0.31 | 0.27 | 0.23 | 0.28 | 0.27 | 0.26 | 0.26 | 0.30034 | 0.32 | 0.28 | 0.27 | 0.25 | 0.216 |  |  | #N/A
| 1 | 3 | 0.2703 | 0.27735 | 0.2271 | 0.33 | 0.304 | 0.27 | 0.23 | 0.28 | 0.27 | 0.27 | 0.25 | 0.30047 | 0.33 | 0.28 | 0.25 | 0.25 | 0.218 |  |  | 16
| 1 | 4 | 0.2725 | 0.27109 | 0.22923 | 0.33 | 0.306 | 0.27 | 0.23 | 0.28 | 0.27 | 0.24 | 0.26 | 0.29821 | 0.32 | 0.28 | 0.26 | 0.25 | 0.195 |  |  | 0.2747
| 1 | 5 | 0.2612 | 0.26618 | 0.22371 | 0.34 | 0.302 | 0.27 | 0.23 | 0.28 | 0.27 | 0.26 | 0.25 | 0.29849 | 0.33 | 0.28 | 0.25 | 0.26 | 0.193 |  |  | 93
| 1 | 6 | 0.2734 | 0.2719 | 0.22659 | 0.32 | 0.306 | 0.28 | 0.23 | 0.29 | 0.27 | 0.27 | 0.25 | 0.29965 | 0.33 | 0.28 | 0.25 | 0.25 | 0.216
| Mean |  | 0.27015 | 0.271 | 0.22843 | 0.33 | 0.30633 | 0.27167 | 0.23167 | 0.28333 | 0.27 | 0.26 | 0.255 | 0.29909 | 0.32667 | 0.28167 | 0.25667 | 0.25167 | 0.20833
| Median |  | 0.27175 | 0.27118 | 0.22817 | 0.33 | 0.306 | 0.27 | 0.23 | 0.28 | 0.27 | 0.26 | 0.255 | 0.29907 | 0.33 | 0.28 | 0.255 | 0.25 | 0.214
| Std Dev. |  | 0.0045152 | 0.0038014 | 0.0032919 | 0.0063246 | 0.0032042 | 0.0040825 | 0.0040825 | 0.005164 | 0 | 0.010954 | 0.0054772 | 0.0012504 | 0.005164 | 0.0040825 | 0.008165 | 0.0040825 | 0.01129
| Rel.Std.Dev. |  | 1.671% | 1.403% | 1.441% | 1.917% | 1.046% | 1.503% | 1.762% | 1.823% | 0% | 4.213% | 2.148% | 0.4181% | 1.581% | 1.449% | 3.181% | 1.622% | 5.419%
| PDM3 |  | -1.655% | -1.345% | -16.84% | 20.13% | 11.52% | -1.103% | -15.66% | 3.145% | -1.709% | -5.35% | -7.17% | 8.88% | 18.92% | 2.538% | -6.563% | -8.383% | -24.16%
| Z-Score (Absolute) |  | 0 | 0.04 | 1.86 | 2.66 | 1.61 | 0.07 | 1.71 | 0.59 | 0.01 | 0.45 | 0.67 | 1.29 | 2.52 | 0.51 | 0.6 | 0.82 | 2.75
| Analytical results for La in OREAS 266 (Certified Value 25.2 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 24.473 | 24.3423 | 26.9 | 27.4 | 26 | 24.9133 | 26 | 24 | 25.6 | 27.077 | 20.4 | 27.7 | 27 | 24.4 | 23.7 |  |  |  |  | 1
| 1 | 2 | 24.494 | 24.6626 | 26.7 | 25.6 | 24.6 | 24.23 | 26.1 | 24 | 26.7 | 27.8074 | 20.5 | 25.6 | 28 | 24.1 | 24.1 |  |  |  |  | 21
| 1 | 3 | 24.518 | 25.2909 | 27.1 | 26.1 | 25.2 | 24.3367 | 25.1 | 25 | 25.3 | 28.0321 | 20.8 | 26.3 | 25.5 | 23.9 | 24.1 |  |  |  |  | 16
| 1 | 4 | 24.751 | 24.5598 | 27.4 | 26.7 | 25.3 | 24.71 | 24 | 22 | 26 | 27.0429 | 20.9 | 27.1 | 28.3 | 23.9 | 23.6 |  |  |  |  | 25.2001
| 1 | 5 | 24.582 | 22.0822 | 27.3 | 27.4 | 25.8 | 24.88 | 25.3 | 24 | 26.2 | 27.3963 | 20.4 | 28.9 | 24.9 | 23.5 | 23.9 |  |  |  |  | 94
| 1 | 6 | 25.088 | 24.6055 | 27.1 | 25.3 | 25.6 | 24.8367 | 25.6 | 25 | 25.2 | 27.8858 | 21.5 | 25.9 | 26.2 | 24.6 | 24.5
| Mean |  | 24.651 | 24.2572 | 27.0833 | 26.4167 | 25.4167 | 24.6511 | 25.35 | 24 | 25.8333 | 27.5403 | 20.75 | 26.9167 | 26.65 | 24.0667 | 23.9833
| Median |  | 24.55 | 24.5827 | 27.1 | 26.4 | 25.45 | 24.7733 | 25.45 | 24 | 25.8 | 27.6019 | 20.65 | 26.7 | 26.6 | 24 | 24
| Std Dev. |  | 0.23654 | 1.1121 | 0.25626 | 0.89759 | 0.49967 | 0.29505 | 0.76616 | 1.0954 | 0.57504 | 0.42786 | 0.42308 | 1.2432 | 1.3605 | 0.39328 | 0.32506
| Rel.Std.Dev. |  | 0.9595% | 4.585% | 0.9462% | 3.398% | 1.966% | 1.197% | 3.022% | 4.564% | 2.226% | 1.554% | 2.039% | 4.619% | 5.105% | 1.634% | 1.355%
| PDM3 |  | -2.179% | -3.741% | 7.473% | 4.828% | 0.8594% | -2.178% | 0.5949% | -4.762% | 2.513% | 9.286% | -17.66% | 6.812% | 5.754% | -4.498% | -4.828%
| Z-Score (Absolute) |  | 0.37 | 0.57 | 0.91 | 0.56 | 0.04 | 0.37 | 0 | 0.71 | 0.25 | 1.15 | 2.42 | 0.82 | 0.68 | 0.67 | 0.72
| Analytical results for Li in OREAS 266 (Certified Value 6.62 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 6.5 | 6.5681 | 6 | 6.6 | 7.7 | 7.1 | 6.6837 | 6.1 | 6.2 | 3.2658 | 7.6 | 6.7 | 6.1 | 6.2 | 8 |  |  |  |  | 1
| 1 | 2 | 6.6 | 6.3508 | 7 | 6.7 | 7.2 | 6.9 | 6.3969 | 6 | 6.3 | 5.2556 | 8.2 | 6.5 | 6.2 | 5.9 | 8 |  |  |  |  | #N/A
| 1 | 3 | 6.6 | 6.7629 | 7 | 6.8 | 7 | 6.9 | 6.4675 | 6 | 5.8 | 4.7775 | 8.1 | 6.6 | 5.7 | 6.1 | 8 |  |  |  |  | 16
| 1 | 4 | 6.6 | 6.5531 | 7 | 6.9 | 7.5 | 6.9 | 6.6008 | 6 | 6 | 3.3393 | 8.1 | 6.6 | 6.1 | 6 | 7 |  |  |  |  | 6.6225
| 1 | 5 | 6.7 | 6.1252 | 7 | 6.9 | 7.5 | 7.1 | 6.6082 | 6 | 5.9 | 5.1862 | 8.1 | 6.6 | 5.7 | 6.5 | 7 |  |  |  |  | 95
| 1 | 6 | 6.5 | 6.5192 | 6 | 6.8 | 7.2 | 6.7 | 6.6809 | 6 | 5.8 | 3.7256 | 8.5 | 6.5 | 5.8 | 6.1 | 8
| Mean |  | 6.5833 | 6.4799 | 6.6667 | 6.7833 | 7.35 | 6.9333 | 6.573 | 6.0167 | 6 | 4.2583 | 8.1 | 6.5833 | 5.9333 | 6.1333 | 7.6667
| Median |  | 6.6 | 6.5362 | 7 | 6.8 | 7.35 | 6.9 | 6.6045 | 6 | 5.95 | 4.2515 | 8.1 | 6.6 | 5.95 | 6.1 | 8
| Std Dev. |  | 0.075277 | 0.2178 | 0.5164 | 0.1169 | 0.25884 | 0.15055 | 0.11667 | 0.040825 | 0.20976 | 0.92071 | 0.28983 | 0.075277 | 0.22509 | 0.20656 | 0.5164
| Rel.Std.Dev. |  | 1.143% | 3.361% | 7.746% | 1.723% | 3.522% | 2.171% | 1.775% | 0.6785% | 3.496% | 21.62% | 3.578% | 1.143% | 3.794% | 3.368% | 6.736%
| PDM3 |  | -0.5909% | -2.153% | 0.6675% | 2.429% | 10.99% | 4.694% | -0.7469% | -9.148% | -9.399% | -35.7% | 22.31% | -0.5909% | -10.41% | -7.386% | 15.77%
| Z-Score (Absolute) |  | 0.02 | 0.14 | NA | 0.32 | 1.19 | 0.55 | 0 | 0.85 | 0.88 | 3.55 | 2.34 | 0.02 | 0.98 | 0.67 | NA
Results from laboratories 1.06 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 266 (Indicative Value 0.1 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.22 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [15 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.25g | [15 empty] | 2
| 1 | 1 | 0.102 | 0.10376 | 0.1 | 0.11 | [15 empty] | 1
| 1 | 2 | 0.099 | 0.10256 | < 0.1 | 0.11 | [15 empty] | 4
| 1 | 3 | 0.095 | 0.10773 | < 0.1 | 0.1 | [15 empty] | 16
| 1 | 4 | 0.101 | 0.10767 | < 0.1 | 0.1 | [15 empty] | 0.10409
| 1 | 5 | 0.105 | 0.065564 | < 0.1 | 0.1 | [15 empty] | 10
| 1 | 6 | 0.104 | 0.10969 | < 0.1 | 0.11
| Mean |  | 0.101 | 0.099497 | 0.1 | 0.105
| Median |  | 0.1015 | 0.10572 | 0.1 | 0.105
| Std Dev. |  | 0.0036332 | 0.016838 |  | 0.0054772
| Rel.Std.Dev. |  | 3.597% | 16.92% |  | 5.216%
| PDM3 |  | -2.973% | -4.417% | -3.933% | 0.87%
| Z-Score (Absolute) |  | 0.18 | 0.18 | 10.27 | 1.16
| Analytical results for Mg in OREAS 266 (Certified Value 0.965 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.96 | 1.166 | 0.96448 | 0.85084 | 1.14 | 1.01 | 1.21 | 0.89 | 0.94 | 0.91 | 0.94 | 0.88 | 1.0083 | 1.03 | 0.97 | 0.91 | 0.87 | 0.859 |  | 1
| 1 | 2 | 0.96 | 1.1244 | 0.9568 | 0.84376 | 1.17 | 1 | 1.09 | 0.85 | 0.92 | 0.9 | 0.92 | 0.9 | 1.0305 | 1.03 | 0.94 | 0.93 | 0.88 | 0.876 |  | #N/A
| 1 | 3 | 0.96 | 1.1167 | 0.97886 | 0.84656 | 1.16 | 1 | 1.11 | 0.85 | 0.96 | 0.9 | 0.95 | 0.86 | 1.0312 | 1.04 | 0.95 | 0.88 | 0.85 | 0.882 |  | 16
| 1 | 4 | 0.96 | 1.1397 | 0.96786 | 0.84073 | 1.15 | 1.01 | 1.12 | 0.85 | 0.96 | 0.9 | 0.88 | 0.86 | 0.99016 | 1.04 | 0.95 | 0.91 | 0.85 | 0.868 |  | 0.96536
| 1 | 5 | 0.92 | 1.1625 | 0.96761 | 0.83338 | 1.2 | 1.01 | 1.16 | 0.87 | 0.96 | 0.9 | 0.96 | 0.88 | 1.0339 | 1.04 | 0.95 | 0.86 | 0.91 | 0.829 |  | 96
| 1 | 6 | 0.96 | 1.1731 | 0.97215 | 0.8396 | 1.14 | 1.03 | 1.11 | 0.86 | 0.97 | 0.91 | 0.96 | 0.86 | 0.9959 | 1.04 | 0.94 | 0.87 | 0.87 | 0.88
| Mean |  | 0.95333 | 1.1471 | 0.96796 | 0.84248 | 1.16 | 1.01 | 1.1333 | 0.86167 | 0.95167 | 0.90333 | 0.935 | 0.87333 | 1.015 | 1.0367 | 0.95 | 0.89333 | 0.87167 | 0.86567
| Median |  | 0.96 | 1.1511 | 0.96773 | 0.84224 | 1.155 | 1.01 | 1.115 | 0.855 | 0.96 | 0.9 | 0.945 | 0.87 | 1.0194 | 1.04 | 0.95 | 0.895 | 0.87 | 0.872
| Std Dev. |  | 0.01633 | 0.023505 | 0.0073905 | 0.0060383 | 0.022804 | 0.010954 | 0.044121 | 0.016021 | 0.018348 | 0.005164 | 0.030822 | 0.01633 | 0.019423 | 0.005164 | 0.010954 | 0.027325 | 0.022286 | 0.019866
| Rel.Std.Dev. |  | 1.713% | 2.049% | 0.7635% | 0.7167% | 1.966% | 1.085% | 3.893% | 1.859% | 1.928% | 0.5717% | 3.296% | 1.87% | 1.914% | 0.4981% | 1.153% | 3.059% | 2.557% | 2.295%
| PDM3 |  | -1.246% | 18.82% | 0.2691% | -12.73% | 20.16% | 4.624% | 17.4% | -10.74% | -1.419% | -6.425% | -3.145% | -9.533% | 5.141% | 7.386% | -1.591% | -7.461% | -9.706% | -10.33%
| Z-Score (Absolute) |  | 0.02 | 1.87 | 0.16 | 1.03 | 1.99 | 0.56 | 1.74 | 0.85 | 0.01 | 0.45 | 0.15 | 0.74 | 0.61 | 0.82 | 0.01 | 0.55 | 0.75 | 0.81
| Analytical results for Mn in OREAS 266 (Certified Value 0.031 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.031 | 0.033371 | 0.031377 | 0.025929 | 0.02 | 0.0334 | 0.035 | 0.0293 | 0.0336 | 0.0304 | 0.0293 | 0.0289 | 0.033228 | 0.0341 | 0.0327 | 0.0296 | 0.0288 | 0.0262 |  | 1
| 1 | 2 | 0.0312 | 0.033628 | 0.031195 | 0.025918 | 0.02 | 0.0333 | 0.0318 | 0.028 | 0.0332 | 0.0301 | 0.0289 | 0.0296 | 0.034545 | 0.0354 | 0.0312 | 0.0304 | 0.0285 | 0.0266 |  | 23
| 1 | 3 | 0.031 | 0.033723 | 0.031802 | 0.025833 | 0.02 | 0.0331 | 0.0327 | 0.028 | 0.0331 | 0.0308 | 0.0299 | 0.0282 | 0.034431 | 0.0355 | 0.0313 | 0.0287 | 0.0277 | 0.0268 |  | 16
| 1 | 4 | 0.0312 | 0.03313 | 0.031304 | 0.025623 | 0.02 | 0.0335 | 0.0329 | 0.0281 | 0.0335 | 0.0308 | 0.0274 | 0.0284 | 0.033387 | 0.0354 | 0.032 | 0.0298 | 0.0277 | 0.0245 |  | 0.030743
| 1 | 5 | 0.03 | 0.032617 | 0.028669 | 0.025289 | 0.02 | 0.0331 | 0.0343 | 0.0286 | 0.0335 | 0.031 | 0.03 | 0.0284 | 0.033669 | 0.0346 | 0.0322 | 0.028 | 0.028 | 0.0242 |  | 97
| 1 | 6 | 0.0313 | 0.03366 | 0.031577 | 0.025445 | 0.02 | 0.0336 | 0.0317 | 0.0284 | 0.0335 | 0.0309 | 0.0299 | 0.0279 | 0.033763 | 0.0366 | 0.0309 | 0.0284 | 0.0283 | 0.0267
| Mean |  | 0.03095 | 0.033355 | 0.030988 | 0.025673 | 0.02 | 0.033333 | 0.033067 | 0.0284 | 0.0334 | 0.030667 | 0.029233 | 0.028567 | 0.033837 | 0.035267 | 0.031717 | 0.02915 | 0.028167 | 0.025833
| Median |  | 0.0311 | 0.033499 | 0.031341 | 0.025728 | 0.02 | 0.03335 | 0.0328 | 0.02825 | 0.0335 | 0.0308 | 0.0296 | 0.0284 | 0.033716 | 0.0354 | 0.03165 | 0.02915 | 0.02815 | 0.0264
| Std Dev. |  | 0.00048062 | 0.00042387 | 0.0011562 | 0.00026575 | 0 | 0.00020656 | 0.0013337 | 0.000502 | 0.0002 | 0.00034448 | 0.00099532 | 0.00060222 | 0.00054069 | 0.00085713 | 0.00069113 | 0.00092466 | 0.00044572 | 0.0011708
| Rel.Std.Dev. |  | 1.553% | 1.271% | 3.731% | 1.035% | 0% | 0.6197% | 4.033% | 1.768% | 0.5988% | 1.123% | 3.405% | 2.108% | 1.598% | 2.43% | 2.179% | 3.172% | 1.582% | 4.532%
| PDM3 |  | 0.6749% | 8.497% | 0.7969% | -16.49% | -34.94% | 8.427% | 7.56% | -7.62% | 8.644% | -0.2467% | -4.909% | -7.078% | 10.07% | 14.72% | 3.169% | -5.18% | -8.379% | -15.97%
| Z-Score (Absolute) |  | 0.04 | 0.7 | 0.05 | 1.4 | 2.95 | 0.69 | 0.62 | 0.66 | 0.71 | 0.04 | 0.43 | 0.61 | 0.83 | 1.22 | 0.25 | 0.45 | 0.72 | 1.36
| Analytical results for Mo in OREAS 266 (Certified Value 8.31 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 9 | 8.2383 | < 10 | 8.8 | 8.1 | 8.49 | 8.0033 | 7.25 | 7.8 | 9.953 | 7.48 | 7.5752 | 8.13 | 9.26 | 8.56 | 7.74 | 8.56 |  |  | 1
| 1 | 2 | 8.9 | 8.1451 | < 10 | 8.7 | 7.34 | 8.44 | 7.8667 | 7.39 | 7.9 | 9.647 | 7.69 | 7.864 | 8.47 | 8.18 | 9.24 | 7.56 | 8.77 |  |  | 24
| 1 | 3 | 8.8 | 8.4752 | < 10 | 8.9 | 7.5 | 8.43 | 7.88 | 7.94 | 8.7 | 9.925 | 8.03 | 7.8062 | 8.16 | 8.41 | 8.62 | 7.25 | 8.56 |  |  | 16
| 1 | 4 | 9 | 8.4207 | < 10 | 9.1 | 7.61 | 8.59 | 7.9367 | 7.64 | 8.3 | 9.561 | 7.35 | 7.5199 | 8.37 | 9.15 | 8.71 | 7.39 | 8.54 |  |  | 8.3065
| 1 | 5 | 8.9 | 8.2931 | < 10 | 8.7 | 7.88 | 8.49 | 7.9767 | 7.84 | 8.3 | 9.836 | 8.05 | 7.796 | 8.18 | 9.73 | 8.39 | 7.93 | 8.65 |  |  | 98
| 1 | 6 | 8.9 | 8.4434 | < 10 | 8.8 | 7.13 | 8.56 | 7.9333 | 7.21 | 8.3 | 9.655 | 7.89 | 7.5061 | 8.31 | 8.09 | 8.62 | 7.64 | 8.54
| Mean |  | 8.9167 | 8.336 |  | 8.8333 | 7.5933 | 8.5 | 7.9328 | 7.545 | 8.2167 | 9.7628 | 7.7483 | 7.6779 | 8.27 | 8.8033 | 8.69 | 7.585 | 8.6033
| Median |  | 8.9 | 8.3569 |  | 8.8 | 7.555 | 8.49 | 7.935 | 7.515 | 8.3 | 9.7455 | 7.79 | 7.6856 | 8.245 | 8.78 | 8.62 | 7.6 | 8.56
| Std Dev. |  | 0.075277 | 0.1311 |  | 0.15055 | 0.35404 | 0.063875 | 0.053058 | 0.30833 | 0.32506 | 0.16346 | 0.29137 | 0.16127 | 0.1352 | 0.66926 | 0.28969 | 0.2437 | 0.09136
| Rel.Std.Dev. |  | 0.8442% | 1.573% |  | 1.704% | 4.663% | 0.7515% | 0.6688% | 4.087% | 3.956% | 1.674% | 3.76% | 2.101% | 1.635% | 7.602% | 3.334% | 3.213% | 1.062%
| PDM3 |  | 7.345% | 0.3542% |  | 6.342% | -8.586% | 2.329% | -4.499% | -9.168% | -1.082% | 17.53% | -6.72% | -7.568% | -0.4398% | 5.981% | 4.617% | -8.686% | 3.573%
| Z-Score (Absolute) |  | 0.82 | 0.08 | 4.13 | 0.71 | 0.86 | 0.29 | 0.43 | 0.92 | 0.07 | 1.89 | 0.66 | 0.75 | 0 | 0.67 | 0.53 | 0.87 | 0.42
| Analytical results for Na in OREAS 266 (Certified Value 0.291 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.31 | 0.29629 | 0.29672 | 1.17 | 0.3 | 0.314 | 0.28 | 0.29 | 0.29 | 0.288 | 0.28 | 0.27104 | 0.33 | 0.33 | 0.3 | 0.28 | 0.207 |  |  | 1
| 1 | 2 | 0.31 | 0.29547 | 0.29366 | 1.2 | 0.3 | 0.29 | 0.27 | 0.28 | 0.29 | 0.287 | 0.29 | 0.29441 | 0.33 | 0.28 | 0.31 | 0.28 | 0.213 |  |  | #N/A
| 1 | 3 | 0.31 | 0.30477 | 0.28876 | 1.19 | 0.3 | 0.295 | 0.27 | 0.28 | 0.3 | 0.296 | 0.27 | 0.29348 | 0.34 | 0.28 | 0.29 | 0.27 | 0.214 |  |  | 16
| 1 | 4 | 0.31 | 0.2961 | 0.29111 | 1.19 | 0.3 | 0.303 | 0.27 | 0.29 | 0.3 | 0.27 | 0.28 | 0.28572 | 0.33 | 0.32 | 0.3 | 0.27 | 0.185 |  |  | 0.29105
| 1 | 5 | 0.3 | 0.29614 | 0.285 | 1.21 | 0.29 | 0.301 | 0.27 | 0.29 | 0.3 | 0.294 | 0.27 | 0.29369 | 0.33 | 0.32 | 0.28 | 0.3 | 0.183 |  |  | 99
| 1 | 6 | 0.31 | 0.29867 | 0.28756 | 1.18 | 0.29 | 0.291 | 0.27 | 0.29 | 0.29 | 0.296 | 0.27 | 0.27215 | 0.33 | 0.27 | 0.29 | 0.28 | 0.211
| Mean |  | 0.30833 | 0.29791 | 0.29047 | 1.19 | 0.29667 | 0.299 | 0.27167 | 0.28667 | 0.295 | 0.2885 | 0.27667 | 0.28508 | 0.33167 | 0.3 | 0.295 | 0.28 | 0.20217
| Median |  | 0.31 | 0.29621 | 0.28993 | 1.19 | 0.3 | 0.298 | 0.27 | 0.29 | 0.295 | 0.291 | 0.275 | 0.2896 | 0.33 | 0.3 | 0.295 | 0.28 | 0.209
| Std Dev. |  | 0.0040825 | 0.0035405 | 0.0042677 | 0.014142 | 0.005164 | 0.0090111 | 0.0040825 | 0.005164 | 0.0054772 | 0.0098742 | 0.008165 | 0.010922 | 0.0040825 | 0.026077 | 0.010488 | 0.010954 | 0.014289
| Rel.Std.Dev. |  | 1.324% | 1.188% | 1.469% | 1.188% | 1.741% | 3.014% | 1.503% | 1.801% | 1.857% | 3.423% | 2.951% | 3.831% | 1.231% | 8.692% | 3.555% | 3.912% | 7.068%
| PDM3 |  | 5.939% | 2.357% | -0.1989% | 308.9% | 1.931% | 2.733% | -6.659% | -1.505% | 1.358% | -0.8751% | -4.941% | -2.049% | 13.96% | 3.076% | 1.358% | -3.796% | -30.54%
| Z-Score (Absolute) |  | 1.08 | 0.24 | 0.37 | 72.42 | 0.13 | 0.32 | 1.89 | 0.67 | 0 | 0.53 | 1.48 | 0.8 | 2.97 | 0.4 | 0 | 1.21 | 7.51
| Analytical results for Nb in OREAS 266 (Certified Value 0.62 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.76 | 0.52744 | 13 | 1.1 | 0.5 | < 1 | < 5 | 0.64 | 0.54 | 1.165 | 0.7 | 0.7 | 0.77 | 0.53 | 0.4 |  |  |  |  | 1
| 1 | 2 | 0.77 | 0.5236 | 15 | 1 | 0.4 | < 1 | < 5 | 0.69 | 0.48 | 1.4358 | 0.9 | 0.67 | 0.92 | 0.52 | 0.43 |  |  |  |  | 25
| 1 | 3 | 0.74 | 0.56526 | 16 | 1.1 | 0.4 | < 1 | < 5 | 0.76 | 0.45 | 1.5899 | 0.7 | 0.67 | 0.96 | 0.48 | 0.41 |  |  |  |  | 16
| 1 | 4 | 0.72 | 0.53149 | 14 | 1.1 | 0.5 | < 1 | < 5 | 0.67 | 0.56 | 1.3412 | 0.8 | 0.76 | 0.82 | 0.49 | 0.42 |  |  |  |  | 0.62243
| 1 | 5 | 0.7 | 0.44086 | 16 | 1.1 | 0.4 | < 1 | < 5 | 0.78 | 0.51 | 1.4588 | 0.8 | 0.79 | 0.83 | 0.55 | 0.41 |  |  |  |  | 100
| 1 | 6 | 0.71 | 0.54088 | 31 | 1 | 0.5 | < 1 | < 5 | 0.72 | 0.49 | 1.2498 | 0.8 | 0.68 | 0.87 | 0.47 | 0.41
| Mean |  | 0.73333 | 0.52159 | 17.5 | 1.0667 | 0.45 |  |  | 0.71 | 0.505 | 1.3734 | 0.78333 | 0.71167 | 0.86167 | 0.50667 | 0.41333
| Median |  | 0.73 | 0.52947 | 15.5 | 1.1 | 0.45 |  |  | 0.705 | 0.5 | 1.3885 | 0.8 | 0.69 | 0.85 | 0.505 | 0.41
| Std Dev. |  | 0.028048 | 0.042269 | 6.7157 | 0.05164 | 0.054772 |  |  | 0.053666 | 0.040373 | 0.15349 | 0.075277 | 0.051153 | 0.06969 | 0.031411 | 0.010328
| Rel.Std.Dev. |  | 3.825% | 8.104% | 38.38% | 4.841% | 12.17% |  |  | 7.559% | 7.995% | 11.18% | 9.61% | 7.188% | 8.088% | 6.2% | 2.499%
| PDM3 |  | 17.82% | -16.2% | 2712% | 71.37% | -27.7% |  |  | 14.07% | -18.87% | 120.7% | 25.85% | 14.34% | 38.44% | -18.6% | -33.59%
| Z-Score (Absolute) |  | 0.07 | 0.62 | 55.22 | NA | NA | 0.69 | 5.89 | 0 | 0.68 | 2.18 | NA | 0 | 0.5 | 0.67 | 0.98
Results from laboratories 1.07, 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 266 (Indicative Value 24.8 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.09 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 0.5g | [17 empty] | 1
| 1 | 1 | 24.961 | 25 | [17 empty] | 1
| 1 | 2 | 25.564 | 23.7 | [17 empty] | 5
| 1 | 3 | 25.255 | 24.7 | [17 empty] | 16
| 1 | 4 | 25.243 | 24.6 | [17 empty] | 24.8433
| 1 | 5 | 25.21 | 25.1 | [17 empty] | 11
| 1 | 6 | 24.987 | 23.8
| Mean |  | 25.2033 | 24.4833
| Median |  | 25.2265 | 24.65
| Std Dev. |  | 0.21904 | 0.59805
| Rel.Std.Dev. |  | 0.8691% | 2.443%
| PDM3 |  | 1.449% | -1.449%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 266 (Certified Value 80 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 1
| 1 | 1 | 79.9 | 84.888 | 79.0971 | 75.796 | 65 | 77 | 84.9 | 83.3 | 76.6277 | 80.5 | 71.5 | 82.6 | 85.1522 | 84.4 | 86.9 | 79.9 | 78.1 | 60.2 |  | 1
| 1 | 2 | 80.5 | 81.6696 | 78.0951 | 76.839 | 68 | 77 | 86.4 | 79.3 | 75.7323 | 80.1 | 71.4 | 84.3 | 88.5112 | 86.8 | 80.8 | 85.2 | 76.9 | 59.6 |  | 27
| 1 | 3 | 80.1 | 82.3672 | 79.0017 | 75.47 | 68 | 77 | 84.1 | 79.8 | 75.8987 | 77.1 | 79.8 | 81.5 | 86.4365 | 87.4 | 82.9 | 81 | 76.5 | 59.4 |  | 16
| 1 | 4 | 80.5 | 82.0584 | 77.6036 | 73.621 | 68 | 76 | 84.4 | 81.1 | 76.1523 | 74.3 | 73.4 | 82.9 | 82.774 | 86.3 | 85.6 | 82 | 77.8 | 58.1 |  | 79.6198
| 1 | 5 | 77.4 | 82.35 | 72.48 | 72.597 | 71 | 77 | 85.9 | 81.4 | 76.4037 | 77.1 | 73 | 81.5 | 86.7415 | 88.8 | 89.6 | 78.5 | 79.5 | 58.6 |  | 101
| 1 | 6 | 81.1 | 84.6936 | 78.2155 | 74.739 | 66 | 76 | 85.8 | 81.5 | 76.5603 | 79.2 | 75.2 | 81.7 | 84.2088 | 89.2 | 81.6 | 80.9 | 79.3 | 60.5
| Mean |  | 79.9167 | 83.0045 | 77.4155 | 74.8437 | 67.6667 | 76.6667 | 85.25 | 81.0667 | 76.2292 | 78.05 | 74.05 | 82.4167 | 85.6374 | 87.15 | 84.5667 | 81.25 | 78.0167 | 59.4
| Median |  | 80.3 | 82.3586 | 78.1553 | 75.1045 | 68 | 77 | 85.35 | 81.25 | 76.278 | 78.15 | 73.2 | 82.15 | 85.7943 | 87.1 | 84.25 | 80.95 | 77.95 | 59.5
| Std Dev. |  | 1.2999 | 1.408 | 2.4836 | 1.5384 | 2.0656 | 0.5164 | 0.91815 | 1.4152 | 0.36357 | 2.3407 | 3.1444 | 1.0962 | 2.0294 | 1.7547 | 3.3945 | 2.2687 | 1.2205 | 0.91869
| Rel.Std.Dev. |  | 1.627% | 1.696% | 3.208% | 2.056% | 3.053% | 0.6736% | 1.077% | 1.746% | 0.4769% | 2.999% | 4.246% | 1.33% | 2.37% | 2.013% | 4.014% | 2.792% | 1.564% | 1.547%
| PDM3 |  | 0.3728% | 4.251% | -2.769% | -5.999% | -15.01% | -3.709% | 7.071% | 1.817% | -4.259% | -1.972% | -6.996% | 3.513% | 7.558% | 9.458% | 6.213% | 2.047% | -2.014% | -25.4%
| Z-Score (Absolute) |  | 0.17 | 0.73 | 0.28 | 0.75 | 2.05 | 0.42 | 1.13 | 0.38 | 0.5 | 0.17 | 0.89 | 0.62 | 1.2 | 1.48 | 1.01 | 0.41 | 0.17 | 3.54
| Analytical results for P in OREAS 266 (Certified Value 0.064 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0621 | 0.067272 | 0.063701 | 0.07 | 0.0684 | 0.066 | 0.0627 | 0.0593 | 0.065 | 0.062 | 0.062 | 0.070525 | 0.07 | 0.065 | 0.064 | 0.062 | 0.057 |  |  | 1
| 1 | 2 | 0.0624 | 0.068032 | 0.062819 | 0.07 | 0.0688 | 0.065 | 0.063 | 0.0578 | 0.064 | 0.061 | 0.062 | 0.072718 | 0.07 | 0.062 | 0.067 | 0.063 | 0.058 |  |  | #N/A
| 1 | 3 | 0.0619 | 0.068658 | 0.063775 | 0.07 | 0.0691 | 0.064 | 0.0626 | 0.0591 | 0.064 | 0.061 | 0.06 | 0.071924 | 0.07 | 0.062 | 0.064 | 0.061 | 0.059 |  |  | 16
| 1 | 4 | 0.0625 | 0.067556 | 0.062282 | 0.07 | 0.0681 | 0.066 | 0.0633 | 0.0585 | 0.062 | 0.058 | 0.062 | 0.070779 | 0.07 | 0.064 | 0.064 | 0.06 | 0.056 |  |  | 0.063826
| 1 | 5 | 0.0597 | 0.067293 | 0.061242 | 0.07 | 0.0685 | 0.066 | 0.063 | 0.059 | 0.064 | 0.063 | 0.062 | 0.07316 | 0.07 | 0.064 | 0.062 | 0.06 | 0.055 |  |  | 102
| 1 | 6 | 0.0627 | 0.067466 | 0.063247 | 0.07 | 0.0695 | 0.067 | 0.0626 | 0.0587 | 0.064 | 0.063 | 0.061 | 0.070779 | 0.07 | 0.063 | 0.063 | 0.062 | 0.057
| Mean |  | 0.061883 | 0.067713 | 0.062844 | 0.07 | 0.068733 | 0.065667 | 0.062867 | 0.058733 | 0.063833 | 0.061333 | 0.0615 | 0.071648 | 0.07 | 0.063333 | 0.064 | 0.061333 | 0.057
| Median |  | 0.06225 | 0.067511 | 0.063033 | 0.07 | 0.06865 | 0.066 | 0.06285 | 0.05885 | 0.064 | 0.0615 | 0.062 | 0.071352 | 0.07 | 0.0635 | 0.064 | 0.0615 | 0.057
| Std Dev. |  | 0.0011071 | 0.00053874 | 0.00096403 | 0 | 0.00050859 | 0.0010328 | 0.00028048 | 0.00053914 | 0.00098319 | 0.0018619 | 0.00083666 | 0.0011206 | 0 | 0.0012111 | 0.0016733 | 0.0012111 | 0.0014142
| Rel.Std.Dev. |  | 1.789% | 0.7956% | 1.534% | 0% | 0.74% | 1.573% | 0.4461% | 0.9179% | 1.54% | 3.036% | 1.36% | 1.564% | 0% | 1.912% | 2.615% | 1.975% | 2.481%
| PDM3 |  | -3.043% | 6.091% | -1.538% | 9.674% | 7.689% | 2.884% | -1.503% | -7.979% | 0.01201% | -3.905% | -3.644% | 12.26% | 9.674% | -0.7714% | 0.2731% | -3.905% | -10.69%
| Z-Score (Absolute) |  | 0.49 | 1.48 | 0.16 | 2.25 | 1.82 | 0.79 | 0.16 | 1.55 | 0.17 | 0.67 | 0.62 | 2.8 | 2.25 | 0 | 0.22 | 0.67 | 2.14
| Analytical results for Pb in OREAS 266 (Certified Value 9.95 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 10.1 | 10.0856 | < 100 | 19 | 10.6 | 11.1 | 9.6 | 10.1725 | 8.9 | 8.8 | 6.0001 | 9.7 | 10.9477 | 11.7 | 9.6 | 10.3 | 9.8 |  |  | 1
| 1 | 2 | 10.1 | 9.867 | < 100 | 16 | 10.2 | 9.7 | 9.3 | 9.7973 | 9 | 8.6 | 5.9932 | 8.8 | 11.093 | 12.2 | 9.9 | 10.5 | 9.8 |  |  | 28
| 1 | 3 | 10.1 | 10.3967 | < 100 | 13 | 10.3 | 10 | 9.4 | 10.0155 | 9.2 | 9.7 | 5.9933 | 8.3 | 11.4082 | 11.6 | 10.1 | 11.8 | 9.2 |  |  | 16
| 1 | 4 | 10.2 | 9.7205 | < 100 | 12 | 10.3 | 10.2 | 9.3 | 10.06 | 8.6 | 8.7 | 6.0176 | 9 | 11.0976 | 11.7 | 9.6 | 10.5 | 9.5 |  |  | 9.9482
| 1 | 5 | 10.1 | 9.3111 | < 100 | 14 | 10.4 | 10.3 | 9.1 | 9.9233 | 8.9 | 9 | 6.0662 | 8.4 | 10.8805 | 11.5 | 10.2 | 10.2 | 9.2 |  |  | 103
| 1 | 6 | 9.9 | 10.0106 | < 100 | 13 | 10.4 | 9.8 | 9.2 | 9.905 | 9 | 9.3 | 6.0975 | 8.4 | 11.5141 | 12 | 10.2 | 10.2 | 9.6
| Mean |  | 10.0833 | 9.8986 |  | 14.5 | 10.3667 | 10.1833 | 9.3167 | 9.9789 | 8.9333 | 9.0167 | 6.028 | 8.7667 | 11.1568 | 11.7833 | 9.9333 | 10.5833 | 9.5167
| Median |  | 10.1 | 9.9388 |  | 13.5 | 10.35 | 10.1 | 9.3 | 9.9694 | 8.95 | 8.9 | 6.0088 | 8.6 | 11.0953 | 11.7 | 10 | 10.4 | 9.55
| Std Dev. |  | 0.098319 | 0.36701 |  | 2.5884 | 0.13663 | 0.50365 | 0.17224 | 0.13187 | 0.19664 | 0.41673 | 0.04378 | 0.53166 | 0.25238 | 0.26394 | 0.28048 | 0.61128 | 0.27142
| Rel.Std.Dev. |  | 0.9751% | 3.708% |  | 17.85% | 1.318% | 4.946% | 1.849% | 1.321% | 2.201% | 4.622% | 0.7263% | 6.065% | 2.262% | 2.24% | 2.824% | 5.776% | 2.852%
| PDM3 |  | 1.359% | -0.4985% |  | 45.76% | 4.207% | 2.364% | -6.348% | 0.3093% | -10.2% | -9.364% | -39.41% | -11.88% | 12.15% | 18.45% | -0.1491% | 6.385% | -4.337%
| Z-Score (Absolute) |  | 0.11 | 0.08 | 40.75 | 4.6 | 0.39 | 0.21 | 0.67 | 0 | 1.06 | 0.98 | 4.02 | 1.23 | 1.2 | 1.84 | 0.05 | 0.62 | 0.47
| Analytical results for Pd in OREAS 266 (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 01g | [17 empty] | 1
| 1 | 1 | < 10 | < 10 | [17 empty] | 1
| 1 | 2 | < 10 | < 10 | [17 empty] | 6
| 1 | 3 | < 10 | < 10 | [17 empty] | 16
| 1 | 4 | < 10 | < 10 | [17 empty] | < 10
| 1 | 5 | < 10 | < 10 | [17 empty] | 12
| 1 | 6 | < 10 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Pr in OREAS 266 (Indicative Value 6.36 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 6.442 | 6.0432 | 7.1 | [16 empty] | 1
| 1 | 2 | 6.501 | 5.9106 | 6.3 | [16 empty] | 7
| 1 | 3 | 6.404 | 6.0812 | 6.6 | [16 empty] | 16
| 1 | 4 | 6.453 | 6.1721 | 6.6 | [16 empty] | 6.3606
| 1 | 5 | 6.355 | 3.7386 | 6.7 | [16 empty] | 13
| 1 | 6 | 6.287 | 6.1671 | 6.3
| Mean |  | 6.407 | 5.6855 | 6.6
| Median |  | 6.423 | 6.0622 | 6.6
| Std Dev. |  | 0.076538 | 0.95859 | 0.29665
| Rel.Std.Dev. |  | 1.195% | 16.86% | 4.495%
| PDM3 |  | 0.7293% | -10.61% | 3.764%
| Z-Score (Absolute) |  | 0 | 2.52 | 0.67
| Analytical results for Pt in OREAS 266 (Indicative Value < 5 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 25g | 01g | [17 empty] | 2
| 1 | 1 | < 5 | < 5 | [17 empty] | 1
| 1 | 2 | < 5 | < 5 | [17 empty] | 3
| 1 | 3 | < 5 | < 5 | [17 empty] | 16
| 1 | 4 | < 5 | < 5 | [17 empty] | < 5
| 1 | 5 | < 5 | < 5 | [17 empty] | 9
| 1 | 6 | < 5 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Rb in OREAS 266 (Certified Value 15.3 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 16.73 | 15.7936 | 14 | 16 | 16.1 | 13.6 | 15.7 | 15.5278 | 19.4 | 15.5 | 15.7 | 14.5 | 13.31 | [6 empty] | 1
| 1 | 2 | 16.6 | 15.875 | 14.1 | 14.8 | 15.8 | 13.8 | 16.8 | 16.121 | 20.2 | 15.4 | 16.4 | 14.2 | 13.46 | [6 empty] | 30
| 1 | 3 | 16.29 | 16.4463 | 14.6 | 14.6 | 15.6 | 14.7 | 14.5 | 15.8623 | 20.3 | 15.6 | 15.2 | 14.1 | 13.55 | [6 empty] | 16
| 1 | 4 | 16.56 | 16.2091 | 14.6 | 15.3 | 15.6 | 13.9 | 15.8 | 15.4597 | 19.9 | 15.2 | 16 | 14.1 | 13.07 | [6 empty] | 15.3084
| 1 | 5 | 16.18 | 16.1412 | 14.1 | 15.5 | 15.8 | 14.6 | 15 | 15.9949 | 19.5 | 16.2 | 14.6 | 13.8 | 13.01 | [6 empty] | 104
| 1 | 6 | 16.01 | 16.297 | 14.7 | 14.7 | 15.9 | 13.7 | 14.6 | 15.6536 | 20.5 | 15.6 | 15.5 | 14.5 | 13.26
| Mean |  | 16.395 | 16.127 | 14.35 | 15.15 | 15.8 | 14.05 | 15.4 | 15.7699 | 19.9667 | 15.5833 | 15.5667 | 14.2 | 13.2767
| Median |  | 16.425 | 16.1751 | 14.35 | 15.05 | 15.8 | 13.85 | 15.35 | 15.758 | 20.05 | 15.55 | 15.6 | 14.15 | 13.285
| Std Dev. |  | 0.27819 | 0.24999 | 0.31464 | 0.54681 | 0.18974 | 0.47645 | 0.87407 | 0.26488 | 0.44572 | 0.33714 | 0.62823 | 0.26833 | 0.21144
| Rel.Std.Dev. |  | 1.697% | 1.55% | 2.193% | 3.609% | 1.201% | 3.391% | 5.676% | 1.68% | 2.232% | 2.163% | 4.036% | 1.89% | 1.593%
| PDM3 |  | 7.098% | 5.348% | -6.26% | -1.034% | 3.212% | -8.22% | 0.5987% | 3.015% | 30.43% | 1.796% | 1.687% | -7.24% | -13.27%
| Z-Score (Absolute) |  | 1 | 0.67 | 1.46 | 0.5 | 0.28 | 1.83 | 0.2 | 0.24 | 5.29 | 0.02 | 0 | 1.64 | 2.76
| Analytical results for Re in OREAS 266 (Certified Value < 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.001 | 0.001 | [9 empty] | 1
| 1 | 2 | < 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 31
| 1 | 3 | < 0.001 | < 0.05 | 0.001 | < 0.002 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 16
| 1 | 4 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.001 | [9 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 105
| 1 | 6 | < 0.001 | < 0.05 | 0.002 | < 0.002 | < 0.001 | 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.001
| Mean |  |  |  | 0.0013333 |  | 0.001 | 0.001 |  | 0.001 |  | 0.001
| Median |  |  |  | 0.001 |  | 0.001 | 0.001 |  | 0.001 |  | 0.001
| Std Dev. |  |  |  | 0.00057735 |  | 0 | 0 |  | 0
| Rel.Std.Dev. |  |  |  | 43.3% |  | 0% | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.94 | 78.35 | 0.4 | 0.67 | 0.13 | 0.13 | 78.35 | 0.13 | 0.94 | 0.67
| Analytical results for S in OREAS 266 (Certified Value 0.03 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0285 | 0.028576 |  | 0.53 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.027225 | 0.03 | 0.032089 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 |  | 1
| 1 | 2 | 0.0287 | 0.029381 |  | 0.55 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.027225 | 0.03 | 0.031567 | 0.03 | 0.02 | 0.04 | 0.03 | 0.03 |  | 16
| 1 | 3 | 0.0285 | 0.02918 |  | 0.55 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.026725 | 0.03 | 0.031313 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 |  | 16
| 1 | 4 | 0.0287 | 0.028771 |  | 0.53 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.026324 | 0.03 | 0.030696 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 |  | 0.029922
| 1 | 5 | 0.0278 | 0.028101 |  | 0.56 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.027022 | 0.03 | 0.031871 | 0.03 | 0.04 | 0.03 | 0.05 | 0.03 |  | 106
| 1 | 6 | 0.029 | 0.028882 |  | 0.54 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.026724 | 0.03 | 0.030569 | 0.03 | 0.02 | 0.03 | 0.03 | 0.03
| Mean |  | 0.028533 | 0.028815 |  | 0.54333 | 0.03 |  | 0.03 | 0.03 | 0.03 |  | 0.026874 | 0.03 | 0.031351 | 0.03 | 0.031667 | 0.031667 | 0.033333 | 0.03
| Median |  | 0.0286 | 0.028826 |  | 0.545 | 0.03 |  | 0.03 | 0.03 | 0.03 |  | 0.026874 | 0.03 | 0.03144 | 0.03 | 0.035 | 0.03 | 0.03 | 0.03
| Std Dev. |  | 0.00040332 | 0.00045289 |  | 0.012111 | 0 |  | 0 | 0 | 0 |  | 0.00035082 | 0 | 0.00061697 | 0 | 0.0098319 | 0.0040825 | 0.008165 | 0
| Rel.Std.Dev. |  | 1.414% | 1.572% |  | 2.229% | 0% |  | 0% | 0% | 0% |  | 1.305% | 0% | 1.968% | 0% | 31.05% | 12.89% | 24.49% | 0%
| PDM3 |  | -4.641% | -3.699% |  | 1716% | 0.2609% |  | 0.2609% | 0.2609% | 0.2609% |  | -10.18% | 0.2609% | 4.775% | 0.2609% | 5.831% | 5.831% | 11.4% | 0.2609%
| Z-Score (Absolute) |  | 0.73 | 0.59 | NA | 256.25 | 0 | 14.73 | 0 | 0 | 0 | 2.5 | 1.56 | 0 | 0.67 | 0 | 0.83 | 0.83 | 1.66 | 0
| Analytical results for Sb in OREAS 266 (Certified Value 7.02 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 8.9 | 6.0911 | < 10 | 8.5 | 6.96 | 6.87 | 5.71 | 6.9733 | 6.52 | 6.8 | 7.87 | 5.86 | 7.9886 | 8.54 | 7.27 | 6.96 | 6.78 | 6.52 |  | 1
| 1 | 2 | 8.73 | 6.0249 | < 10 | 7.6 | 6.97 | 6.06 | 5.41 | 6.7933 | 6.65 | 6.8 | 7.839 | 5.98 | 7.554 | 8.77 | 6.43 | 7.65 | 6.6 | 6.66 |  | 32
| 1 | 3 | 8.78 | 6.039 | < 10 | 5.6 | 7.06 | 6.31 | 5.3 | 6.5867 | 7.15 | 7.4 | 7.785 | 5.61 | 7.8499 | 8.65 | 6.8 | 7.85 | 6.03 | 6.63 |  | 16
| 1 | 4 | 8.92 | 6.1055 | < 10 | 8 | 7.12 | 6.69 | 5.14 | 6.4267 | 6.85 | 7.1 | 7.575 | 6.03 | 8.1414 | 8.88 | 7.27 | 6.8 | 6.14 | 6.46 |  | 7.0221
| 1 | 5 | 8.8 | 5.4897 | < 10 | 8.2 | 7.07 | 6.75 | 5.46 | 6.6767 | 7.06 | 7 | 7.891 | 5.67 | 8.0523 | 8.65 | 7.87 | 7.43 | 6.9 | 6.6 |  | 107
| 1 | 6 | 8.67 | 5.9876 | < 10 | 15.2 | 6.74 | 6.41 | 5.36 | 6.3367 | 6.54 | 7 | 7.727 | 5.64 | 7.5939 | 8.88 | 6.88 | 7.38 | 6.25 | 6.44
| Mean |  | 8.8 | 5.9563 |  | 8.85 | 6.9867 | 6.515 | 5.3967 | 6.6322 | 6.795 | 7.0167 | 7.7812 | 5.7983 | 7.8633 | 8.7283 | 7.0867 | 7.345 | 6.45 | 6.5517
| Median |  | 8.79 | 6.0319 |  | 8.1 | 7.015 | 6.55 | 5.385 | 6.6317 | 6.75 | 7 | 7.812 | 5.765 | 7.9193 | 8.71 | 7.075 | 7.405 | 6.425 | 6.56
| Std Dev. |  | 0.09654 | 0.23265 |  | 3.2777 | 0.1356 | 0.30723 | 0.18917 | 0.23493 | 0.26868 | 0.22286 | 0.11726 | 0.18302 | 0.2438 | 0.13819 | 0.49738 | 0.40044 | 0.35956 | 0.091742
| Rel.Std.Dev. |  | 1.097% | 3.906% |  | 37.04% | 1.941% | 4.716% | 3.505% | 3.542% | 3.954% | 3.176% | 1.507% | 3.156% | 3.101% | 1.583% | 7.019% | 5.452% | 5.575% | 1.4%
| PDM3 |  | 25.32% | -15.18% |  | 26.03% | -0.5051% | -7.222% | -23.15% | -5.553% | -3.235% | -0.07792% | 10.81% | -17.43% | 11.98% | 24.3% | 0.9189% | 4.598% | -8.148% | -6.7%
| Z-Score (Absolute) |  | 1.92 | 0.94 | 1.9 | 1.97 | 0.1 | 0.38 | 1.5 | 0.26 | 0.1 | 0.13 | 0.89 | 1.1 | 0.98 | 1.84 | 0.2 | 0.46 | 0.44 | 0.34
| Analytical results for Sc in OREAS 266 (Certified Value 3.87 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 4.1 | 3.8357 | 4 | 4.4 | 3.9 | < 5 | 3.7 | 3.9 | 3.6 | 4.1039 | 4.6 | 4.1 | 3.9 | 3.8 | 4 |  |  |  |  | 1
| 1 | 2 | 4.1 | 3.8095 | 4 | 3.9 | 3.7 | < 5 | 3.8 | 3.8 | 3.8 | 3.7553 | 4.9 | 4.2 | 4 | 3.7 | 4 |  |  |  |  | 33
| 1 | 3 | 4.1 | 3.9176 | 4 | 4 | 3.7 | < 5 | 3.7 | 4 | 3.5 | 3.8341 | 4.9 | 4.3 | 3.6 | 3.6 | 4.1 |  |  |  |  | 16
| 1 | 4 | 4.1 | 3.845 | 4 | 4.1 | 3.7 | < 5 | 3.5 | 3.6 | 3.6 | 3.8612 | 4.8 | 4 | 4 | 3.7 | 4 |  |  |  |  | 3.8711
| 1 | 5 | 3.9 | 3.4805 | 4 | 4.2 | 3.7 | < 5 | 3.6 | 3.9 | 3.5 | 3.8015 | 4.8 | 4.3 | 3.5 | 3.5 | 4 |  |  |  |  | 108
| 1 | 6 | 4.1 | 3.8243 | 4 | 3.9 | 3.7 | < 5 | 3.7 | 4 | 3.5 | 4.086 | 5 | 4.2 | 3.7 | 3.9 | 4.1
| Mean |  | 4.0667 | 3.7854 | 4 | 4.0833 | 3.7333 |  | 3.6667 | 3.8667 | 3.5833 | 3.907 | 4.8333 | 4.1833 | 3.7833 | 3.7 | 4.0333
| Median |  | 4.1 | 3.83 | 4 | 4.05 | 3.7 |  | 3.7 | 3.9 | 3.55 | 3.8476 | 4.85 | 4.2 | 3.8 | 3.7 | 4
| Std Dev. |  | 0.08165 | 0.154 | 0 | 0.19408 | 0.08165 |  | 0.10328 | 0.15055 | 0.1169 | 0.14991 | 0.13663 | 0.1169 | 0.2137 | 0.14142 | 0.05164
| Rel.Std.Dev. |  | 2.008% | 4.068% | 0% | 4.753% | 2.187% |  | 2.817% | 3.894% | 3.262% | 3.837% | 2.827% | 2.795% | 5.648% | 3.822% | 1.28%
| PDM3 |  | 5.052% | -2.213% | 3.329% | 5.482% | -3.559% |  | -5.281% | -0.115% | -7.434% | 0.9266% | 24.86% | 8.065% | -2.268% | -4.42% | 4.19%
| Z-Score (Absolute) |  | 0.88 | 0.15 | NA | 0.95 | 0.34 | 4.88 | 0.59 | 0.15 | 0.89 | 0.3 | 3.7 | 1.31 | 0.16 | 0.46 | 0.76
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 266 (Indicative Value 0.55 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1.5 | < 1 | 0.8 | 0.8 | < 5 | 0.6 | < 0.5 | 0.403 | 0.8 | < 1 | 1.6 | 0.5 | 0.2 | 0.5 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1.5 | < 1 | 0.8 | 0.8 | < 5 | 0.7 | 0.9 | 0.362 | 0.7 | < 1 | 1.5 | 0.4 | 0.3 | 0.4 |  |  |  |  | 34
| 1 | 3 | < 1 | < 1.5 | < 1 | 0.7 | 0.7 | < 5 | 0.6 | 0.8 | 0.394 | 0.5 | < 1 | 1.9 | 0.4 | 0.3 | 0.4 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1.5 | < 1 | 0.7 | 0.7 | < 5 | 0.4 | 0.8 | 0.364 | 0.7 | < 1 | 1.6 | 0.6 | < 0.2 | 0.3 |  |  |  |  | 0.55448
| 1 | 5 | < 1 | < 1.5 | < 1 | 0.6 | 0.8 | < 5 | < 0.2 | 0.7 | 0.416 | 0.4 | < 1 | 1.7 | 0.6 | < 0.2 | 0.5 |  |  |  |  | 10
| 1 | 6 | < 1 | < 1.5 | < 1 | 0.6 | 0.8 | < 5 | 0.4 | 0.7 | 0.383 | 0.8 | < 1 | 1.8 | 0.4 | < 0.2 | 0.4
| Mean |  |  |  |  | 0.7 | 0.76667 |  | 0.54 | 0.78 | 0.387 | 0.65 |  | 1.6833 | 0.48333 | 0.26667 | 0.41667
| Median |  |  |  |  | 0.7 | 0.8 |  | 0.6 | 0.8 | 0.3885 | 0.7 |  | 1.65 | 0.45 | 0.3 | 0.4
| Std Dev. |  |  |  |  | 0.089443 | 0.05164 |  | 0.13416 | 0.083666 | 0.021522 | 0.16432 |  | 0.1472 | 0.098319 | 0.057735 | 0.075277
| Rel.Std.Dev. |  |  |  |  | 12.78% | 6.736% |  | 24.85% | 10.73% | 5.561% | 25.28% |  | 8.744% | 20.34% | 21.65% | 18.07%
| PDM3 |  |  |  |  | 26.24% | 38.27% |  | -2.612% | 40.67% | -30.21% | 17.23% |  | 203.6% | -12.83% | -51.91% | -24.85%
| Z-Score (Absolute) |  | 0 | 1.12 | 0 | 0.9 | 1.2 | 8.99 | 0.15 | 0.86 | 0.51 | 0.67 | 0 | 5.32 | 0.07 | 1.42 | 0.37
| Analytical results for Sm in OREAS 266 (Indicative Value 4.45 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 4.416 | 4.3525 | 5.6 | [16 empty] | 1
| 1 | 2 | 4.355 | 4.3175 | 4.3 | [16 empty] | 5
| 1 | 3 | 4.307 | 4.3582 | 4.2 | [16 empty] | 16
| 1 | 4 | 4.504 | 4.4232 | 4.4 | [16 empty] | 4.4516
| 1 | 5 | 4.433 | 2.655 | 4.6 | [16 empty] | 11
| 1 | 6 | 4.346 | 4.4385 | 4.4
| Mean |  | 4.3935 | 4.0908 | 4.5833
| Median |  | 4.3855 | 4.3554 | 4.4
| Std Dev. |  | 0.071411 | 0.7049 | 0.51543
| Rel.Std.Dev. |  | 1.625% | 17.23% | 11.25%
| PDM3 |  | -1.305% | -8.105% | 2.959%
| Z-Score (Absolute) |  | 0 | 1.08 | 0.67
| Analytical results for Sn in OREAS 266 (Certified Value 1.56 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 1.6 | 1.5189 | 1 | 1.8 | 1.91 | 1.6 | < 10 | 1.5 | 1.3 | 1.3313 | 1.9 | 1.7 | 1.6 | 1.3 | 1.4 |  |  |  |  | 1
| 1 | 2 | 1.52 | 1.4919 | 1 | 1.7 | 1.73 | 1.6 | < 10 | 1.5 | 1.4 | 1.5634 | 2 | 1.4 | 1.5 | 1.3 | 1.4 |  |  |  |  | 19
| 1 | 3 | 1.54 | 1.514 | 4 | 1.9 | 1.73 | 1.5 | < 10 | 1.5 | 1.3 | 1.6468 | 2 | 1.4 | 1.5 | 1.3 | 1.4 |  |  |  |  | 16
| 1 | 4 | 1.63 | 1.5189 | 3 | 1.9 | 1.82 | 1.5 | < 10 | 1.4 | 1.3 | 1.2139 | 1.9 | 1.6 | 1.6 | 1.3 | 1.4 |  |  |  |  | 1.558
| 1 | 5 | 1.56 | 1.2648 | 6 | 1.9 | 1.91 | 1.6 | < 10 | 1.4 | 1.3 | 1.5458 | 2 | 1.8 | 1.5 | 1.3 | 1.4 |  |  |  |  | 109
| 1 | 6 | 1.6 | 1.5721 | < 1 | 1.9 | 1.74 | 1.6 | < 10 | 1.4 | 1.3 | 1.1912 | 2 | 1.4 | 1.5 | 1.3 | 1.4
| Mean |  | 1.575 | 1.4801 | 3 | 1.85 | 1.8067 | 1.5667 |  | 1.45 | 1.3167 | 1.4154 | 1.9667 | 1.55 | 1.5333 | 1.3 | 1.4
| Median |  | 1.58 | 1.5164 | 3 | 1.9 | 1.78 | 1.6 |  | 1.45 | 1.3 | 1.4385 | 2 | 1.5 | 1.5 | 1.3 | 1.4
| Std Dev. |  | 0.041833 | 0.10874 | 2.1213 | 0.083666 | 0.086872 | 0.05164 |  | 0.054772 | 0.040825 | 0.19513 | 0.05164 | 0.17607 | 0.05164 | 0 | 2.4324e-16
| Rel.Std.Dev. |  | 2.656% | 7.347% | 70.71% | 4.522% | 4.808% | 3.296% |  | 3.777% | 3.101% | 13.79% | 2.626% | 11.36% | 3.368% | 0% | 1.737e-14%
| PDM3 |  | 1.094% | -4.999% | 92.56% | 18.74% | 15.96% | 0.5586% |  | -6.93% | -15.49% | -9.152% | 26.23% | -0.5111% | -1.581% | -16.56% | -10.14%
| Z-Score (Absolute) |  | 0.17 | 0.31 | NA | 1.55 | 1.33 | 0.13 | 17.41 | 0.46 | 1.13 | 0.64 | 2.14 | 0.04 | 0.04 | 1.22 | 0.71
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 266 (Certified Value 66 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  | 0
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 70.03 | 65.0784 | 218 | 64.6 | 86.6 | 67 | 68.6467 | 66.6 | 63.8 | 66.2 | 64.0127 | 65.3 | 70.9 | 67.2 | 65.6 | 48.6 |  |  |  | 1
| 1 | 2 | 68.54 | 64.9671 | 224 | 64.7 | 78.3 | 65 | 66.22 | 66.4 | 62.7 | 69.2 | 67.2991 | 68.4 | 64.7 | 69 | 63.5 | 49.7 |  |  |  | 20
| 1 | 3 | 68.49 | 67.2409 | 227 | 63 | 81 | 66 | 66.82 | 64.5 | 65.8 | 64.6 | 65.9152 | 68.7 | 66.2 | 65.7 | 60.9 | 50.4 |  |  |  | 16
| 1 | 4 | 69.36 | 65.3428 | 219 | 64.8 | 81.9 | 65 | 68.0867 | 62.5 | 59.2 | 66.4 | 64.6657 | 66.9 | 69.2 | 69.5 | 62 | 45.3 |  |  |  | 66.0081
| 1 | 5 | 68.24 | 60.1599 | 229 | 62.9 | 85.3 | 67 | 68.38 | 63.7 | 64.8 | 66.3 | 66.849 | 67.7 | 73.2 | 64.2 | 59.7 | 48.7 |  |  |  | 110
| 1 | 6 | 66.94 | 65.1055 | 218 | 64.9 | 81.9 | 66 | 68.4033 | 66.5 | 65.5 | 63.7 | 64.6532 | 68.7 | 65.8 | 67.3 | 64.2 | 49.6
| Mean |  | 68.6 | 64.6491 | 222.5 | 64.15 | 82.5 | 66 | 67.7594 | 65.0333 | 63.6333 | 66.0667 | 65.5658 | 67.6167 | 68.3333 | 67.15 | 62.65 | 48.7167
| Median |  | 68.515 | 65.0919 | 221.5 | 64.65 | 81.9 | 66 | 68.2333 | 65.45 | 64.3 | 66.25 | 65.2904 | 68.05 | 67.7 | 67.25 | 62.75 | 49.15
| Std Dev. |  | 1.051 | 2.3599 | 4.8477 | 0.93541 | 3.0087 | 0.89443 | 0.99462 | 1.7294 | 2.4533 | 1.8843 | 1.3288 | 1.3303 | 3.3224 | 1.9887 | 2.1916 | 1.8038
| Rel.Std.Dev. |  | 1.532% | 3.65% | 2.179% | 1.458% | 3.647% | 1.355% | 1.468% | 2.659% | 3.855% | 2.852% | 2.027% | 1.967% | 4.862% | 2.962% | 3.498% | 3.703%
| PDM3 |  | 3.927% | -2.059% | 237.1% | -2.815% | 24.98% | -0.0123% | 2.653% | -1.477% | -3.598% | 0.0887% | -0.6701% | 2.437% | 3.523% | 1.73% | -5.087% | -26.2%
| Z-Score (Absolute) |  | 0.96 | 0.52 | 58.46 | 0.7 | 6.15 | 0.01 | 0.64 | 0.37 | 0.9 | 0.01 | 0.17 | 0.59 | 0.86 | 0.42 | 1.26 | 6.47
| Analytical results for Ta in OREAS 266 (Certified Value < 0.01 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 21
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 111
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01
| Mean | [6 empty] | 0.01 |  |  | 0.3 |  | 0.01
| Median | [6 empty] | 0.01 |  |  | 0.3 |  | 0.01
| Std Dev. | [6 empty] | 0 |  |  | 0 |  | 0
| Rel.Std.Dev. | [6 empty] | 0% |  |  | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 7.42 | 7.42 | 0.67 | 1008.77 | 0 | 0.67 | 7.42 | 118.68 | 0.67 | 0 | 0.67
| Analytical results for Tb in OREAS 266 (Certified Value 0.5 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.479 | 0.48706 | 0.5 | 0.5 | 0.54 | [14 empty] | 1
| 1 | 2 | 0.473 | 0.47118 | 0.5 | 0.5 | 0.53 | [14 empty] | 22
| 1 | 3 | 0.471 | 0.4856 | 0.5 | 0.5 | 0.53 | [14 empty] | 16
| 1 | 4 | 0.481 | 0.49251 | 0.5 | 0.6 | 0.51 | [14 empty] | 0.50396
| 1 | 5 | 0.479 | 0.30407 | 0.5 | 0.5 | 0.51 | [14 empty] | 112
| 1 | 6 | 0.463 | 0.49104 | 0.5 | 0.6 | 0.54
| Mean |  | 0.47433 | 0.45524 | 0.5 | 0.53333 | 0.52667
| Median |  | 0.476 | 0.48633 | 0.5 | 0.5 | 0.53
| Std Dev. |  | 0.0067725 | 0.074445 | 0 | 0.05164 | 0.013663
| Rel.Std.Dev. |  | 1.428% | 16.35% | 0% | 9.682% | 2.594%
| PDM3 |  | -5.879% | -9.667% | -0.7862% | 5.828% | 4.505%
| Z-Score (Absolute) |  | 0.65 | 1.13 | 0 | 0.84 | 0.67
| Analytical results for Te in OREAS 266 (Certified Value 0.058 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.1 | < 0.1 | 0.06 | 0.07 | 0.07 | < 5 | 0.05 | < 0.2 | 0.054 | 0.06 | < 0.05 | 0.12 | 0.06 | 0.05 | 0.05 |  |  |  |  | 1
| 1 | 2 | < 0.1 | 0.12131 | 0.05 | 0.09 | 0.06 | < 5 | 0.07 | < 0.2 | 0.057 | 0.07 | < 0.05 | 0.13 | 0.06 | 0.03 | 0.05 |  |  |  |  | 23
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | 0.05 | 0.06 | < 5 | 0.05 | < 0.2 | 0.053 | 0.06 | < 0.05 | 0.1 | 0.05 | 0.06 | 0.05 |  |  |  |  | 16
| 1 | 4 | < 0.1 | < 0.1 | 0.05 | 0.07 | 0.06 | < 5 | 0.04 | < 0.2 | 0.053 | 0.05 | < 0.05 | 0.07 | 0.06 | 0.05 | 0.05 |  |  |  |  | 0.057689
| 1 | 5 | < 0.1 | < 0.1 | 0.05 | 0.06 | 0.06 | < 5 | 0.07 | < 0.2 | 0.053 | 0.05 | < 0.05 | 0.1 | 0.08 | 0.06 | 0.05 |  |  |  |  | 113
| 1 | 6 | < 0.1 | < 0.1 | 0.07 | 0.09 | 0.07 | < 5 | 0.06 | < 0.2 | 0.053 | 0.05 | < 0.05 | 0.06 | 0.04 | 0.06 | 0.06
| Mean |  |  | 0.12131 | 0.056 | 0.071667 | 0.063333 |  | 0.056667 |  | 0.053833 | 0.056667 |  | 0.096667 | 0.058333 | 0.051667 | 0.051667
| Median |  |  | 0.12131 | 0.05 | 0.07 | 0.06 |  | 0.055 |  | 0.053 | 0.055 |  | 0.1 | 0.06 | 0.055 | 0.05
| Std Dev. |  |  |  | 0.0089443 | 0.016021 | 0.005164 |  | 0.012111 |  | 0.0016021 | 0.008165 |  | 0.027325 | 0.013292 | 0.01169 | 0.0040825
| Rel.Std.Dev. |  |  |  | 15.97% | 22.35% | 8.154% |  | 21.37% |  | 2.976% | 14.41% |  | 28.27% | 22.79% | 22.63% | 7.902%
| PDM3 |  |  | 110.3% | -2.928% | 24.23% | 9.784% |  | -1.772% |  | -6.683% | -1.772% |  | 67.57% | 1.117% | -10.44% | -10.44%
| Z-Score (Absolute) |  | 0.77 | 0.6 | 0.67 | 1.73 | 0.77 | 282.44 | 0 | 5.01 | 0.33 | 0 | 3.66 | 4.62 | 0.19 | 0.58 | 0.58
| Analytical results for Th in OREAS 266 (Certified Value 6.75 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7.14 | 6.7104 | < 100 | 6.87 | 7.6 | 6.5 | 6.6 | 6.6 | 6 | 8.2441 | 6.54 | 7.2 | 7.4 | 6.9 | 6.79 |  |  |  |  | 1
| 1 | 2 | 7.22 | 6.5579 | < 100 | 6.76 | 7 | 6.2 | 6.6 | 6.4 | 6.7 | 8.2884 | 6.3 | 6.5 | 7.7 | 6.9 | 6.75 |  |  |  |  | 24
| 1 | 3 | 7.26 | 6.7772 | < 100 | 6.94 | 7.1 | 6.3 | 7 | 7.1 | 6 | 8.3019 | 6.53 | 6.9 | 6.9 | 6.4 | 6.87 |  |  |  |  | 16
| 1 | 4 | 7.23 | 6.743 | < 100 | 6.91 | 7.2 | 6.3 | 6.6 | 6.5 | 5.9 | 8.2441 | 6.68 | 7.1 | 7.6 | 6.6 | 6.81 |  |  |  |  | 6.7478
| 1 | 5 | 7.11 | 6.5218 | < 100 | 7.02 | 7.4 | 6.3 | 6.9 | 6.5 | 5.8 | 8.254 | 6.43 | 7.6 | 7 | 5.9 | 6.75 |  |  |  |  | 114
| 1 | 6 | 7.02 | 6.9298 | < 100 | 6.81 | 7 | 6.4 | 6.6 | 6.9 | 5.6 | 8.2859 | 6.73 | 6.9 | 7.2 | 6.7 | 6.8
| Mean |  | 7.1633 | 6.7067 |  | 6.885 | 7.2167 | 6.3333 | 6.7167 | 6.6667 | 6 | 8.2698 | 6.535 | 7.0333 | 7.3 | 6.5667 | 6.795
| Median |  | 7.18 | 6.7267 |  | 6.89 | 7.15 | 6.3 | 6.6 | 6.55 | 5.95 | 8.27 | 6.535 | 7 | 7.3 | 6.65 | 6.795
| Std Dev. |  | 0.09048 | 0.14997 |  | 0.093113 | 0.24014 | 0.10328 | 0.18348 | 0.27325 | 0.37417 | 0.025313 | 0.15834 | 0.36697 | 0.32249 | 0.37771 | 0.044609
| Rel.Std.Dev. |  | 1.263% | 2.236% |  | 1.352% | 3.328% | 1.631% | 2.732% | 4.099% | 6.236% | 0.3061% | 2.423% | 5.218% | 4.418% | 5.752% | 0.6565%
| PDM3 |  | 6.158% | -0.6096% |  | 2.033% | 6.948% | -6.143% | -0.4617% | -1.203% | -11.08% | 22.55% | -3.154% | 4.231% | 8.183% | -2.685% | 0.6992%
| Z-Score (Absolute) |  | 0.96 | 0.23 | 112.05 | 0.23 | 1.09 | 1.2 | 0.2 | 0.33 | 2.06 | 3.82 | 0.67 | 0.62 | 1.31 | 0.59 | 0
| Analytical results for Ti in OREAS 266 (Certified Value 0.142 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.1542 | 0.15158 | 0.11389 | 0.13 | 0.195 | 0.2 | 0.128 | 0.14 | 0.143 | 0.136 | 0.151 | 0.26227 | 0.21 | 0.17 | 0.148 | 0.144 | 0.108 |  |  | 1
| 1 | 2 | 0.1526 | 0.15326 | 0.11185 | 0.15 | 0.197 | 0.2 | 0.125 | 0.14 | 0.146 | 0.135 | 0.152 | 0.25238 | 0.2 | 0.158 | 0.152 | 0.143 | 0.109 |  |  | 25
| 1 | 3 | 0.1506 | 0.15871 | 0.10604 | 0.16 | 0.198 | 0.2 | 0.118 | 0.14 | 0.148 | 0.142 | 0.137 | 0.26657 | 0.21 | 0.161 | 0.144 | 0.137 | 0.111 |  |  | 16
| 1 | 4 | 0.1535 | 0.15306 | 0.11045 | 0.14 | 0.199 | 0.21 | 0.119 | 0.15 | 0.147 | 0.126 | 0.148 | 0.25769 | 0.21 | 0.168 | 0.15 | 0.139 | 0.1 |  |  | 0.14219
| 1 | 5 | 0.1458 | 0.1526 | 0.10591 | 0.17 | 0.199 | 0.2 | 0.124 | 0.14 | 0.149 | 0.138 | 0.143 | 0.25377 | 0.21 | 0.166 | 0.14 | 0.143 | 0.096 |  |  | 115
| 1 | 6 | 0.1528 | 0.15509 | 0.10772 | 0.13 | 0.193 | 0.21 | 0.122 | 0.14 | 0.15 | 0.141 | 0.142 | 0.2658 | 0.21 | 0.162 | 0.142 | 0.14 | 0.109
| Mean |  | 0.15158 | 0.15405 | 0.10931 | 0.14667 | 0.19683 | 0.20333 | 0.12267 | 0.14167 | 0.14717 | 0.13633 | 0.1455 | 0.25975 | 0.20833 | 0.16417 | 0.146 | 0.141 | 0.1055
| Median |  | 0.1527 | 0.15316 | 0.10909 | 0.145 | 0.1975 | 0.2 | 0.123 | 0.14 | 0.1475 | 0.137 | 0.1455 | 0.25998 | 0.21 | 0.164 | 0.146 | 0.1415 | 0.1085
| Std Dev. |  | 0.0030805 | 0.0025532 | 0.0032704 | 0.01633 | 0.0024014 | 0.005164 | 0.0037771 | 0.0040825 | 0.0024833 | 0.0057504 | 0.0058224 | 0.006061 | 0.0040825 | 0.0045789 | 0.0047329 | 0.0027568 | 0.0060249
| Rel.Std.Dev. |  | 2.032% | 1.657% | 2.992% | 11.13% | 1.22% | 2.54% | 3.079% | 2.882% | 1.687% | 4.218% | 4.002% | 2.333% | 1.96% | 2.789% | 3.242% | 1.955% | 5.711%
| PDM3 |  | 6.603% | 8.338% | -23.13% | 3.145% | 38.43% | 43% | -13.73% | -0.3713% | 3.497% | -4.122% | 2.325% | 82.67% | 46.51% | 15.45% | 2.676% | -0.8401% | -25.81%
| Z-Score (Absolute) |  | 0.32 | 0.48 | 2.44 | 0 | 3.27 | 3.7 | 1.57 | 0.33 | 0.03 | 0.67 | 0.08 | 7.38 | 4.02 | 1.14 | 0.04 | 0.37 | 2.69
| Analytical results for Tl in OREAS 266 (Certified Value 0.12 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.12 | 0.11582 | 0.13 | 0.14 | 0.14 | < 10 | 0.12 | < 0.1 | 0.117 | 0.11 | 0.11685 | 0.16 | 0.11 | 0.12 | 0.11 |  |  |  |  | 1
| 1 | 2 | 0.12 | 0.11102 | 0.13 | 0.13 | 0.12 | < 10 | 0.11 | < 0.1 | 0.114 | 0.11 | 0.1191 | 0.17 | 0.11 | 0.13 | 0.1 |  |  |  |  | 26
| 1 | 3 | 0.12 | 0.11698 | 0.13 | 0.14 | 0.15 | < 10 | 0.12 | 0.1 | 0.114 | 0.1 | 0.12294 | 0.16 | 0.11 | 0.12 | 0.11 |  |  |  |  | 16
| 1 | 4 | 0.12 | 0.11676 | 0.13 | 0.14 | 0.15 | < 10 | 0.11 | < 0.1 | 0.111 | 0.1 | 0.1156 | 0.16 | 0.11 | 0.13 | 0.11 |  |  |  |  | 0.11894
| 1 | 5 | 0.13 | < 0.1 | 0.13 | 0.14 | 0.14 | < 10 | 0.12 | < 0.1 | 0.116 | 0.1 | 0.12376 | 0.16 | 0.12 | 0.12 | 0.12 |  |  |  |  | 116
| 1 | 6 | 0.12 | 0.12112 | 0.13 | 0.14 | 0.15 | < 10 | 0.11 | < 0.1 | 0.114 | 0.1 | 0.12514 | 0.17 | 0.11 | 0.12 | 0.11
| Mean |  | 0.12167 | 0.11634 | 0.13 | 0.13833 | 0.14167 |  | 0.115 | 0.1 | 0.11433 | 0.10333 | 0.12057 | 0.16333 | 0.11167 | 0.12333 | 0.11
| Median |  | 0.12 | 0.11676 | 0.13 | 0.14 | 0.145 |  | 0.115 | 0.1 | 0.114 | 0.1 | 0.12102 | 0.16 | 0.11 | 0.12 | 0.11
| Std Dev. |  | 0.0040825 | 0.0036059 | 0 | 0.0040825 | 0.01169 |  | 0.0054772 |  | 0.0020656 | 0.005164 | 0.0039342 | 0.005164 | 0.0040825 | 0.005164 | 0.0063246
| Rel.Std.Dev. |  | 3.355% | 3.099% | 0% | 2.951% | 8.252% |  | 4.763% |  | 1.807% | 4.997% | 3.263% | 3.162% | 3.656% | 4.187% | 5.75%
| PDM3 |  | 2.291% | -2.188% | 9.298% | 16.3% | 19.11% |  | -3.314% | -15.92% | -3.874% | -13.12% | 1.366% | 37.32% | -6.116% | 3.693% | -7.517%
| Z-Score (Absolute) |  | 0.07 | 0.98 | 0.6 | 1.13 | 1.35 | 311.41 | 0.36 | 3.97 | 0.4 | 1.1 | 0 | 2.73 | 0.57 | 0.18 | 0.67
| Analytical results for Tm in OREAS 266 (Indicative Value 0.12 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.137 | 0.11283 | 0.1 | [16 empty] | 1
| 1 | 2 | 0.132 | 0.10987 | 0.1 | [16 empty] | 6
| 1 | 3 | 0.13 | 0.11722 | 0.1 | [16 empty] | 16
| 1 | 4 | 0.135 | 0.12024 | 0.1 | [16 empty] | 0.11673
| 1 | 5 | 0.134 | 0.070726 | 0.1 | [16 empty] | 12
| 1 | 6 | 0.131 | 0.12501 | 0.1
| Mean |  | 0.13317 | 0.10932 | 0.1
| Median |  | 0.133 | 0.11502 | 0.1
| Std Dev. |  | 0.0026394 | 0.019647 | 1.5202e-17
| Rel.Std.Dev. |  | 1.982% | 17.97% | 1.52e-14%
| PDM3 |  | 14.08% | -6.354% | -14.33%
| Z-Score (Absolute) |  | 1.73 | 0 | 0.67
| Analytical results for U in OREAS 266 (Certified Value 1.05 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 1.06 | 1.0953 | 1.02 | 1.1 | 1.06 | 1.05 | 1 | 1.103 | 0.94 | 1.1234 | 1.2 | 1.02 | 1.11 | 1 | 1.08 |  |  |  |  | 1
| 1 | 2 | 1.06 | 1.0761 | 1.04 | 1 | 1 | 1.06 | 1 | 1.071 | 0.95 | 1.1006 | 1.19 | 0.96 | 1.12 | 1.07 | 1.09 |  |  |  |  | 28
| 1 | 3 | 1.06 | 1.0421 | 1.05 | 1 | 1.03 | 1.06 | 1.1 | 1.079 | 0.86 | 1.1285 | 1.26 | 0.99 | 1.06 | 0.94 | 1.09 |  |  |  |  | 16
| 1 | 4 | 1.08 | 1.0356 | 1.04 | 1 | 1.03 | 0.96 | 1 | 1.051 | 0.95 | 1.0643 | 1.19 | 1 | 1.15 | 0.98 | 1.08 |  |  |  |  | 1.0492
| 1 | 5 | 1.06 | 1.1045 | 1.04 | 1 | 1.05 | 1.01 | 1 | 1.093 | 0.88 | 1.1 | 1.18 | 1.07 | 1.08 | 0.91 | 1.08 |  |  |  |  | 117
| 1 | 6 | 1.04 | 1.049 | 1.03 | 1 | 1.02 | 1.05 | 1.1 | 1.071 | 0.88 | 1.0677 | 1.23 | 0.98 | 1.1 | 0.98 | 1.1
| Mean |  | 1.06 | 1.0671 | 1.0367 | 1.0167 | 1.0317 | 1.0317 | 1.0333 | 1.078 | 0.91 | 1.0974 | 1.2083 | 1.0033 | 1.1033 | 0.98 | 1.0867
| Median |  | 1.06 | 1.0626 | 1.04 | 1 | 1.03 | 1.05 | 1 | 1.075 | 0.91 | 1.1003 | 1.195 | 0.995 | 1.105 | 0.98 | 1.085
| Std Dev. |  | 0.012649 | 0.029051 | 0.010328 | 0.040825 | 0.02137 | 0.039707 | 0.05164 | 0.018319 | 0.040988 | 0.026987 | 0.030605 | 0.038297 | 0.031411 | 0.054772 | 0.008165
| Rel.Std.Dev. |  | 1.193% | 2.722% | 0.9963% | 4.016% | 2.071% | 3.849% | 4.997% | 1.699% | 4.504% | 2.459% | 2.533% | 3.817% | 2.847% | 5.589% | 0.7514%
| PDM3 |  | 1.025% | 1.702% | -1.199% | -3.105% | -1.675% | -1.675% | -1.517% | 2.74% | -13.27% | 4.592% | 15.16% | -4.376% | 5.155% | -6.6% | 3.566%
| Z-Score (Absolute) |  | 0.47 | 0.62 | 0 | 0.4 | 0.1 | 0.1 | 0.07 | 0.84 | 2.56 | 1.23 | 3.47 | 0.67 | 1.35 | 1.15 | 1.01
| Analytical results for V in OREAS 266 (Certified Value 34.5 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 36 | 37.0874 | 33.052 | 32 | 39 | 34 | 31 | 37.5755 | 35 | 35 | 33 | 36.4931 | 38 | 37 | 34 | 33 | 30 |  |  | 1
| 1 | 2 | 36 | 37.02 | 33.195 | 33 | 40 | 30 | 29 | 36.9316 | 35 | 35 | 34 | 37.9486 | 39 | 35 | 35 | 33 | 28 |  |  | 29
| 1 | 3 | 36 | 37.0177 | 32.73 | 34 | 40 | 31 | 30 | 36.9273 | 35 | 36 | 32 | 38.0671 | 39 | 35 | 33 | 32 | 28 |  |  | 16
| 1 | 4 | 36 | 36.1525 | 32.66 | 32 | 40 | 32 | 30 | 37.4293 | 34 | 33 | 33 | 37.0269 | 39 | 36 | 34 | 32 | 29 |  |  | 34.5107
| 1 | 5 | 35 | 35.4962 | 31.727 | 35 | 39 | 32 | 31 | 37.3888 | 35 | 36 | 33 | 37.0343 | 38 | 36 | 32 | 32 | 28 |  |  | 118
| 1 | 6 | 37 | 36.7805 | 33.155 | 31 | 40 | 31 | 30 | 37.6087 | 35 | 36 | 32 | 37.0352 | 39 | 35 | 33 | 33 | 30
| Mean |  | 36 | 36.5924 | 32.7532 | 32.8333 | 39.6667 | 31.6667 | 30.1667 | 37.3102 | 34.8333 | 35.1667 | 32.8333 | 37.2675 | 38.6667 | 35.6667 | 33.5 | 32.5 | 28.8333
| Median |  | 36 | 36.8991 | 32.891 | 32.5 | 40 | 31.5 | 30 | 37.409 | 35 | 35.5 | 33 | 37.0347 | 39 | 35.5 | 33.5 | 32.5 | 28.5
| Std Dev. |  | 0.63246 | 0.63864 | 0.54922 | 1.472 | 0.5164 | 1.3663 | 0.75277 | 0.3065 | 0.40825 | 1.169 | 0.75277 | 0.61141 | 0.5164 | 0.8165 | 1.0488 | 0.54772 | 0.98319
| Rel.Std.Dev. |  | 1.757% | 1.745% | 1.677% | 4.483% | 1.302% | 4.315% | 2.495% | 0.8215% | 1.172% | 3.324% | 2.293% | 1.641% | 1.336% | 2.289% | 3.131% | 1.685% | 3.41%
| PDM3 |  | 4.316% | 6.032% | -5.093% | -4.86% | 14.94% | -8.241% | -12.59% | 8.112% | 0.935% | 1.901% | -4.86% | 7.989% | 12.04% | 3.35% | -2.929% | -5.826% | -16.45%
| Z-Score (Absolute) |  | 0.38 | 0.57 | 0.67 | 0.65 | 1.57 | 1.03 | 1.51 | 0.8 | 0 | 0.11 | 0.65 | 0.79 | 1.24 | 0.27 | 0.43 | 0.76 | 1.94
| Analytical results for W in OREAS 266 (Certified Value 11.7 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 11.85 | 13.1713 | 18 | 8.5 | 8.2 | < 20 | 17.96 | 12.2 | 8.9 | 11.9 | 12.7293 | 13.1 | 13.75 | 13.4 | 11.2 |  |  |  |  | 1
| 1 | 2 | 11.91 | 13.567 | 18 | 8.1 | 6.5 | < 20 | 19.2967 | 12.4 | 8.9 | 11.1 | 12.5165 | 12.3 | 13.55 | 14.5 | 10.6 |  |  |  |  | 30
| 1 | 3 | 11.72 | 12.7787 | 20 | 8.2 | 7.1 | < 20 | 17.86 | 14.5 | 9.5 | 12.9 | 12.3968 | 12.4 | 14.25 | 15.55 | 10.85 |  |  |  |  | 16
| 1 | 4 | 12.13 | 13.4651 | 18 | 8 | 8.3 | < 20 | 17.8033 | 14.15 | 9.6 | 11.75 | 12.8999 | 13.1 | 14.05 | 13.2 | 10.8 |  |  |  |  | 11.7105
| 1 | 5 | 11.7 | 13.3067 | 18 | 7.9 | 8 | < 20 | 17.97 | 14.9 | 9.2 | 13.25 | 12.5654 | 13.2 | 14.85 | 14.45 | 12.2 |  |  |  |  | 119
| 1 | 6 | 11.55 | 12.7233 | 19 | 8 | 7.9 | < 20 | 19.2933 | 12.4 | 10 | 13.2 | 12.7988 | 12.1 | 14.5 | 14.4 | 10.95
| Mean |  | 11.81 | 13.1687 | 18.5 | 8.1167 | 7.6667 |  | 18.3639 | 13.425 | 9.35 | 12.35 | 12.6511 | 12.7 | 14.1583 | 14.25 | 11.1
| Median |  | 11.785 | 13.239 | 18 | 8.05 | 7.95 |  | 17.965 | 13.275 | 9.35 | 12.4 | 12.6473 | 12.75 | 14.15 | 14.425 | 10.9
| Std Dev. |  | 0.2009 | 0.35102 | 0.83666 | 0.2137 | 0.71181 |  | 0.72393 | 1.2214 | 0.43243 | 0.88994 | 0.18971 | 0.4858 | 0.48002 | 0.85323 | 0.57359
| Rel.Std.Dev. |  | 1.701% | 2.666% | 4.522% | 2.633% | 9.284% |  | 3.942% | 9.098% | 4.625% | 7.206% | 1.5% | 3.825% | 3.39% | 5.988% | 5.167%
| PDM3 |  | 0.8493% | 12.45% | 57.98% | -30.69% | -34.53% |  | 56.82% | 14.64% | -20.16% | 5.461% | 8.032% | 8.449% | 20.9% | 21.69% | -5.214%
| Z-Score (Absolute) |  | 0.3 | 0.29 | NA | 1.93 | 2.13 | 1.1 | 2.59 | 0.41 | 1.39 | 0.07 | 0.07 | 0.09 | 0.73 | 0.77 | 0.62
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 266 (Certified Value 10.1 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 10.48 | 10.1753 | 10 | 9.54 | 9.97 | 9.89 | 10.3533 | 9.93 | 10.35 | 10.4054 | 10.2 | 10.7 | 10.9 | 9.95 | 9.72 |  |  |  |  | 1
| 1 | 2 | 10.34 | 10.1932 | 11 | 9.66 | 9.18 | 9.91 | 10.1367 | 10 | 10.9 | 10.604 | 10.4 | 10.15 | 11.05 | 9.77 | 9.78 |  |  |  |  | 31
| 1 | 3 | 10.51 | 10.5239 | 11 | 9.87 | 9.24 | 9.84 | 10.1767 | 10.3 | 9.87 | 10.4787 | 10.5 | 10.4 | 10.2 | 9.73 | 9.91 |  |  |  |  | 16
| 1 | 4 | 10.5 | 10.2946 | 11 | 9.79 | 9.54 | 9.97 | 10.2767 | 9.79 | 10.65 | 10.4594 | 10.5 | 10.3 | 11.05 | 9.67 | 9.81 |  |  |  |  | 10.1425
| 1 | 5 | 10.41 | 10.2838 | 11 | 9.57 | 9.9 | 10.03 | 10.3333 | 10.1 | 9.9 | 10.5703 | 10.3 | 11.3 | 10.2 | 9.02 | 9.72 |  |  |  |  | 120
| 1 | 6 | 10.32 | 10.3188 | 10 | 9.81 | 9.15 | 9.95 | 10.3367 | 9.82 | 9.91 | 10.4787 | 10.8 | 10.2 | 10.45 | 10.1 | 10.03
| Mean |  | 10.4267 | 10.2983 | 10.6667 | 9.7067 | 9.4967 | 9.9317 | 10.2689 | 9.99 | 10.2633 | 10.4994 | 10.45 | 10.5083 | 10.6417 | 9.7067 | 9.8283
| Median |  | 10.445 | 10.2892 | 11 | 9.725 | 9.39 | 9.93 | 10.305 | 9.965 | 10.13 | 10.4787 | 10.45 | 10.35 | 10.675 | 9.75 | 9.795
| Std Dev. |  | 0.082865 | 0.12465 | 0.5164 | 0.13633 | 0.36741 | 0.066458 | 0.091571 | 0.19016 | 0.44135 | 0.073818 | 0.20736 | 0.43407 | 0.40671 | 0.37184 | 0.12123
| Rel.Std.Dev. |  | 0.7947% | 1.21% | 4.841% | 1.405% | 3.869% | 0.6692% | 0.8917% | 1.903% | 4.3% | 0.7031% | 1.984% | 4.131% | 3.822% | 3.831% | 1.233%
| PDM3 |  | 2.801% | 1.535% | 5.168% | -4.297% | -6.368% | -2.079% | 1.246% | -1.504% | 1.191% | 3.519% | 3.031% | 3.607% | 4.921% | -4.297% | -3.098%
| Z-Score (Absolute) |  | 0.42 | 0.08 | NA | 1.46 | 2 | 0.87 | 0.01 | 0.72 | 0.01 | 0.61 | 0.48 | 0.63 | 0.98 | 1.46 | 1.14
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 266 (Certified Value 0.74 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.765 | 0.68172 | 0.7 | 0.8 | 0.8 | [14 empty] | 1
| 1 | 2 | 0.718 | 0.6691 | 0.7 | 0.8 | 0.8 | [14 empty] | 32
| 1 | 3 | 0.728 | 0.71076 | 0.7 | 0.8 | 0.8 | [14 empty] | 16
| 1 | 4 | 0.751 | 0.70939 | 0.7 | 0.8 | 0.8 | [14 empty] | 0.74358
| 1 | 5 | 0.753 | 0.41984 | 0.7 | 0.7 | 0.8 | [14 empty] | 121
| 1 | 6 | 0.722 | 0.7043 | 0.7 | 0.8 | 0.8
| Mean |  | 0.7395 | 0.64918 | 0.7 | 0.78333 | 0.8
| Median |  | 0.7395 | 0.69301 | 0.7 | 0.8 | 0.8
| Std Dev. |  | 0.019316 | 0.11358 | 1.2162e-16 | 0.040825 | 1.2162e-16
| Rel.Std.Dev. |  | 2.612% | 17.5% | 1.737e-14% | 5.212% | 1.52e-14%
| PDM3 |  | -0.5483% | -12.69% | -5.86% | 5.347% | 7.588%
| Z-Score (Absolute) |  | 0 | 1.39 | 0.61 | 0.67 | 0.93
| Analytical results for Zn in OREAS 266 (Certified Value 58 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 60 | 68.0832 | 61.0109 | 50.873 | 49 | 63 | 64.2 | 54 | 53.2633 | 60 | 57 | 54.366 | 58 | 64.9629 | 65.5 | 62 | 58 | 55 | 53 | 1
| 1 | 2 | 60 | 65.772 | 60.6757 | 49.675 | 57 | 64 | 65.7 | 54 | 52.87 | 60 | 57 | 54.286 | 58 | 66.5464 | 67.8 | 60 | 60 | 56 | 50 | 33
| 1 | 3 | 60 | 64.8216 | 61.7785 | 49.698 | 50 | 63 | 64.7 | 54 | 52.5233 | 60 | 58 | 54.26 | 56 | 65.8313 | 68.1 | 61 | 59 | 54 | 49 | 16
| 1 | 4 | 60 | 66.2256 | 60.3706 | 48.779 | 51 | 64 | 65.2 | 51 | 53.13 | 60 | 53 | 54.718 | 56 | 63.2379 | 67.3 | 61 | 58 | 54 | 47 | 58.2965
| 1 | 5 | 60 | 66.7008 | 61.0454 | 48.772 | 55 | 63 | 65.1 | 52 | 53.55 | 60 | 59 | 54.211 | 57 | 63.6206 | 67.2 | 61 | 56 | 61 | 51 | 122
| 1 | 6 | 59 | 68.7204 | 60.9468 | 49.587 | 50 | 64 | 66 | 51 | 52.9133 | 61 | 58 | 54.395 | 55 | 62.6859 | 69.5 | 60 | 56 | 55 | 48
| Mean |  | 59.8333 | 66.7206 | 60.9713 | 49.564 | 52 | 63.5 | 65.15 | 52.6667 | 53.0417 | 60.1667 | 57 | 54.3727 | 56.6667 | 64.4808 | 67.5667 | 60.8333 | 57.8333 | 55.8333 | 49.6667
| Median |  | 60 | 66.4632 | 60.9788 | 49.631 | 50.5 | 63.5 | 65.15 | 53 | 53.0217 | 60 | 57.5 | 54.326 | 56.5 | 64.2917 | 67.55 | 61 | 58 | 55 | 49.5
| Std Dev. |  | 0.40825 | 1.4565 | 0.47063 | 0.77302 | 3.2249 | 0.54772 | 0.65345 | 1.5055 | 0.35495 | 0.40825 | 2.0976 | 0.18229 | 1.2111 | 1.5381 | 1.3079 | 0.75277 | 1.6021 | 2.6394 | 2.1602
| Rel.Std.Dev. |  | 0.6823% | 2.183% | 0.7719% | 1.56% | 6.202% | 0.8626% | 1.003% | 2.859% | 0.6692% | 0.6785% | 3.68% | 0.3353% | 2.137% | 2.385% | 1.936% | 1.237% | 2.77% | 4.727% | 4.349%
| PDM3 |  | 2.636% | 14.45% | 4.588% | -14.98% | -10.8% | 8.926% | 11.76% | -9.657% | -9.014% | 3.208% | -2.224% | -6.731% | -2.796% | 10.61% | 15.9% | 4.352% | -0.7946% | -4.225% | -14.8%
| Z-Score (Absolute) |  | 0.28 | 1.25 | 0.44 | 1.16 | 0.82 | 0.8 | 1.03 | 0.73 | 0.67 | 0.33 | 0.12 | 0.49 | 0.16 | 0.94 | 1.37 | 0.42 | 0 | 0.28 | 1.15
| Analytical results for Zr in OREAS 266 (Indicative Value 32.4 ppm) | [20 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 12.2 | 42.3872 | 80 | 41.8 | 5.1 | 29 | 27.7613 | 27.8 | 28.3 | 54.3883 | 47.2 | 25.8 | 27.2 | 22.8 | 46.6 |  |  |  |  | 1
| 1 | 2 | 14 | 44.3984 | 83 | 41 | 3.9 | 28.5 | 27.7011 | 31.3 | 19.4 | 56.6341 | 47 | 22 | 36.8 | 21 | 47.6 |  |  |  |  | 7
| 1 | 3 | 12.9 | 45.3318 | 99 | 42.2 | 5.2 | 29.3 | 27.9399 | 39.3 | 21.8 | 57.2037 | 49.3 | 19.8 | 40.2 | 20.6 | 48.8 |  |  |  |  | 16
| 1 | 4 | 12.9 | 42.7513 | 83 | 40.9 | 5.6 | 28.8 | 27.2905 | 41.4 | 31.4 | 50.3863 | 47.7 | 27.7 | 25.4 | 19.4 | 48.1 |  |  |  |  | 32.3824
| 1 | 5 | 13.1 | 42.8148 | 87 | 39.4 | 5.6 | 29.4 | 26.8652 | 48.1 | 30.3 | 57.8253 | 46 | 29.8 | 36.5 | 25.8 | 46.8 |  |  |  |  | 13
| 1 | 6 | 12.8 | 43.7635 | 81 | 41.8 | 5.3 | 28.7 | 27.384 | 32.7 | 29.6 | 51.8512 | 49.9 | 27.7 | 37.7 | 15.7 | 47.5
| Mean |  | 12.9833 | 43.5745 | 85.5 | 41.1833 | 5.1167 | 28.95 | 27.4903 | 36.7667 | 26.8 | 54.7148 | 47.85 | 25.4667 | 33.9667 | 20.8833 | 47.5667
| Median |  | 12.9 | 43.2891 | 83 | 41.4 | 5.25 | 28.9 | 27.5425 | 36 | 28.95 | 55.5112 | 47.45 | 26.75 | 36.65 | 20.8 | 47.55
| Std Dev. |  | 0.58452 | 1.1373 | 7.0356 | 1.0088 | 0.63061 | 0.35071 | 0.39038 | 7.5253 | 4.9651 | 3.0528 | 1.4761 | 3.8208 | 6.106 | 3.3766 | 0.8165
| Rel.Std.Dev. |  | 4.502% | 2.61% | 8.229% | 2.45% | 12.32% | 1.211% | 1.42% | 20.47% | 18.53% | 5.579% | 3.085% | 15% | 17.98% | 16.17% | 1.717%
| PDM3 |  | -59.91% | 34.56% | 164% | 27.18% | -84.2% | -10.6% | -15.11% | 13.54% | -17.24% | 68.96% | 47.77% | -21.36% | 4.893% | -35.51% | 46.89%
| Z-Score (Absolute) |  | 1.47 | 0.67 | 3.62 | 0.51 | 2.02 | 0.35 | 0.45 | 0.2 | 0.5 | 1.46 | 0.97 | 0.6 | 0 | 0.92 | 0.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 266 (Indicative Value 12.74 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.75 | 1
| 1 | 2 | 12.73 | 9
| Mean |  | 12.74 | 16
| Median |  | 12.74 | 12.74
| Std Dev. |  | 0.014142 | 15
| Rel.Std.Dev. |  | 0.111%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 266 (Indicative Value 2.99 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.99 | 1
| 1 | 2 | 2.99 | 10
| Mean |  | 2.99 | 16
| Median |  | 2.99 | 2.99
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 266 (Indicative Value 7.45 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 7.46 | 1
| 1 | 2 | 7.44 | 11
| Mean |  | 7.45 | 16
| Median |  | 7.45 | 7.45
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.1898%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 266 (Indicative Value 1.84 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.84 | 1
| 1 | 2 | 1.84 | 12
| Mean |  | 1.84 | 16
| Median |  | 1.84 | 1.84
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 266 (Indicative Value 3.06 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.06 | 1
| 1 | 2 | 3.06 | 9
| Mean |  | 3.06 | 16
| Median |  | 3.06 | 3.06
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 266 (Indicative Value 0.068 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.068 | 1
| 1 | 2 | 0.067 | 10
| Mean |  | 0.0675 | 16
| Median |  | 0.0675 | 0.0675
| Std Dev. |  | 0.00070711 | 16
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 266 (Indicative Value 1.57 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.56 | 1
| 1 | 2 | 1.57 | 11
| Mean |  | 1.565 | 16
| Median |  | 1.565 | 1.565
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 266 (Indicative Value 0.241 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.241 | 12
| Mean |  | 0.2405 | 16
| Median |  | 0.2405 | 0.2405
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.294%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 266 (Indicative Value 65.89 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 65.92 | 1
| 1 | 2 | 65.86 | 9
| Mean |  | 65.89 | 16
| Median |  | 65.89 | 65.89
| Std Dev. |  | 0.042426 | 15
| Rel.Std.Dev. |  | 0.06439%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 266 (Indicative Value 0.081 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.081 | 10
| Mean |  | 0.0805 | 16
| Median |  | 0.0805 | 0.0805
| Std Dev. |  | 0.00070711 | 16
| Rel.Std.Dev. |  | 0.8784%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 266 (Indicative Value 1.2 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.199 | 1
| 1 | 2 | 1.201 | 11
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0.0014142 | 17
| Rel.Std.Dev. |  | 0.1179%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 266 (Indicative Value 2.62 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.64 | 14
| Mean |  | 2.62 | 16
| Median |  | 2.62 | 2.62
| Std Dev. |  | 0.028284 | 20
| Rel.Std.Dev. |  | 1.08%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 266 (Indicative Value 0.11 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.11 | 16
| Mean |  | 0.11 | 16
| Median |  | 0.11 | 0.11
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 266 (Indicative Value 0.01 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | 0.01 | 16
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.01
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 266 (Indicative Value 4.45 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.5 | 18
| Mean |  | 4.45 | 16
| Median |  | 4.45 | 4.45
| Std Dev. |  | 0.070711 | 24
| Rel.Std.Dev. |  | 1.589%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 266 (Indicative Value 181 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 182 | 1
| 1 | 2 | 179 | 19
| Mean |  | 180.5 | 16
| Median |  | 180.5 | 180.5
| Std Dev. |  | 2.1213 | 25
| Rel.Std.Dev. |  | 1.175%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 266 (Indicative Value 459 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 460 | 1
| 1 | 2 | 458 | 20
| Mean |  | 459 | 16
| Median |  | 459 | 459
| Std Dev. |  | 1.4142 | 26
| Rel.Std.Dev. |  | 0.3081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 266 (Indicative Value 2.8 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.8 | 21
| Mean |  | 2.8 | 16
| Median |  | 2.8 | 2.8
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 266 (Indicative Value 3.18 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.16 | 1
| 1 | 2 | 3.2 | 22
| Mean |  | 3.18 | 16
| Median |  | 3.18 | 3.18
| Std Dev. |  | 0.028284 | 28
| Rel.Std.Dev. |  | 0.8894%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 266 (Indicative Value 0.13 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | 0.2 | 23
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.125
| Std Dev. |  |  | 29
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 266 (Indicative Value 82 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 82.6 | 1
| 1 | 2 | 81.1 | 24
| Mean |  | 81.85 | 16
| Median |  | 81.85 | 81.85
| Std Dev. |  | 1.0607 | 30
| Rel.Std.Dev. |  | 1.296%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 266 (Indicative Value 23.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 23.1 | 25
| Mean |  | 23.05 | 16
| Median |  | 23.05 | 23.05
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 0.3068%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 266 (Indicative Value 151 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 150 | 1
| 1 | 2 | 151 | 26
| Mean |  | 150.5 | 16
| Median |  | 150.5 | 150.5
| Std Dev. |  | 0.70711 | 32
| Rel.Std.Dev. |  | 0.4698%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 266 (Indicative Value 4.38 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.36 | 27
| Mean |  | 4.38 | 16
| Median |  | 4.38 | 4.38
| Std Dev. |  | 0.028284 | 33
| Rel.Std.Dev. |  | 0.6458%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 266 (Indicative Value 35 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36 | 1
| 1 | 2 | 34 | 28
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 1.4142 | 34
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 266 (Indicative Value 5.24 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.4 | 1
| 1 | 2 | 5.07 | 3
| Mean |  | 5.235 | 16
| Median |  | 5.235 | 5.235
| Std Dev. |  | 0.23335 | 35
| Rel.Std.Dev. |  | 4.457%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 266 (Indicative Value 2.68 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.67 | 1
| 1 | 2 | 2.69 | 4
| Mean |  | 2.68 | 16
| Median |  | 2.68 | 2.68
| Std Dev. |  | 0.014142 | 36
| Rel.Std.Dev. |  | 0.5277%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 266 (Indicative Value 1.84 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.84 | 1
| 1 | 2 | 1.83 | 5
| Mean |  | 1.835 | 16
| Median |  | 1.835 | 1.835
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 0.3853%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 266 (Indicative Value 18.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.8 | 1
| 1 | 2 | 18.8 | 32
| Mean |  | 18.8 | 16
| Median |  | 18.8 | 18.8
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 266 (Indicative Value 6.19 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.22 | 1
| 1 | 2 | 6.16 | 6
| Mean |  | 6.19 | 16
| Median |  | 6.19 | 6.19
| Std Dev. |  | 0.042426 | 39
| Rel.Std.Dev. |  | 0.6854%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 266 (Indicative Value 1.28 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.2 | 1
| Mean |  | 1.275 | 16
| Median |  | 1.275 | 1.275
| Std Dev. |  | 0.10607 | 40
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 266 (Indicative Value 7.14 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.05 | 1
| 1 | 2 | 7.23 | 7
| Mean |  | 7.14 | 16
| Median |  | 7.14 | 7.14
| Std Dev. |  | 0.12728 | 24
| Rel.Std.Dev. |  | 1.783%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 266 (Indicative Value 0.99 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.98 | 1
| 1 | 2 | 0.99 | 3
| Mean |  | 0.985 | 16
| Median |  | 0.985 | 0.985
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.7179%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 266 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 20
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 266 (Indicative Value 45.1 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45.3 | 1
| 1 | 2 | 44.9 | 21
| Mean |  | 45.1 | 16
| Median |  | 45.1 | 45.1
| Std Dev. |  | 0.28284 | 27
| Rel.Std.Dev. |  | 0.6271%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 266 (Indicative Value 0.32 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.31 | 4
| Mean |  | 0.315 | 16
| Median |  | 0.315 | 0.315
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 2.245%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 266 (Indicative Value 0.055 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0556 | 1
| 1 | 2 | 0.0552 | 23
| Mean |  | 0.0554 | 16
| Median |  | 0.0554 | 0.0554
| Std Dev. |  | 0.00028284 | 29
| Rel.Std.Dev. |  | 0.5105%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 266 (Indicative Value 9.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.6 | 1
| 1 | 2 | 9.6 | 24
| Mean |  | 9.6 | 16
| Median |  | 9.6 | 9.6
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 266 (Indicative Value 34.2 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34.4 | 1
| 1 | 2 | 34 | 25
| Mean |  | 34.2 | 16
| Median |  | 34.2 | 34.2
| Std Dev. |  | 0.28284 | 31
| Rel.Std.Dev. |  | 0.827%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 266 (Indicative Value 38.9 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.3 | 1
| 1 | 2 | 38.5 | 5
| Mean |  | 38.9 | 16
| Median |  | 38.9 | 38.9
| Std Dev. |  | 0.56569 | 32
| Rel.Std.Dev. |  | 1.454%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 266 (Indicative Value 101 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 102 | 1
| 1 | 2 | 100 | 27
| Mean |  | 101 | 16
| Median |  | 101 | 101
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 266 (Indicative Value 16 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16 | 1
| 1 | 2 | 16 | 28
| Mean |  | 16 | 16
| Median |  | 16 | 16
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 266 (Indicative Value 10.5 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.5 | 1
| 1 | 2 | 10.4 | 7
| Mean |  | 10.45 | 16
| Median |  | 10.45 | 10.45
| Std Dev. |  | 0.070711 | 35
| Rel.Std.Dev. |  | 0.6767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 266 (Indicative Value 75 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75.7 | 1
| 1 | 2 | 74.4 | 30
| Mean |  | 75.05 | 16
| Median |  | 75.05 | 75.05
| Std Dev. |  | 0.91924 | 36
| Rel.Std.Dev. |  | 1.225%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 266 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 31
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 37
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 266 (Indicative Value 11 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 11 | 32
| Mean |  | 11 | 16
| Median |  | 11 | 11
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 266 (Indicative Value 12.5 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.5 | 1
| 1 | 2 | 12.4 | 33
| Mean |  | 12.45 | 16
| Median |  | 12.45 | 12.45
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.568%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 266 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 34
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 266 (Indicative Value 7.57 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.68 | 1
| 1 | 2 | 7.46 | 5
| Mean |  | 7.57 | 16
| Median |  | 7.57 | 7.57
| Std Dev. |  | 0.15556 | 24
| Rel.Std.Dev. |  | 2.055%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 266 (Indicative Value 7.9 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8 | 1
| 1 | 2 | 7.8 | 19
| Mean |  | 7.9 | 16
| Median |  | 7.9 | 7.9
| Std Dev. |  | 0.14142 | 25
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 266 (Indicative Value 327 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 327 | 1
| 1 | 2 | 326 | 20
| Mean |  | 326.5 | 16
| Median |  | 326.5 | 326.5
| Std Dev. |  | 0.70711 | 26
| Rel.Std.Dev. |  | 0.2166%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 266 (Indicative Value 2.29 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.27 | 1
| 1 | 2 | 2.31 | 21
| Mean |  | 2.29 | 16
| Median |  | 2.29 | 2.29
| Std Dev. |  | 0.028284 | 27
| Rel.Std.Dev. |  | 1.235%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 266 (Indicative Value 0.95 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.94 | 1
| 1 | 2 | 0.95 | 22
| Mean |  | 0.945 | 16
| Median |  | 0.945 | 0.945
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 266 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 266 (Indicative Value 11.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.3 | 1
| 1 | 2 | 11.1 | 24
| Mean |  | 11.2 | 16
| Median |  | 11.2 | 11.2
| Std Dev. |  | 0.14142 | 30
| Rel.Std.Dev. |  | 1.263%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 266 (Indicative Value 0.726 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.728 | 1
| 1 | 2 | 0.724 | 25
| Mean |  | 0.726 | 16
| Median |  | 0.726 | 0.726
| Std Dev. |  | 0.0028284 | 31
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 266 (Indicative Value 0.3 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.4 | 26
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 266 (Indicative Value 0.37 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.37 | 6
| Mean |  | 0.37 | 16
| Median |  | 0.37 | 0.37
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 266 (Indicative Value 2.28 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.26 | 1
| 1 | 2 | 2.3 | 28
| Mean |  | 2.28 | 16
| Median |  | 2.28 | 2.28
| Std Dev. |  | 0.028284 | 34
| Rel.Std.Dev. |  | 1.241%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 266 (Indicative Value 99 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 99.9 | 1
| 1 | 2 | 98.2 | 29
| Mean |  | 99.05 | 16
| Median |  | 99.05 | 99.05
| Std Dev. |  | 1.2021 | 35
| Rel.Std.Dev. |  | 1.214%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 266 (Indicative Value 33.8 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.5 | 1
| 1 | 2 | 34 | 30
| Mean |  | 33.75 | 16
| Median |  | 33.75 | 33.75
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 266 (Indicative Value 25.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 24.5 | 31
| Mean |  | 25.25 | 16
| Median |  | 25.25 | 25.25
| Std Dev. |  | 1.0607 | 37
| Rel.Std.Dev. |  | 4.201%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 266 (Indicative Value 2.26 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.28 | 1
| 1 | 2 | 2.23 | 32
| Mean |  | 2.255 | 16
| Median |  | 2.255 | 2.255
| Std Dev. |  | 0.035355 | 38
| Rel.Std.Dev. |  | 1.568%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 266 (Indicative Value 95 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 95 | 1
| 1 | 2 | 95 | 33
| Mean |  | 95 | 16
| Median |  | 95 | 95
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 266 (Indicative Value 270 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 269 | 1
| 1 | 2 | 270 | 7
| Mean |  | 269.5 | 16
| Median |  | 269.5 | 269.5
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 0.2624%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.