## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 231b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.55644 | 0.54995 | 0.56293 | 0.55174 | 0.56114
PhotonAssay (recommended gross mass* 358±15 g)
Au, Gold (ppm) | 0.55617 | 0.54882 | 0.56352 | 0.55489 | 0.55746
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 0.52588 | 0.51657 | 0.53518 | 0.52101 | 0.53074
Cyanide Leach
Au, Gold (ppm) | 0.52369 | 0.51541 | 0.53197 | 0.52198 | 0.5254
4-Acid Digestion
Ag, Silver (ppm) | 0.18226 | 0.17033 | 0.19419 | 0.16627 | 0.19826
Al, Aluminium (wt.%) | 6.8854 | 6.7044 | 7.0664 | 6.7685 | 7.0023
As, Arsenic (ppm) | 33.976 | 31.8795 | 36.0725 | 32.1968 | 35.7552
Ba, Barium (ppm) | 299.2542 | 287.2454 | 311.2631 | 288.9029 | 309.6056
Be, Beryllium (ppm) | 0.36832 | 0.3362 | 0.40044 | 0.34897 | 0.38766
Bi, Bismuth (ppm) | 0.028909 | 0.021422 | 0.036396 | IND | IND
Ca, Calcium (wt.%) | 6.4334 | 6.2623 | 6.6044 | 6.3272 | 6.5396
Cd, Cadmium (ppm) | 0.39122 | 0.35603 | 0.4264 | 0.36501 | 0.41742
Ce, Cerium (ppm) | 11.3417 | 10.7081 | 11.9753 | 10.9712 | 11.7123
Co, Cobalt (ppm) | 46.563 | 44.8152 | 48.3109 | 45.5209 | 47.6052
Cr, Chromium (ppm) | 153.7205 | 145.607 | 161.8339 | 148.6301 | 158.8109
Cs, Caesium (ppm) | 0.58443 | 0.55253 | 0.61634 | 0.55892 | 0.60995
Cu, Copper (ppm) | 155.6637 | 149.867 | 161.4604 | 152.6985 | 158.6289
Dy, Dysprosium (ppm) | 3.7696 | 3.5485 | 3.9908 | 3.6253 | 3.914
Er, Erbium (ppm) | 2.315 | 2.1602 | 2.4698 | 2.2121 | 2.4179
Eu, Europium (ppm) | 0.92279 | 0.86118 | 0.98439 | 0.89156 | 0.95401
Fe, Iron (wt.%) | 8.3054 | 8.0732 | 8.5376 | 8.1664 | 8.4444
Ga, Gallium (ppm) | 15.3744 | 14.6138 | 16.1349 | 14.7841 | 15.9646
Gd, Gadolinium (ppm) | 3.0421 | 2.7665 | 3.3177 | 2.968 | 3.1162
Hf, Hafnium (ppm) | 1.5439 | 1.4002 | 1.6877 | 1.4574 | 1.6304
Ho, Holmium (ppm) | 0.77451 | 0.7271 | 0.82192 | 0.72608 | 0.82294
In, Indium (ppm) | 0.071179 | 0.062399 | 0.079959 | 0.062938 | 0.079419
K, Potassium (wt.%) | 0.47693 | 0.45865 | 0.49521 | 0.46693 | 0.48692
La, Lanthanum (ppm) | 4.5248 | 4.2956 | 4.7539 | 4.3604 | 4.6891
Li, Lithium (ppm) | 11.9618 | 11.1089 | 12.8146 | 11.5644 | 12.3591
Lu, Lutetium (ppm) | 0.33032 | 0.29033 | 0.37031 | 0.32209 | 0.33855
Mg, Magnesium (wt.%) | 4.1414 | 4.0391 | 4.2437 | 4.0574 | 4.2254
Mn, Manganese (wt.%) | 0.15794 | 0.15338 | 0.1625 | 0.15541 | 0.16046
Mo, Molybdenum (ppm) | 1.4618 | 1.3791 | 1.5444 | 1.3699 | 1.5537
Na, Sodium (wt.%) | 2.0387 | 1.9858 | 2.0916 | 1.9972 | 2.0802
Nb, Niobium (ppm) | 3.6631 | 3.3906 | 3.9356 | 3.4904 | 3.8357
Nd, Neodymium (ppm) | 7.9799 | 7.6422 | 8.3176 | 7.8329 | 8.1269
Ni, Nickel (ppm) | 104.8458 | 101.1578 | 108.5338 | 102.8294 | 106.8622
P, Phosphorus (wt.%) | 0.043809 | 0.042145 | 0.045473 | 0.042645 | 0.044973
Pb, Lead (ppm) | 13.5667 | 12.8277 | 14.3058 | 12.9711 | 14.1624
Pr, Praseodymium (ppm) | 1.642 | 1.5212 | 1.7628 | 1.5754 | 1.7086
Rb, Rubidium (ppm) | 10.3618 | 9.7407 | 10.9829 | 10.0092 | 10.7144
Re, Rhenium (ppm) | 0.0025729 | 0.001629 | 0.0035168 | IND | IND
S, Sulphur (wt.%) | 0.26428 | 0.25009 | 0.27848 | 0.25681 | 0.27176
Sb, Antimony (ppm) | 0.82095 | 0.75105 | 0.89086 | 0.77211 | 0.8698
Sc, Scandium (ppm) | 39.4319 | 37.8068 | 41.0571 | 38.228 | 40.6359
Sm, Samarium (ppm) | 2.44 | 2.2063 | 2.6737 | 2.3244 | 2.5556
Sn, Tin (ppm) | 0.93501 | 0.81848 | 1.0515 | 0.81811 | 1.0519
Sr, Strontium (ppm) | 203.2353 | 196.6729 | 209.7978 | 199.6378 | 206.8329
Ta, Tantalum (ppm) | 0.24188 | 0.20943 | 0.27433 | 0.21083 | 0.27293
Tb, Terbium (ppm) | 0.55959 | 0.51768 | 0.60149 | 0.53857 | 0.58061
Te, Tellurium (ppm) | 0.073508 | 0.054035 | 0.092981 | IND | IND
Th, Thorium (ppm) | 0.65239 | 0.6087 | 0.69607 | 0.62062 | 0.68416
Ti, Titanium (wt.%) | 0.63779 | 0.6225 | 0.65308 | 0.62076 | 0.65483
Tl, Thallium (ppm) | 0.12208 | 0.10863 | 0.13553 | IND | IND
Tm, Thulium (ppm) | 0.33216 | 0.30545 | 0.35888 | 0.31458 | 0.34975
U, Uranium (ppm) | 0.20272 | 0.19967 | 0.20576 | 0.19316 | 0.21228
V, Vanadium (ppm) | 282.8955 | 274.8909 | 290.9001 | 277.7437 | 288.0473
W, Tungsten (ppm) | 16.9595 | 16.1723 | 17.7467 | 16.5358 | 17.3832
Y, Yttrium (ppm) | 20.4727 | 19.3527 | 21.5928 | 19.8604 | 21.085
Yb, Ytterbium (ppm) | 2.1304 | 1.959 | 2.3017 | 2.0249 | 2.2358
Zn, Zinc (ppm) | 117.7622 | 113.4194 | 122.1049 | 114.7801 | 120.7443
Zr, Zirconium (ppm) | 45.4343 | 41.7363 | 49.1322 | 42.9209 | 47.9476
Aqua Regia Digestion
Ag, Silver (ppm) | 0.18226 | 0.16916 | 0.19535 | 0.16735 | 0.19716
Al, Aluminium (wt.%) | 3.1318 | 2.9952 | 3.2684 | 3.0684 | 3.1951
As, Arsenic (ppm) | 33.3234 | 32.08 | 34.5668 | 32.3843 | 34.2625
B, Boron (ppm) | 28.5899 | 22.7947 | 34.385 | 26.8833 | 30.2965
Ba, Barium (ppm) | 35.8332 | 33.3628 | 38.3036 | 34.5416 | 37.1247
Be, Beryllium (ppm) | 0.17019 | 0.154 | 0.18639 | IND | IND
Bi, Bismuth (ppm) | 0.027289 | 0.019147 | 0.035431 | IND | IND
Ca, Calcium (wt.%) | 2.1067 | 1.9695 | 2.2439 | 2.0538 | 2.1597
Cd, Cadmium (ppm) | 0.34391 | 0.32209 | 0.36572 | 0.3223 | 0.36551
Ce, Cerium (ppm) | 7.936 | 7.651 | 8.2209 | 7.7614 | 8.1106
Co, Cobalt (ppm) | 33.3786 | 32.1778 | 34.5794 | 32.5405 | 34.2168
Cr, Chromium (ppm) | 23.8965 | 22.7086 | 25.0844 | 23.0687 | 24.7244
Cs, Caesium (ppm) | 0.4401 | 0.41571 | 0.46449 | 0.42484 | 0.45536
Cu, Copper (ppm) | 154.8415 | 150.7379 | 158.9451 | 151.9967 | 157.6863
Dy, Dysprosium (ppm) | 2.115 | 1.7695 | 2.4605 | 1.99 | 2.24
Er, Erbium (ppm) | 1.2177 | 1.0143 | 1.421 | 1.1659 | 1.2695
Eu, Europium (ppm) | 0.44405 | 0.34268 | 0.54542 | 0.43005 | 0.45805
Fe, Iron (wt.%) | 5.9057 | 5.7829 | 6.0284 | 5.7772 | 6.0342
Ga, Gallium (ppm) | 10.2022 | 9.6265 | 10.7779 | 10.0013 | 10.4031
Gd, Gadolinium (ppm) | 1.8169 | 1.5697 | 2.0641 | 1.6645 | 1.9694
Ge, Germanium (ppm) | 0.13 | 0.10467 | 0.15533 | IND | IND
Hf, Hafnium (ppm) | 0.39831 | 0.36298 | 0.43365 | 0.37421 | 0.42242
Ho, Holmium (ppm) | 0.42389 | 0.36304 | 0.48475 | 0.37995 | 0.46784
In, Indium (ppm) | 0.024324 | 0.021726 | 0.026922 | 0.021434 | 0.027213
K, Potassium (wt.%) | 0.10389 | 0.099379 | 0.1084 | 0.10038 | 0.1074
La, Lanthanum (ppm) | 3.3424 | 3.2268 | 3.4579 | 3.2334 | 3.4513
Li, Lithium (ppm) | 10.5797 | 10.1073 | 11.0521 | 10.3539 | 10.8055
Lu, Lutetium (ppm) | 0.12997 | 0.10741 | 0.15254 | IND | IND
Mg, Magnesium (wt.%) | 1.9568 | 1.9066 | 2.007 | 1.9162 | 1.9974
Mn, Manganese (wt.%) | 0.084178 | 0.081242 | 0.087114 | 0.082853 | 0.085503
Mo, Molybdenum (ppm) | 1.3948 | 1.3165 | 1.473 | 1.3223 | 1.4673
Na, Sodium (wt.%) | 0.14683 | 0.13999 | 0.15367 | 0.13724 | 0.15642
Nd, Neodymium (ppm) | 5.3144 | 4.844 | 5.7848 | 5.155 | 5.4738
Ni, Nickel (ppm) | 72.8669 | 70.548 | 75.1858 | 71.2776 | 74.4562
P, Phosphorus (wt.%) | 0.043109 | 0.041621 | 0.044598 | 0.042047 | 0.044172
Pb, Lead (ppm) | 13.3381 | 12.7818 | 13.8944 | 12.8003 | 13.876
Pr, Praseodymium (ppm) | 1.0724 | 0.96011 | 1.1847 | 1.0241 | 1.1206
Rb, Rubidium (ppm) | 3.9903 | 3.7641 | 4.2164 | 3.8615 | 4.119
Re, Rhenium (ppm) | 0.0019444 | 0.0011354 | 0.0027535 | IND | IND
S, Sulphur (wt.%) | 0.2615 | 0.2515 | 0.2715 | 0.25269 | 0.27031
Sb, Antimony (ppm) | 0.50231 | 0.44362 | 0.561 | 0.4677 | 0.53692
Sc, Scandium (ppm) | 4.4618 | 4.2203 | 4.7034 | 4.2765 | 4.6472
Sm, Samarium (ppm) | 1.5312 | 1.3665 | 1.6959 | 1.4815 | 1.5809
Sn, Tin (ppm) | 0.52088 | 0.45645 | 0.58532 | 0.47503 | 0.56674
Sr, Strontium (ppm) | 32.046 | 28.9719 | 35.1202 | 30.2054 | 33.8867
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.3066 | 0.2851 | 0.32809 | 0.2886 | 0.32459
Te, Tellurium (ppm) | 0.059653 | 0.047003 | 0.072303 | IND | IND
Th, Thorium (ppm) | 0.45898 | 0.40061 | 0.51735 | 0.42625 | 0.4917
Ti, Titanium (wt.%) | 0.32883 | 0.30066 | 0.35699 | 0.31907 | 0.33859
Tl, Thallium (ppm) | 0.064382 | 0.056635 | 0.072129 | IND | IND
Tm, Thulium (ppm) | 0.16578 | 0.13901 | 0.19254 | IND | IND
U, Uranium (ppm) | 0.13068 | 0.1168 | 0.14455 | 0.1192 | 0.14215
V, Vanadium (ppm) | 140.6777 | 131.4506 | 149.9049 | 137.175 | 144.1804
W, Tungsten (ppm) | 11.8727 | 10.806 | 12.9394 | 11.4207 | 12.3247
Y, Yttrium (ppm) | 10.777 | 10.3361 | 11.2179 | 10.5114 | 11.0426
Yb, Ytterbium (ppm) | 0.98121 | 0.83787 | 1.1246 | 0.93397 | 1.0285
Zn, Zinc (ppm) | 102.0616 | 98.8398 | 105.2835 | 99.8704 | 104.2528
Zr, Zirconium (ppm) | 12.1224 | 10.953 | 13.2919 | 11.4793 | 12.7656
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 231b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 20.7858 | Pt | ppb | 10.285
4-Acid Digestion
B | ppm | 32.6325 | Hg | ppm | 0.1885
Ge | ppm | 0.15995 | Se | ppm | 0.99918
Aqua Regia Digestion
Hg | ppm | 0.035994 | Pd | ppb | 16.4333 | Si | wt.% | 0.080817
Ir | ppm | < 0.003 | Pt | ppb | 8.7733
Nb | ppm | 0.10883 | Se | ppm | 0.52439
Borate Fusion XRF
Al2O3 | wt.% | 13.145 | MgO | wt.% | 7.045 | S | wt.% | 0.21905
CaO | wt.% | 9.375 | MnO | wt.% | 0.192 | SiO2 | wt.% | 50.17
Fe2O3 | wt.% | 12.305 | Na2O | wt.% | 2.735 | TiO2 | wt.% | 1.1125
K2O | wt.% | 0.55 | P2O5 | wt.% | 0.0965
Thermogravimetry
LOI1000 | wt.% | 2.975
Infrared Combustion
C | wt.% | 0.1 | S | wt.% | 0.245
Laser Ablation ICP-MS
Ag | ppm | 0.2 | Hf | ppm | 1.88 | Sm | ppm | 2.605
As | ppm | 33.2 | Ho | ppm | 0.835 | Sn | ppm | 0.8
Ba | ppm | 294.5 | In | ppm | 0.05 | Sr | ppm | 196.5
Be | ppm | 0.6 | La | ppm | 4.865 | Ta | ppm | 0.205
Bi | ppm | 0.04 | Lu | ppm | 0.35 | Tb | ppm | 0.575
Cd | ppm | 0.4 | Mn | wt.% | 0.168 | Te | ppm | < 0.2
Ce | ppm | 11 | Mo | ppm | 1.3 | Th | ppm | 0.645
Co | ppm | 47.8 | Nb | ppm | 3.59 | Ti | wt.% | 0.6665
Cr | ppm | 189.5 | Nd | ppm | 8.24 | Tl | ppm | < 0.2
Cs | ppm | 0.545 | Ni | ppm | 115 | Tm | ppm | 0.35
Cu | ppm | 160 | Pb | ppm | 13.5 | U | ppm | 0.17
Dy | ppm | 3.86 | Pr | ppm | 1.735 | V | ppm | 297.5
Er | ppm | 2.455 | Rb | ppm | 9.9 | W | ppm | 16.5
Eu | ppm | 0.885 | Re | ppm | 0.0175 | Y | ppm | 21.5
Ga | ppm | 15.05 | Sb | ppm | 0.8 | Yb | ppm | 2.44
Gd | ppm | 3.21 | Sc | ppm | 41.25 | Zn | ppm | 120
Ge | ppm | 1.525 | Se | ppm | < 5 | Zr | ppm | 64.75
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 231b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.55644 | 0.017285 | 0.52187 | 0.59101 | 0.50458 | 0.60829 | 3.106% | 6.213% | 9.319% | 0.52862 | 0.58426
PhotonAssay
Au, ppm | 0.55617 | 0.030441 | 0.49529 | 0.61705 | 0.46485 | 0.6475 | 5.473% | 10.95% | 16.42% | 0.52836 | 0.58398
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.52588 | 0.023781 | 0.47832 | 0.57344 | 0.45454 | 0.59722 | 4.522% | 9.044% | 13.57% | 0.49958 | 0.55217
Cyanide Leach
Au, ppm | 0.52369 | 0.023875 | 0.47594 | 0.57144 | 0.45206 | 0.59531 | 4.559% | 9.118% | 13.68% | 0.4975 | 0.54987
4-Acid Digestion
Ag, ppm | 0.18226 | 0.011496 | 0.15927 | 0.20526 | 0.14778 | 0.21675 | 6.307% | 12.61% | 18.92% | 0.17315 | 0.19138
Al, wt.% | 6.8854 | 0.23431 | 6.4168 | 7.354 | 6.1825 | 7.5884 | 3.403% | 6.806% | 10.21% | 6.5411 | 7.2297
As, ppm | 33.976 | 2.6371 | 28.7017 | 39.2503 | 26.0646 | 41.8874 | 7.762% | 15.52% | 23.29% | 32.2772 | 35.6748
Ba, ppm | 299.2542 | 14.2538 | 270.7467 | 327.7618 | 256.4929 | 342.0156 | 4.763% | 9.526% | 14.29% | 284.2915 | 314.217
Be, ppm | 0.36832 | 0.046329 | 0.27566 | 0.46098 | 0.22933 | 0.5073 | 12.58% | 25.16% | 37.74% | 0.3499 | 0.38673
Bi, ppm | 0.028909 | 0.004718 | 0.019473 | 0.038345 | 0.014755 | 0.043063 | 16.32% | 32.64% | 48.96% | 0.027464 | 0.030355
Ca, wt.% | 6.4334 | 0.26715 | 5.8991 | 6.9677 | 5.6319 | 7.2348 | 4.153% | 8.305% | 12.46% | 6.1117 | 6.755
Cd, ppm | 0.39122 | 0.037382 | 0.31645 | 0.46598 | 0.27907 | 0.50336 | 9.555% | 19.11% | 28.67% | 0.37165 | 0.41078
Ce, ppm | 11.3417 | 0.5621 | 10.2175 | 12.4659 | 9.6554 | 13.028 | 4.956% | 9.912% | 14.87% | 10.7746 | 11.9088
Co, ppm | 46.563 | 1.6627 | 43.2376 | 49.8884 | 41.5749 | 51.5511 | 3.571% | 7.142% | 10.71% | 44.2349 | 48.8912
Cr, ppm | 153.7205 | 13.3601 | 127.0002 | 180.4407 | 113.6401 | 193.8008 | 8.691% | 17.38% | 26.07% | 146.0344 | 161.4065
Cs, ppm | 0.58443 | 0.042853 | 0.49873 | 0.67014 | 0.45587 | 0.71299 | 7.332% | 14.66% | 22% | 0.55521 | 0.61366
Cu, ppm | 155.6637 | 6.2859 | 143.0918 | 168.2356 | 136.8059 | 174.5215 | 4.038% | 8.076% | 12.11% | 147.8805 | 163.4469
Dy, ppm | 3.7696 | 0.1292 | 3.5112 | 4.028 | 3.382 | 4.1572 | 3.427% | 6.855% | 10.28% | 3.5811 | 3.9581
Er, ppm | 2.315 | 0.093178 | 2.1286 | 2.5014 | 2.0355 | 2.5945 | 4.025% | 8.05% | 12.07% | 2.1993 | 2.4308
Eu, ppm | 0.92279 | 0.049219 | 0.82435 | 1.0212 | 0.77513 | 1.0704 | 5.334% | 10.67% | 16% | 0.87665 | 0.96892
Fe, wt.% | 8.3054 | 0.37124 | 7.5629 | 9.0478 | 7.1917 | 9.4191 | 4.47% | 8.94% | 13.41% | 7.8901 | 8.7206
Ga, ppm | 15.3744 | 1.0519 | 13.2705 | 17.4782 | 12.2186 | 18.5301 | 6.842% | 13.68% | 20.53% | 14.6056 | 16.1431
Gd, ppm | 3.0421 | 0.23408 | 2.5739 | 3.5102 | 2.3399 | 3.7443 | 7.695% | 15.39% | 23.08% | 2.89 | 3.1942
Hf, ppm | 1.5439 | 0.20986 | 1.1242 | 1.9636 | 0.91434 | 2.1735 | 13.59% | 27.19% | 40.78% | 1.4667 | 1.6211
Ho, ppm | 0.77451 | 0.042961 | 0.68858 | 0.86043 | 0.64562 | 0.90339 | 5.547% | 11.09% | 16.64% | 0.73578 | 0.81323
In, ppm | 0.071179 | 0.00877 | 0.053639 | 0.088719 | 0.044869 | 0.097489 | 12.32% | 24.64% | 36.96% | 0.06762 | 0.074738
K, wt.% | 0.47693 | 0.023778 | 0.42937 | 0.52448 | 0.4056 | 0.54826 | 4.986% | 9.971% | 14.96% | 0.45308 | 0.50078
La, ppm | 4.5248 | 0.18867 | 4.1474 | 4.9021 | 3.9587 | 5.0908 | 4.17% | 8.339% | 12.51% | 4.2985 | 4.751
Li, ppm | 11.9618 | 1.3573 | 9.2472 | 14.6763 | 7.8899 | 16.0336 | 11.35% | 22.69% | 34.04% | 11.3637 | 12.5598
Lu, ppm | 0.33032 | 0.027819 | 0.27468 | 0.38596 | 0.24686 | 0.41378 | 8.422% | 16.84% | 25.27% | 0.3138 | 0.34683
Mg, wt.% | 4.1414 | 0.12087 | 3.8996 | 4.3831 | 3.7788 | 4.504 | 2.919% | 5.837% | 8.756% | 3.9343 | 4.3485
Mn, wt.% | 0.15794 | 0.0057654 | 0.14641 | 0.16947 | 0.14064 | 0.17524 | 3.65% | 7.301% | 10.95% | 0.15004 | 0.16584
Mo, ppm | 1.4618 | 0.089263 | 1.2832 | 1.6403 | 1.194 | 1.7296 | 6.106% | 12.21% | 18.32% | 1.3887 | 1.5349
Na, wt.% | 2.0387 | 0.078365 | 1.882 | 2.1955 | 1.8036 | 2.2738 | 3.844% | 7.688% | 11.53% | 1.9368 | 2.1407
Nb, ppm | 3.6631 | 0.34381 | 2.9755 | 4.3507 | 2.6317 | 4.6945 | 9.386% | 18.77% | 28.16% | 3.4799 | 3.8462
Nd, ppm | 7.9799 | 0.28195 | 7.416 | 8.5438 | 7.134 | 8.8258 | 3.533% | 7.067% | 10.6% | 7.5809 | 8.3789
Ni, ppm | 104.8458 | 4.9164 | 95.013 | 114.6786 | 90.0967 | 119.5949 | 4.689% | 9.378% | 14.07% | 99.6035 | 110.0881
P, wt.% | 0.043809 | 0.0016997 | 0.040409 | 0.047208 | 0.03871 | 0.048908 | 3.88% | 7.76% | 11.64% | 0.041618 | 0.045999
Pb, ppm | 13.5667 | 1.1618 | 11.2431 | 15.8904 | 10.0813 | 17.0522 | 8.564% | 17.13% | 25.69% | 12.8884 | 14.2451
Pr, ppm | 1.642 | 0.1079 | 1.4262 | 1.8578 | 1.3183 | 1.9657 | 6.571% | 13.14% | 19.71% | 1.5599 | 1.7241
Rb, ppm | 10.3618 | 0.57065 | 9.2205 | 11.5031 | 8.6499 | 12.0738 | 5.507% | 11.01% | 16.52% | 9.8437 | 10.8799
Re, ppm | 0.0025729 | 0.00059802 | 0.0013769 | 0.003769 | 0.00077886 | 0.004367 | 23.24% | 46.49% | 69.73% | 0.0024443 | 0.0027016
S, wt.% | 0.26428 | 0.013975 | 0.23633 | 0.29223 | 0.22236 | 0.30621 | 5.288% | 10.58% | 15.86% | 0.25107 | 0.2775
Sb, ppm | 0.82095 | 0.062104 | 0.69675 | 0.94516 | 0.63464 | 1.0073 | 7.565% | 15.13% | 22.69% | 0.77991 | 0.862
Sc, ppm | 39.4319 | 2.2528 | 34.9264 | 43.9375 | 32.6736 | 46.1903 | 5.713% | 11.43% | 17.14% | 37.4603 | 41.4035
Sm, ppm | 2.44 | 0.18756 | 2.0649 | 2.8151 | 1.8773 | 3.0027 | 7.687% | 15.37% | 23.06% | 2.318 | 2.562
Sn, ppm | 0.93501 | 0.13501 | 0.665 | 1.205 | 0.52999 | 1.34 | 14.44% | 28.88% | 43.32% | 0.88826 | 0.98176
Sr, ppm | 203.2353 | 8.6025 | 186.0304 | 220.4403 | 177.4279 | 229.0428 | 4.233% | 8.466% | 12.7% | 193.0736 | 213.3971
Ta, ppm | 0.24188 | 0.049696 | 0.14249 | 0.34127 | 0.092794 | 0.39097 | 20.55% | 41.09% | 61.64% | 0.22979 | 0.25398
Tb, ppm | 0.55959 | 0.040555 | 0.47848 | 0.6407 | 0.43792 | 0.68125 | 7.247% | 14.49% | 21.74% | 0.53161 | 0.58757
Te, ppm | 0.073508 | 0.018313 | 0.036882 | 0.11013 | 0.018569 | 0.12845 | 24.91% | 49.83% | 74.74% | 0.069833 | 0.077184
Th, ppm | 0.65239 | 0.045704 | 0.56098 | 0.7438 | 0.51528 | 0.7895 | 7.006% | 14.01% | 21.02% | 0.61977 | 0.68501
Ti, wt.% | 0.63779 | 0.018035 | 0.60172 | 0.67386 | 0.58369 | 0.6919 | 2.828% | 5.656% | 8.483% | 0.6059 | 0.66968
Tl, ppm | 0.12208 | 0.012309 | 0.097464 | 0.1467 | 0.085155 | 0.15901 | 10.08% | 20.17% | 30.25% | 0.11598 | 0.12819
Tm, ppm | 0.33216 | 0.023555 | 0.28505 | 0.37927 | 0.2615 | 0.40283 | 7.092% | 14.18% | 21.27% | 0.31555 | 0.34877
U, ppm | 0.20272 | 0.0073575 | 0.188 | 0.21743 | 0.18065 | 0.22479 | 3.629% | 7.259% | 10.89% | 0.19258 | 0.21285
V, ppm | 282.8955 | 9.8642 | 263.1671 | 302.6239 | 253.3029 | 312.4881 | 3.487% | 6.974% | 10.46% | 268.7507 | 297.0403
W, ppm | 16.9595 | 1.2864 | 14.3867 | 19.5323 | 13.1004 | 20.8186 | 7.585% | 15.17% | 22.76% | 16.1115 | 17.8075
Y, ppm | 20.4727 | 1.3398 | 17.7932 | 23.1523 | 16.4534 | 24.4921 | 6.544% | 13.09% | 19.63% | 19.4491 | 21.4964
Yb, ppm | 2.1304 | 0.17135 | 1.7877 | 2.4731 | 1.6163 | 2.6444 | 8.043% | 16.09% | 24.13% | 2.0239 | 2.2369
Zn, ppm | 117.7622 | 6.7112 | 104.3398 | 131.1846 | 97.6286 | 137.8957 | 5.699% | 11.4% | 17.1% | 111.8741 | 123.6503
Zr, ppm | 45.4343 | 6.4793 | 32.4757 | 58.3928 | 25.9965 | 64.872 | 14.26% | 28.52% | 42.78% | 43.1625 | 47.706
Aqua Regia Digestion
Ag, ppm | 0.18226 | 0.014711 | 0.15283 | 0.21168 | 0.13812 | 0.22639 | 8.072% | 16.14% | 24.21% | 0.17314 | 0.19137
Al, wt.% | 3.1318 | 0.24765 | 2.6365 | 3.6271 | 2.3888 | 3.8747 | 7.908% | 15.82% | 23.72% | 2.9752 | 3.2884
As, ppm | 33.3234 | 1.637 | 30.0495 | 36.5974 | 28.4125 | 38.2344 | 4.912% | 9.825% | 14.74% | 31.6573 | 34.9896
B, ppm | 28.5899 | 8.0586 | 12.4727 | 44.7071 | 4.4141 | 52.7657 | 28.19% | 56.37% | 84.56% | 27.1604 | 30.0194
Ba, ppm | 35.8332 | 4.3907 | 27.0517 | 44.6146 | 22.661 | 49.0054 | 12.25% | 24.51% | 36.76% | 34.0415 | 37.6248
Be, ppm | 0.17019 | 0.023815 | 0.12256 | 0.21782 | 0.098748 | 0.24164 | 13.99% | 27.99% | 41.98% | 0.16168 | 0.1787
Bi, ppm | 0.027289 | 0.0048926 | 0.017504 | 0.037074 | 0.012611 | 0.041967 | 17.93% | 35.86% | 53.79% | 0.025925 | 0.028653
Ca, wt.% | 2.1067 | 0.32196 | 1.4628 | 2.7506 | 1.1409 | 3.0726 | 15.28% | 30.56% | 45.85% | 2.0014 | 2.2121
Cd, ppm | 0.34391 | 0.027883 | 0.28814 | 0.39967 | 0.26026 | 0.42755 | 8.108% | 16.22% | 24.32% | 0.32671 | 0.3611
Ce, ppm | 7.936 | 0.42056 | 7.0948 | 8.7771 | 6.6743 | 9.1976 | 5.299% | 10.6% | 15.9% | 7.5392 | 8.3328
Co, ppm | 33.3786 | 2.1513 | 29.076 | 37.6813 | 26.9247 | 39.8326 | 6.445% | 12.89% | 19.34% | 31.7097 | 35.0476
Cr, ppm | 23.8965 | 2.2873 | 19.3219 | 28.4712 | 17.0345 | 30.7585 | 9.572% | 19.14% | 28.72% | 22.7017 | 25.0913
Cs, ppm | 0.4401 | 0.029885 | 0.38033 | 0.49987 | 0.35044 | 0.52975 | 6.79% | 13.58% | 20.37% | 0.41809 | 0.4621
Cu, ppm | 154.8415 | 5.1696 | 144.5024 | 165.1806 | 139.3328 | 170.3502 | 3.339% | 6.677% | 10.02% | 147.0994 | 162.5836
Dy, ppm | 2.115 | 0.29085 | 1.5333 | 2.6967 | 1.2425 | 2.9875 | 13.75% | 27.5% | 41.26% | 2.0093 | 2.2208
Er, ppm | 1.2177 | 0.17628 | 0.86511 | 1.5702 | 0.68883 | 1.7465 | 14.48% | 28.95% | 43.43% | 1.1568 | 1.2786
Eu, ppm | 0.44405 | 0.091029 | 0.26199 | 0.62611 | 0.17096 | 0.71713 | 20.5% | 41% | 61.5% | 0.42185 | 0.46625
Fe, wt.% | 5.9057 | 0.17165 | 5.5624 | 6.249 | 5.3907 | 6.4206 | 2.906% | 5.813% | 8.719% | 5.6104 | 6.201
Ga, ppm | 10.2022 | 0.91299 | 8.3762 | 12.0282 | 7.4633 | 12.9412 | 8.949% | 17.9% | 26.85% | 9.6921 | 10.7123
Gd, ppm | 1.8169 | 0.19758 | 1.4218 | 2.2121 | 1.2242 | 2.4097 | 10.87% | 21.75% | 32.62% | 1.7261 | 1.9078
Ge, ppm | 0.13 | 0.02774 | 0.074521 | 0.18548 | 0.046781 | 0.21322 | 21.34% | 42.68% | 64.01% | 0.1235 | 0.1365
Hf, ppm | 0.39831 | 0.037818 | 0.32268 | 0.47395 | 0.28486 | 0.51177 | 9.494% | 18.99% | 28.48% | 0.3784 | 0.41823
Ho, ppm | 0.42389 | 0.04213 | 0.33963 | 0.50816 | 0.2975 | 0.55029 | 9.939% | 19.88% | 29.82% | 0.4027 | 0.44509
In, ppm | 0.024324 | 0.0038291 | 0.016665 | 0.031982 | 0.012836 | 0.035811 | 15.74% | 31.48% | 47.23% | 0.023107 | 0.02554
K, wt.% | 0.10389 | 0.0060441 | 0.091801 | 0.11598 | 0.085757 | 0.12202 | 5.818% | 11.64% | 17.45% | 0.098695 | 0.10908
La, ppm | 3.3424 | 0.17001 | 3.0023 | 3.6824 | 2.8323 | 3.8524 | 5.087% | 10.17% | 15.26% | 3.1752 | 3.5095
Li, ppm | 10.5797 | 0.65027 | 9.2792 | 11.8802 | 8.6289 | 12.5305 | 6.146% | 12.29% | 18.44% | 10.0507 | 11.1087
Lu, ppm | 0.12997 | 0.0272 | 0.075572 | 0.18437 | 0.048371 | 0.21157 | 20.93% | 41.86% | 62.78% | 0.12347 | 0.13647
Mg, wt.% | 1.9568 | 0.089371 | 1.778 | 2.1355 | 1.6887 | 2.2249 | 4.567% | 9.135% | 13.7% | 1.8589 | 2.0546
Mn, wt.% | 0.084178 | 0.0059831 | 0.072212 | 0.096144 | 0.066229 | 0.10213 | 7.108% | 14.22% | 21.32% | 0.079969 | 0.088387
Mo, ppm | 1.3948 | 0.09455 | 1.2057 | 1.5839 | 1.1111 | 1.6784 | 6.779% | 13.56% | 20.34% | 1.325 | 1.4645
Na, wt.% | 0.14683 | 0.011189 | 0.12445 | 0.16921 | 0.11326 | 0.1804 | 7.62% | 15.24% | 22.86% | 0.13949 | 0.15417
Nd, ppm | 5.3144 | 0.36565 | 4.5831 | 6.0457 | 4.2174 | 6.4113 | 6.88% | 13.76% | 20.64% | 5.0487 | 5.5801
Ni, ppm | 72.8669 | 3.3314 | 66.204 | 79.5298 | 62.8725 | 82.8612 | 4.572% | 9.144% | 13.72% | 69.2235 | 76.5102
P, wt.% | 0.043109 | 0.0023731 | 0.038363 | 0.047856 | 0.03599 | 0.050229 | 5.505% | 11.01% | 16.51% | 0.040954 | 0.045265
Pb, ppm | 13.3381 | 0.58127 | 12.1756 | 14.5007 | 11.5943 | 15.0819 | 4.358% | 8.716% | 13.07% | 12.6712 | 14.005
Pr, ppm | 1.0724 | 0.072834 | 0.92672 | 1.2181 | 0.85389 | 1.2909 | 6.792% | 13.58% | 20.38% | 1.0188 | 1.126
Rb, ppm | 3.9903 | 0.40049 | 3.1893 | 4.7912 | 2.7888 | 5.1917 | 10.04% | 20.07% | 30.11% | 3.7907 | 4.1898
Re, ppm | 0.0019444 | 0.00049672 | 0.000951 | 0.0029379 | 0.00045427 | 0.0034346 | 25.55% | 51.09% | 76.64% | 0.0018472 | 0.0020417
S, wt.% | 0.2615 | 0.012902 | 0.2357 | 0.28731 | 0.2228 | 0.30021 | 4.934% | 9.868% | 14.8% | 0.24843 | 0.27458
Sb, ppm | 0.50231 | 0.1094 | 0.28351 | 0.72111 | 0.17411 | 0.83051 | 21.78% | 43.56% | 65.34% | 0.47719 | 0.52742
Sc, ppm | 4.4618 | 0.29324 | 3.8754 | 5.0483 | 3.5821 | 5.3415 | 6.572% | 13.14% | 19.72% | 4.2387 | 4.6849
Sm, ppm | 1.5312 | 0.1243 | 1.2826 | 1.7798 | 1.1583 | 1.9041 | 8.118% | 16.24% | 24.35% | 1.4546 | 1.6078
Sn, ppm | 0.52088 | 0.053294 | 0.4143 | 0.62747 | 0.361 | 0.68077 | 10.23% | 20.46% | 30.69% | 0.49484 | 0.54693
Sr, ppm | 32.046 | 6.5305 | 18.985 | 45.1071 | 12.4545 | 51.6376 | 20.38% | 40.76% | 61.14% | 30.4437 | 33.6484
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.3066 | 0.010977 | 0.28464 | 0.32855 | 0.27366 | 0.33953 | 3.58% | 7.161% | 10.74% | 0.29127 | 0.32193
Te, ppm | 0.059653 | 0.0077324 | 0.044188 | 0.075118 | 0.036456 | 0.08285 | 12.96% | 25.92% | 38.89% | 0.05667 | 0.062635
Th, ppm | 0.45898 | 0.052249 | 0.35448 | 0.56348 | 0.30223 | 0.61573 | 11.38% | 22.77% | 34.15% | 0.43603 | 0.48193
Ti, wt.% | 0.32883 | 0.059143 | 0.21054 | 0.44711 | 0.1514 | 0.50625 | 17.99% | 35.97% | 53.96% | 0.31239 | 0.34527
Tl, ppm | 0.064382 | 0.0062632 | 0.051856 | 0.076908 | 0.045592 | 0.083172 | 9.728% | 19.46% | 29.18% | 0.061163 | 0.067601
Tm, ppm | 0.16578 | 0.019349 | 0.12708 | 0.20447 | 0.10773 | 0.22382 | 11.67% | 23.34% | 35.01% | 0.15749 | 0.17407
U, ppm | 0.13068 | 0.010963 | 0.10875 | 0.1526 | 0.097787 | 0.16356 | 8.389% | 16.78% | 25.17% | 0.12414 | 0.13721
V, ppm | 140.6777 | 20.258 | 100.1618 | 181.1937 | 79.9038 | 201.4516 | 14.4% | 28.8% | 43.2% | 133.6438 | 147.7116
W, ppm | 11.8727 | 1.9514 | 7.9699 | 15.7755 | 6.0186 | 17.7268 | 16.44% | 32.87% | 49.31% | 11.2791 | 12.4663
Y, ppm | 10.777 | 0.64293 | 9.4911 | 12.0628 | 8.8482 | 12.7058 | 5.966% | 11.93% | 17.9% | 10.2381 | 11.3158
Yb, ppm | 0.98121 | 0.13804 | 0.70513 | 1.2573 | 0.56709 | 1.3953 | 14.07% | 28.14% | 42.2% | 0.93215 | 1.0303
Zn, ppm | 102.0616 | 5.7356 | 90.5904 | 113.5329 | 84.8548 | 119.2685 | 5.62% | 11.24% | 16.86% | 96.9586 | 107.1647
Zr, ppm | 12.1224 | 1.8305 | 8.4614 | 15.7835 | 6.6309 | 17.614 | 15.1% | 30.2% | 45.3% | 11.5163 | 12.7286
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 231b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
N.A. | Not Applicable (Lab 1.27)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 231b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
Nagrom, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 231b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 3 | 3 | 238517 | 0.088314 | 0.58215 | 0.57786
2 | 1 | ANSLu | 1 | 5 | 5 | 238518 | 0.085717 | 0.59777 | 0.57869
2 | 1 | ANSLu | 1 | 1 | 1 | 238519 | 0.083765 | 0.53229 | 0.57519
2 | 1 | ANSLu | 1 | 15 | 15 | 238520 | 0.086151 | 0.57987 | 0.57773
2 | 1 | ANSLu | 1 | 17 | 17 | 238521 | 0.08574 | 0.55763 | 0.57654
2 | 1 | ANSLu | 1 | 9 | 9 | 238522 | 0.087255 | 0.55038 | 0.57616
2 | 1 | ANSLu | 1 | 6 | 6 | 238523 | 0.086271 | 0.53701 | 0.57544
2 | 1 | ANSLu | 1 | 14 | 14 | 238524 | 0.08533 | 0.56893 | 0.57715
2 | 1 | ANSLu | 1 | 10 | 10 | 238525 | 0.082889 | 0.52829 | 0.57497
2 | 1 | ANSLu | 1 | 16 | 16 | 238526 | 0.084395 | 0.58093 | 0.57779
2 | 1 | ANSLu | 1 | 4 | 4 | 238527 | 0.084648 | 0.58928 | 0.57824
2 | 1 | ANSLu | 1 | 13 | 13 | 238528 | 0.085756 | 0.57244 | 0.57734
2 | 1 | ANSLu | 1 | 11 | 11 | 238529 | 0.088808 | 0.61881 | 0.57982
2 | 1 | ANSLu | 1 | 8 | 8 | 238530 | 0.085914 | 0.59283 | 0.57843
2 | 1 | ANSLu | 1 | 7 | 7 | 238531 | 0.087071 | 0.55543 | 0.57643
2 | 1 | ANSLu | 1 | 12 | 12 | 238532 | 0.086159 | 0.62089 | 0.57993
2 | 1 | ANSLu | 1 | 20 | 20 | 238533 | 0.082772 | 0.57003 | 0.57721
2 | 1 | ANSLu | 1 | 18 | 18 | 238534 | 0.088494 | 0.59315 | 0.57844
2 | 1 | ANSLu | 1 | 2 | 2 | 238535 | 0.085701 | 0.63202 | 0.58052
2 | 1 | ANSLu | 1 | 19 | 19 | 238536 | 0.085664 | 0.59213 | 0.57839
Mean | [6 empty] | 0.085841 | 0.57761 | 0.57761
Median | [7 empty] | 0.5804 | 0.57776
SD | [7 empty] | 0.028765 | 0.0015387
RSD | [7 empty] | 4.98% | 0.2664%
RSD Upscale Factor | [7 empty] | 0.053492
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 231b (Certified Value 0.556 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 30g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 50g | 50g | 25g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 15g | 30g | 30g | 25g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 0.53229 | 0.557 | 0.575 | 0.584 | 0.5603 | 0.56264 | 0.52 | 0.574 | 0.56 | 0.5 | 0.56 | 0.551 | 0.6 | 0.629 | 0.56 | 0.559 | 0.592 | 0.544 | 0.535 | 0.494 | 0.55 | 0.585 | 0.522 | 0.536 | 0.58485 | 0.547 | 0.5556 | 0.5569 | 0.559 | 0.567 | 0.558 | 0.579 | 1
| 1 | 2 | 0.63202 | 0.546 | 0.57 | 0.554 | 0.55555 | 0.57777 | 0.54 | 0.564 | 0.558 | 0.58 | 0.57 | 0.547 | 0.6 | 0.588 | 0.555 | 0.543 | 0.557 | 0.558 | 0.528 | 0.55 | 0.54 | 0.56 | 0.516 | 0.552 | 0.5754 | 0.56 | 0.53426 | 0.561 | 0.544 | 0.561 | 0.562 | 0.587 | #N/A
| 1 | 3 | 0.58215 | 0.57 | 0.566 | 0.546 | 0.5565 | 0.56901 | 0.54 | 0.548 | 0.568 | 0.48 | 0.56 | 0.55 | 0.6 | 0.561 | 0.547 | 0.551 | 0.558 | 0.554 | 0.541 | 0.518 | 0.55 | 0.572 | 0.52 | 0.552 | 0.57435 | 0.556 | 0.53366 | 0.563 | 0.488 | 0.556 | 0.566 | 0.588 | 16
| 1 | 4 | 0.58928 | 0.562 | 0.569 | 0.582 | 0.55935 | 0.58014 | 0.52 | 0.563 | 0.547 | 0.51 | 0.57 | 0.555 | 0.6 | 0.556 | 0.545 | 0.553 | 0.563 | 0.571 | 0.523 | 0.528 | 0.55 | 0.586 | 0.539 | 0.561 | 0.5796 | 0.55 | 0.56682 | 0.5435 | 0.562 | 0.547 | 0.556 | 0.578 | 0.55644
| 1 | 5 | 0.59777 | 0.563 | 0.568 | 0.492 | 0.5413 | 0.58491 | 0.52 | 0.558 | 0.577 | 0.51 | 0.57 | 0.556 | 0.614 | 0.6 | 0.556 | 0.546 | 0.559 | 0.573 | 0.533 | 0.539 | 0.54 | 0.581 | 0.544 | 0.528 | 0.5733 | 0.545 | 0.54945 | 0.5552 | 0.544 | 0.572 | 0.55 | 0.577 | 7
| 1 | 6 | 0.53701 | 0.557 | 0.571 | 0.588 | 0.55745 | 0.57554 | 0.52 | 0.556 | 0.573 | 0.55 | 0.56 | 0.561 | 0.6 | 0.581 | 0.552 | 0.547 | 0.544 | 0.57 | 0.533 | 0.516 | 0.55 | 0.55 | 0.535 | 0.517 | 0.57015 | 0.538 | 0.55191 | 0.5591 | 0.569 | 0.566 | 0.544 | 0.599
|  | 7 | 0.55543
|  | 8 | 0.59283
|  | 9 | 0.55038
|  | 10 | 0.52829
|  | 11 | 0.61881
|  | 12 | 0.62089
|  | 13 | 0.57244
|  | 14 | 0.56893
|  | 15 | 0.57987
|  | 16 | 0.58093
|  | 17 | 0.55763
|  | 18 | 0.59315
|  | 19 | 0.59213
|  | 20 | 0.57003
| Mean |  | 0.57761 | 0.55917 | 0.56983 | 0.55767 | 0.55507 | 0.575 | 0.52667 | 0.5605 | 0.56383 | 0.52167 | 0.565 | 0.55333 | 0.60233 | 0.58583 | 0.5525 | 0.54983 | 0.56217 | 0.56167 | 0.53217 | 0.52417 | 0.54667 | 0.57233 | 0.52933 | 0.541 | 0.57627 | 0.54933 | 0.54862 | 0.55645 | 0.54433 | 0.5615 | 0.556 | 0.58467
| Median |  | 0.5804 | 0.5595 | 0.5695 | 0.568 | 0.55697 | 0.57666 | 0.52 | 0.5605 | 0.564 | 0.51 | 0.565 | 0.553 | 0.6 | 0.5845 | 0.5535 | 0.549 | 0.5585 | 0.564 | 0.533 | 0.523 | 0.55 | 0.5765 | 0.5285 | 0.544 | 0.57488 | 0.5485 | 0.55068 | 0.558 | 0.5515 | 0.5635 | 0.557 | 0.583
| Std Dev. |  | 0.028765 | 0.0080353 | 0.0030605 | 0.036517 | 0.0069746 | 0.0080172 | 0.010328 | 0.0087579 | 0.011017 | 0.03656 | 0.0054772 | 0.0050067 | 0.0057155 | 0.026828 | 0.0056833 | 0.0057417 | 0.015968 | 0.01157 | 0.0061455 | 0.01958 | 0.005164 | 0.014652 | 0.011483 | 0.016805 | 0.0052043 | 0.0078909 | 0.012818 | 0.0069296 | 0.029358 | 0.008961 | 0.008 | 0.0084538
| Rel.Std.Dev. |  | 4.98% | 1.437% | 0.5371% | 6.548% | 1.257% | 1.394% | 1.961% | 1.563% | 1.954% | 7.008% | 0.9694% | 0.9048% | 0.9489% | 4.58% | 1.029% | 1.044% | 2.84% | 2.06% | 1.155% | 3.735% | 0.9446% | 2.56% | 2.169% | 3.106% | 0.9031% | 1.436% | 2.336% | 1.245% | 5.393% | 1.596% | 1.439% | 1.446%
| PDM3 |  | 3.805% | 0.49% | 2.407% | 0.2205% | -0.2453% | 3.336% | -5.351% | 0.7297% | 1.329% | -6.249% | 1.538% | -0.5583% | 8.248% | 5.282% | -0.708% | -1.187% | 1.029% | 0.9393% | -4.362% | -5.8% | -1.756% | 2.856% | -4.871% | -2.775% | 3.565% | -1.277% | -1.406% | 0.001827% | -2.176% | 0.9094% | -0.07904% | 5.073%
| Z-Score (Absolute) |  | NA | 0.21 | 1.06 | 0.1 | 0.11 | 1.46 | 2.35 | 0.32 | 0.58 | 2.74 | 0.67 | 0.25 | 3.62 | 2.32 | 0.31 | 0.52 | 0.45 | 0.41 | 1.92 | 2.55 | 0.77 | 1.25 | 2.14 | 1.22 | 1.56 | 0.56 | 0.62 | 0 | 0.96 | 0.4 | 0.04 | 2.23
| Analytical results for Pd in OREAS 231b (Indicative Value 20.8 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 22.2174 | [31 empty] | 1
| 1 | 2 | 20.0811 | [31 empty] | 1
| 1 | 3 | 21.1875 | [31 empty] | 16
| 1 | 4 | 21.2447 | [31 empty] | 20.7858
| 1 | 5 | 20.0003 | [31 empty] | 7
| 1 | 6 | 19.9839
| Mean |  | 20.7858
| Median |  | 20.6343
| Std Dev. |  | 0.91411
| Rel.Std.Dev. |  | 4.398%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 231b (Indicative Value 10.3 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 10.6336 | [31 empty] | 1
| 1 | 2 | 9.9895 | [31 empty] | 1
| 1 | 3 | 9.8636 | [31 empty] | 16
| 1 | 4 | 10.6708 | [31 empty] | 10.285
| 1 | 5 | 10.2086 | [31 empty] | 7
| 1 | 6 | 10.344
| Mean |  | 10.285
| Median |  | 10.2763
| Std Dev. |  | 0.33002
| Rel.Std.Dev. |  | 3.209%
| PDM3 |  | -3.331e-13%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 231b (Certified Value 0.556 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 6.01 | 6.02 | 6.03 | 6.04 | ppm
|  |  | INAA | Raw*PA | Various [a] | Raw*PA | Various [b] | Various [c] | Raw*PA | Various [d] | Raw*PA | Raw*PA | Raw*PA | Various [e] | 3
|  |  | 0.085g | 400g | see footer | 400g | see footer | see footer | 400g | see footer | 450g | 400g | 450g | see footer | 3
| 1 | 1 | 0.53229 | 0.59 | 0.55 |  | 0.56383 | 0.525 | 0.50946 | 0.554 | 0.528 | 0.556 | 0.536 | 0.586 | 1
| 1 | 2 | 0.63202 | 0.57 | 0.56 |  | 0.56181 | 0.532 | 0.55338 | 0.554 | 0.514 | 0.571 | 0.521 | 0.566 | #N/A
| 1 | 3 | 0.58215 | 0.54 | 0.57 |  | 0.5471 | 0.532 | 0.57485 | 0.553 | 0.528 | 0.535 | 0.583 | 0.561 | 16
| 1 | 4 | 0.58928 | 0.57 | 0.57 |  | 0.52559 | 0.511 | 0.48208 | 0.495 | 0.625 | 0.59 | 0.51 | 0.569 | 0.55617
| 1 | 5 | 0.59777 | 0.57 | 0.61 |  | 0.57401 | 0.519 | 0.56361 | 0.587 | 0.554 | 0.589 | 0.521 | 0.623 | 9
| 1 | 6 | 0.53701 | 0.57 | 0.54 |  | 0.54863 | 0.571 | 0.5012 | 0.553 | 0.521 | 0.546 | 0.568 | 0.57
| 2 | 7 | 0.55543 | 0.57 | 0.55 |  | 0.52 | 0.53 | 0.596 | 0.494 | 0.548 | 0.627 | 0.55345 | 0.529
| 2 | 8 | 0.59283 | 0.56 | 0.62 |  | 0.53 | 0.53 | 0.511 | 0.501 | 0.615 | 0.552 | 0.58482 | 0.56
| 2 | 9 | 0.55038 | 0.56 | 0.64 |  | 0.58 | 0.61 | 0.513 | 0.495 | 0.567 | 0.557 | 0.61852 | 0.534
| 2 | 10 | 0.52829 | 0.54 | 0.61 |  | 0.6 | 0.54 | 0.532 | 0.56 | 0.584 | 0.578 | 0.58398 | 0.587
| 2 | 11 | 0.61881 | 0.55 | 0.55 |  | 0.58 | 0.61 | 0.538 | 0.552 | 0.56 | 0.557 | 0.58469 | 0.548
| 2 | 12 | 0.62089 | 0.57 | 0.61 |  | 0.62 | 0.58 | 0.54 | 0.456 | 0.585 | 0.545 | 0.56236 | 0.514
| 3 | 13 | 0.57244 | 0.66 | 0.58
| 3 | 14 | 0.56893 | 0.58 | 0.53
| 3 | 15 | 0.57987 | 0.61 | 0.59
| 3 | 16 | 0.58093 | 0.55 | 0.54
| 3 | 17 | 0.55763 | 0.56 | 0.52
| 3 | 18 | 0.59315 | 0.6 | 0.59
| 4 | 19 | 0.59213 | 0.58 | 0.59 | 0.55
| 4 | 20 | 0.57003 | 0.54 | 0.52 | 0.55
| 4 | 21 |  | 0.56 | 0.58 | 0.57
| 4 | 22 |  | 0.53 | 0.56 | 0.58
| 4 | 23 |  | 0.52 | 0.52
| 4 | 24 |  | 0.55 | 0.58
| 5 | 25 |  |  | 0.52 | 0.52
| 5 | 26 |  |  | 0.55 | 0.52
| 5 | 27 |  |  | 0.54 | 0.52
| 5 | 28 |  |  | 0.56 | 0.52
| 5 | 29 |  |  | 0.55
| 5 | 30 |  |  | 0.59
| 6 | 31 |  |  | 0.54 | 0.55
| 6 | 32 |  |  | 0.57 | 0.54
| 6 | 33 |  |  | 0.6 | 0.54
| 6 | 34 |  |  | 0.57 | 0.54
| 6 | 35 |  |  | 0.55
| 6 | 36 |  |  | 0.61
| Mean |  | 0.57761 | 0.56667 | 0.5675 | 0.54167 | 0.56258 | 0.54917 | 0.53455 | 0.5295 | 0.56075 | 0.56692 | 0.56057 | 0.56225
| Median |  | 0.5804 | 0.565 | 0.565 | 0.54 | 0.56282 | 0.532 | 0.535 | 0.5525 | 0.557 | 0.557 | 0.56518 | 0.5635
| Std Dev. |  | 0.028765 | 0.028993 | 0.031294 | 0.019924 | 0.030322 | 0.034625 | 0.033249 | 0.0392 | 0.036129 | 0.025664 | 0.033003 | 0.029277
| Rel.Std.Dev. |  | 4.98% | 5.116% | 5.514% | 3.678% | 5.39% | 6.305% | 6.22% | 7.403% | 6.443% | 4.527% | 5.887% | 5.207%
| PDM3 |  | 3.855% | 1.887% | 2.037% | -2.608% | 1.152% | -1.26% | -3.888% | -4.796% | 0.8232% | 1.932% | 0.7906% | 1.093%
| Z-Score (Absolute) |  | NA | 0.65 | 0.74 | 2.09 | 0.2 | 1.27 | 2.87 | 3.42 | 0 | 0.67 | 0.02 | 0.16
[a]: Lab 5.02: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (350g), Rnd 03 - Raw*PA (350g), Rnd 04 - Raw*PA (400g), Rnd 05 - Raw*PA (400g), Rnd 06 - Raw*PA (400g)
[b]: Lab 5.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[c]: Lab 5.05: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (300g)
[d]: Lab 5.07: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[e]: Lab 6.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 231b (Certified Value 0.526 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.33 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.085g | 30g | 25g | 25g | 30g | 50g | 30g | 15g | 50g | 30g | 10g | 30g | 25g | 25g | 25g | 25g | 15g | 25g | 15g | 50g | 25g | 25g | 10g | 20g | 3
| 1 | 1 | 0.53229 | 0.65 | 0.61 | 0.5 | 0.533 | 0.51 | 0.516 | 0.545 | 0.5 | 0.52 | 0.513 | 0.54316 | 0.513 | 0.521 | 0.489 | 0.528 | 0.5353 | 0.535 | 0.4997 | 0.5358 | 0.575 | 0.52 | 0.543 | 0.539 | 1
| 1 | 2 | 0.63202 | 0.52 | 0.6 | 0.52 | 0.543 | 0.5 | 0.497 | 0.572 | 0.46 | 0.53 | 0.516 | 0.526 | 0.526 | 0.507 | 0.514 | 0.537 | 0.5353 | 0.498 | 0.5086 | 0.5374 | 0.57 | 0.51 | 0.545 | 0.5395 | #N/A
| 1 | 3 | 0.58215 | 0.53 | 0.61 | 0.56 | 0.543 | 0.5 | 0.5 | 0.546 | 0.5 | 0.52 | 0.507 | 0.508 | 0.496 | 0.515 | 0.491 | 0.547 | 0.5252 | 0.493 | 0.4942 | 0.555 | 0.575 | 0.526 | 0.549 | 0.5435 | 16
| 1 | 4 | 0.58928 | 0.56 | 0.62 | 0.6 | 0.543 | 0.5 | 0.504 | 0.565 | 0.48 | 0.52 | 0.532 | 0.52 | 0.526 | 0.528 | 0.485 | 0.526 | 0.5252 | 0.544 | 0.5152 | 0.5491 | 0.55 | 0.514 | 0.53 | 0.5344 | 0.52588
| 1 | 5 | 0.59777 | 0.57 | 0.59 | 0.52 | 0.55 | 0.5 | 0.505 | 0.569 | 0.51 | 0.52 | 0.505 | 0.52 | 0.538 | 0.507 | 0.503 | 0.505 | 0.5454 | 0.518 | 0.5029 | 0.5759 | 0.55 | 0.5 | 0.543 | 0.5332 | 11
| 1 | 6 | 0.53701 | 0.47 | 0.59 | 0.62 | 0.543 | 0.49 | 0.485 | 0.575 | 0.5 | 0.53 | 0.534 | 0.538 | 0.521 | 0.525 | 0.517 | 0.524 | 0.5252 | 0.517 | 0.4932 | 0.5472 | 0.57 | 0.54 | 0.547 | 0.5252
|  | 7 | 0.55543
|  | 8 | 0.59283
|  | 9 | 0.55038
|  | 10 | 0.52829
|  | 11 | 0.61881
|  | 12 | 0.62089
|  | 13 | 0.57244
|  | 14 | 0.56893
|  | 15 | 0.57987
|  | 16 | 0.58093
|  | 17 | 0.55763
|  | 18 | 0.59315
|  | 19 | 0.59213
|  | 20 | 0.57003
| Mean |  | 0.57761 | 0.55 | 0.60333 | 0.55333 | 0.5425 | 0.5 | 0.50117 | 0.562 | 0.49167 | 0.52333 | 0.51783 | 0.52586 | 0.52 | 0.51717 | 0.49983 | 0.52783 | 0.53193 | 0.5175 | 0.5023 | 0.55007 | 0.565 | 0.51833 | 0.54283 | 0.5358
| Median |  | 0.5804 | 0.545 | 0.605 | 0.54 | 0.543 | 0.5 | 0.502 | 0.567 | 0.5 | 0.52 | 0.5145 | 0.523 | 0.5235 | 0.518 | 0.497 | 0.527 | 0.53025 | 0.5175 | 0.5013 | 0.54815 | 0.57 | 0.517 | 0.544 | 0.5367
| Std Dev. |  | 0.028765 | 0.060332 | 0.012111 | 0.048442 | 0.0054314 | 0.0063246 | 0.010226 | 0.013206 | 0.018348 | 0.005164 | 0.012416 | 0.012917 | 0.014297 | 0.0089981 | 0.013571 | 0.014077 | 0.0082466 | 0.019947 | 0.0085046 | 0.014593 | 0.011832 | 0.013823 | 0.0067057 | 0.0063922
| Rel.Std.Dev. |  | 4.98% | 10.97% | 2.007% | 8.755% | 1.001% | 1.265% | 2.04% | 2.35% | 3.732% | 0.9867% | 2.398% | 2.456% | 2.749% | 1.74% | 2.715% | 2.667% | 1.55% | 3.855% | 1.693% | 2.653% | 2.094% | 2.667% | 1.235% | 1.193%
| PDM3 |  | 9.838% | 4.587% | 14.73% | 5.221% | 3.161% | -4.921% | -4.699% | 6.869% | -6.505% | -0.4837% | -1.53% | -0.003291% | -1.118% | -1.656% | -4.952% | 0.372% | 1.152% | -1.593% | -4.483% | 4.6% | 7.44% | -1.434% | 3.224% | 1.887%
| Z-Score (Absolute) |  | NA | 0.96 | 3.08 | 1.09 | 0.66 | 1.03 | 0.98 | 1.44 | 1.36 | 0.1 | 0.32 | 0 | 0.23 | 0.35 | 1.03 | 0.08 | 0.24 | 0.33 | 0.94 | 0.96 | 1.55 | 0.3 | 0.67 | 0.39
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 231b (Certified Value 0.524 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.30 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 05g | 50g | 60g | 30g | 30g | 30g | 200g | 30g | 200g | 200g | 20g | 30g | 05g | 3
| 1 | 1 | 0.53229 | 0.495 | 0.53 | 0.575 | 0.512 | 0.52289 | 0.577 | 0.502 | 0.56 | 0.5 | 0.49397 | 0.54 | 0.59 | 0.43 | 0.55 | 0.56 | 0.53 | 0.513 | 0.53429 | 0.51 | 0.5296 | 0.54 | 1
| 1 | 2 | 0.63202 | 0.438 | 0.53 | 0.55 | 0.518 | 0.50823 | 0.511 | 0.51 | 0.53 | 0.5 | 0.512 | 0.56 | 0.59 | 0.4 | 0.55 | 0.544 | 0.54 | 0.509 | 0.53204 | 0.51 | 0.5255 | 0.53 | #N/A
| 1 | 3 | 0.58215 | 0.448 | 0.51 | 0.575 | 0.512 | 0.52111 | 0.534 | 0.518 | 0.52737 | 0.5 | 0.49597 | 0.54 | 0.6 | 0.48 | 0.54 | 0.562 | 0.53 | 0.509 | 0.52995 | 0.51 | 0.5284 | 0.53 | 16
| 1 | 4 | 0.58928 | 0.519 | 0.54 | 0.575 | 0.512 | 0.50863 | 0.474 | 0.503 | 0.56917 | 0.5 | 0.512 | 0.55 | 0.59 | 0.47 | 0.53 | 0.56 | 0.53 | 0.508 | 0.54614 | 0.52 | 0.5207 | 0.55 | 0.52369
| 1 | 5 | 0.59777 | 0.532 | 0.54 | 0.5 | 0.516 | 0.54272 | 0.474 | 0.513 | 0.52914 | 0.5 | 0.51 | 0.55 | 0.59 | 0.46 | 0.45 | 0.53 | 0.54 | 0.507 | 0.54326 | 0.51 | 0.5198 | 0.54 | 13
| 1 | 6 | 0.53701 | 0.505 | 0.53 | 0.525 | 0.528 | 0.51002 | 0.488 | 0.496 | 0.55168 | 0.5 | 0.52133 | 0.53 | 0.6 | 0.44 | 0.47 | 0.546 | 0.53 | 0.504 | 0.53877 | 0.52 | 0.527 | 0.54
|  | 7 | 0.55543
|  | 8 | 0.59283
|  | 9 | 0.55038
|  | 10 | 0.52829
|  | 11 | 0.61881
|  | 12 | 0.62089
|  | 13 | 0.57244
|  | 14 | 0.56893
|  | 15 | 0.57987
|  | 16 | 0.58093
|  | 17 | 0.55763
|  | 18 | 0.59315
|  | 19 | 0.59213
|  | 20 | 0.57003
| Mean |  | 0.57761 | 0.4895 | 0.53 | 0.55 | 0.51633 | 0.51893 | 0.50967 | 0.507 | 0.54456 | 0.5 | 0.50754 | 0.545 | 0.59333 | 0.44667 | 0.515 | 0.55033 | 0.53333 | 0.50833 | 0.53741 | 0.51333 | 0.52517 | 0.53833
| Median |  | 0.5804 | 0.5 | 0.53 | 0.5625 | 0.514 | 0.51556 | 0.4995 | 0.5065 | 0.54084 | 0.5 | 0.511 | 0.545 | 0.59 | 0.45 | 0.535 | 0.553 | 0.53 | 0.5085 | 0.53653 | 0.51 | 0.52625 | 0.54
| Std Dev. |  | 0.028765 | 0.038266 | 0.010954 | 0.031623 | 0.0062503 | 0.013312 | 0.040352 | 0.0080994 | 0.01811 | 0 | 0.010526 | 0.010488 | 0.005164 | 0.029439 | 0.043704 | 0.012612 | 0.005164 | 0.0029439 | 0.0064273 | 0.005164 | 0.0040579 | 0.0075277
| Rel.Std.Dev. |  | 4.98% | 7.817% | 2.067% | 5.75% | 1.211% | 2.565% | 7.917% | 1.598% | 3.326% | 0% | 2.074% | 1.924% | 0.8703% | 6.591% | 8.486% | 2.292% | 0.9682% | 0.5791% | 1.196% | 1.006% | 0.7727% | 1.398%
| PDM3 |  | 10.3% | -6.528% | 1.205% | 5.024% | -1.404% | -0.9082% | -2.677% | -3.187% | 3.986% | -4.523% | -3.083% | 4.07% | 13.3% | -14.71% | -1.659% | 5.088% | 1.842% | -2.932% | 2.62% | -1.977% | 0.2823% | 2.796%
| Z-Score (Absolute) |  | NA | 1.38 | 0.52 | 1.45 | 0.12 | 0 | 0.43 | 0.56 | 1.2 | 0.89 | 0.53 | 1.22 | 3.48 | 3.38 | 0.18 | 1.47 | 0.67 | 0.5 | 0.87 | 0.26 | 0.29 | 0.91
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 231b (Certified Value 0.182 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.25 | 0.19 | 0.1799 | < 0.3 | 0.41 | 0.2 | 0.16 | < 0.5 | 0.2 | 0.19 | 0.19 | 0.19 | 0.18 | < 0.2 | 0.2 | 0.1 | 0.2 | < 0.2 | < 1 | 0.17 | < 0.3 | 0.19 | 0.2 | 0.22 | 0.18 |  |  |  |  | 1
| 1 | 2 | 0.24 | 0.19 | 0.18464 | < 0.3 | 0.46 | 0.2 | 0.17 | < 0.5 | 0.2 | 0.19 | 0.2 | 0.18 | 0.21 | < 0.2 | 0.2 | 0.2 | 0.19 | < 0.2 | < 1 | 0.2 | < 0.3 | 0.18 | 0.17 | 0.21 | 0.17 |  |  |  |  | 19
| 1 | 3 | 0.26 | 0.2 | 0.17913 | < 0.3 | 0.43 | 0.2 | 0.17 | < 0.5 | 0.2 | 0.19 | 0.19 | 0.19 | 0.18 | < 0.2 | 0.2 | 0.2 | 0.19 | 0.22 | 5 | 0.17 | < 0.3 | 0.18 | 0.18 | 0.21 | 0.17 |  |  |  |  | 16
| 1 | 4 | 0.25 | 0.2 | 0.1789 | < 0.3 | 0.35 | 0.2 | 0.17 | < 0.5 | 0.2 | 0.19 | 0.18 | 0.19 | 0.18 | < 0.2 | 0.2 | 0.2 | 0.2 | 0.24 | < 1 | 0.18 | < 0.3 | 0.18 | 0.17 | 0.21 | 0.16 |  |  |  |  | 0.18226
| 1 | 5 | 0.26 | 0.19 | 0.19092 | < 0.3 | 0.33 | 0.2 | 0.16 | < 0.5 | 0.2 | 0.17 | 0.17 | 0.19 | 0.17 | < 0.2 | 0.2 | 0.2 | 0.19 | 0.21 | < 1 | 0.19 | < 0.3 | 0.19 | 0.19 | 0.22 | 0.16 |  |  |  |  | 15
| 1 | 6 | 0.25 | 0.19 | 0.17993 | < 0.3 | 0.32 | 0.2 | 0.18 | < 0.5 | 0.2 | 0.19 | 0.18 | 0.18 | 0.17 | < 0.2 | 0.2 | 0.2 | 0.23 | 0.61 | < 1 | 0.21 | < 0.3 | 0.17 | 0.18 | 0.21 | 0.16
| Mean |  | 0.25167 | 0.19333 | 0.18224 |  | 0.38333 | 0.2 | 0.16833 |  | 0.2 | 0.18667 | 0.185 | 0.18667 | 0.18167 |  | 0.2 | 0.18333 | 0.2 | 0.32 | 5 | 0.18667 |  | 0.18167 | 0.18167 | 0.21333 | 0.16667
| Median |  | 0.25 | 0.19 | 0.17992 |  | 0.38 | 0.2 | 0.17 |  | 0.2 | 0.19 | 0.185 | 0.19 | 0.18 |  | 0.2 | 0.2 | 0.195 | 0.23 | 5 | 0.185 |  | 0.18 | 0.18 | 0.21 | 0.165
| Std Dev. |  | 0.0075277 | 0.005164 | 0.0047483 |  | 0.05785 | 3.0405e-17 | 0.0075277 |  | 3.0405e-17 | 0.008165 | 0.010488 | 0.005164 | 0.01472 |  | 3.0405e-17 | 0.040825 | 0.015492 | 0.19374 |  | 0.01633 |  | 0.0075277 | 0.01169 | 0.005164 | 0.008165
| Rel.Std.Dev. |  | 2.991% | 2.671% | 2.606% |  | 15.09% | 1.52e-14% | 4.472% |  | 1.52e-14% | 4.374% | 5.669% | 2.766% | 8.103% |  | 1.52e-14% | 22.27% | 7.746% | 60.54% |  | 8.748% |  | 4.144% | 6.435% | 2.421% | 4.899%
| PDM3 |  | 38.08% | 6.073% | -0.01367% |  | 110.3% | 9.731% | -7.643% |  | 9.731% | 2.415% | 1.501% | 2.415% | -0.3278% |  | 9.731% | 0.5866% | 9.731% | 75.57% | 2643% | 2.415% |  | -0.3278% | -0.3278% | 17.05% | -8.558%
| Z-Score (Absolute) |  | 2.8 | 0.32 | 0.15 | 1.53 | 8.41 | NA | 0.75 | 2.73 | NA | 0.04 | 0.04 | 0.04 | 0.18 | 3.66 | NA | NA | 0.6 | 2.59 | NA | 0.04 | 1.53 | 0.18 | 0.18 | 1.17 | 0.82
Results from laboratories 1.10, 1.13, 1.22 and 1.23 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Al in OREAS 231b (Certified Value 6.89 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.84 | 6.7864 | 7.1604 | 6.8718 | 6.5 | 7.568 | 7.06 | 6.8329 | 6.66 | 6.98 | 7.28 | 6.7 | 6.69 | 6.6 | 6.95 | 7.001 | 6.97 | 6.75 | 4.09 | 6.5505 | 6.854 | 6.99 | 6.064 | 6.95 | 7.16 | 7 | 6.8 |  |  | 1
| 1 | 2 | 7.39 | 6.6798 | 7.1568 | 6.9332 | 6.67 | 7.621 | 7.01 | 6.7603 | 6.66 | 6.79 | 7.33 | 6.87 | 6.78 | 7.1 | 6.85 | 7.066 | 6.93 | 6.7 | 3.87 | 6.8302 | 6.832 | 6.81 | 6.395 | 6.88 | 7.42 | 7.08 | 7.6 |  |  | #N/A
| 1 | 3 | 6.93 | 6.9442 | 7.1436 | 6.8804 | 6.73 | 7.674 | 6.98 | 6.8378 | 6.69 | 6.78 | 7.35 | 6.64 | 6.77 | 6.52 | 7 | 6.991 | 6.86 | 6.77 | 3.5 | 6.7841 | 6.819 | 6.96 | 6.296 | 7.01 | 6.91 | 7.14 | 6.7 |  |  | 16
| 1 | 4 | 7.23 | 6.6165 | 7.19 | 6.8885 | 6.64 | 7.727 | 7.12 | 6.8878 | 6.68 | 6.7 | 7.41 | 6.64 | 6.61 | 6.58 | 6.86 | 7.005 | 6.95 | 7.09 | 4.2 | 6.6579 | 6.789 | 6.85 | 6.381 | 6.96 | 7.17 | 7.08 | 7.4 |  |  | 6.8854
| 1 | 5 | 6.93 | 6.834 | 7.1998 | 6.9204 | 6.69 | 7.727 | 7.1 | 6.6708 | 6.67 | 6.58 | 6.87 | 6.49 | 6.61 | 6.8 | 6.98 | 7.024 | 7 | 7.15 | 3.82 | 6.5074 | 6.863 | 6.8 | 6.606 | 6.94 | 7.19 | 7.14 | 7.1 |  |  | 16
| 1 | 6 | 7.32 | 6.8046 | 7.1582 | 6.9826 | 6.84 | 7.621 | 6.82 | 6.7806 | 6.7 | 6.57 | 7.22 | 6.8 | 6.83 | 6.67 | 6.81 | 7.052 | 6.92 | 6.73 | 3.52 | 6.4726 | 6.849 | 7.05 | 6.59 | 7.01 | 7.33 | 7.12 | 7
| Mean |  | 7.1067 | 6.7776 | 7.1681 | 6.9128 | 6.6783 | 7.6563 | 7.015 | 6.795 | 6.6767 | 6.7333 | 7.2433 | 6.69 | 6.715 | 6.7117 | 6.9083 | 7.0232 | 6.9383 | 6.865 | 3.8333 | 6.6338 | 6.8343 | 6.91 | 6.3887 | 6.9583 | 7.1967 | 7.0933 | 7.1
| Median |  | 7.08 | 6.7955 | 7.1593 | 6.9044 | 6.68 | 7.6475 | 7.035 | 6.8068 | 6.675 | 6.74 | 7.305 | 6.67 | 6.73 | 6.635 | 6.905 | 7.0145 | 6.94 | 6.76 | 3.845 | 6.6042 | 6.8405 | 6.905 | 6.388 | 6.955 | 7.18 | 7.1 | 7.05
| Std Dev. |  | 0.23432 | 0.11598 | 0.021775 | 0.041622 | 0.11161 | 0.064186 | 0.10913 | 0.075785 | 0.01633 | 0.15332 | 0.19387 | 0.13387 | 0.092898 | 0.21302 | 0.078337 | 0.030076 | 0.047924 | 0.19977 | 0.28675 | 0.14876 | 0.027274 | 0.10412 | 0.20093 | 0.048751 | 0.17432 | 0.053166 | 0.34641
| Rel.Std.Dev. |  | 3.297% | 1.711% | 0.3038% | 0.6021% | 1.671% | 0.8383% | 1.556% | 1.115% | 0.2446% | 2.277% | 2.677% | 2.001% | 1.383% | 3.174% | 1.134% | 0.4282% | 0.6907% | 2.91% | 7.48% | 2.242% | 0.3991% | 1.507% | 3.145% | 0.7006% | 2.422% | 0.7495% | 4.879%
| PDM3 |  | 3.213% | -1.566% | 4.106% | 0.3976% | -3.008% | 11.2% | 1.882% | -1.313% | -3.032% | -2.209% | 5.198% | -2.838% | -2.475% | -2.523% | 0.3328% | 2.001% | 0.7685% | -0.2965% | -44.33% | -3.655% | -0.7419% | 0.357% | -7.215% | 1.059% | 4.52% | 3.02% | 3.116%
| Z-Score (Absolute) |  | 0.7 | 0.46 | 0.91 | 0.02 | 0.81 | 2.63 | 0.38 | 0.4 | 0.82 | 0.62 | 1.18 | 0.77 | 0.68 | 0.69 | 0 | 0.4 | 0.11 | 0.15 | 10.82 | 0.97 | 0.26 | 0.01 | 1.83 | 0.18 | 1.01 | 0.65 | 0.67
| Analytical results for As in OREAS 231b (Certified Value 34 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 38 | 33.6 | 38.7653 | 11 | 37 | 35 | 32 | 28.9 | 35 | 31.7 | 33 | 33.4 | 34 | 33 | 32 | 34 | 32.199 | 33 | 35.6 | 39.5764 | 33.8 | 32 | 36.1 | 32.7 | [5 empty] | 1
| 1 | 2 | 36 | 34.4 | 38.751 | 30.6 | 36 | 34.2 | 32 | 29.4 | 35.4 | 33 | 35.3 | 36.8 | 33 | 34 | 31 | 33 | 26.34 | 32 | 35.6 | 36.7007 | 35.6 | 34 | 33.9 | 31.4 | [5 empty] | 20
| 1 | 3 | 39 | 35.9 | 38.845 | 31.6 | 36 | 34.3 | 32 | 27.7 | 36.2 | 30.9 | 35.7 | 35.3 | 34 | 36 | 32 | 33 | 35.01 | 33 | 35.7 | 36.5462 | 33.9 | 32 | 34.4 | 31.9 | [5 empty] | 16
| 1 | 4 | 39 | 34.3 | 38.4492 | 23 | 35 | 34.4 | 31 | 32.2 | 35.6 | 31.1 | 35.1 | 35 | 33 | 34 | 32 | 35 | 28.55 | 30 | 35.3 | 36.8178 | 35.8 | 33 | 33.6 | 31 | [5 empty] | 33.976
| 1 | 5 | 40 | 34.7 | 38.756 | 21 | 36 | 33.7 | 31 | 27.4 | 34.6 | 30.1 | 33.5 | 35.7 | 33 | 33 | 31 | 35 | 32.42 | 30 | 35.4 | 36.5462 | 36.3 | 34 | 35.2 | 32.6 | [5 empty] | 17
| 1 | 6 | 40 | 36 | 38.9049 | 31.3 | 34 | 35.2 | 32 | 29.2 | 35.6 | 32.2 | 36.5 | 34.6 | 33 | 34 | 31 | 34 | 29.39 | 31 | 36.3 | 34.0214 | 35.6 | 31 | 34.3 | 30.3
| Mean |  | 38.6667 | 34.8167 | 38.7452 | 24.75 | 35.6667 | 34.4667 | 31.6667 | 29.1333 | 35.4 | 31.5 | 34.85 | 35.1333 | 33.3333 | 34 | 31.5 | 34 | 30.6515 | 31.5 | 35.65 | 36.7015 | 35.1667 | 32.6667 | 34.5833 | 31.65
| Median |  | 39 | 34.55 | 38.7607 | 26.8 | 36 | 34.35 | 32 | 29.05 | 35.5 | 31.4 | 35.2 | 35.15 | 33 | 34 | 31.5 | 34 | 30.7945 | 31.5 | 35.6 | 36.6234 | 35.6 | 32.5 | 34.35 | 31.65
| Std Dev. |  | 1.5055 | 0.94956 | 0.15726 | 8.1269 | 1.0328 | 0.55015 | 0.5164 | 1.7084 | 0.55136 | 1.0257 | 1.3383 | 1.1343 | 0.5164 | 1.0954 | 0.54772 | 0.89443 | 3.1342 | 1.3784 | 0.35071 | 1.7612 | 1.052 | 1.2111 | 0.9196 | 0.93541
| Rel.Std.Dev. |  | 3.894% | 2.727% | 0.4059% | 32.84% | 2.896% | 1.596% | 1.631% | 5.864% | 1.558% | 3.256% | 3.84% | 3.229% | 1.549% | 3.222% | 1.739% | 2.631% | 10.23% | 4.376% | 0.9838% | 4.799% | 2.991% | 3.707% | 2.659% | 2.955%
| PDM3 |  | 13.81% | 2.474% | 14.04% | -27.15% | 4.976% | 1.444% | -6.797% | -14.25% | 4.191% | -7.288% | 2.572% | 3.406% | -1.892% | 0.07061% | -7.288% | 0.07061% | -9.785% | -7.288% | 4.927% | 8.022% | 3.504% | -3.854% | 1.788% | -6.846%
| Z-Score (Absolute) |  | 1.99 | 0.26 | 2.03 | 4.26 | 0.64 | 0.1 | 1.15 | 2.29 | 0.52 | 1.23 | 0.28 | 0.4 | 0.4 | 0.1 | 1.23 | 0.1 | 1.61 | 1.23 | 0.64 | 1.11 | 0.42 | 0.7 | 0.16 | 1.16
| Analytical results for B in OREAS 231b (Indicative Value 32.6 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.03 | 1.25 | 1.27 | [26 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 1
1
| 1 | 1 | 58.51 | 13407.54 | 5 | [26 empty] | 1
| 1 | 2 | 56.08 | 11198.89 | 10 | [26 empty] | 3
| 1 | 3 | 57.47 | 17296.76 | 5 | [26 empty] | 16
| 1 | 4 | 58.73 | 12028.18 | 5 | [26 empty] | 32.6325
| 1 | 5 | 61.44 | 15611.79 | 10 | [26 empty] | 9
| 1 | 6 | 54.36 | 10953.92 | 10
| Mean |  | 57.765 | 13416.18 | 7.5
| Median |  | 57.99 | 12717.86 | 7.5
| Std Dev. |  | 2.4295 | 2560.6766 | 2.7386
| Rel.Std.Dev. |  | 4.206% | 19.09% | 36.51%
| PDM3 |  | 77.02% | 41013% | -77.02%
| Z-Score (Absolute) |  | 0 | 179.2 | 0.67
| Analytical results for Ba in OREAS 231b (Certified Value 299 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 302 | 296.2 | 315.798 | 294.8225 | 265 | 339 | 315 | 299.9 | 300 | 304 | 330 | 310 | 290 | 300 | 241 | 280 | 276 | 276 | 585.07 | 281.187 | 295 | 300.4772 | 310 | 331.3984 | 304 | 332 | 288 | 291 |  | 1
| 1 | 2 | 299 | 300.1 | 310.6583 | 298.2595 | 278 | 337 | 302 | 296.9 | 302 | 294 | 330 | 320 | 290 | 320 | 234 | 304 | 280 | 272 | 1064.78 | 289.26 | 300 | 308.5858 | 300 | 331.0372 | 302 | 337 | 288 | 276 |  | 21
| 1 | 3 | 308 | 311.4 | 314.7635 | 295.982 | 295 | 339 | 303 | 305.4 | 296 | 296 | 330 | 300 | 300 | 290 | 239 | 302 | 279 | 274 | 876.51 | 289.917 | 295 | 301.4451 | 310 | 339.1931 | 307 | 321 | 290 | 280 |  | 16
| 1 | 4 | 300 | 301.3 | 317.0237 | 298.045 | 285 | 345 | 310 | 304.8 | 292 | 298 | 330 | 300 | 290 | 300 | 236 | 309 | 276 | 287 | 1016.98 | 284.247 | 295 | 301.3553 | 310 | 340.6016 | 304 | 324 | 288 | 269 |  | 299.2542
| 1 | 5 | 307 | 307.2 | 314.1294 | 299.1179 | 271 | 344 | 308 | 295.7 | 299 | 300 | 310 | 300 | 290 | 310 | 236 | 304 | 282 | 289 | 1125.02 | 277.56 | 300 | 297.9995 | 310 | 342.504 | 304 | 323 | 292 | 289 |  | 18
| 1 | 6 | 310 | 312.5 | 319.3 | 299.6802 | 289 | 341 | 297 | 307.3 | 301 | 301 | 320 | 310 | 290 | 300 | 235 | 284 | 276 | 273 | 1124.13 | 274.815 | 300 | 306.483 | 310 | 330.8809 | 307 | 331 | 291 | 280
| Mean |  | 304.3333 | 304.7833 | 315.2788 | 297.6512 | 280.5 | 340.8333 | 305.8333 | 301.6667 | 298.3333 | 298.8333 | 325 | 306.6667 | 291.6667 | 303.3333 | 236.8333 | 297.1667 | 278.1667 | 278.5 | 965.415 | 282.831 | 297.5 | 302.7243 | 308.3333 | 335.9359 | 304.6667 | 328 | 289.5 | 280.8333
| Median |  | 304.5 | 304.25 | 315.2808 | 298.1522 | 281.5 | 340 | 305.5 | 302.35 | 299.5 | 299 | 330 | 305 | 290 | 300 | 236 | 303 | 277.5 | 275 | 1040.88 | 282.717 | 297.5 | 301.4002 | 310 | 335.2957 | 304 | 327.5 | 289 | 280
| Std Dev. |  | 4.5898 | 6.5865 | 2.9127 | 1.8752 | 11.3093 | 3.1252 | 6.4317 | 4.836 | 3.7238 | 3.6009 | 8.3666 | 8.165 | 4.0825 | 10.328 | 2.6394 | 12.0402 | 2.5626 | 7.5033 | 207.732 | 6.1366 | 2.7386 | 3.9842 | 4.0825 | 5.3974 | 1.9664 | 6.261 | 1.7607 | 8.1833
| Rel.Std.Dev. |  | 1.508% | 2.161% | 0.9239% | 0.63% | 4.032% | 0.9169% | 2.103% | 1.603% | 1.248% | 1.205% | 2.574% | 2.662% | 1.4% | 3.405% | 1.114% | 4.052% | 0.9212% | 2.694% | 21.52% | 2.17% | 0.9205% | 1.316% | 1.324% | 1.607% | 0.6454% | 1.909% | 0.6082% | 2.914%
| PDM3 |  | 1.697% | 1.848% | 5.355% | -0.5357% | -6.267% | 13.89% | 2.198% | 0.8061% | -0.3077% | -0.1407% | 8.603% | 2.477% | -2.535% | 1.363% | -20.86% | -0.6976% | -7.047% | -6.935% | 222.6% | -5.488% | -0.5862% | 1.16% | 3.034% | 12.26% | 1.809% | 9.606% | -3.26% | -6.156%
| Z-Score (Absolute) |  | 0.17 | 0.21 | 1.06 | 0.37 | 1.76 | 3.13 | 0.29 | 0.04 | 0.31 | 0.27 | 1.85 | 0.36 | 0.85 | 0.09 | 5.29 | 0.41 | 1.94 | 1.92 | 53.67 | 1.57 | 0.38 | 0.04 | 0.5 | 2.73 | 0.2 | 2.09 | 1.03 | 1.73
| Analytical results for Be in OREAS 231b (Certified Value 0.37 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.4 | 0.41 | 0.21 | 0.36076 | < 0.5 | 0.4 | 0.4 | 0.5 | 0.41 | < 5 | < 1 | 0.4 | 0.4 | 0.39 | 0.34 | < 0.5 | < 0.5 | < 1 | 0.36 | < 0.5 | 0.3375 | 0.35 | 0.26558 | < 0.05 | 0.4 | 0.47 | 0.45 |  |  | 1
| 1 | 2 | 0.4 | 0.41 | 0.21 | 0.32959 | < 0.5 | 0.4 | 0.4 | < 0.5 | 0.4 | < 5 | < 1 | 0.39 | 0.41 | 0.35 | 0.38 | < 0.5 | < 0.5 | < 1 | 0.3 | < 0.5 | 0.31946 | 0.34 | 0.27578 | < 0.05 | 0.5 | 0.41 | 0.36 |  |  | 22
| 1 | 3 | 0.4 | 0.41 | 0.2 | 0.329 | < 0.5 | 0.4 | 0.4 | 0.6 | 0.4 | < 5 | < 1 | 0.4 | 0.39 | 0.38 | 0.35 | < 0.5 | < 0.5 | < 1 | 0.38 | < 0.5 | 0.30842 | 0.35 | 0.27574 | < 0.05 | 0.5 | 0.41 | 0.45 |  |  | 16
| 1 | 4 | 0.4 | 0.41 | 0.2 | 0.35213 | < 0.5 | 0.4 | 0.4 | < 0.5 | 0.41 | < 5 | < 1 | 0.42 | 0.4 | 0.39 | 0.35 | < 0.5 | < 0.5 | < 1 | 0.33 | < 0.5 | 0.30985 | 0.34 | 0.28372 | < 0.05 | 0.4 | 0.39 | 0.32 |  |  | 0.36832
| 1 | 5 | 0.4 | 0.42 | 0.2 | 0.33738 | < 0.5 | 0.4 | 0.4 | < 0.5 | 0.41 | < 5 | < 1 | 0.38 | 0.39 | 0.38 | 0.36 | < 0.5 | < 0.5 | < 1 | 0.36 | < 0.5 | 0.3163 | 0.35 | 0.28854 | < 0.05 | 0.5 | 0.42 | 0.46 |  |  | 19
| 1 | 6 | 0.3 | 0.41 | 0.2 | 0.35173 | < 0.5 | 0.4 | 0.4 | < 0.5 | 0.45 | < 5 | < 1 | 0.4 | 0.41 | 0.37 | 0.35 | < 0.5 | < 0.5 | < 1 | 0.35 | < 0.5 | 0.3097 | 0.35 | 0.28639 | < 0.05 | 0.5 | 0.43 | 0.33
| Mean |  | 0.38333 | 0.41167 | 0.20333 | 0.34343 |  | 0.4 | 0.4 | 0.55 | 0.41333 |  |  | 0.39833 | 0.4 | 0.37667 | 0.355 |  |  |  | 0.34667 |  | 0.31687 | 0.34667 | 0.27929 |  | 0.46667 | 0.42167 | 0.395
| Median |  | 0.4 | 0.41 | 0.2 | 0.34456 |  | 0.4 | 0.4 | 0.55 | 0.41 |  |  | 0.4 | 0.4 | 0.38 | 0.35 |  |  |  | 0.355 |  | 0.31307 | 0.35 | 0.27975 |  | 0.5 | 0.415 | 0.405
| Std Dev. |  | 0.040825 | 0.0040825 | 0.005164 | 0.013275 |  | 6.0809e-17 | 6.0809e-17 | 0.070711 | 0.018619 |  |  | 0.013292 | 0.0089443 | 0.015055 | 0.013784 |  |  |  | 0.028048 |  | 0.010997 | 0.005164 | 0.0085853 |  | 0.05164 | 0.027142 | 0.065345
| Rel.Std.Dev. |  | 10.65% | 0.9917% | 2.54% | 3.865% |  | 1.52e-14% | 1.52e-14% | 12.86% | 4.505% |  |  | 3.337% | 2.236% | 3.997% | 3.883% |  |  |  | 8.091% |  | 3.471% | 1.49% | 3.074% |  | 11.07% | 6.437% | 16.54%
| PDM3 |  | 4.077% | 11.77% | -44.79% | -6.757% |  | 8.602% | 8.602% | 49.33% | 12.22% |  |  | 8.149% | 8.602% | 2.267% | -3.616% |  |  |  | -5.878% |  | -13.97% | -5.878% | -24.17% |  | 26.7% | 14.48% | 7.244%
| Z-Score (Absolute) |  | NA | 0.62 | 1.49 | 0.07 | 1.02 | NA | NA | NA | 0.63 | 21.74 | 1.51 | 0.48 | 0.5 | 0.26 | 0.04 | 1.02 | 1.02 | 1.51 | 0.04 | 1.02 | 0.34 | 0.04 | 0.72 | 3.3 | NA | 0.72 | 0.45
Results from laboratories 1.01, 1.07, 1.08, 1.10 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 231b (Certified Value 0.029 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | < 0.04 | 0.03 | 0.080296 | 0.03 | 0.7 | 0.05 | 0.03 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.02 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | < 0.1 | < 0.11 | 0.03 | 0.0108 | 0.1 | < 0.04 | 0.04 | 0.02 |  |  |  |  | 1
| 1 | 2 | 0.06 | 0.02 | 0.068582 | 0.02 | 0.2 | 0.05 | 0.03 | < 0.1 | 0.02 | 0.03 | 0.03 | 0.03 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | < 0.1 | < 0.11 | 0.03 | < 0.01 | 0.03 | < 0.04 | 0.04 | < 0.01 |  |  |  |  | 23
| 1 | 3 | < 0.04 | 0.03 | 0.052625 | 0.03 | 0.2 | 0.05 | 0.03 | < 0.1 | 0.02 | 0.03 | 0.03 | 0.03 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | 0.5 | < 0.11 | 0.03 | < 0.01 | 0.03 | < 0.04 | 0.04 | < 0.01 |  |  |  |  | 16
| 1 | 4 | 0.07 | 0.03 | 0.056336 | 0.03 | 0.2 | < 0.05 | 0.03 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.02 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | < 0.1 | < 0.11 | 0.03 | < 0.01 | 0.03 | < 0.04 | 0.03 | 0.01 |  |  |  |  | 0.028909
| 1 | 5 | < 0.04 | 0.03 | 0.060867 | 0.02 | 0.2 | < 0.05 | 0.03 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.02 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | < 0.1 | < 0.11 | 0.02 | < 0.01 | 0.02 | < 0.04 | 0.03 | 0.01 |  |  |  |  | 20
| 1 | 6 | 0.04 | 0.03 | 0.058508 | 0.03 | 0.2 | 0.05 | 0.03 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.02 | < 5 | < 0.05 | < 0.1 | 0.03 | < 0.1 | < 0.1 | < 0.11 | 0.03 | 0.169 | 0.03 | < 0.04 | 0.04 | 0.02
| Mean |  | 0.056667 | 0.028333 | 0.062869 | 0.026667 | 0.28333 | 0.05 | 0.03 |  | 0.026667 | 0.03 | 0.03 | 0.023333 |  |  |  | 0.03 |  | 0.5 |  | 0.028333 | 0.0899 | 0.04 |  | 0.036667 | 0.015
| Median |  | 0.06 | 0.03 | 0.059687 | 0.03 | 0.2 | 0.05 | 0.03 |  | 0.03 | 0.03 | 0.03 | 0.02 |  |  |  | 0.03 |  | 0.5 |  | 0.03 | 0.0899 | 0.03 |  | 0.04 | 0.015
| Std Dev. |  | 0.015275 | 0.0040825 | 0.01007 | 0.005164 | 0.20412 | 0 | 0 |  | 0.005164 | 0 | 0 | 0.005164 |  |  |  | 0 |  |  |  | 0.0040825 | 0.11186 | 0.029665 |  | 0.005164 | 0.0057735
| Rel.Std.Dev. |  | 26.96% | 14.41% | 16.02% | 19.36% | 72.04% | 0% | 0% |  | 19.36% | 0% | 0% | 22.13% |  |  |  | 0% |  |  |  | 14.41% | 124.4% | 74.16% |  | 14.08% | 38.49%
| PDM3 |  | 96.02% | -1.992% | 117.5% | -7.757% | 880.1% | 72.96% | 3.774% |  | -7.757% | 3.774% | 3.774% | -19.29% |  |  |  | 3.774% |  | 1630% |  | -1.992% | 211% | 38.36% |  | 26.83% | -48.11%
| Z-Score (Absolute) |  | 0.6 | 0.3 | 2.81 | 0.45 | NA | 0.9 | 0.15 | 1.65 | 0.45 | 0.15 | 0.15 | 0.75 | 221.92 | 0.6 | 1.65 | 0.15 | 1.65 | 8.39 | 2.1 | 0.3 | 0.15 | 0.75 | 1.05 | 0.45 | 1.8
Results from laboratory 1.08 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 231b (Certified Value 6.43 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.24 | 6.8031 | 6.5537 | 6.4517 | 5.86 | 7.218 | 6.5 | 6.9605 | 6.3 | 6.35 | 6.83 | 6.3 | 6.36 | 6.27 | 6.26 | 6.659 | 6.29 | 6.413 | 4.6 | 6.0295 | 6.512 | 6.67 | 6.4877 | 6.31 | 5.92 | 6.14 | 6.3 |  |  | 1
| 1 | 2 | 6.17 | 6.7283 | 6.5938 | 6.4906 | 6.07 | 7.218 | 6.49 | 6.7393 | 6.35 | 6.27 | 6.77 | 6.53 | 6.43 | 6.72 | 6.19 | 6.726 | 6.22 | 6.34 | 4.44 | 6.1074 | 6.521 | 6.5 | 6.6427 | 6.24 | 6.04 | 6.18 | 7.2 |  |  | #N/A
| 1 | 3 | 6.35 | 6.9655 | 6.597 | 6.4602 | 6.38 | 7.29 | 6.43 | 6.8342 | 6.28 | 6.34 | 6.84 | 6.34 | 6.39 | 6.28 | 6.32 | 6.621 | 6.18 | 6.454 | 3.99 | 6.1029 | 6.595 | 6.69 | 6.7177 | 6.38 | 5.69 | 6.23 | 6.3 |  |  | 16
| 1 | 4 | 6.22 | 6.7001 | 6.5559 | 6.438 | 6.4 | 7.29 | 6.47 | 6.8328 | 6.27 | 6.35 | 6.93 | 6.28 | 6.31 | 6.27 | 6.25 | 6.647 | 6.26 | 6.73 | 4.69 | 6.1699 | 6.524 | 6.64 | 6.8637 | 6.31 | 5.98 | 6.17 | 6.9 |  |  | 6.4334
| 1 | 5 | 6.26 | 6.7637 | 6.5832 | 6.4837 | 5.95 | 7.361 | 6.54 | 6.5737 | 6.25 | 6.36 | 6.47 | 6.2 | 6.3 | 6.5 | 6.22 | 6.689 | 6.26 | 6.791 | 4.51 | 6.039 | 6.578 | 6.72 | 6.9687 | 6.28 | 5.77 | 6.24 | 6.7 |  |  | 21
| 1 | 6 | 6.31 | 6.8317 | 6.5538 | 6.5167 | 6.24 | 7.29 | 6.35 | 6.7904 | 6.31 | 6.32 | 6.81 | 6.42 | 6.52 | 6.36 | 6.15 | 6.709 | 6.23 | 6.387 | 4.46 | 5.9833 | 6.571 | 6.79 | 6.9007 | 6.33 | 5.96 | 6.22 | 6.7
| Mean |  | 6.2583 | 6.7987 | 6.5729 | 6.4735 | 6.15 | 7.2778 | 6.4633 | 6.7885 | 6.2933 | 6.3317 | 6.775 | 6.345 | 6.385 | 6.4 | 6.2317 | 6.6752 | 6.24 | 6.5192 | 4.4483 | 6.072 | 6.5502 | 6.6683 | 6.7635 | 6.3083 | 5.8933 | 6.1967 | 6.6833
| Median |  | 6.25 | 6.7834 | 6.5695 | 6.472 | 6.155 | 7.29 | 6.48 | 6.8116 | 6.29 | 6.345 | 6.82 | 6.32 | 6.375 | 6.32 | 6.235 | 6.674 | 6.245 | 6.4335 | 4.485 | 6.0709 | 6.5475 | 6.68 | 6.7907 | 6.31 | 5.94 | 6.2 | 6.7
| Std Dev. |  | 0.064317 | 0.094693 | 0.020705 | 0.028908 | 0.22539 | 0.05389 | 0.066232 | 0.12822 | 0.035024 | 0.033116 | 0.15846 | 0.11589 | 0.082158 | 0.18011 | 0.059133 | 0.039771 | 0.038471 | 0.19154 | 0.2431 | 0.067164 | 0.035244 | 0.09704 | 0.18095 | 0.047081 | 0.13471 | 0.039328 | 0.34881
| Rel.Std.Dev. |  | 1.028% | 1.393% | 0.315% | 0.4466% | 3.665% | 0.7405% | 1.025% | 1.889% | 0.5565% | 0.523% | 2.339% | 1.826% | 1.287% | 2.814% | 0.9489% | 0.5958% | 0.6165% | 2.938% | 5.465% | 1.106% | 0.5381% | 1.455% | 2.675% | 0.7463% | 2.286% | 0.6347% | 5.219%
| PDM3 |  | -2.721% | 5.679% | 2.169% | 0.6236% | -4.404% | 13.13% | 0.466% | 5.52% | -2.177% | -1.581% | 5.31% | -1.373% | -0.7517% | -0.5185% | -3.135% | 3.759% | -3.006% | 1.334% | -30.86% | -5.617% | 1.816% | 3.652% | 5.132% | -1.943% | -8.394% | -3.679% | 3.886%
| Z-Score (Absolute) |  | 0.55 | 1.56 | 0.67 | 0.29 | 0.98 | 3.42 | 0.25 | 1.52 | 0.42 | 0.27 | 1.46 | 0.21 | 0.06 | 0 | 0.66 | 1.07 | 0.62 | 0.46 | 7.61 | 1.28 | 0.59 | 1.05 | 1.42 | 0.36 | 1.98 | 0.79 | 1.11
| Analytical results for Cd in OREAS 231b (Certified Value 0.39 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.45 | 0.38 | 0.23 | 0.43589 | 0.3 | 0.5 | 0.43 | 0.36 | 0.7 | 0.4 | 0.39 | 0.38 | 0.34 | 0.36 | < 1 | 0.33 | 0.3 | 0.37 | < 0.1 | < 0.5 | 0.34009 | 0.39 | 0.4116 | 0.38 | 0.42 | 0.45 | 0.37 |  |  | 1
| 1 | 2 | 0.44 | 0.38 | 0.25 | 0.43259 | 0.4 | 0.4 | 0.47 | 0.36 | 0.8 | 0.4 | 0.4 | 0.4 | 0.37 | 0.38 | < 1 | 0.37 | 0.3 | 0.34 | < 0.1 | < 0.5 | 0.36763 | 0.39 | 0.4094 | 0.36 | 0.43 | 0.45 | 0.36 |  |  | 24
| 1 | 3 | 0.45 | 0.43 | 0.23 | 0.40578 | < 0.3 | 0.4 | 0.44 | 0.43 | 0.8 | 0.5 | 0.38 | 0.37 | 0.4 | 0.4 | < 1 | 0.38 | 0.4 | 0.34 | < 0.1 | < 0.5 | 0.34392 | 0.37 | 0.4174 | 0.36 | 0.39 | 0.41 | 0.35 |  |  | 16
| 1 | 4 | 0.46 | 0.35 | 0.24 | 0.4412 | < 0.3 | 0.5 | 0.5 | 0.4 | 0.8 | 0.4 | 0.39 | 0.37 | 0.35 | 0.41 | < 1 | 0.4 | 0.3 | 0.36 | < 0.1 | < 0.5 | 0.34799 | 0.36 | 0.4139 | 0.38 | 0.4 | 0.39 | 0.34 |  |  | 0.39122
| 1 | 5 | 0.45 | 0.42 | 0.23 | 0.42293 | 0.4 | 0.4 | 0.35 | 0.39 | 0.8 | 0.4 | 0.38 | 0.35 | 0.35 | 0.38 | < 1 | 0.41 | 0.3 | 0.34 | < 0.1 | < 0.5 | 0.37457 | 0.38 | 0.4505 | 0.33 | 0.39 | 0.45 | 0.37 |  |  | 22
| 1 | 6 | 0.44 | 0.34 | 0.23 | 0.40726 | 0.3 | 0.4 | 0.47 | 0.41 | 0.8 | 0.5 | 0.4 | 0.37 | 0.38 | 0.35 | < 1 | 0.35 | 0.3 | 0.38 | < 0.1 | < 0.5 | 0.35874 | 0.36 | 0.4049 | 0.35 | 0.44 | 0.46 | 0.35
| Mean |  | 0.44833 | 0.38333 | 0.235 | 0.42428 | 0.35 | 0.43333 | 0.44333 | 0.39167 | 0.78333 | 0.43333 | 0.39 | 0.37333 | 0.365 | 0.38 |  | 0.37333 | 0.31667 | 0.355 |  |  | 0.35549 | 0.375 | 0.41795 | 0.36 | 0.41167 | 0.435 | 0.35667
| Median |  | 0.45 | 0.38 | 0.23 | 0.42776 | 0.35 | 0.4 | 0.455 | 0.395 | 0.8 | 0.4 | 0.39 | 0.37 | 0.36 | 0.38 |  | 0.375 | 0.3 | 0.35 |  |  | 0.35337 | 0.375 | 0.41275 | 0.36 | 0.41 | 0.45 | 0.355
| Std Dev. |  | 0.0075277 | 0.036148 | 0.0083666 | 0.014991 | 0.057735 | 0.05164 | 0.052026 | 0.027869 | 0.040825 | 0.05164 | 0.0089443 | 0.01633 | 0.022583 | 0.022804 |  | 0.030111 | 0.040825 | 0.017607 |  |  | 0.013779 | 0.013784 | 0.016492 | 0.018974 | 0.02137 | 0.028107 | 0.012111
| Rel.Std.Dev. |  | 1.679% | 9.43% | 3.56% | 3.533% | 16.5% | 11.92% | 11.74% | 7.115% | 5.212% | 11.92% | 2.293% | 4.374% | 6.187% | 6.001% |  | 8.065% | 12.89% | 4.96% |  |  | 3.876% | 3.676% | 3.946% | 5.27% | 5.191% | 6.461% | 3.395%
| PDM3 |  | 14.6% | -2.015% | -39.93% | 8.451% | -10.54% | 10.77% | 13.32% | 0.1154% | 100.2% | 10.77% | -0.3106% | -4.571% | -6.701% | -2.867% |  | -4.571% | -19.06% | -9.257% |  |  | -9.132% | -4.145% | 6.834% | -7.979% | 5.228% | 11.19% | -8.831%
| Z-Score (Absolute) |  | 2.15 | 0.18 | 4.32 | 1.42 | NA | NA | 1.99 | 0.43 | NA | NA | 0.38 | 0.13 | 0.38 | 0.08 | 3.71 | 0.13 | NA | 0.68 | 9.92 | 3.86 | 0.67 | 0.08 | 1.23 | 0.53 | 1.04 | 1.74 | 0.63
Results from laboratories 1.07, 1.08, 1.12, 1.13 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 231b (Certified Value 11.3 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 10.6 | 11.35 | 10.27 | 10.8266 | 11.339 | 10.4 | 14 | 11.46 | 12 | 12.1 | 12.15 | 10.25 | 10.5 | < 10 | 11 | 11.17 | 12 | 11.5498 | 12.2 | 6.052 | 13.1 | 11.43 | 13.29 | 11.4 | [5 empty] | 1
| 1 | 2 | 10.75 | 11.25 | 10.84 | 10.6663 | 11.4222 | 10.7 | 14 | 11.02 | 11 | 11.5 | 12.8 | 11 | 11.45 | < 10 | 11 | 11.15 | 13 | 11.9121 | 11.5 | 6.2803 | 13 | 11.83 | 13.19 | 10.7 | [5 empty] | 25
| 1 | 3 | 10.67 | 11.61 | 11.07 | 10.573 | 11.6224 | 11.3 | 14 | 11.3 | 11 | 11.9 | 12.15 | 11.6 | 11.15 | < 10 | 10 | 11.02 | 14.5 | 11.7338 | 11.9 | 6.4383 | 13 | 11.13 | 13.34 | 10.9 | [5 empty] | 16
| 1 | 4 | 10.78 | 11.22 | 10.54 | 10.7782 | 11.5614 | 11.1 | 14 | 11.05 | 12 | 11.7 | 12 | 11.15 | 11 | < 10 | 11 | 10.86 | 12 | 11.5561 | 11.5 | 6.5749 | 13.2 | 11.41 | 12.61 | 10.5 | [5 empty] | 11.3417
| 1 | 5 | 11.57 | 11.68 | 10.58 | 10.727 | 11.5768 | 10.6 | 14 | 11.37 | 11 | 11.45 | 11.6 | 10.6 | 11.35 | < 10 | 11 | 10.97 | 12.5 | 11.479 | 11.45 | 6.5429 | 13.7 | 11.34 | 13.2 | 11.3 | [5 empty] | 23
| 1 | 6 | 11.95 | 12.02 | 10.55 | 10.8496 | 11.7296 | 11.3 | 14 | 11.49 | 12 | 11.95 | 12.45 | 11.25 | 11.15 | < 10 | 11 | 11.14 | 12.5 | 11.5983 | 12.2 | 6.1901 | 13.6 | 11.64 | 12.97 | 10.8
| Mean |  | 11.0533 | 11.5217 | 10.6417 | 10.7368 | 11.5419 | 10.9 | 14 | 11.2817 | 11.5 | 11.7667 | 12.1917 | 10.975 | 11.1 |  | 10.8333 | 11.0517 | 12.75 | 11.6382 | 11.7917 | 6.3464 | 13.2667 | 11.4633 | 13.1 | 10.9333
| Median |  | 10.765 | 11.48 | 10.565 | 10.7526 | 11.5691 | 10.9 | 14 | 11.335 | 11.5 | 11.8 | 12.15 | 11.075 | 11.15 |  | 11 | 11.08 | 12.5 | 11.5772 | 11.7 | 6.3593 | 13.15 | 11.42 | 13.195 | 10.85
| Std Dev. |  | 0.56394 | 0.30838 | 0.27694 | 0.1043 | 0.14055 | 0.38471 | 0 | 0.20272 | 0.54772 | 0.26013 | 0.40671 | 0.48244 | 0.33466 |  | 0.40825 | 0.12319 | 0.93541 | 0.15853 | 0.35555 | 0.20724 | 0.30768 | 0.24328 | 0.27159 | 0.35024
| Rel.Std.Dev. |  | 5.102% | 2.677% | 2.602% | 0.9715% | 1.218% | 3.529% | 0% | 1.797% | 4.763% | 2.211% | 3.336% | 4.396% | 3.015% |  | 3.768% | 1.115% | 7.337% | 1.362% | 3.015% | 3.265% | 2.319% | 2.122% | 2.073% | 3.203%
| PDM3 |  | -2.543% | 1.587% | -6.172% | -5.334% | 1.765% | -3.895% | 23.44% | -0.5295% | 1.396% | 3.747% | 7.494% | -3.233% | -2.131% |  | -4.482% | -2.557% | 12.42% | 2.614% | 3.967% | -44.04% | 16.97% | 1.072% | 15.5% | -3.601%
| Z-Score (Absolute) |  | 0.4 | 0.42 | 1.13 | 0.96 | 0.46 | 0.67 | NA | 0 | NA | 0.86 | 1.61 | 0.54 | 0.32 | 11.1 | NA | 0.41 | 2.59 | 0.63 | 0.9 | 8.72 | 3.51 | 0.32 | 3.21 | 0.62
Results from laboratories 1.08, 1.13 and 1.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 231b (Certified Value 46.6 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 49.8 | 46.9 | 45.8 | 39.3911 | 45.8544 | 44 | 54.7 | 48 | 48.4 | 35 | 46.6 | 48.1 | 45 | 46.1 | 44.3 | 45 | 44.6 | 45 | 48.2 | 40.65 | 40 | 46.4895 | 47.5 | 45.6235 | 53.7 | 48 | 55.4 | 46.4 |  | 1
| 1 | 2 | 48.7 | 46.8 | 45.48 | 39.0402 | 45.8404 | 46 | 54 | 46 | 46.8 | 36 | 45.4 | 47.9 | 46.8 | 47.8 | 46.8 | 45 | 46.2 | 43.7 | 44.2 | 39.78 | 40 | 49.2968 | 45.4 | 46.1308 | 53.4 | 49.3 | 54.7 | 43.5 |  | 26
| 1 | 3 | 51.2 | 48.8 | 45.32 | 39.8078 | 45.9949 | 45 | 53.1 | 47 | 47.9 | 36 | 46 | 48.4 | 44.7 | 48.9 | 46.4 | 46 | 45.2 | 44.8 | 47.3 | 41.64 | 40 | 47.2586 | 47.9 | 47.0622 | 52.4 | 48.6 | 55.4 | 44.6 |  | 16
| 1 | 4 | 48.8 | 47.9 | 45.33 | 39.7655 | 45.879 | 45 | 54.2 | 48 | 48 | 35 | 46.5 | 47.3 | 44.1 | 46.5 | 43.9 | 45 | 44.2 | 43.7 | 49.3 | 40.02 | 40 | 47.2627 | 46.8 | 47.7886 | 53.8 | 47.1 | 52.7 | 42.4 |  | 46.563
| 1 | 5 | 49.1 | 47.7 | 45.45 | 39.4485 | 45.8419 | 45 | 52.3 | 48 | 46.6 | 36 | 45.8 | 47.1 | 42.9 | 44.7 | 46.6 | 45 | 46.9 | 45.2 | 48.6 | 40.18 | 40 | 47.9424 | 47.1 | 47.5494 | 53.9 | 47.3 | 54.9 | 45.7 |  | 24
| 1 | 6 | 50.4 | 48.1 | 45.41 | 39.4703 | 46.2646 | 46 | 51.9 | 46 | 48.1 | 36 | 46.6 | 49.2 | 46.2 | 50.2 | 45.9 | 44 | 45.6 | 45 | 47.7 | 38.57 | 40 | 47.2741 | 48.4 | 52.9104 | 52.7 | 48.2 | 53.8 | 43.7
| Mean |  | 49.6667 | 47.7 | 45.465 | 39.4872 | 45.9459 | 45.1667 | 53.3667 | 47.1667 | 47.6333 | 35.6667 | 46.15 | 48 | 44.95 | 47.3667 | 45.65 | 45 | 45.45 | 44.5667 | 47.55 | 40.14 | 40 | 47.5873 | 47.1833 | 47.8441 | 53.3167 | 48.0833 | 54.4833 | 44.3833
| Median |  | 49.45 | 47.8 | 45.43 | 39.4594 | 45.8667 | 45 | 53.55 | 47.5 | 47.95 | 36 | 46.25 | 48 | 44.85 | 47.15 | 46.15 | 45 | 45.4 | 44.9 | 47.95 | 40.1 | 40 | 47.2684 | 47.3 | 47.3058 | 53.55 | 48.1 | 54.8 | 44.15
| Std Dev. |  | 0.9913 | 0.75631 | 0.17604 | 0.2797 | 0.16658 | 0.75277 | 1.1165 | 0.98319 | 0.74476 | 0.5164 | 0.49699 | 0.7642 | 1.4124 | 2.0017 | 1.2438 | 0.63246 | 1.0035 | 0.68313 | 1.783 | 1.012 | 0 | 0.95551 | 1.042 | 2.6167 | 0.62423 | 0.81833 | 1.0534 | 1.4851
| Rel.Std.Dev. |  | 1.996% | 1.586% | 0.3872% | 0.7083% | 0.3626% | 1.667% | 2.092% | 2.085% | 1.564% | 1.448% | 1.077% | 1.592% | 3.142% | 4.226% | 2.725% | 1.405% | 2.208% | 1.533% | 3.75% | 2.521% | 0% | 2.008% | 2.208% | 5.469% | 1.171% | 1.702% | 1.933% | 3.346%
| PDM3 |  | 6.665% | 2.442% | -2.358% | -15.2% | -1.325% | -2.999% | 14.61% | 1.296% | 2.299% | -23.4% | -0.8871% | 3.086% | -3.464% | 1.726% | -1.961% | -3.357% | -2.39% | -4.287% | 2.12% | -13.79% | -14.09% | 2.2% | 1.332% | 2.751% | 14.5% | 3.265% | 17.01% | -4.681%
| Z-Score (Absolute) |  | 1.47 | 0.51 | 0.58 | 3.5 | 0.35 | 0.73 | 3.27 | 0.25 | 0.48 | 5.36 | 0.25 | 0.65 | 0.83 | 0.35 | 0.49 | 0.81 | 0.59 | 1.02 | 0.43 | 3.18 | 3.25 | 0.45 | 0.26 | 0.58 | 3.25 | 0.69 | 3.81 | 1.11
| Analytical results for Cr in OREAS 231b (Certified Value 154 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 157 | 188 | 144.7796 | 97 | 219 | 156 | 134 | 162 | 168 | 150 | 147 | 148 | 137 | 157 | 111 | 150 | 154 | 156.85 | 168.534 | 165 | 188.7111 | 152 | 155.5836 | 148 | 164 | 169 | 148 |  |  | 1
| 1 | 2 | 161 | 186 | 143.3966 | 107 | 217 | 149 | 136 | 159 | 166 | 147 | 151 | 143 | 153 | 156 | 126 | 149 | 151 | 150.15 | 163.107 | 165 | 193.8634 | 147 | 158.6792 | 153 | 167 | 167 | 147 |  |  | 27
| 1 | 3 | 165 | 183 | 141.1514 | 113 | 221 | 151 | 145 | 162 | 165 | 152 | 147 | 147 | 137 | 158 | 129 | 152 | 148 | 154.53 | 173.844 | 165 | 188.792 | 150 | 159.176 | 151 | 160 | 168 | 141 |  |  | 16
| 1 | 4 | 164 | 178 | 141.4802 | 115 | 223 | 164 | 143 | 163 | 165 | 156 | 140 | 146 | 137 | 156 | 119 | 154 | 165 | 148.6 | 173.547 | 170 | 190.149 | 148 | 163.3757 | 146 | 164 | 161 | 143 |  |  | 153.7205
| 1 | 5 | 120 | 179 | 141.5821 | 102 | 220 | 164 | 134 | 157 | 164 | 143 | 145 | 144 | 141 | 157 | 119 | 152 | 157 | 147.73 | 170.352 | 170 | 192.8602 | 154 | 164.0709 | 144 | 165 | 167 | 146 |  |  | 25
| 1 | 6 | 129 | 190 | 144.0151 | 108 | 225 | 156 | 152 | 163 | 154 | 155 | 151 | 146 | 141 | 159 | 113 | 152 | 169 | 147.39 | 158.949 | 165 | 187.4372 | 152 | 162.274 | 139 | 169 | 165 | 127
| Mean |  | 149.3333 | 184 | 142.7342 | 107 | 220.8333 | 156.6667 | 140.6667 | 161 | 163.6667 | 150.5 | 146.8333 | 145.6667 | 141 | 157.1667 | 119.5 | 151.5 | 157.3333 | 150.875 | 168.0555 | 166.6667 | 190.3021 | 150.5 | 160.5266 | 146.8333 | 164.8333 | 166.1667 | 142
| Median |  | 159 | 184.5 | 142.4894 | 107.5 | 220.5 | 156 | 139.5 | 162 | 165 | 151 | 147 | 146 | 139 | 157 | 119 | 152 | 155.5 | 149.375 | 169.443 | 165 | 189.4705 | 151 | 160.725 | 147 | 164.5 | 167 | 144.5
| Std Dev. |  | 19.6435 | 4.858 | 1.5276 | 6.7231 | 2.8577 | 6.314 | 7.2572 | 2.4495 | 4.9261 | 4.9295 | 4.1191 | 1.8619 | 6.1968 | 1.169 | 7.0356 | 1.7607 | 8.165 | 3.9191 | 5.9428 | 2.582 | 2.5405 | 2.6646 | 3.2681 | 5.0365 | 3.0605 | 2.8577 | 7.7974
| Rel.Std.Dev. |  | 13.15% | 2.64% | 1.07% | 6.283% | 1.294% | 4.03% | 5.159% | 1.521% | 3.01% | 3.275% | 2.805% | 1.278% | 4.395% | 0.7438% | 5.888% | 1.162% | 5.19% | 2.598% | 3.536% | 1.549% | 1.335% | 1.77% | 2.036% | 3.43% | 1.857% | 1.72% | 5.491%
| PDM3 |  | -2.854% | 19.7% | -7.147% | -30.39% | 43.66% | 1.917% | -8.492% | 4.736% | 6.47% | -2.095% | -4.48% | -5.239% | -8.275% | 2.242% | -22.26% | -1.444% | 2.35% | -1.851% | 9.325% | 8.422% | 23.8% | -2.095% | 4.428% | -4.48% | 7.229% | 8.097% | -7.625%
| Z-Score (Absolute) |  | 0.15 | 2.31 | 0.62 | 3.16 | 4.92 | 0.37 | 0.77 | 0.67 | 0.86 | 0.07 | 0.33 | 0.41 | 0.75 | 0.4 | 2.27 | 0 | 0.41 | 0.04 | 1.18 | 1.08 | 2.75 | 0.07 | 0.64 | 0.33 | 0.95 | 1.04 | 0.67
| Analytical results for Cs in OREAS 231b (Certified Value 0.58 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 0.58 | 0.51864 | 0.57 | 0.6 | 0.5 | 0.62 | 0.8 | 0.59 | 0.64 | 0.55 | 0.57 | 0.6 | 0.59 | 0.5 | 0.52498 | 0.62 | 0.6418 | 0.53 | 0.82 | 0.66 | 0.54 | [7 empty] | 1
| 1 | 2 | < 5 | 0.58 | 0.53201 | 0.55 | 0.6 | 0.6 | 0.61 | 0.6 | 0.6 | 0.66 | 0.54 | 0.6 | 0.6 | 0.55 | 0.5 | 0.53024 | 0.6 | 0.6237 | 0.54 | 0.77 | 0.65 | 0.52 | [7 empty] | 28
| 1 | 3 | < 5 | 0.6 | 0.51794 | 0.59 | 0.6 | 0.5 | 0.62 | 0.5 | 0.59 | 0.61 | 0.6 | 0.57 | 0.6 | 0.57 | 1 | 0.50822 | 0.6 | 0.6515 | 0.56 | 0.71 | 0.66 | 0.57 | [7 empty] | 16
| 1 | 4 | < 5 | 0.59 | 0.54576 | 0.59 | 0.6 | 0.6 | 0.61 | 0.7 | 0.58 | 0.63 | 0.58 | 0.59 | 0.6 | 0.57 | 0.5 | 0.51641 | 0.62 | 0.6478 | 0.56 | 0.73 | 0.64 | 0.52 | [7 empty] | 0.58443
| 1 | 5 | < 5 | 0.58 | 0.50771 | 0.57 | 0.6 | 0.6 | 0.58 | 0.7 | 0.56 | 0.59 | 0.57 | 0.56 | 0.6 | 0.58 | 0.5 | 0.53197 | 0.6 | 0.6679 | 0.55 | 0.7 | 0.67 | 0.54 | [7 empty] | 26
| 1 | 6 | < 5 | 0.61 | 0.53393 | 0.6 | 0.6 | 0.6 | 0.62 | 0.7 | 0.58 | 0.64 | 0.56 | 0.59 | 0.6 | 0.59 | 0.5 | 0.52191 | 0.63 | 0.6507 | 0.55 | 0.75 | 0.66 | 0.53
| Mean |  |  | 0.59 | 0.526 | 0.57833 | 0.6 | 0.56667 | 0.61 | 0.66667 | 0.58333 | 0.62833 | 0.56667 | 0.58 | 0.6 | 0.575 | 0.58333 | 0.52229 | 0.61167 | 0.64723 | 0.54833 | 0.74667 | 0.65667 | 0.53667
| Median |  |  | 0.585 | 0.52532 | 0.58 | 0.6 | 0.6 | 0.615 | 0.7 | 0.585 | 0.635 | 0.565 | 0.58 | 0.6 | 0.575 | 0.5 | 0.52344 | 0.61 | 0.64925 | 0.55 | 0.74 | 0.66 | 0.535
| Std Dev. |  |  | 0.012649 | 0.013735 | 0.018348 | 0 | 0.05164 | 0.015492 | 0.10328 | 0.013663 | 0.024833 | 0.021602 | 0.015492 | 0 | 0.015166 | 0.20412 | 0.0089087 | 0.013292 | 0.014429 | 0.01169 | 0.044121 | 0.010328 | 0.018619
| Rel.Std.Dev. |  |  | 2.144% | 2.611% | 3.173% | 0% | 9.113% | 2.54% | 15.49% | 2.342% | 3.952% | 3.812% | 2.671% | 0% | 2.638% | 34.99% | 1.706% | 2.173% | 2.229% | 2.132% | 5.909% | 1.573% | 3.469%
| PDM3 |  |  | 0.9523% | -9.999% | -1.044% | 2.663% | -3.04% | 4.374% | 14.07% | -0.1884% | 7.511% | -3.04% | -0.7588% | 2.663% | -1.614% | -0.1884% | -10.63% | 4.66% | 10.75% | -6.177% | 27.76% | 12.36% | -8.173%
| Z-Score (Absolute) |  | 36.93 | 0.13 | 1.1 | 0.1 | NA | NA | 0.51 | NA | 0 | 0.87 | 0.32 | 0.06 | NA | 0.16 | NA | 1.18 | 0.55 | 1.23 | 0.67 | 3.15 | 1.41 | 0.9
Results from laboratories 1.08, 1.10, 1.13, 1.22 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 231b (Certified Value 156 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 174 | 155.1 | 145.71 | 167.69 | 154.3047 | 154 | 170 | 159 | 157.1 | 154 | 154.4 | 165.5 | 153.5 | 150 | 154.5 | 164 | 159 | 142.1 | 157 | 11.84 | 148.194 | 150 | 144.8339 | 159 | 144.2526 | 170 | 153 | 154.6 | 154 | 1
| 1 | 2 | 163.1 | 154.8 | 153.44 | 165.26 | 155.4504 | 157 | 170 | 153 | 152.8 | 154 | 148 | 165 | 161 | 155 | 167 | 158 | 158 | 139.5 | 155 | 16.07 | 154.503 | 150 | 153.7211 | 155.5 | 139.4336 | 169 | 158.7 | 156.2 | 153 | 29
| 1 | 3 | 161.7 | 158 | 145.59 | 162.86 | 153.9135 | 157 | 166 | 151 | 155.9 | 154 | 151.4 | 166.5 | 153.5 | 158 | 153 | 165 | 159 | 142.8 | 160 | 16.8 | 152.847 | 150 | 147.0103 | 157 | 142.4664 | 169 | 151.8 | 156.3 | 148 | 16
| 1 | 4 | 160.9 | 155 | 144.96 | 167.1 | 154.9224 | 160 | 170 | 151 | 157 | 156 | 152.9 | 168.5 | 154 | 155.5 | 156 | 160 | 156 | 139.5 | 168 | 18.17 | 150.21 | 150 | 147.8199 | 155 | 146.0582 | 168 | 155 | 154.3 | 153 | 155.6637
| 1 | 5 | 159.4 | 160.1 | 147.34 | 165.87 | 155.1435 | 158 | 166 | 150 | 149.7 | 155 | 147.4 | 157.5 | 146.5 | 151 | 161 | 159 | 159 | 142.1 | 168 | 16.61 | 148.194 | 145 | 152.6945 | 156.5 | 141.9537 | 169 | 156.3 | 156.5 | 152 | 27
| 1 | 6 | 162 | 157.2 | 148.92 | 170.3 | 155.3117 | 161 | 161 | 157 | 154.5 | 153 | 153.2 | 165 | 156.5 | 160 | 157 | 159 | 156 | 140.8 | 157 | 15.25 | 145.602 | 150 | 147.3433 | 157.5 | 242.8339 | 168 | 153.4 | 156.2 | 147
| Mean |  | 163.5167 | 156.7 | 147.66 | 166.5133 | 154.841 | 157.8333 | 167.1667 | 153.5 | 154.5 | 154.3333 | 151.2167 | 164.6667 | 154.1667 | 154.9167 | 158.0833 | 160.8333 | 157.8333 | 141.1333 | 160.8333 | 15.79 | 149.925 | 149.1667 | 148.9038 | 156.75 | 159.4997 | 168.8333 | 154.7 | 155.6833 | 151.1667
| Median |  | 161.85 | 156.15 | 146.525 | 166.485 | 155.0329 | 157.5 | 168 | 152 | 155.2 | 154 | 152.15 | 165.25 | 153.75 | 155.25 | 156.5 | 159.5 | 158.5 | 141.45 | 158.5 | 16.34 | 149.202 | 150 | 147.5816 | 156.75 | 143.3595 | 169 | 154.2 | 156.2 | 152.5
| Std Dev. |  | 5.2814 | 2.1241 | 3.1773 | 2.506 | 0.60648 | 2.4833 | 3.6009 | 3.6742 | 2.8601 | 1.0328 | 2.893 | 3.7506 | 4.7293 | 3.8784 | 5.1421 | 2.9269 | 1.472 | 1.4208 | 5.7764 | 2.1601 | 3.2942 | 2.0412 | 3.5024 | 1.4405 | 40.8861 | 0.75277 | 2.5156 | 0.96626 | 2.9269
| Rel.Std.Dev. |  | 3.23% | 1.356% | 2.152% | 1.505% | 0.3917% | 1.573% | 2.154% | 2.394% | 1.851% | 0.6692% | 1.913% | 2.278% | 3.068% | 2.504% | 3.253% | 1.82% | 0.9326% | 1.007% | 3.592% | 13.68% | 2.197% | 1.368% | 2.352% | 0.919% | 25.63% | 0.4459% | 1.626% | 0.6207% | 1.936%
| PDM3 |  | 5.045% | 0.6657% | -5.142% | 6.97% | -0.5285% | 1.394% | 7.39% | -1.39% | -0.7476% | -0.8547% | -2.857% | 5.784% | -0.9617% | -0.4799% | 1.554% | 3.321% | 1.394% | -9.334% | 3.321% | -89.86% | -3.687% | -4.174% | -4.343% | 0.6978% | 2.464% | 8.46% | -0.6191% | 0.0126% | -2.889%
| Z-Score (Absolute) |  | 1.55 | 0.32 | 1.3 | 2.09 | 0.01 | 0.52 | 2.2 | 0.25 | 0.07 | 0.1 | 0.67 | 1.75 | 0.13 | 0 | 0.57 | 1.06 | 0.52 | 2.48 | 1.06 | 25.02 | 0.9 | 1.03 | 1.08 | 0.33 | 0.82 | 2.5 | 0.04 | 0.14 | 0.67
| Analytical results for Dy in OREAS 231b (Certified Value 3.77 ppm) | [30 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.71 | 3.46 | 3.5193 | 3.4 | 3.79 | 3.74 | 3.7 | 3.8759 | [21 empty] | 1
| 1 | 2 | 3.76 | 3.69 | 3.4823 | 3.5 | 3.85 | 3.68 | 4 | 3.9208 | [21 empty] | 30
| 1 | 3 | 3.87 | 3.84 | 3.4254 | 3.6 | 3.79 | 3.84 | 3.9 | 3.8485 | [21 empty] | 16
| 1 | 4 | 3.82 | 3.45 | 3.4857 | 3.7 | 3.88 | 3.75 | 3.8 | 3.9082 | [21 empty] | 3.7696
| 1 | 5 | 3.94 | 3.49 | 3.492 | 3.5 | 3.79 | 3.73 | 3.7 | 3.7373 | [21 empty] | 28
| 1 | 6 | 3.85 | 3.49 | 3.5334 | 3.7 | 3.64 | 3.8 | 3.9 | 3.7857
| Mean |  | 3.825 | 3.57 | 3.4897 | 3.5667 | 3.79 | 3.7567 | 3.8333 | 3.8461
| Median |  | 3.835 | 3.49 | 3.4889 | 3.55 | 3.79 | 3.745 | 3.85 | 3.8622
| Std Dev. |  | 0.08167 | 0.15912 | 0.037427 | 0.12111 | 0.082704 | 0.056095 | 0.12111 | 0.071843
| Rel.Std.Dev. |  | 2.135% | 4.457% | 1.072% | 3.395% | 2.182% | 1.493% | 3.159% | 1.868%
| PDM3 |  | 1.469% | -5.296% | -7.426% | -5.384% | 0.5406% | -0.3436% | 1.69% | 2.028%
| Z-Score (Absolute) |  | 0.52 | 2.07 | 2.88 | 2.1 | 0.17 | 0.17 | 0.61 | 0.74
| Analytical results for Er in OREAS 231b (Certified Value 2.32 ppm) | [30 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.3 | 2.04 | 2.2225 | 2.1 | 2.33 | 2.3 | 2.3 | 2.4383 | [21 empty] | 1
| 1 | 2 | 2.26 | 2.18 | 2.2131 | 2 | 2.35 | 2.29 | 2.4 | 2.4791 | [21 empty] | 31
| 1 | 3 | 2.36 | 2.27 | 2.1697 | 2.3 | 2.31 | 2.28 | 2.5 | 2.4849 | [21 empty] | 16
| 1 | 4 | 2.29 | 2.04 | 2.2041 | 2.2 | 2.4 | 2.2 | 2.3 | 2.4825 | [21 empty] | 2.315
| 1 | 5 | 2.36 | 2.08 | 2.2536 | 2.1 | 2.32 | 2.2 | 2.3 | 2.4288 | [21 empty] | 29
| 1 | 6 | 2.36 | 2.08 | 2.2403 | 2.3 | 2.2 | 2.28 | 2.3 | 2.4431
| Mean |  | 2.3217 | 2.115 | 2.2172 | 2.1667 | 2.3183 | 2.2583 | 2.35 | 2.4595
| Median |  | 2.33 | 2.08 | 2.2178 | 2.15 | 2.325 | 2.28 | 2.3 | 2.4611
| Std Dev. |  | 0.044008 | 0.091597 | 0.029408 | 0.12111 | 0.066156 | 0.045789 | 0.083666 | 0.025382
| Rel.Std.Dev. |  | 1.896% | 4.331% | 1.326% | 5.59% | 2.854% | 2.028% | 3.56% | 1.032%
| PDM3 |  | 0.2878% | -8.639% | -4.224% | -6.408% | 0.1438% | -2.448% | 1.512% | 6.24%
| Z-Score (Absolute) |  | 0.43 | 2.24 | 0.92 | NA | 0.39 | 0.39 | NA | 2.21
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 231b (Certified Value 0.92 ppm) | [30 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.9 | 0.88 | 1.3318 | 0.82 | 0.97 | 0.99 | 0.9 | 0.9183 | [21 empty] | 1
| 1 | 2 | 0.9 | 0.93 | 1.2833 | 0.8 | 0.95 | 0.98 | 0.9 | 0.93914 | [21 empty] | 32
| 1 | 3 | 0.93 | 0.95 | 1.3174 | 0.89 | 0.96 | 0.98 | 1.1 | 0.92561 | [21 empty] | 16
| 1 | 4 | 0.92 | 0.87 | 1.3545 | 0.87 | 0.95 | 0.98 | 0.9 | 0.94406 | [21 empty] | 0.92279
| 1 | 5 | 0.91 | 0.88 | 1.3026 | 0.82 | 0.95 | 0.98 | 0.9 | 0.91147 | [21 empty] | 30
| 1 | 6 | 0.95 | 0.87 | 1.3418 | 0.92 | 0.91 | 1 | 1.1 | 0.92569
| Mean |  | 0.91833 | 0.89667 | 1.3219 | 0.85333 | 0.94833 | 0.985 | 0.96667 | 0.92738
| Median |  | 0.915 | 0.88 | 1.3246 | 0.845 | 0.95 | 0.98 | 0.9 | 0.92565
| Std Dev. |  | 0.019408 | 0.034448 | 0.026218 | 0.047188 | 0.020412 | 0.0083666 | 0.10328 | 0.012313
| Rel.Std.Dev. |  | 2.113% | 3.842% | 1.983% | 5.53% | 2.152% | 0.8494% | 10.68% | 1.328%
| PDM3 |  | -0.4824% | -2.83% | 43.25% | -7.526% | 2.769% | 6.742% | 4.755% | 0.4977%
| Z-Score (Absolute) |  | 0.2 | 0.67 | 8.66 | 1.63 | 0.46 | 1.27 | NA | 0
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 231b (Certified Value 8.31 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 8.46 | 8.48 | 8.552 | 8.1878 | 7.75 | 9.093 | 8.48 | 8.49 | 8.45 | 7.93 | 8.84 | 8.33 | 8.32 | 8.29 | 7.79 | 8.75 | 7.76 | 8.14 | 8.5 | 7.6902 | 8.44 | 8.57 | 7.9433 | 8.17 | 7.69 | 8.67 | 8.4 |  |  | 1
| 1 | 2 | 8.45 | 8.38 | 8.5681 | 8.2739 | 7.8 | 9.163 | 8.46 | 8.18 | 8.49 | 7.82 | 8.86 | 8.58 | 8.03 | 8.9 | 7.74 | 8.92 | 7.65 | 8.04 | 8.3 | 7.9798 | 8.395 | 8.14 | 8.1213 | 8.1 | 7.87 | 8.76 | 9.6 |  |  | #N/A
| 1 | 3 | 8.62 | 8.76 | 8.5265 | 8.2037 | 8.51 | 9.163 | 8.41 | 8.34 | 8.45 | 7.87 | 8.91 | 8.33 | 8.41 | 8.17 | 7.86 | 8.78 | 7.63 | 8.12 | 7.93 | 8.0048 | 8.465 | 8.52 | 8.1803 | 8.26 | 7.51 | 8.79 | 8.3 |  |  | 16
| 1 | 4 | 8.49 | 8.57 | 8.591 | 8.2221 | 8.28 | 9.233 | 8.58 | 8.53 | 8.42 | 7.94 | 9.01 | 8.31 | 8.22 | 8.28 | 7.8 | 8.83 | 7.75 | 8.56 | 8.29 | 7.8368 | 8.395 | 8.41 | 8.3133 | 8.18 | 7.58 | 8.74 | 9.3 |  |  | 8.3054
| 1 | 5 | 8.48 | 8.57 | 8.5454 | 8.2543 | 7.8 | 9.233 | 8.55 | 8.25 | 8.45 | 7.9 | 8.35 | 8.09 | 8.19 | 8.54 | 7.78 | 8.86 | 7.73 | 8.62 | 7.76 | 7.6274 | 8.49 | 8.48 | 8.4503 | 8.14 | 7.6 | 8.8 | 9 |  |  | 31
| 1 | 6 | 8.65 | 8.65 | 8.5479 | 8.3254 | 8.28 | 9.163 | 8.26 | 8.43 | 8.44 | 7.89 | 8.76 | 8.47 | 8.11 | 8.37 | 7.65 | 8.87 | 7.62 | 8.08 | 7.82 | 7.5482 | 8.46 | 8.65 | 8.4083 | 8.23 | 7.68 | 8.82 | 9
| Mean |  | 8.525 | 8.5683 | 8.5551 | 8.2445 | 8.07 | 9.1747 | 8.4567 | 8.37 | 8.45 | 7.8917 | 8.7883 | 8.3517 | 8.2133 | 8.425 | 7.77 | 8.835 | 7.69 | 8.26 | 8.1 | 7.7812 | 8.4408 | 8.4617 | 8.2361 | 8.18 | 7.655 | 8.7633 | 8.9333
| Median |  | 8.485 | 8.57 | 8.5499 | 8.2382 | 8.04 | 9.163 | 8.47 | 8.385 | 8.45 | 7.895 | 8.85 | 8.33 | 8.205 | 8.33 | 7.785 | 8.845 | 7.69 | 8.13 | 8.11 | 7.7635 | 8.45 | 8.5 | 8.2468 | 8.175 | 7.64 | 8.775 | 9
| Std Dev. |  | 0.086891 | 0.13167 | 0.022032 | 0.050773 | 0.32558 | 0.052694 | 0.11431 | 0.13784 | 0.022804 | 0.043551 | 0.23008 | 0.1657 | 0.13779 | 0.26312 | 0.070427 | 0.062209 | 0.063561 | 0.25861 | 0.30299 | 0.18908 | 0.038912 | 0.17725 | 0.19157 | 0.05831 | 0.1247 | 0.053914 | 0.50465
| Rel.Std.Dev. |  | 1.019% | 1.537% | 0.2575% | 0.6158% | 4.034% | 0.5743% | 1.352% | 1.647% | 0.2699% | 0.5519% | 2.618% | 1.984% | 1.678% | 3.123% | 0.9064% | 0.7041% | 0.8265% | 3.131% | 3.741% | 2.43% | 0.461% | 2.095% | 2.326% | 0.7128% | 1.629% | 0.6152% | 5.649%
| PDM3 |  | 2.644% | 3.166% | 3.007% | -0.7327% | -2.834% | 10.47% | 1.822% | 0.7781% | 1.741% | -4.981% | 5.815% | 0.5573% | -1.108% | 1.44% | -6.446% | 6.377% | -7.409% | -0.5464% | -2.473% | -6.311% | 1.631% | 1.882% | -0.834% | -1.51% | -7.831% | 5.514% | 7.561%
| Z-Score (Absolute) |  | 0.55 | 0.7 | 0.66 | 0.45 | 1.06 | 2.85 | 0.31 | 0 | 0.28 | 1.7 | 1.48 | 0.07 | 0.56 | 0.2 | 2.13 | 1.65 | 2.41 | 0.39 | 0.96 | 2.09 | 0.25 | 0.33 | 0.48 | 0.67 | 2.54 | 1.4 | 2
| Analytical results for Ga in OREAS 231b (Certified Value 15.4 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 16.5 | 15.27 | 14.4238 | 13.3402 | 12.3 | 14.9 | 16 | 16.1 | 19 | 16 | 15.95 | 16.15 | 14.35 | 14.3 | 11 | 14.2 | 15.98 | 11.27 | 12 | 12.42 | 16.3 | 16.1354 | 19.2 | 15.7 | 16.93 | 14.6 |  |  |  | 1
| 1 | 2 | 16.8 | 15.36 | 14.2126 | 13.4941 | 12.6 | 15 | 15.3 | 16.09 | 19 | 15 | 16.1 | 15.85 | 15 | 15.9 | 10 | 16.4 | 14.94 | 12.08 | 12 | 12.381 | 14.85 | 16.3212 | 19.2 | 16 | 16.22 | 13.9 |  |  |  | 33
| 1 | 3 | 16.8 | 15.75 | 14.3054 | 13.3412 | 13.4 | 14.8 | 15.9 | 16.2 | 19 | 15 | 16.3 | 15.3 | 15.95 | 15.25 | 10 | 17.3 | 15.71 | 11.86 | 12 | 12.2305 | 16.2 | 16.5574 | 19 | 15.2 | 16.25 | 14.2 |  |  |  | 16
| 1 | 4 | 16.2 | 15.53 | 14.731 | 13.5286 | 14.4 | 15.1 | 17 | 16.82 | 19 | 16 | 15.75 | 15.15 | 15.2 | 14.95 | 10 | 15.9 | 15.8 | 12.71 | 12 | 12.2452 | 15.85 | 16.9562 | 19.1 | 15.5 | 15.67 | 13.7 |  |  |  | 15.3744
| 1 | 5 | 16.4 | 15.69 | 14.7803 | 14.1331 | 13.2 | 14.6 | 15.8 | 15.87 | 19 | 15 | 15.7 | 14.35 | 14.6 | 15.55 | 10 | 13.1 | 15.72 | 11.37 | 12 | 12.4536 | 16.15 | 16.8475 | 19.1 | 15.4 | 16.37 | 14.7 |  |  |  | 32
| 1 | 6 | 16.9 | 15.58 | 14.8933 | 13.8608 | 13.7 | 14 | 15.6 | 16.51 | 20 | 15 | 16.3 | 15.4 | 15.15 | 15.5 | 10 | 14.9 | 16.3 | 11.64 | 12 | 12.4243 | 16.7 | 16.844 | 18.6 | 16.1 | 16.07 | 14.2
| Mean |  | 16.6 | 15.53 | 14.5577 | 13.6163 | 13.2667 | 14.7333 | 15.9333 | 16.265 | 19.1667 | 15.3333 | 16.0167 | 15.3667 | 15.0417 | 15.2417 | 10.1667 | 15.3 | 15.7417 | 11.8217 | 12 | 12.3591 | 16.0083 | 16.6103 | 19.0333 | 15.65 | 16.2517 | 14.2167
| Median |  | 16.65 | 15.555 | 14.5774 | 13.5114 | 13.3 | 14.85 | 15.85 | 16.15 | 19 | 15 | 16.025 | 15.35 | 15.075 | 15.375 | 10 | 15.4 | 15.76 | 11.75 | 12 | 12.4005 | 16.175 | 16.7007 | 19.1 | 15.6 | 16.235 | 14.2
| Std Dev. |  | 0.27568 | 0.18601 | 0.28029 | 0.31664 | 0.75807 | 0.39833 | 0.5785 | 0.34239 | 0.40825 | 0.5164 | 0.26204 | 0.62183 | 0.55445 | 0.55984 | 0.40825 | 1.5349 | 0.4508 | 0.5289 | 0 | 0.096819 | 0.63041 | 0.32937 | 0.22509 | 0.35071 | 0.41146 | 0.38687
| Rel.Std.Dev. |  | 1.661% | 1.198% | 1.925% | 2.325% | 5.714% | 2.704% | 3.631% | 2.105% | 2.13% | 3.368% | 1.636% | 4.047% | 3.686% | 3.673% | 4.016% | 10.03% | 2.864% | 4.474% | 0% | 0.7834% | 3.938% | 1.983% | 1.183% | 2.241% | 2.532% | 2.721%
| PDM3 |  | 7.972% | 1.012% | -5.312% | -11.43% | -13.71% | -4.169% | 3.636% | 5.793% | 24.67% | -0.2668% | 4.178% | -0.04998% | -2.164% | -0.863% | -33.87% | -0.4836% | 2.389% | -23.11% | -21.95% | -19.61% | 4.124% | 8.039% | 23.8% | 1.793% | 5.706% | -7.53%
| Z-Score (Absolute) |  | 1.02 | 0.07 | 0.79 | 1.63 | 1.94 | 0.64 | 0.43 | 0.73 | NA | NA | 0.5 | 0.07 | 0.36 | 0.18 | NA | 0.13 | 0.26 | 3.22 | NA | 2.74 | 0.5 | 1.03 | 3.18 | 0.18 | 0.71 | 1.09
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 231b (Certified Value 3.04 ppm) | [30 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.24 | 2.78 | 3.2578 | 2.7 | 2.9 | 3.28 | 3 | 0.57468 | [21 empty] | 1
| 1 | 2 | 3.23 | 2.97 | 3.2112 | 2.8 | 2.8 | 3.3 | 3 | 0.59293 | [21 empty] | 34
| 1 | 3 | 3.32 | 3.06 | 3.2919 | 2.8 | 2.8 | 3.28 | 3.2 | 0.57037 | [21 empty] | 16
| 1 | 4 | 3.15 | 2.76 | 3.2449 | 2.9 | 2.8 | 3.27 | 3.1 | 0.57834 | [21 empty] | 3.0421
| 1 | 5 | 3.36 | 2.81 | 3.242 | 2.7 | 2.7 | 3.28 | 3.1 | 0.57258 | [21 empty] | 33
| 1 | 6 | 3.34 | 2.79 | 3.2559 | 2.9 | 2.6 | 3.45 | 3.2 | 0.57882
| Mean |  | 3.2733 | 2.8617 | 3.2506 | 2.8 | 2.7667 | 3.31 | 3.1 | 0.57795
| Median |  | 3.28 | 2.8 | 3.2504 | 2.8 | 2.8 | 3.28 | 3.1 | 0.57651
| Std Dev. |  | 0.080416 | 0.12319 | 0.026249 | 0.089443 | 0.10328 | 0.069282 | 0.089443 | 0.0080286
| Rel.Std.Dev. |  | 2.457% | 4.305% | 0.8075% | 3.194% | 3.733% | 2.093% | 2.885% | 1.389%
| PDM3 |  | 7.602% | -5.931% | 6.855% | -7.958% | -9.054% | 8.807% | 1.904% | -81%
| Z-Score (Absolute) |  | 0.82 | 0.33 | 0.75 | 0.5 | 0.6 | 0.92 | 0.33 | 6.7
| Analytical results for Ge in OREAS 231b (Indicative Value 0.16 ppm) | [30 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 0.1 | 1.7 | 0.13048 | < 0.2 | 0.3 | 1.3 | 0.8 | 3.1 | 0.09 | 0.1 | 0.1 | 0.07 | < 10 | 0.6 | < 5 | 0.08 | 0.2805 | 0.11 | 0.3 | < 0.05 | 0.11 | [8 empty] | 1
| 1 | 2 | 0.1 | 1.7 | 0.12015 | < 0.2 | 0.4 | 1.1 | 1 | 3 | 0.09 | 0.11 | 0.06 | 0.08 | < 10 | 0.7 | < 5 | 0.42 | 0.2256 | 0.16 | 0.4 | < 0.05 | 0.09 | [8 empty] | 4
| 1 | 3 | 0.2 | 1.7 | 0.13126 | < 0.2 | 0.4 | 1.3 | 1.1 | 2.9 | 0.11 | 0.09 | 0.16 | 0.08 | < 10 | 0.7 | < 5 | 0.09 | 0.226 | 0.14 | 0.4 | < 0.05 | 0.1 | [8 empty] | 16
| 1 | 4 | 0.5 | 1.7 | 0.13465 | < 0.2 | 0.2 | 1.5 | 1.2 | 3.1 | 0.08 | 0.1 | 0.09 | 0.09 | < 10 | 0.7 | < 5 | 0.09 | 0.2628 | 0.12 | 0.4 | < 0.05 | 0.1 | [8 empty] | 0.15995
| 1 | 5 | < 0.1 | 1.6 | 0.13012 | < 0.2 | 0.4 | 1.4 | 1.1 | 3 | 0.08 | 0.1 | 0.08 | 0.07 | < 10 | 0.6 | < 5 | 0.1 | 0.2474 | 0.14 | 0.4 | 0.05 | 0.11 | [8 empty] | 10
| 1 | 6 | < 0.1 | 1.7 | 0.13111 | < 0.2 | 0.4 | 1.6 | 1 | 3 | 0.08 | 0.11 | 0.08 | 0.06 | < 10 | 0.6 | < 5 | 0.1 | 0.311 | 0.13 | 0.4 | 0.06 | 0.11
| Mean |  | 0.225 | 1.6833 | 0.12963 |  | 0.35 | 1.3667 | 1.0333 | 3.0167 | 0.088333 | 0.10167 | 0.095 | 0.075 |  | 0.65 |  | 0.14667 | 0.25888 | 0.13333 | 0.38333 | 0.055 | 0.10333
| Median |  | 0.15 | 1.7 | 0.1308 |  | 0.4 | 1.35 | 1.05 | 3 | 0.085 | 0.1 | 0.085 | 0.075 |  | 0.65 |  | 0.095 | 0.2551 | 0.135 | 0.4 | 0.055 | 0.105
| Std Dev. |  | 0.1893 | 0.040825 | 0.0049176 |  | 0.083666 | 0.17512 | 0.13663 | 0.075277 | 0.01169 | 0.0075277 | 0.034496 | 0.010488 |  | 0.054772 |  | 0.13411 | 0.033226 | 0.017512 | 0.040825 | 0.0070711 | 0.008165
| Rel.Std.Dev. |  | 84.13% | 2.425% | 3.794% |  | 23.9% | 12.81% | 13.22% | 2.495% | 13.23% | 7.404% | 36.31% | 13.98% |  | 8.427% |  | 91.44% | 12.83% | 13.13% | 10.65% | 12.86% | 7.902%
| PDM3 |  | 40.67% | 952.4% | -18.96% |  | 118.8% | 754.4% | 546% | 1786% | -44.78% | -36.44% | -40.61% | -53.11% |  | 306.4% |  | -8.307% | 61.85% | -16.64% | 139.7% | -65.62% | -35.4%
| Z-Score (Absolute) |  | 0 | 11.09 | 0.27 | 0.49 | 1.34 | 8.77 | 6.34 | 20.84 | 0.57 | 0.48 | 0.52 | 0.67 | 35.34 | 3.53 | 17.06 | 0.15 | 0.67 | 0.24 | 1.58 | 0.96 | 0.46
| Analytical results for Hf in OREAS 231b (Certified Value 1.54 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.59 | 1.41 | 0.77603 | 0.9 | 1.5 | 1.8 | 1.54 | 1.9 | 1.5 | 1.6 | 1.4 | 1.4 | 1.4 | 1.3 | 1.62 | 2.5 | 1.876 | 1.7 | 1.384 | 2 | 2.4 | 1.8 | 1.36 | [6 empty] | 1
| 1 | 2 | 2.03 | 1.34 | 0.76221 | 1.5 | 1.5 | 1.8 | 1.5 | 1.7 | 1.4 | 1.6 | 1.3 | 1.5 | 1.5 | 1.3 | 1.65 | 2 | 1.8774 | 1.4 | 1.4116 | 2.3 | 2.27 | 2 | 1.27 | [6 empty] | 19
| 1 | 3 | 1.56 | 1.45 | 0.79705 | 1.5 | 1.5 | 1.6 | 1.61 | 1.7 | 1.5 | 1.6 | 1.4 | 1.4 | 1.5 | 1.3 | 1.69 | 2 | 1.8706 | 1.5 | 1.4174 | 2.3 | 2.13 | 2 | 1.27 | [6 empty] | 16
| 1 | 4 | 1.8 | 1.4 | 0.78976 | 0.5 | 1.5 | 1.5 | 1.59 | 1.7 | 1.5 | 1.6 | 1.6 | 1.4 | 1.6 | 1.3 | 1.62 | 1.5 | 1.8637 | 1.5 | 1.4367 | 1.9 | 2.18 | 1.9 | 1.27 | [6 empty] | 1.5439
| 1 | 5 | 1.53 | 1.47 | 0.73955 | 0.9 | 1.5 | 1.5 | 1.5 | 1.7 | 1.4 | 1.5 | 1.4 | 1.5 | 1.4 | 1.3 | 1.65 | 1.5 | 1.808 | 1.5 | 1.4292 | 2 | 2.13 | 1.9 | 1.32 | [6 empty] | 34
| 1 | 6 | 1.84 | 1.43 | 0.77466 | 1.4 | 1.4 | 1.8 | 1.53 | 1.8 | 1.5 | 1.6 | 1.6 | 1.3 | 1.4 | 1.3 | 1.63 | 1.5 | 1.8419 | 1.5 | 1.3654 | 2 | 2.14 | 2 | 1.32
| Mean |  | 1.725 | 1.4167 | 0.77321 | 1.1167 | 1.4833 | 1.6667 | 1.545 | 1.75 | 1.4667 | 1.5833 | 1.45 | 1.4167 | 1.4667 | 1.3 | 1.6433 | 1.8333 | 1.8563 | 1.5167 | 1.4074 | 2.0833 | 2.2083 | 1.9333 | 1.3017
| Median |  | 1.695 | 1.42 | 0.77535 | 1.15 | 1.5 | 1.7 | 1.535 | 1.7 | 1.5 | 1.6 | 1.4 | 1.4 | 1.45 | 1.3 | 1.64 | 1.75 | 1.8671 | 1.5 | 1.4145 | 2 | 2.16 | 1.95 | 1.295
| Std Dev. |  | 0.19766 | 0.045461 | 0.020522 | 0.41191 | 0.040825 | 0.15055 | 0.045935 | 0.083666 | 0.05164 | 0.040825 | 0.12247 | 0.075277 | 0.08165 | 0 | 0.026583 | 0.40825 | 0.026962 | 0.098319 | 0.027433 | 0.17224 | 0.10797 | 0.08165 | 0.037639
| Rel.Std.Dev. |  | 11.46% | 3.209% | 2.654% | 36.89% | 2.752% | 9.033% | 2.973% | 4.781% | 3.521% | 2.578% | 8.447% | 5.314% | 5.567% | 0% | 1.618% | 22.27% | 1.452% | 6.483% | 1.949% | 8.268% | 4.889% | 4.223% | 2.892%
| PDM3 |  | 11.73% | -8.242% | -49.92% | -27.67% | -3.924% | 7.951% | 0.07029% | 13.35% | -5.003% | 2.553% | -6.083% | -8.242% | -5.003% | -15.8% | 6.439% | 18.75% | 20.23% | -1.765% | -8.843% | 34.94% | 43.03% | 25.22% | -15.69%
| Z-Score (Absolute) |  | 0.94 | 0.45 | 3.34 | 1.8 | 0.15 | 0.67 | 0.13 | 1.05 | 0.22 | 0.3 | 0.3 | 0.45 | 0.22 | 0.97 | 0.57 | 1.42 | 1.53 | 0 | 0.49 | 2.55 | 3.11 | 1.87 | 0.97
| Analytical results for Hg in OREAS 231b (Indicative Value 0.19 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.05 | 1.25 | 1.30 | [26 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 2
2
| 1 | 1 | < 1 | 0.199 | 0.1168 | [26 empty] | 1
| 1 | 2 | < 1 | 0.497 | 0.0433 | [26 empty] | 6
| 1 | 3 | < 1 | < 0.1 | 0.0183 | [26 empty] | 16
| 1 | 4 | < 1 | 0.3396 | < 0.01 | [26 empty] | 0.1885
| 1 | 5 | < 1 | 0.5302 | < 0.0101 | [26 empty] | 9
| 1 | 6 | < 1 | 0.5443 | < 0.0265
| Mean |  |  | 0.42202 | 0.059467
| Median |  |  | 0.497 | 0.0433
| Std Dev. |  |  | 0.149 | 0.051201
| Rel.Std.Dev. |  |  | 35.31% | 86.1%
| PDM3 |  |  | 123.9% | -68.45%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.57
| Analytical results for Ho in OREAS 231b (Certified Value 0.77 ppm) | [30 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.8 | 0.72 | 0.77306 | 0.7 | 0.74 | 0.79 | 0.75 | 0.82772 | [21 empty] | 1
| 1 | 2 | 0.77 | 0.75 | 0.74698 | 0.7 | 0.74 | 0.81 | 0.8 | 0.84696 | [21 empty] | 20
| 1 | 3 | 0.82 | 0.79 | 0.7324 | 0.8 | 0.73 | 0.8 | 0.75 | 0.83844 | [21 empty] | 16
| 1 | 4 | 0.83 | 0.71 | 0.75499 | 0.7 | 0.74 | 0.82 | 0.8 | 0.83914 | [21 empty] | 0.77451
| 1 | 5 | 0.84 | 0.72 | 0.73853 | 0.7 | 0.72 | 0.8 | 0.75 | 0.80731 | [21 empty] | 35
| 1 | 6 | 0.81 | 0.72 | 0.77671 | 0.8 | 0.7 | 0.83 | 0.75 | 0.81304
| Mean |  | 0.81167 | 0.735 | 0.75378 | 0.73333 | 0.72833 | 0.80833 | 0.76667 | 0.82877
| Median |  | 0.815 | 0.72 | 0.75099 | 0.7 | 0.735 | 0.805 | 0.75 | 0.83308
| Std Dev. |  | 0.024833 | 0.030166 | 0.018083 | 0.05164 | 0.016021 | 0.01472 | 0.02582 | 0.015754
| Rel.Std.Dev. |  | 3.059% | 4.104% | 2.399% | 7.042% | 2.2% | 1.821% | 3.368% | 1.901%
| PDM3 |  | 4.798% | -5.101% | -2.676% | -5.316% | -5.962% | 4.367% | -1.012% | 7.006%
| Z-Score (Absolute) |  | 0.79 | 0.56 | 0.23 | NA | 0.67 | 0.73 | 0 | 1.09
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 231b (Certified Value 0.071 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.057386 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.076 | 0.082 | 0.066 | 0.082 | < 10 | 0.06 | 0.06 | 0.071 | 34.9 | < 0.1 | 0.06411 | 0.04 | 0.0907 | 0.08 | 0.07 | 0.071 | 0.065 | [5 empty] | 1
| 1 | 2 | 0.07 | 0.07 | 0.05664 | < 0.1 | 0.08 | 0.07 | 0.07 | 0.075 | 0.088 | 0.067 | 0.095 | < 10 | 0.07 | 0.06 | 0.068 | 35.26 | < 0.1 | 0.065 | 0.072 | 0.0794 | 0.07 | 0.07 | 0.073 | 0.059 | [5 empty] | 21
| 1 | 3 | 0.08 | 0.07 | 0.055398 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.075 | 0.083 | 0.075 | 0.08 | < 10 | 0.08 | 0.07 | 0.07 | 27.6 | < 0.1 | 0.06277 | 0.055 | 0.0856 | 0.07 | 0.07 | 0.07 | 0.062 | [5 empty] | 16
| 1 | 4 | 0.08 | 0.06 | 0.056019 | < 0.1 | 0.06 | 0.07 | 0.09 | 0.077 | 0.082 | 0.071 | 0.081 | < 10 | 0.07 | 0.06 | 0.071 | 36.32 | < 0.1 | 0.06278 | 0.038 | 0.0849 | 0.06 | 0.07 | 0.074 | 0.067 | [5 empty] | 0.071179
| 1 | 5 | 0.08 | 0.07 | 0.056268 | < 0.1 | 0.07 | 0.07 | 0.09 | 0.076 | 0.072 | 0.073 | 0.081 | < 10 | 0.05 | 0.07 | 0.07 | 32.47 | < 0.1 | 0.06447 | 0.05 | 0.0879 | 0.072 | 0.07 | 0.079 | 0.064 | [5 empty] | 36
| 1 | 6 | 0.08 | 0.07 | 0.056641 | < 0.1 | 0.07 | 0.07 | 0.08 | 0.078 | 0.082 | 0.083 | 0.084 | < 10 | 0.07 | 0.07 | 0.074 | 31.33 | < 0.1 | 0.06517 | 0.056 | 0.0728 | 0.066 | 0.08 | 0.076 | 0.064
| Mean |  | 0.078333 | 0.07 | 0.056392 |  | 0.07 | 0.07 | 0.078333 | 0.076167 | 0.0815 | 0.0725 | 0.083833 |  | 0.066667 | 0.065 | 0.070667 | 32.98 |  | 0.06405 | 0.051833 | 0.08355 | 0.069667 | 0.071667 | 0.073833 | 0.0635
| Median |  | 0.08 | 0.07 | 0.056454 |  | 0.07 | 0.07 | 0.075 | 0.076 | 0.082 | 0.072 | 0.0815 |  | 0.07 | 0.065 | 0.0705 | 33.685 |  | 0.06429 | 0.0525 | 0.08525 | 0.07 | 0.07 | 0.0735 | 0.064
| Std Dev. |  | 0.0040825 | 0.0063246 | 0.00067131 |  | 0.0063246 | 0 | 0.0098319 | 0.001169 | 0.0052058 | 0.0061887 | 0.0056362 |  | 0.010328 | 0.0054772 | 0.0019664 | 3.2239 |  | 0.0010572 | 0.0124 | 0.006463 | 0.0066232 | 0.0040825 | 0.0033116 | 0.0027386
| Rel.Std.Dev. |  | 5.212% | 9.035% | 1.19% |  | 9.035% | 0% | 12.55% | 1.535% | 6.387% | 8.536% | 6.723% |  | 15.49% | 8.427% | 2.783% | 9.775% |  | 1.651% | 23.92% | 7.736% | 9.507% | 5.696% | 4.485% | 4.313%
| PDM3 |  | 10.05% | -1.656% | -20.77% |  | -1.656% | -1.656% | 10.05% | 7.008% | 14.5% | 1.856% | 17.78% |  | -6.339% | -8.681% | -0.7194% | 46234% |  | -10.02% | -27.18% | 17.38% | -2.124% | 0.6855% | 3.729% | -10.79%
| Z-Score (Absolute) |  | 0.82 | 0.03 | 1.43 | 2.09 | 0.03 | 0.03 | 0.82 | 0.6 | 1.15 | 0.22 | 1.39 | 506.85 | 0.38 | 0.55 | 0.03 | 3383.68 | 2.09 | 0.65 | 1.9 | 1.36 | 0.07 | 0.14 | 0.36 | 0.7
| Analytical results for K in OREAS 231b (Certified Value 0.477 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.49 | 0.4665 | 0.46503 | 0.4639 | 0.44 | 0.5047 | 0.458 | 0.4586 | 0.48 | 0.51 | 0.5 | 0.46 | 0.45 | 0.45 | 0.48 | 0.477 | 0.43 | 0.51 | 0.595 | 0.46267 | 0.455 | 0.48 | 0.4948 | 0.47 | 0.51 | 0.5 | 0.49 |  |  | 1
| 1 | 2 | 0.48 | 0.4653 | 0.4602 | 0.46754 | 0.45 | 0.5081 | 0.454 | 0.451 | 0.49 | 0.52 | 0.5 | 0.47 | 0.46 | 0.48 | 0.47 | 0.479 | 0.43 | 0.49 | 0.602 | 0.47808 | 0.455 | 0.45 | 0.5037 | 0.47 | 0.54 | 0.5 | 0.48 |  |  | #N/A
| 1 | 3 | 0.51 | 0.4796 | 0.46114 | 0.46978 | 0.48 | 0.5114 | 0.453 | 0.4554 | 0.48 | 0.52 | 0.49 | 0.46 | 0.46 | 0.44 | 0.48 | 0.474 | 0.44 | 0.5 | 0.604 | 0.47968 | 0.455 | 0.47 | 0.5088 | 0.47 | 0.49 | 0.5 | 0.47 |  |  | 16
| 1 | 4 | 0.51 | 0.4636 | 0.46331 | 0.4692 | 0.48 | 0.5105 | 0.46 | 0.4561 | 0.48 | 0.51 | 0.5 | 0.46 | 0.44 | 0.45 | 0.47 | 0.474 | 0.44 | 0.52 | 0.637 | 0.47344 | 0.455 | 0.47 | 0.5225 | 0.47 | 0.51 | 0.5 | 0.47 |  |  | 0.47693
| 1 | 5 | 0.5 | 0.4766 | 0.46468 | 0.4698 | 0.44 | 0.513 | 0.461 | 0.444 | 0.48 | 0.51 | 0.47 | 0.44 | 0.44 | 0.46 | 0.47 | 0.475 | 0.44 | 0.53 | 0.661 | 0.46497 | 0.455 | 0.47 | 0.52949 | 0.47 | 0.51 | 0.51 | 0.48 |  |  | 37
| 1 | 6 | 0.51 | 0.471 | 0.4639 | 0.4721 | 0.48 | 0.5064 | 0.445 | 0.4487 | 0.49 | 0.5 | 0.48 | 0.47 | 0.46 | 0.46 | 0.47 | 0.478 | 0.44 | 0.5 | 0.775 | 0.46119 | 0.455 | 0.48 | 0.52259 | 0.48 | 0.53 | 0.5 | 0.46
| Mean |  | 0.5 | 0.47043 | 0.46304 | 0.46872 | 0.46167 | 0.50902 | 0.45517 | 0.4523 | 0.48333 | 0.51167 | 0.49 | 0.46 | 0.45167 | 0.45667 | 0.47333 | 0.47617 | 0.43667 | 0.50833 | 0.64567 | 0.47001 | 0.455 | 0.47 | 0.51365 | 0.47167 | 0.515 | 0.50167 | 0.475
| Median |  | 0.505 | 0.46875 | 0.4636 | 0.46949 | 0.465 | 0.5093 | 0.456 | 0.4532 | 0.48 | 0.51 | 0.495 | 0.46 | 0.455 | 0.455 | 0.47 | 0.476 | 0.44 | 0.505 | 0.6205 | 0.46921 | 0.455 | 0.47 | 0.51565 | 0.47 | 0.51 | 0.5 | 0.475
| Std Dev. |  | 0.012649 | 0.0064945 | 0.0019562 | 0.0027762 | 0.020412 | 0.0031644 | 0.0059133 | 0.0054273 | 0.005164 | 0.0075277 | 0.012649 | 0.010954 | 0.0098319 | 0.013663 | 0.005164 | 0.002137 | 0.005164 | 0.01472 | 0.068187 | 0.0080938 | 0 | 0.010954 | 0.013321 | 0.0040825 | 0.017607 | 0.0040825 | 0.010488
| Rel.Std.Dev. |  | 2.53% | 1.381% | 0.4225% | 0.5923% | 4.421% | 0.6217% | 1.299% | 1.2% | 1.068% | 1.471% | 2.581% | 2.381% | 2.177% | 2.992% | 1.091% | 0.4488% | 1.183% | 2.896% | 10.56% | 1.722% | 0% | 2.331% | 2.593% | 0.8655% | 3.419% | 0.8138% | 2.208%
| PDM3 |  | 4.837% | -1.362% | -2.912% | -1.722% | -3.2% | 6.728% | -4.563% | -5.164% | 1.343% | 7.283% | 2.741% | -3.55% | -5.297% | -4.249% | -0.754% | -0.16% | -8.442% | 6.585% | 35.38% | -1.452% | -4.598% | -1.453% | 7.699% | -1.103% | 7.982% | 5.187% | -0.4046%
| Z-Score (Absolute) |  | 1.16 | 0.05 | 0.35 | 0.12 | 0.41 | 1.53 | 0.67 | 0.79 | 0.48 | 1.63 | 0.75 | 0.48 | 0.82 | 0.61 | 0.07 | 0.18 | 1.43 | 1.5 | 7.11 | 0.07 | 0.68 | 0.07 | 1.72 | 0 | 1.77 | 1.23 | 0.14
| Analytical results for La in OREAS 231b (Certified Value 4.52 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 4.4 | 4.5 | 4.09 | 4.3632 | 4.4138 | 4.2 | 5 | 4.65 | < 5 | 8 | 4.5 | 4.9 | 4.2 | 4.3 | 5 | 4.2 | 4.4 | 5 | 4.6127 | 4.8 | 4.7063 | 5.4 | 4.7 | 5.4 | 4.7 |  |  |  |  | 1
| 1 | 2 | 4.6 | 4.51 | 4.35 | 4.3043 | 4.4989 | 4.3 | 5 | 4.44 | < 5 | 8 | 4.6 | 4.9 | 4.5 | 4.7 | 5 | 4.3 | 4.5 | 5 | 4.7455 | 4.6 | 4.8135 | 5.4 | 4.9 | 5.3 | 4.4 |  |  |  |  | 22
| 1 | 3 | 4.5 | 4.52 | 4.45 | 4.3504 | 4.4574 | 4.6 | 5 | 4.58 | < 5 | 7 | 4.7 | 4.6 | 4.9 | 4.5 | 5 | 4.1 | 4.4 | 7 | 4.6741 | 4.7 | 4.8511 | 5.4 | 4.6 | 5.4 | 4.5 |  |  |  |  | 16
| 1 | 4 | 4.5 | 4.41 | 4.27 | 4.3376 | 4.4913 | 4.5 | 5 | 4.5 | < 5 | 8 | 4.6 | 4.6 | 4.6 | 4.5 | 5 | 4.2 | 4.5 | 5 | 4.5787 | 4.7 | 4.8598 | 5.4 | 4.6 | 5.1 | 4.3 |  |  |  |  | 4.5248
| 1 | 5 | 5.4 | 4.47 | 4.25 | 4.2992 | 4.5277 | 4.3 | 5 | 4.58 | < 5 | 8 | 4.5 | 4.6 | 4.4 | 4.7 | 5 | 4.2 | 4.7 | 5 | 4.5435 | 4.5 | 4.8999 | 5.6 | 4.5 | 5.4 | 4.7 |  |  |  |  | 38
| 1 | 6 | 5.6 | 4.47 | 4.24 | 4.3667 | 4.57 | 4.6 | 5 | 4.57 | 5 | 8 | 4.6 | 4.8 | 4.6 | 4.6 | 5 | 4.2 | 4.4 | 5 | 4.5664 | 4.9 | 4.6915 | 5.7 | 4.7 | 5.3 | 4.5
| Mean |  | 4.8333 | 4.48 | 4.275 | 4.3369 | 4.4932 | 4.4167 | 5 | 4.5533 | 5 | 7.8333 | 4.5833 | 4.7333 | 4.5333 | 4.55 | 5 | 4.2 | 4.4833 | 5.3333 | 4.6201 | 4.7 | 4.8037 | 5.4833 | 4.6667 | 5.3167 | 4.5167
| Median |  | 4.55 | 4.485 | 4.26 | 4.344 | 4.4951 | 4.4 | 5 | 4.575 | 5 | 8 | 4.6 | 4.7 | 4.55 | 4.55 | 5 | 4.2 | 4.45 | 5 | 4.5957 | 4.7 | 4.8323 | 5.4 | 4.65 | 5.35 | 4.5
| Std Dev. |  | 0.52409 | 0.04 | 0.12029 | 0.029143 | 0.054247 | 0.17224 | 0 | 0.073121 |  | 0.40825 | 0.075277 | 0.15055 | 0.23381 | 0.15166 | 0 | 0.063246 | 0.1169 | 0.8165 | 0.076365 | 0.14142 | 0.085814 | 0.13292 | 0.13663 | 0.1169 | 0.16021
| Rel.Std.Dev. |  | 10.84% | 0.8929% | 2.814% | 0.672% | 1.207% | 3.9% | 0% | 1.606% |  | 5.212% | 1.642% | 3.181% | 5.158% | 3.333% | 0% | 1.506% | 2.608% | 15.31% | 1.653% | 3.009% | 1.786% | 2.424% | 2.928% | 2.199% | 3.547%
| PDM3 |  | 6.82% | -0.9891% | -5.52% | -4.152% | -0.6981% | -2.389% | 10.5% | 0.6317% | 10.5% | 73.12% | 1.295% | 4.61% | 0.1896% | 0.558% | 10.5% | -7.177% | -0.9154% | 17.87% | 2.108% | 3.873% | 6.165% | 21.19% | 3.136% | 17.5% | -0.1787%
| Z-Score (Absolute) |  | 1.52 | 0.39 | 1.49 | 1.16 | 0.32 | 0.73 | NA | 0.01 | NA | NA | 0.17 | 0.98 | 0.1 | 0.01 | NA | 1.9 | 0.37 | NA | 0.37 | 0.8 | 1.36 | 5.03 | 0.62 | 4.13 | 0.19
Results from laboratories 1.08, 1.12, 1.13 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 231b (Certified Value 12 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | < 1 | 12.2 | 11.7 | 9.5241 | 14.2929 | 11 | 6 | 14.9 | 12.2 | 12 | 14.9 | 14.4 | 11.9 | 12.6 | 12.4 | < 5 | 11 | 10.3 | 11.9 | 13 | 9.6906 | 13 | 9.867 | 12.4 | 16 | 13.6 | 12.7 |  |  | 1
| 1 | 2 | < 1 | 11.9 | 11.12 | 9.4918 | 13.9433 | 12 | 6 | 14 | 11.8 | 12 | 13.4 | 14.4 | 12.3 | 12.5 | 13.4 | < 5 | 12.2 | 10.2 | 10.5 | 12 | 9.5859 | 12.4 | 10.0854 | 12.5 | 16 | 12.9 | 11.5 |  |  | #N/A
| 1 | 3 | < 1 | 12.4 | 11.04 | 9.5645 | 13.5528 | 12 | 6 | 13.8 | 12.1 | 12 | 13.8 | 14.6 | 11.7 | 13 | 12.2 | < 5 | 13.1 | 10.3 | 11.3 | 13 | 9.4218 | 13 | 10.2414 | 12.5 | 14 | 12.9 | 11.9 |  |  | 16
| 1 | 4 | < 1 | 12 | 10.78 | 9.406 | 13.4364 | 11 | 6 | 13.6 | 11.9 | 11 | 15 | 14.6 | 11.9 | 12.6 | 12.4 | < 5 | 12.2 | 10.4 | 11.8 | 12 | 9.535 | 12.8 | 10.4208 | 12.4 | 15 | 12.5 | 11.8 |  |  | 11.9618
| 1 | 5 | < 1 | 12.1 | 10.76 | 8.9852 | 13.3261 | 12 | 6 | 14 | 11.9 | 12 | 16.9 | 13.2 | 11.5 | 12.4 | 12.7 | < 5 | 12.5 | 10.6 | 11.1 | 12 | 9.6866 | 12.8 | 10.5612 | 12.3 | 15 | 12.9 | 12 |  |  | 39
| 1 | 6 | < 1 | 12.5 | 10.48 | 9.358 | 13.3626 | 12 | 6 | 13.1 | 12.1 | 12 | 15.4 | 14.4 | 12 | 12.8 | 12.3 | < 5 | 10.9 | 10.6 | 11.5 | 12 | 9.6266 | 13.1 | 10.53 | 12.3 | 15 | 12.8 | 12.2
| Mean |  |  | 12.1833 | 10.98 | 9.3883 | 13.6523 | 11.6667 | 6 | 13.9 | 12 | 11.8333 | 14.9 | 14.2667 | 11.8833 | 12.65 | 12.5667 |  | 11.9833 | 10.4 | 11.35 | 12.3333 | 9.5911 | 12.85 | 10.2843 | 12.4 | 15.1667 | 12.9333 | 12.0167
| Median |  |  | 12.15 | 10.91 | 9.4489 | 13.4946 | 12 | 6 | 13.9 | 12 | 12 | 14.95 | 14.4 | 11.9 | 12.6 | 12.4 |  | 12.2 | 10.35 | 11.4 | 12 | 9.6063 | 12.9 | 10.3311 | 12.4 | 15 | 12.9 | 11.95
| Std Dev. |  |  | 0.23166 | 0.41952 | 0.2116 | 0.38519 | 0.5164 | 0 | 0.5933 | 0.15492 | 0.40825 | 1.2426 | 0.53166 | 0.27142 | 0.21679 | 0.44121 |  | 0.86583 | 0.16733 | 0.51284 | 0.5164 | 0.10211 | 0.251 | 0.27248 | 0.089443 | 0.75277 | 0.36148 | 0.40702
| Rel.Std.Dev. |  |  | 1.901% | 3.821% | 2.254% | 2.821% | 4.426% | 0% | 4.268% | 1.291% | 3.45% | 8.339% | 3.727% | 2.284% | 1.714% | 3.511% |  | 7.225% | 1.609% | 4.518% | 4.187% | 1.065% | 1.953% | 2.649% | 0.7213% | 4.963% | 2.795% | 3.387%
| PDM3 |  |  | 1.852% | -8.207% | -21.51% | 14.13% | -2.467% | -49.84% | 16.2% | 0.3198% | -1.074% | 24.56% | 19.27% | -0.6556% | 5.754% | 5.057% |  | 0.1804% | -13.06% | -5.114% | 3.106% | -19.82% | 7.426% | -14.02% | 3.664% | 26.79% | 8.122% | 0.4591%
| Z-Score (Absolute) |  | 7.95 | 0.12 | 0.71 | 1.81 | 1.14 | NA | NA | 1.31 | 0.01 | NA | 2 | 1.56 | 0.09 | 0.44 | 0.39 | 6.56 | 0.02 | 1.11 | 0.45 | NA | 1.67 | 0.58 | 1.19 | 0.27 | NA | 0.64 | 0.01
Results from laboratories 1.07, 1.08, 1.12, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 231b (Certified Value 0.33 ppm) | [30 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | 1.32 | [20 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [20 empty] | 2
3
| 1 | 1 | 0.34 | 0.32 | 0.28 | 4.2752 | 0.3 | 0.34 | 0.5 | 0.35459 | 0.31 | [20 empty] | 1
| 1 | 2 | 0.35 | 0.32 | 0.3 | 4.2478 | 0.3 | 0.34 | 0.5 | 0.3566 | 0.32 | [20 empty] | 23
| 1 | 3 | 0.35 | 0.33 | 0.31 | 4.2612 | 0.3 | 0.35 | 0.6 | 0.35332 | 0.29 | [20 empty] | 16
| 1 | 4 | 0.36 | 0.32 | 0.27 | 4.3008 | 0.4 | 0.35 | 0.5 | 0.35481 | 0.31 | [20 empty] | 0.33032
| 1 | 5 | 0.37 | 0.33 | 0.29 | 4.2887 | 0.3 | 0.34 | 0.5 | 0.33756 | 0.3 | [20 empty] | 40
| 1 | 6 | 0.38 | 0.35 | 0.29 | 4.2947 | 0.3 | 0.36 | 0.5 | 0.34857 | 0.3
| Mean |  | 0.35833 | 0.32833 | 0.29 | 4.2781 | 0.31667 | 0.34667 | 0.51667 | 0.35091 | 0.305
| Median |  | 0.355 | 0.325 | 0.29 | 4.2819 | 0.3 | 0.345 | 0.5 | 0.35396 | 0.305
| Std Dev. |  | 0.01472 | 0.01169 | 0.014142 | 0.020561 | 0.040825 | 0.008165 | 0.040825 | 0.00708 | 0.010488
| Rel.Std.Dev. |  | 4.108% | 3.561% | 4.877% | 0.4806% | 12.89% | 2.355% | 7.902% | 2.018% | 3.439%
| PDM3 |  | 8.481% | -0.601% | -12.21% | 1195% | -4.133% | 4.949% | 56.41% | 6.233% | -7.665%
| Z-Score (Absolute) |  | 0.43 | 0.67 | 2.08 | 144.6 | NA | 0 | NA | 0.16 | 1.53
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 231b (Certified Value 4.14 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4.04 | 4.077 | 4.2468 | 4.0081 | 3.64 | 4.4506 | 4.16 | 4.207 | 3.97 | 3.71 | 4.51 | 4.14 | 4.07 | 4.08 | 4.2 | 4.235 | 4.27 | 4.023 | 1.649 | 3.3449 | 4.3505 | 4.28 | 3.9526 | 4.23 | 3.13 | 4.13 | 3.9 |  |  | 1
| 1 | 2 | 3.99 | 4.0607 | 4.2238 | 4.0447 | 3.82 | 4.4687 | 4.18 | 4.1385 | 3.94 | 3.67 | 4.51 | 4.28 | 4.02 | 4.4 | 4.12 | 4.231 | 4.19 | 3.966 | 1.3 | 3.4824 | 4.355 | 4.1 | 4.0388 | 4.2 | 3.29 | 4.17 | 4.5 |  |  | #N/A
| 1 | 3 | 4.09 | 4.1333 | 4.2166 | 4.013 | 4.11 | 4.4928 | 4.17 | 4.2294 | 3.98 | 3.69 | 4.55 | 4.14 | 4.13 | 4.03 | 4.19 | 4.257 | 4.19 | 4.03 | 0.628 | 3.5057 | 4.364 | 4.25 | 4.1051 | 4.27 | 3.13 | 4.18 | 3.9 |  |  | 16
| 1 | 4 | 4 | 3.9895 | 4.219 | 4.0148 | 4.05 | 4.523 | 4.22 | 4.2123 | 3.95 | 3.66 | 4.61 | 4.14 | 4.03 | 4.08 | 4.13 | 4.269 | 4.24 | 4.238 | 1.59 | 3.4325 | 4.336 | 4.18 | 4.1868 | 4.22 | 3.16 | 4.17 | 4.3 |  |  | 4.1414
| 1 | 5 | 4.06 | 4.1368 | 4.2357 | 4.0292 | 3.81 | 4.5169 | 4.25 | 4.0997 | 3.99 | 3.66 | 4.27 | 4.03 | 4.02 | 4.2 | 4.13 | 4.289 | 4.25 | 4.247 | 1.21 | 3.3391 | 4.3725 | 4.22 | 4.2287 | 4.21 | 3.17 | 4.19 | 4.2 |  |  | 41
| 1 | 6 | 4.09 | 4.0794 | 4.273 | 4.0494 | 4.05 | 4.4868 | 4.08 | 4.1735 | 4.03 | 3.62 | 4.49 | 4.21 | 4.08 | 4.12 | 4.07 | 4.291 | 4.2 | 3.99 | 1.01 | 3.3548 | 4.3695 | 4.32 | 4.2383 | 4.25 | 3.25 | 4.19 | 4.1
| Mean |  | 4.045 | 4.0795 | 4.2358 | 4.0265 | 3.9133 | 4.4898 | 4.1767 | 4.1767 | 3.9767 | 3.6683 | 4.49 | 4.1567 | 4.0583 | 4.1517 | 4.14 | 4.262 | 4.2233 | 4.0823 | 1.2312 | 3.4099 | 4.3579 | 4.225 | 4.1251 | 4.23 | 3.1883 | 4.1717 | 4.15
| Median |  | 4.05 | 4.0782 | 4.2298 | 4.022 | 3.935 | 4.4898 | 4.175 | 4.1903 | 3.975 | 3.665 | 4.51 | 4.14 | 4.05 | 4.1 | 4.13 | 4.263 | 4.22 | 4.0265 | 1.255 | 3.3936 | 4.3595 | 4.235 | 4.146 | 4.225 | 3.165 | 4.175 | 4.15
| Std Dev. |  | 0.043243 | 0.054106 | 0.021458 | 0.017436 | 0.18446 | 0.0277 | 0.058195 | 0.049722 | 0.032042 | 0.030605 | 0.11593 | 0.083586 | 0.043551 | 0.13423 | 0.048166 | 0.025822 | 0.034448 | 0.12623 | 0.3798 | 0.073773 | 0.013621 | 0.07791 | 0.11398 | 0.026077 | 0.066458 | 0.022286 | 0.23452
| Rel.Std.Dev. |  | 1.069% | 1.326% | 0.5066% | 0.433% | 4.714% | 0.617% | 1.393% | 1.19% | 0.8057% | 0.8343% | 2.582% | 2.011% | 1.073% | 3.233% | 1.163% | 0.6059% | 0.8157% | 3.092% | 30.85% | 2.164% | 0.3126% | 1.844% | 2.763% | 0.6165% | 2.084% | 0.5342% | 5.651%
| PDM3 |  | -2.327% | -1.496% | 2.28% | -2.774% | -5.507% | 8.413% | 0.8518% | 0.8534% | -3.977% | -11.42% | 8.418% | 0.3689% | -2.006% | 0.2481% | -0.03357% | 2.912% | 1.979% | -1.426% | -70.27% | -17.66% | 5.228% | 2.019% | -0.394% | 2.14% | -23.01% | 0.7311% | 0.2079%
| Z-Score (Absolute) |  | 0.82 | 0.55 | 0.67 | 0.97 | 1.86 | 2.67 | 0.21 | 0.21 | 1.36 | 3.78 | 2.67 | 0.05 | 0.72 | 0.01 | 0.08 | 0.88 | 0.58 | 0.53 | 22.93 | 5.81 | 1.63 | 0.59 | 0.2 | 0.63 | 7.56 | 0.17 | 0
| Analytical results for Mn in OREAS 231b (Certified Value 0.158 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.17 | 0.1638 | 0.16696 | 0.15519 | 0.154 | 0.1758 | 0.157 | 0.1629 | 0.1442 | 0.0892 | 0.163 | 0.157 | 0.1595 | 0.154 | 0.1587 | 0.1606 | 0.1635 | 0.1568 | 0.106 | 0.1453 | 0.157 | 0.1565 | 0.13313 | 0.164 | 0.1523 | 0.1538 | 0.157 |  |  | 1
| 1 | 2 | 0.16 | 0.1637 | 0.16427 | 0.15663 | 0.154 | 0.1766 | 0.156 | 0.159 | 0.1449 | 0.0883 | 0.164 | 0.163 | 0.1565 | 0.1665 | 0.156 | 0.1633 | 0.1605 | 0.1547 | 0.099 | 0.15205 | 0.156 | 0.153 | 0.1354 | 0.162 | 0.1528 | 0.1544 | 0.154 |  |  | 24
| 1 | 3 | 0.17 | 0.1658 | 0.16801 | 0.15496 | 0.168 | 0.1781 | 0.156 | 0.1629 | 0.1448 | 0.0777 | 0.164 | 0.1575 | 0.1605 | 0.152 | 0.1589 | 0.161 | 0.1603 | 0.1567 | 0.089 | 0.15223 | 0.157 | 0.1555 | 0.13691 | 0.165 | 0.1527 | 0.1547 | 0.149 |  |  | 16
| 1 | 4 | 0.16 | 0.164 | 0.1672 | 0.15516 | 0.165 | 0.1789 | 0.158 | 0.162 | 0.1448 | 0.082 | 0.1655 | 0.156 | 0.1565 | 0.1545 | 0.1571 | 0.1617 | 0.1624 | 0.164 | 0.101 | 0.14781 | 0.1565 | 0.1535 | 0.13961 | 0.163 | 0.1521 | 0.1542 | 0.152 |  |  | 0.15794
| 1 | 5 | 0.16 | 0.164 | 0.16649 | 0.15528 | 0.155 | 0.1797 | 0.159 | 0.1578 | 0.1451 | 0.0853 | 0.155 | 0.154 | 0.1575 | 0.1605 | 0.1572 | 0.1634 | 0.1621 | 0.1646 | 0.096 | 0.14591 | 0.158 | 0.1555 | 0.14069 | 0.163 | 0.153 | 0.1558 | 0.156 |  |  | 42
| 1 | 6 | 0.16 | 0.164 | 0.16959 | 0.15623 | 0.164 | 0.1774 | 0.153 | 0.1655 | 0.1448 | 0.0774 | 0.163 | 0.159 | 0.159 | 0.1555 | 0.1549 | 0.1624 | 0.1602 | 0.1552 | 0.092 | 0.14425 | 0.1575 | 0.1575 | 0.1395 | 0.165 | 0.1563 | 0.1551 | 0.151
| Mean |  | 0.16333 | 0.16422 | 0.16709 | 0.15558 | 0.16 | 0.17775 | 0.1565 | 0.16168 | 0.14477 | 0.083317 | 0.16242 | 0.15775 | 0.15825 | 0.15717 | 0.15713 | 0.16207 | 0.1615 | 0.15867 | 0.097167 | 0.14792 | 0.157 | 0.15525 | 0.13754 | 0.16367 | 0.1532 | 0.15467 | 0.15317
| Median |  | 0.16 | 0.164 | 0.16708 | 0.15524 | 0.1595 | 0.17775 | 0.1565 | 0.16245 | 0.1448 | 0.08365 | 0.1635 | 0.15725 | 0.15825 | 0.155 | 0.15715 | 0.16205 | 0.1613 | 0.15675 | 0.0975 | 0.14686 | 0.157 | 0.1555 | 0.13821 | 0.1635 | 0.15275 | 0.15455 | 0.153
| Std Dev. |  | 0.005164 | 0.00078592 | 0.0017575 | 0.00068332 | 0.0063561 | 0.0014488 | 0.0020736 | 0.0028252 | 0.00030111 | 0.0051316 | 0.0037472 | 0.0030619 | 0.0016658 | 0.0053821 | 0.0015397 | 0.001169 | 0.0013638 | 0.0044442 | 0.0061779 | 0.0034618 | 0.00070711 | 0.0017248 | 0.0029151 | 0.0012111 | 0.0015543 | 0.00070899 | 0.0030605
| Rel.Std.Dev. |  | 3.162% | 0.4786% | 1.052% | 0.4392% | 3.973% | 0.8151% | 1.325% | 1.747% | 0.208% | 6.159% | 2.307% | 1.941% | 1.053% | 3.424% | 0.9799% | 0.7213% | 0.8445% | 2.801% | 6.358% | 2.34% | 0.4504% | 1.111% | 2.119% | 0.74% | 1.015% | 0.4584% | 1.998%
| PDM3 |  | 3.415% | 3.975% | 5.792% | -1.496% | 1.305% | 12.54% | -0.9113% | 2.371% | -8.34% | -47.25% | 2.835% | -0.1198% | 0.1967% | -0.4892% | -0.5103% | 2.613% | 2.254% | 0.4605% | -38.48% | -6.341% | -0.5947% | -1.703% | -12.92% | 3.626% | -3.001% | -2.072% | -3.022%
| Z-Score (Absolute) |  | 0.96 | 1.1 | 1.54 | 0.25 | 0.44 | 3.2 | 0.1 | 0.7 | 1.93 | 11.49 | 0.82 | 0.09 | 0.17 | 0 | 0.01 | 0.76 | 0.67 | 0.23 | 9.34 | 1.44 | 0.03 | 0.3 | 3.05 | 1.01 | 0.62 | 0.39 | 0.62
| Analytical results for Mo in OREAS 231b (Certified Value 1.46 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.38 | 1.4 | 1.54 | 1.3757 | 1.5151 | < 1 | 1.92 | 1.4 | 1.5 | < 2 | 1.8 | 1.49 | 1.61 | 1.41 | 1.39 | < 2 | 1.1 | 1.2 | 1.5 | 1.92 | 2 | 1.4136 | 1.49 | 1.439 | 1.46 | 1.8 | 1.59 | 1.42 |  | 1
| 1 | 2 | 1.73 | 1.5 | 2.14 | 1.4193 | 1.4585 | 1 | 1.74 | 1.5 | 1.4 | < 2 | 1.5 | 1.52 | 1.57 | 1.49 | 1.5 | < 2 | 1.4 | 1.3 | 1.37 | 1.35 | 2 | 1.506 | 1.47 | 1.4895 | 1.58 | 1.86 | 1.57 | 1.38 |  | 25
| 1 | 3 | 1.78 | 1.4 | 1.38 | 1.3651 | 1.4184 | 1 | 1.71 | 1.4 | 1.6 | < 2 | 1.5 | 1.5 | 1.49 | 1.44 | 1.43 | < 2 | 1.3 | 1.3 | 1.46 | 2.19 | 2 | 1.4496 | 1.54 | 1.5087 | 1.46 | 1.67 | 1.58 | 1.34 |  | 16
| 1 | 4 | 1.75 | 1.4 | 0.91 | 1.3416 | 1.4745 | < 1 | 1.8 | 1.5 | 1.6 | < 2 | 1.5 | 1.5 | 1.52 | 1.44 | 1.41 | < 2 | 1.2 | 1.2 | 1.41 | 1.98 | 1 | 1.4529 | 1.51 | 1.5515 | 1.63 | 1.8 | 1.52 | 1.34 |  | 1.4618
| 1 | 5 | 1.68 | 1.5 | 1.61 | 1.4028 | 1.451 | < 1 | 1.79 | 1.4 | 1.4 | < 2 | 1.3 | 1.42 | 1.51 | 1.34 | 1.46 | < 2 | 1 | 1.3 | 1.43 | 1.78 | 2 | 1.4689 | 1.52 | 1.5381 | 1.56 | 1.73 | 1.52 | 1.38 |  | 43
| 1 | 6 | 1.66 | 1.5 | 2.27 | 1.3695 | 1.4433 | < 1 | 1.64 | 1.5 | 1.6 | < 2 | 1.5 | 1.5 | 1.61 | 1.42 | 1.43 | < 2 | 1.1 | 1.2 | 1.47 | 1.4 | 1 | 1.4596 | 1.51 | 1.4879 | 1.62 | 1.76 | 1.59 | 1.32
| Mean |  | 1.83 | 1.45 | 1.6417 | 1.379 | 1.4601 | 1 | 1.7667 | 1.45 | 1.5167 |  | 1.5167 | 1.4883 | 1.5517 | 1.4233 | 1.4367 |  | 1.1833 | 1.25 | 1.44 | 1.77 | 1.6667 | 1.4585 | 1.5067 | 1.5024 | 1.5517 | 1.77 | 1.5617 | 1.3633
| Median |  | 1.74 | 1.45 | 1.575 | 1.3726 | 1.4547 | 1 | 1.765 | 1.45 | 1.55 |  | 1.5 | 1.5 | 1.545 | 1.43 | 1.43 |  | 1.15 | 1.25 | 1.445 | 1.85 | 2 | 1.4563 | 1.51 | 1.4991 | 1.57 | 1.78 | 1.575 | 1.36
| Std Dev. |  | 0.27306 | 0.054772 | 0.50165 | 0.027863 | 0.032682 | 0 | 0.095009 | 0.054772 | 0.098319 |  | 0.16021 | 0.034881 | 0.052313 | 0.049261 | 0.038816 |  | 0.1472 | 0.054772 | 0.046476 | 0.33359 | 0.5164 | 0.029982 | 0.024221 | 0.040302 | 0.075476 | 0.065727 | 0.033116 | 0.036697
| Rel.Std.Dev. |  | 14.92% | 3.777% | 30.56% | 2.021% | 2.238% | 0% | 5.378% | 3.777% | 6.483% |  | 10.56% | 2.344% | 3.371% | 3.461% | 2.702% |  | 12.44% | 4.382% | 3.227% | 18.85% | 30.98% | 2.056% | 1.608% | 2.682% | 4.864% | 3.713% | 2.121% | 2.692%
| PDM3 |  | 25.19% | -0.8055% | 12.31% | -5.662% | -0.1127% | -31.59% | 20.86% | -0.8055% | 3.755% |  | 3.755% | 1.817% | 6.15% | -2.63% | -1.718% |  | -19.05% | -14.49% | -1.49% | 21.09% | 14.02% | -0.2266% | 3.071% | 2.783% | 6.15% | 21.09% | 6.834% | -6.734%
| Z-Score (Absolute) |  | 3.1 | 0.21 | 1.46 | 0.83 | 0.12 | NA | 2.55 | 0.21 | 0.37 | 4.13 | 0.37 | 0.12 | 0.67 | 0.44 | 0.33 | 4.13 | 2.53 | 1.95 | 0.3 | 2.58 | NA | 0.14 | 0.28 | 0.25 | 0.67 | 2.58 | 0.76 | 0.97
Results from laboratories 1.07 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 231b (Certified Value 2.04 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.08 | 2.0519 | 1.9207 | 1.9162 | 2.025 | 1.82 | 2.263 | 2 | 2.0169 | 1.96 | 2.15 | 2.17 | 2.01 | 2.02 | 1.96 | 2.1 | 2.057 | 2.016 | 1.964 | 1.946 | 2.0016 | 2.105 | 2.1 | 2.44 | 2.08 | 2.48 | 2.16 | 2.15 |  | 1
| 1 | 2 | 2.06 | 2.0525 | 1.9066 | 1.8892 | 2.0431 | 1.87 | 2.233 | 2 | 1.976 | 1.95 | 2.13 | 2.17 | 2.07 | 2.01 | 2.11 | 2.06 | 2.069 | 2.003 | 1.946 | 1.759 | 2.0854 | 2.09 | 2.01 | 2.54 | 2.07 | 2.49 | 2.18 | 2.06 |  | #N/A
| 1 | 3 | 2.11 | 2.1116 | 1.8576 | 1.9347 | 2.0284 | 2.06 | 2.255 | 1.99 | 2.0075 | 1.96 | 2.12 | 2.18 | 2.02 | 2.06 | 1.94 | 2.1 | 2.062 | 1.989 | 1.926 | 1.807 | 2.0818 | 2.1 | 2.1 | 2.559 | 2.11 | 2.31 | 2.19 | 2.04 |  | 16
| 1 | 4 | 2.09 | 2.0289 | 2.0081 | 1.9249 | 2.0249 | 2 | 2.255 | 2.04 | 1.9946 | 1.94 | 2.15 | 2.21 | 2.03 | 2 | 1.97 | 2.05 | 2.06 | 2.018 | 2.003 | 2.021 | 2.0311 | 2.09 | 2.07 | 2.7 | 2.09 | 2.43 | 2.18 | 2.06 |  | 2.0387
| 1 | 5 | 2.09 | 2.1122 | 1.979 | 1.9221 | 2.0252 | 1.89 | 2.285 | 2.02 | 1.9445 | 1.96 | 2.13 | 2.06 | 1.97 | 2.01 | 2.03 | 2.06 | 2.068 | 2.024 | 2.036 | 2.061 | 1.985 | 2.115 | 2.1 | 2.691 | 2.08 | 2.42 | 2.2 | 2.12 |  | 44
| 1 | 6 | 2.11 | 2.08 | 2.0006 | 1.9466 | 2.0465 | 2.04 | 2.263 | 1.95 | 1.9977 | 2.01 | 2.1 | 2.15 | 2.08 | 2.02 | 1.98 | 2.05 | 2.055 | 2 | 1.943 | 2.171 | 1.9541 | 2.11 | 2.13 | 2.694 | 2.09 | 2.5 | 2.19 | 2.06
| Mean |  | 2.09 | 2.0729 | 1.9454 | 1.9223 | 2.0322 | 1.9467 | 2.259 | 2 | 1.9895 | 1.9633 | 2.13 | 2.1567 | 2.03 | 2.02 | 1.9983 | 2.07 | 2.0618 | 2.0083 | 1.9697 | 1.9608 | 2.0232 | 2.1017 | 2.085 | 2.604 | 2.0867 | 2.4383 | 2.1833 | 2.0817
| Median |  | 2.09 | 2.0663 | 1.9499 | 1.9235 | 2.0268 | 1.945 | 2.259 | 2 | 1.9962 | 1.96 | 2.13 | 2.17 | 2.025 | 2.015 | 1.975 | 2.06 | 2.061 | 2.0095 | 1.955 | 1.9835 | 2.0164 | 2.1025 | 2.1 | 2.625 | 2.085 | 2.455 | 2.185 | 2.06
| Std Dev. |  | 0.018974 | 0.03431 | 0.059888 | 0.019411 | 0.0099313 | 0.099532 | 0.016829 | 0.030332 | 0.025982 | 0.024221 | 0.018974 | 0.051251 | 0.040497 | 0.020976 | 0.062423 | 0.023664 | 0.0057067 | 0.013186 | 0.041764 | 0.15649 | 0.053033 | 0.010328 | 0.041352 | 0.10761 | 0.013663 | 0.070828 | 0.013663 | 0.043089
| Rel.Std.Dev. |  | 0.9078% | 1.655% | 3.078% | 1.01% | 0.4887% | 5.113% | 0.745% | 1.517% | 1.306% | 1.234% | 0.8908% | 2.376% | 1.995% | 1.038% | 3.124% | 1.143% | 0.2768% | 0.6566% | 2.12% | 7.981% | 2.621% | 0.4914% | 1.983% | 4.133% | 0.6548% | 2.905% | 0.6258% | 2.07%
| PDM3 |  | 2.515% | 1.674% | -4.576% | -5.712% | -0.3206% | -4.516% | 10.8% | -1.9% | -2.413% | -3.698% | 4.477% | 5.785% | -0.428% | -0.9185% | -1.981% | 1.534% | 1.133% | -1.491% | -3.387% | -3.821% | -0.7634% | 3.087% | 2.27% | 27.73% | 2.351% | 19.6% | 7.093% | 2.106%
| Z-Score (Absolute) |  | 0.56 | 0.34 | 1.33 | 1.63 | 0.19 | 1.31 | 2.77 | 0.61 | 0.75 | 1.09 | 1.08 | 1.43 | 0.22 | 0.35 | 0.64 | 0.3 | 0.19 | 0.5 | 1.01 | 1.12 | 0.31 | 0.71 | 0.5 | 7.27 | 0.52 | 5.11 | 1.78 | 0.45
| Analytical results for Nb in OREAS 231b (Certified Value 3.66 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4.4 | 3.78 | 21.92 | 2.7194 | 4.0395 | 0.6 | 4 | 3.8 | 3.56 | < 5 | 4 | 3.7 | 3.5 | 3.4 | 3.3 | 26 | 3.3 | 3.1 | 3.7 | 6 | 2.9653 | 3.7 | 5.4362 | 3.9 | 4.3 | 3.9 | 3.4 |  |  | 1
| 1 | 2 | 4.4 | 3.76 | 21.53 | 2.6597 | 3.8006 | 3.1 | 3.9 | 3.6 | 3.41 | < 5 | 3.8 | 3.7 | 3.6 | 3.5 | 3.6 | 25 | 3.6 | 3 | 3.6 | 6 | 2.9908 | 3.5 | 5.484 | 4.1 | 4.1 | 4.3 | 3.3 |  |  | 8
| 1 | 3 | 4.6 | 3.78 | 20.97 | 2.5609 | 3.7213 | 3.1 | 3.9 | 3.6 | 3.51 | < 5 | 3.8 | 3.8 | 3.5 | 3.7 | 3.4 | 26 | 3.8 | 3 | 3.8 | 5 | 2.9328 | 3.8 | 5.5128 | 4.2 | 3.8 | 4.2 | 3.3 |  |  | 16
| 1 | 4 | 4.5 | 3.74 | 20.07 | 2.6572 | 3.6997 | 0.1 | 4 | 3.7 | 3.54 | < 5 | 3.9 | 3.7 | 3.5 | 3.5 | 3.3 | 25 | 3.5 | 3 | 3.9 | 5 | 2.9903 | 3.7 | 5.5974 | 4 | 3.8 | 4 | 3.2 |  |  | 3.6631
| 1 | 5 | 3.6 | 3.75 | 22.03 | 2.434 | 3.7611 | 0.1 | 3.8 | 3.8 | 3.48 | < 5 | 3.7 | 3.6 | 3.4 | 3.4 | 3.5 | 26 | 3.4 | 3.1 | 3.6 | 5 | 2.9768 | 3.7 | 5.5744 | 4.2 | 3.8 | 4.1 | 3.4 |  |  | 45
| 1 | 6 | 3.8 | 3.82 | 19.87 | 2.5429 | 3.6708 | 2.2 | 3.8 | 3.6 | 3.71 | < 5 | 3.8 | 3.8 | 3.5 | 3.6 | 3.4 | 25 | 3.6 | 3.1 | 3.6 | 5 | 2.9904 | 3.9 | 5.3656 | 4.2 | 3.9 | 4.1 | 3.3
| Mean |  | 4.2167 | 3.7717 | 21.065 | 2.5957 | 3.7822 | 1.5333 | 3.9 | 3.6833 | 3.535 |  | 3.8333 | 3.7167 | 3.5 | 3.5167 | 3.4167 | 25.5 | 3.5333 | 3.05 | 3.7 | 5.3333 | 2.9744 | 3.7167 | 5.4951 | 4.1 | 3.95 | 4.1 | 3.3167
| Median |  | 4.4 | 3.77 | 21.25 | 2.6091 | 3.7412 | 1.4 | 3.9 | 3.65 | 3.525 |  | 3.8 | 3.7 | 3.5 | 3.5 | 3.4 | 25.5 | 3.55 | 3.05 | 3.65 | 5 | 2.9836 | 3.7 | 5.4984 | 4.15 | 3.85 | 4.1 | 3.3
| Std Dev. |  | 0.41191 | 0.028577 | 0.92796 | 0.10326 | 0.13412 | 1.4376 | 0.089443 | 0.098319 | 0.10055 |  | 0.10328 | 0.075277 | 0.063246 | 0.1169 | 0.1169 | 0.54772 | 0.17512 | 0.054772 | 0.12649 | 0.5164 | 0.022773 | 0.13292 | 0.086521 | 0.12649 | 0.20736 | 0.14142 | 0.075277
| Rel.Std.Dev. |  | 9.769% | 0.7577% | 4.405% | 3.978% | 3.546% | 93.76% | 2.293% | 2.669% | 2.844% |  | 2.694% | 2.025% | 1.807% | 3.324% | 3.422% | 2.148% | 4.956% | 1.796% | 3.419% | 9.682% | 0.7656% | 3.576% | 1.575% | 3.085% | 5.25% | 3.449% | 2.27%
| PDM3 |  | 15.11% | 2.964% | 475.1% | -29.14% | 3.251% | -58.14% | 6.468% | 0.5527% | -3.497% |  | 4.648% | 1.463% | -4.452% | -3.997% | -6.727% | 596.1% | -3.542% | -16.74% | 1.008% | 45.6% | -18.8% | 1.463% | 50.01% | 11.93% | 7.833% | 11.93% | -9.457%
| Z-Score (Absolute) |  | 1.29 | 0.16 | 43.89 | 2.81 | 0.19 | 5.5 | 0.48 | 0.06 | 0.44 | 3.06 | 0.32 | 0.02 | 0.53 | 0.48 | 0.74 | 55.1 | 0.44 | 1.66 | 0.02 | NA | 1.86 | 0.02 | 4.52 | 0.99 | 0.61 | 0.99 | 0.99
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 231b (Certified Value 7.98 ppm) | [30 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 8.06 | 4.61 | 7.6911 | 7 | 8.4 | 7.8 | 8 | 8.3288 | [21 empty] | 1
| 1 | 2 | 8.07 | 5.06 | 7.6302 | 7.3 | 8.5 | 7.9 | 8 | 8.421 | [21 empty] | 27
| 1 | 3 | 8.12 | 5.24 | 7.6292 | 7.9 | 8.3 | 7.8 | 9 | 8.4613 | [21 empty] | 16
| 1 | 4 | 8.07 | 4.55 | 7.7077 | 7.8 | 8.4 | 7.8 | 8 | 8.0855 | [21 empty] | 7.9799
| 1 | 5 | 8.06 | 4.65 | 7.6955 | 7.4 | 8.2 | 7.9 | 8 | 8.0745 | [21 empty] | 46
| 1 | 6 | 8.07 | 4.57 | 7.7171 | 8.1 | 7.9 | 7.8 | 8 | 8.3642
| Mean |  | 8.075 | 4.78 | 7.6784 | 7.5833 | 8.2833 | 7.8333 | 8.1667 | 8.2892
| Median |  | 8.07 | 4.63 | 7.6933 | 7.6 | 8.35 | 7.8 | 8 | 8.3465
| Std Dev. |  | 0.022583 | 0.29421 | 0.038865 | 0.41673 | 0.2137 | 0.05164 | 0.40825 | 0.16839
| Rel.Std.Dev. |  | 0.2797% | 6.155% | 0.5062% | 5.495% | 2.58% | 0.6592% | 4.999% | 2.031%
| PDM3 |  | 1.192% | -40.1% | -3.778% | -4.97% | 3.802% | -1.837% | 2.34% | 3.876%
| Z-Score (Absolute) |  | 0.27 | 7.08 | 0.61 | 0.83 | 0.73 | 0.27 | 0.47 | 0.75
| Analytical results for Ni in OREAS 231b (Certified Value 105 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 103.2 | 106.5 | 99.31 | 104.28 | 110.8162 | 99 | 125 | 105 | 107.4 | 96 | 104.1 | 115.5 | 107.5 | 102 | 102 | 95 | 105 | 103.2 | 101 | 90.27 | 100.2726 | 105 | 107.0992 | 109.5 | 99.9157 | 96.9 | 108.1 | 108.4 | 111 | 1
| 1 | 2 | 102.7 | 106.9 | 99.79 | 104.2 | 111.4491 | 102 | 124 | 104 | 104.9 | 96 | 100.5 | 116 | 112.5 | 106.5 | 109 | 94 | 104 | 102.1 | 101 | 89.22 | 100.8678 | 105 | 112.1235 | 101 | 100.8956 | 97.5 | 114.2 | 109.2 | 108 | 28
| 1 | 3 | 105.2 | 111.1 | 99.39 | 107.9 | 110.4407 | 101 | 122 | 104 | 108.4 | 98 | 101.8 | 117 | 105 | 108.5 | 103 | 96 | 104 | 103.2 | 101 | 92.98 | 101.8536 | 105 | 106.845 | 110 | 102.2649 | 98.9 | 107.3 | 110.2 | 106 | 16
| 1 | 4 | 102.8 | 108.2 | 99.57 | 107.41 | 110.3161 | 102 | 124 | 107 | 107.9 | 95 | 102.9 | 116.5 | 105.5 | 105 | 103 | 94 | 104 | 101.6 | 106 | 88.44 | 101.9001 | 105 | 105.9934 | 107 | 104.7668 | 96.1 | 109.8 | 108.4 | 108 | 104.8458
| 1 | 5 | 105.2 | 111.1 | 99.85 | 106.59 | 110.7236 | 102 | 120 | 106 | 104.9 | 97 | 102.5 | 111 | 103 | 103.5 | 107.5 | 95 | 104 | 103.5 | 106 | 88.17 | 99.0915 | 105 | 110.2381 | 109.5 | 105.4924 | 97.8 | 106.9 | 109.6 | 111 | 47
| 1 | 6 | 105.4 | 109.2 | 102.07 | 108.45 | 111.6037 | 103 | 119 | 102 | 108.6 | 98 | 102.3 | 117 | 107.5 | 109 | 105 | 93 | 104 | 102.4 | 103 | 87.49 | 95.5854 | 105 | 105.7757 | 110.5 | 276.3006 | 94.5 | 111.5 | 109.8 | 107
| Mean |  | 104.0833 | 108.8333 | 99.9967 | 106.4717 | 110.8916 | 101.5 | 122.3333 | 104.6667 | 107.0167 | 96.6667 | 102.35 | 115.5 | 106.8333 | 105.75 | 104.9167 | 94.5 | 104.1667 | 102.6667 | 103 | 89.4283 | 99.9285 | 105 | 108.0125 | 107.9167 | 131.606 | 96.95 | 109.6333 | 109.2667 | 108.5
| Median |  | 104.2 | 108.7 | 99.68 | 107 | 110.7699 | 102 | 123 | 104.5 | 107.65 | 96.5 | 102.4 | 116.25 | 106.5 | 105.75 | 104 | 94.5 | 104 | 102.8 | 102 | 88.83 | 100.5702 | 105 | 106.9721 | 109.5 | 103.5159 | 97.2 | 108.95 | 109.4 | 108
| Std Dev. |  | 1.3091 | 1.9997 | 1.0377 | 1.8337 | 0.52656 | 1.3784 | 2.4221 | 1.7512 | 1.6916 | 1.2111 | 1.1929 | 2.2804 | 3.2506 | 2.7704 | 2.8003 | 1.0488 | 0.40825 | 0.74744 | 2.4495 | 1.9841 | 2.3731 | 0 | 2.5741 | 3.5975 | 70.9184 | 1.5202 | 2.8154 | 0.74476 | 2.0736
| Rel.Std.Dev. |  | 1.258% | 1.837% | 1.038% | 1.722% | 0.4748% | 1.358% | 1.98% | 1.673% | 1.581% | 1.253% | 1.166% | 1.974% | 3.043% | 2.62% | 2.669% | 1.11% | 0.3919% | 0.728% | 2.378% | 2.219% | 2.375% | 0% | 2.383% | 3.334% | 53.89% | 1.568% | 2.568% | 0.6816% | 1.911%
| PDM3 |  | -0.7272% | 3.803% | -4.625% | 1.551% | 5.766% | -3.191% | 16.68% | -0.1709% | 2.071% | -7.801% | -2.38% | 10.16% | 1.896% | 0.8624% | 0.06759% | -9.868% | -0.6477% | -2.078% | -1.76% | -14.7% | -4.69% | 0.1471% | 3.02% | 2.929% | 25.52% | -7.531% | 4.566% | 4.217% | 3.485%
| Z-Score (Absolute) |  | 0.18 | 0.74 | 0.96 | 0.28 | 1.14 | 0.67 | 3.34 | 0.06 | 0.39 | 1.61 | 0.51 | 2.02 | 0.35 | 0.14 | 0.02 | 2.02 | 0.16 | 0.45 | 0.39 | 3 | 0.98 | 0 | 0.58 | 0.56 | 5.13 | 1.55 | 0.89 | 0.82 | 0.67
| Analytical results for P in OREAS 231b (Certified Value 0.044 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.0437 | 0.0472 | 0.042093 | 0.042995 | 0.028 | 0.0537 | 0.046 | 0.0432 | 0.047 | 0.045 | 0.048 | 0.046 | 0.044 | 0.043 | 0.0426 | 0.045 | 0.043 | 0.044 | 0.0415 | 0.046 | 0.040944 | 0.0427 | 0.0455 | 0.0431 | 0.042 |  |  |  |  | 1
| 1 | 2 | 0.0437 | 0.0433 | 0.041859 | 0.043324 | 0.042 | 0.0524 | 0.046 | 0.0433 | 0.0477 | 0.042 | 0.047 | 0.047 | 0.043 | 0.047 | 0.0423 | 0.045 | 0.041 | 0.044 | 0.043 | 0.045 | 0.041547 | 0.0443 | 0.0425 | 0.043 | 0.0418 |  |  |  |  | #N/A
| 1 | 3 | 0.0446 | 0.0463 | 0.041767 | 0.04304 | 0.04 | 0.0532 | 0.046 | 0.0445 | 0.0473 | 0.043 | 0.047 | 0.045 | 0.045 | 0.043 | 0.043 | 0.045 | 0.044 | 0.044 | 0.0435 | 0.046 | 0.041729 | 0.0446 | 0.0419 | 0.0434 | 0.0403 |  |  |  |  | 16
| 1 | 4 | 0.0439 | 0.0455 | 0.041729 | 0.043005 | 0.036 | 0.0532 | 0.046 | 0.0417 | 0.0474 | 0.043 | 0.048 | 0.045 | 0.044 | 0.043 | 0.0421 | 0.045 | 0.043 | 0.047 | 0.043 | 0.045 | 0.043297 | 0.043 | 0.0429 | 0.0433 | 0.0413 |  |  |  |  | 0.043809
| 1 | 5 | 0.0436 | 0.046 | 0.042072 | 0.042937 | 0.036 | 0.0537 | 0.047 | 0.044 | 0.0477 | 0.042 | 0.045 | 0.044 | 0.044 | 0.044 | 0.042 | 0.045 | 0.043 | 0.047 | 0.044 | 0.045 | 0.042962 | 0.0444 | 0.0423 | 0.0437 | 0.0419 |  |  |  |  | 48
| 1 | 6 | 0.0441 | 0.0477 | 0.042135 | 0.043455 | 0.041 | 0.0532 | 0.045 | 0.0419 | 0.0475 | 0.043 | 0.047 | 0.046 | 0.044 | 0.042 | 0.0413 | 0.045 | 0.042 | 0.044 | 0.0425 | 0.046 | 0.042983 | 0.0436 | 0.0445 | 0.0434 | 0.0405
| Mean |  | 0.043933 | 0.046 | 0.041942 | 0.043126 | 0.037167 | 0.053233 | 0.046 | 0.0431 | 0.047433 | 0.043 | 0.047 | 0.0455 | 0.044 | 0.043667 | 0.042217 | 0.045 | 0.042667 | 0.045 | 0.042917 | 0.0455 | 0.042244 | 0.043767 | 0.043267 | 0.043317 | 0.0413
| Median |  | 0.0438 | 0.04615 | 0.041966 | 0.043022 | 0.038 | 0.0532 | 0.046 | 0.04325 | 0.04745 | 0.043 | 0.047 | 0.0455 | 0.044 | 0.043 | 0.0422 | 0.045 | 0.043 | 0.044 | 0.043 | 0.0455 | 0.042346 | 0.04395 | 0.0427 | 0.04335 | 0.04155
| Std Dev. |  | 0.00037238 | 0.0015466 | 0.00017879 | 0.00021079 | 0.0051543 | 0.0004761 | 0.00063246 | 0.0011153 | 0.00026583 | 0.0010954 | 0.0010954 | 0.0010488 | 0.00063246 | 0.0017512 | 0.00057764 | 7.6012e-18 | 0.0010328 | 0.0015492 | 0.0008612 | 0.00054772 | 0.00096053 | 0.00079162 | 0.0014166 | 0.00024833 | 0.00074027
| Rel.Std.Dev. |  | 0.8476% | 3.362% | 0.4263% | 0.4888% | 13.87% | 0.8944% | 1.375% | 2.588% | 0.5604% | 2.548% | 2.331% | 2.305% | 1.437% | 4.01% | 1.368% | 1.689e-14% | 2.421% | 3.443% | 2.007% | 1.204% | 2.274% | 1.809% | 3.274% | 0.5733% | 1.792%
| PDM3 |  | 0.2841% | 5.002% | -4.26% | -1.559% | -15.16% | 21.51% | 5.002% | -1.618% | 8.273% | -1.846% | 7.284% | 3.86% | 0.4362% | -0.3246% | -3.634% | 2.719% | -2.607% | 2.719% | -2.037% | 3.86% | -3.573% | -0.09638% | -1.238% | -1.124% | -5.727%
| Z-Score (Absolute) |  | 0.13 | 1.18 | 0.87 | 0.27 | 3.29 | 4.84 | 1.18 | 0.29 | 1.9 | 0.34 | 1.69 | 0.93 | 0.17 | 0 | 0.73 | 0.67 | 0.51 | 0.67 | 0.38 | 0.93 | 0.72 | 0.05 | 0.2 | 0.18 | 1.2
| Analytical results for Pb in OREAS 231b (Certified Value 13.6 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 12.8 | 14.1 | 11.9 | 11.9989 | 14.9687 | 10 | 15.8 | 15 | 13.9 | 15 | 11.6 | 12.7 | 14.6 | 13 | 12.7 | 7 | 13 | 12.8 | 15 | 22.26 | 14.121 | 14 | 14.1409 | 13.7 | 39.6684 | 13.8 | 14.3 | 15.6 | 13.7 | 1
| 1 | 2 | 12.7 | 14 | 12.6 | 12.1832 | 14.5917 | 10 | 15.8 | 14 | 13.3 | 15 | 10.4 | 13.4 | 15.3 | 13.6 | 14.2 | 7 | 14 | 12.9 | 14.5 | 17.17 | 11.8064 | 14 | 13.8998 | 15 | 46.1979 | 13.7 | 14.6 | 14.9 | 12.8 | 29
| 1 | 3 | 13.2 | 14.3 | 11.5 | 12.6461 | 14.937 | 8 | 15.7 | 15 | 13.7 | 16 | 11.4 | 13.1 | 14.2 | 13.6 | 13.3 | 7 | 15 | 12.7 | 14.6 | 21.54 | 12.2106 | 23 | 13.8966 | 13.7 | 42.7221 | 13.3 | 14.1 | 15 | 12.8 | 16
| 1 | 4 | 13.2 | 13.9 | 10.8 | 12.0289 | 14.6296 | 10 | 15.9 | 15 | 13.5 | 15 | 10.9 | 12.9 | 14.9 | 13.4 | 13.4 | 8 | 14 | 12.7 | 14.7 | 16.85 | 11.5338 | 13 | 13.9129 | 13.8 | 49.4878 | 13.9 | 13.8 | 14.6 | 12.7 | 13.5667
| 1 | 5 | 12.8 | 14.5 | 11.63 | 12.8049 | 14.7189 | 10 | 15.4 | 15 | 13.5 | 15 | 11.4 | 12.8 | 13.5 | 12.7 | 13.6 | 7 | 13 | 12.9 | 14.5 | 23.46 | 14.0906 | 14 | 13.7601 | 13.8 | 43.9119 | 13.9 | 13.5 | 15.3 | 13 | 49
| 1 | 6 | 13.4 | 14.6 | 12.26 | 12.8543 | 14.868 | 9 | 14.8 | 14 | 13.6 | 15 | 10.6 | 13.3 | 15.1 | 13.8 | 13.6 | 7 | 13 | 12.8 | 14.7 | 16.87 | 11.8252 | 14 | 13.9915 | 13.8 | 48.6964 | 13.8 | 14.7 | 15 | 13.3
| Mean |  | 13.0167 | 14.2333 | 11.7817 | 12.4194 | 14.7857 | 9.5 | 15.5667 | 14.6667 | 13.5833 | 15.1667 | 11.05 | 13.0333 | 14.6 | 13.35 | 13.4667 | 7.1667 | 13.6667 | 12.8 | 14.6667 | 19.6917 | 12.5979 | 15.3333 | 13.9336 | 13.9667 | 45.1141 | 13.7333 | 14.1667 | 15.0667 | 13.05
| Median |  | 13 | 14.2 | 11.765 | 12.4146 | 14.7935 | 10 | 15.75 | 15 | 13.55 | 15 | 11.15 | 13 | 14.75 | 13.5 | 13.5 | 7 | 13.5 | 12.8 | 14.65 | 19.355 | 12.0179 | 14 | 13.9064 | 13.8 | 45.0549 | 13.8 | 14.2 | 15 | 12.9
| Std Dev. |  | 0.28577 | 0.28048 | 0.62904 | 0.39351 | 0.161 | 0.83666 | 0.41312 | 0.5164 | 0.20412 | 0.40825 | 0.48888 | 0.28048 | 0.66332 | 0.41833 | 0.48854 | 0.40825 | 0.8165 | 0.089443 | 0.18619 | 3.0532 | 1.1878 | 3.7771 | 0.12597 | 0.50859 | 3.7406 | 0.22509 | 0.46332 | 0.34448 | 0.38341
| Rel.Std.Dev. |  | 2.195% | 1.971% | 5.339% | 3.168% | 1.089% | 8.807% | 2.654% | 3.521% | 1.503% | 2.692% | 4.424% | 2.152% | 4.543% | 3.134% | 3.628% | 5.696% | 5.974% | 0.6988% | 1.269% | 15.5% | 9.428% | 24.63% | 0.9041% | 3.641% | 8.291% | 1.639% | 3.271% | 2.286% | 2.938%
| PDM3 |  | -4.055% | 4.913% | -13.16% | -8.457% | 8.985% | -29.98% | 14.74% | 8.107% | 0.1223% | 11.79% | -18.55% | -3.932% | 7.616% | -1.598% | -0.7377% | -47.17% | 0.7365% | -5.652% | 8.107% | 45.15% | -7.141% | 13.02% | 2.704% | 2.948% | 232.5% | 1.228% | 4.422% | 11.06% | -3.809%
| Z-Score (Absolute) |  | 0.56 | 0.39 | 1.52 | 1.02 | 0.82 | NA | 1.43 | NA | 0.12 | NA | 2.09 | 0.54 | 0.67 | 0.3 | 0.21 | NA | NA | 0.73 | 0.73 | 4.64 | 0.88 | NA | 0.16 | 0.18 | 24.42 | 0 | 0.34 | 1.04 | 0.53
Results from laboratories 1.07, 1.10, 1.12, 1.19, 1.22 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 231b (Certified Value 1.64 ppm) | [30 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 1.63 | 1.47 | 1.5228 | 1.5 | 1.9 | 1.67 | 1.8 | 1.663 | [21 empty] | 1
| 1 | 2 | 1.61 | 1.54 | 1.5198 | 1.6 | 1.8 | 1.62 | 1.6 | 1.7382 | [21 empty] | 30
| 1 | 3 | 1.69 | 1.58 | 1.5258 | 1.7 | 1.8 | 1.63 | 1.9 | 1.6896 | [21 empty] | 16
| 1 | 4 | 1.66 | 1.5 | 1.5203 | 1.7 | 1.8 | 1.62 | 1.7 | 1.6791 | [21 empty] | 1.642
| 1 | 5 | 1.71 | 1.52 | 1.536 | 1.5 | 1.8 | 1.64 | 1.5 | 1.6617 | [21 empty] | 50
| 1 | 6 | 1.74 | 1.51 | 1.5663 | 1.6 | 1.7 | 1.65 | 1.6 | 1.7037
| Mean |  | 1.6733 | 1.52 | 1.5318 | 1.6 | 1.8 | 1.6383 | 1.6833 | 1.6892
| Median |  | 1.675 | 1.515 | 1.5243 | 1.6 | 1.8 | 1.635 | 1.65 | 1.6844
| Std Dev. |  | 0.049261 | 0.037417 | 0.017898 | 0.089443 | 0.063246 | 0.019408 | 0.1472 | 0.028828
| Rel.Std.Dev. |  | 2.944% | 2.462% | 1.168% | 5.59% | 3.514% | 1.185% | 8.744% | 1.707%
| PDM3 |  | 1.908% | -7.43% | -6.71% | -2.558% | 9.622% | -0.2233% | 2.517% | 2.874%
| Z-Score (Absolute) |  | 0.26 | 2.05 | 1.88 | 0.84 | 2.18 | 0.26 | 0.42 | 0.5
| Analytical results for Rb in OREAS 231b (Certified Value 10.4 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 10.6 | 10.09 | 9.6877 | 9.4397 | 9.6 | 10.7 | 11.3 | 10.45 | 10 | 10.8 | 10.4 | 9.2 | 9.9 | 10.3 | 10 | 10.7 | 10.5 | 10.6517 | 11 | 11.2398 | 11.4 | 10.3 | 12.1 | 9.7 | [5 empty] | 1
| 1 | 2 | 10.3 | 10.09 | 9.6578 | 9.473 | 10 | 10.4 | 10.6 | 10.12 | 9.7 | 11.1 | 10.8 | 10 | 10.5 | 10.9 | 10.1 | 9.9 | 10.5 | 10.9453 | 10 | 11.3446 | 11.5 | 10.6 | 11.9 | 9.4 | [5 empty] | 31
| 1 | 3 | 11.2 | 10.09 | 9.7551 | 9.4614 | 10.5 | 10.2 | 10.6 | 10.15 | 10.5 | 10.8 | 10.3 | 9.9 | 10 | 11 | 10.2 | 10.4 | 11 | 10.6465 | 10.9 | 11.5394 | 11.3 | 10 | 12 | 9.5 | [5 empty] | 16
| 1 | 4 | 10.2 | 10.32 | 9.7422 | 9.3819 | 10.4 | 10.6 | 10.8 | 10.62 | 10.2 | 10.7 | 10.1 | 9.6 | 10.1 | 10.5 | 9.9 | 10.8 | 10.5 | 10.5496 | 10.8 | 11.7295 | 11.4 | 10.2 | 11.5 | 9.2 | [5 empty] | 10.3618
| 1 | 5 | 10.7 | 10.48 | 9.7065 | 9.6488 | 9.8 | 10.4 | 10.4 | 10.24 | 9.5 | 10.4 | 9.7 | 9.3 | 10.5 | 10.5 | 10.4 | 10.3 | 10.5 | 10.8236 | 10.1 | 11.6206 | 11.4 | 9.9 | 11.9 | 9.7 | [5 empty] | 51
| 1 | 6 | 11.5 | 10.41 | 9.8676 | 9.639 | 10.4 | 10.2 | 10.3 | 10.47 | 9.2 | 11 | 10.4 | 10.4 | 10.3 | 11 | 10 | 10.4 | 10.5 | 10.8713 | 10.8 | 11.4779 | 11.2 | 10.4 | 11.7 | 9.5
| Mean |  | 10.75 | 10.2467 | 9.7362 | 9.5073 | 10.1167 | 10.4167 | 10.6667 | 10.3417 | 9.85 | 10.8 | 10.2833 | 9.7333 | 10.2167 | 10.7 | 10.1 | 10.4167 | 10.5833 | 10.748 | 10.6 | 11.492 | 11.3667 | 10.2333 | 11.85 | 9.5
| Median |  | 10.65 | 10.205 | 9.7243 | 9.4672 | 10.2 | 10.4 | 10.6 | 10.345 | 9.85 | 10.8 | 10.35 | 9.75 | 10.2 | 10.7 | 10.05 | 10.4 | 10.5 | 10.7377 | 10.8 | 11.5086 | 11.4 | 10.25 | 11.9 | 9.5
| Std Dev. |  | 0.50892 | 0.17896 | 0.07353 | 0.11042 | 0.37103 | 0.20412 | 0.3559 | 0.20094 | 0.47645 | 0.24495 | 0.3656 | 0.45461 | 0.25626 | 0.30332 | 0.17889 | 0.31885 | 0.20412 | 0.15413 | 0.43359 | 0.17929 | 0.10328 | 0.2582 | 0.21679 | 0.18974
| Rel.Std.Dev. |  | 4.734% | 1.747% | 0.7552% | 1.161% | 3.668% | 1.96% | 3.337% | 1.943% | 4.837% | 2.268% | 3.555% | 4.671% | 2.508% | 2.835% | 1.771% | 3.061% | 1.929% | 1.434% | 4.09% | 1.56% | 0.9086% | 2.523% | 1.829% | 1.997%
| PDM3 |  | 3.746% | -1.111% | -6.038% | -8.247% | -2.366% | 0.5294% | 2.942% | -0.1945% | -4.939% | 4.229% | -0.7574% | -6.065% | -1.401% | 3.264% | -2.527% | 0.5294% | 2.138% | 3.727% | 2.299% | 10.91% | 9.698% | -1.24% | 14.36% | -8.317%
| Z-Score (Absolute) |  | 0.82 | 0.29 | 1.43 | 1.93 | 0.58 | 0.08 | 0.64 | 0.08 | 1.17 | 0.93 | 0.21 | 1.43 | 0.36 | 0.71 | 0.62 | 0.08 | 0.45 | 0.82 | 0.49 | 2.47 | 2.19 | 0.32 | 3.26 | 1.95
| Analytical results for Re in OREAS 231b (Certified Value 0.003 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 | 1.31 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.003 | < 0.01 | 0.002 | < 0.05 | < 0.002 | < 0.005 | 0.002 | 0.003 | 0.003 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0127 | 0.005 | 0.006 | 0.005 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.004 | < 0.05 | < 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.005 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.012 | 0.004 | 0.006 | 0.004 | [10 empty] | 32
| 1 | 3 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.003 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0152 | 0.005 | 0.005 | 0.006 | [10 empty] | 16
| 1 | 4 | 0.002 | < 0.01 | 0.004 | < 0.05 | < 0.002 | < 0.005 | 0.002 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0131 | 0.003 | 0.004 | 0.004 | [10 empty] | 0.0025729
| 1 | 5 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.003 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0136 | 0.005 | 0.005 | 0.005 | [10 empty] | 52
| 1 | 6 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.002 | 0.002 | < 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.003 | 0.0102 | 0.003 | 0.004 | 0.004
| Mean |  | 0.0025 |  | 0.003 |  | 0.0026667 |  | 0.0025 | 0.0026667 | 0.0023333 | 0.0032 |  |  |  |  | 0.0021667 | 0.0128 | 0.0041667 | 0.005 | 0.0046667
| Median |  | 0.0025 |  | 0.003 |  | 0.003 |  | 0.0025 | 0.003 | 0.002 | 0.003 |  |  |  |  | 0.002 | 0.0129 | 0.0045 | 0.005 | 0.0045
| Std Dev. |  | 0.00070711 |  | 0.00089443 |  | 0.00057735 |  | 0.00054772 | 0.0005164 | 0.0005164 | 0.0010954 |  |  |  |  | 0.00040825 | 0.0016673 | 0.00098319 | 0.00089443 | 0.0008165
| Rel.Std.Dev. |  | 28.28% |  | 29.81% |  | 21.65% |  | 21.91% | 19.36% | 22.13% | 34.23% |  |  |  |  | 18.84% | 13.03% | 23.6% | 17.89% | 17.5%
| PDM3 |  | -2.834% |  | 16.6% |  | 3.644% |  | -2.834% | 3.644% | -9.312% | 24.37% |  |  |  |  | -15.79% | 397.5% | 61.94% | 94.33% | 81.38%
| Z-Score (Absolute) |  | 0.87 | 1.16 | 0 | 12.72 | 0.67 | 0.29 | 0.29 | 0.19 | 0.39 | 0.1 | 1443.21 | 12.72 | 0.29 | 12.72 | 0.48 | 5.66 | 0.67 | 1.16 | 0.96
| Analytical results for S in OREAS 231b (Certified Value 0.264 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.28 | 0.28 | 0.26108 | 0.27 | 0.23 | 0.295 | 0.264 | 0.26 | 0.27 | 0.25 | 0.27 | 0.27 | 0.26 | 0.25 | 0.25 | 0.269 | 0.2 | 0.274 | 0.2431 | 0.275 | 0.26 | 0.22512 | 0.28 | 0.25 | 0.26 | 0.28 |  |  |  | 1
| 1 | 2 | 0.28 | 0.29 | 0.26748 | 0.27108 | 0.25 | 0.297 | 0.262 | 0.25 | 0.27 | 0.24 | 0.27 | 0.28 | 0.25 | 0.28 | 0.25 | 0.269 | 0.2 | 0.272 | 0.25206 | 0.27 | 0.25 | 0.22949 | 0.28 | 0.27 | 0.26 | 0.27 |  |  |  | 17
| 1 | 3 | 0.28 | 0.29 | 0.26457 | 0.26894 | 0.25 | 0.302 | 0.264 | 0.25 | 0.28 | 0.25 | 0.27 | 0.27 | 0.26 | 0.25 | 0.25 | 0.268 | 0.2 | 0.27 | 0.25255 | 0.27 | 0.26 | 0.23913 | 0.29 | 0.26 | 0.27 | 0.27 |  |  |  | 16
| 1 | 4 | 0.29 | 0.28 | 0.26376 | 0.27181 | 0.24 | 0.307 | 0.268 | 0.27 | 0.27 | 0.25 | 0.28 | 0.26 | 0.26 | 0.26 | 0.25 | 0.272 | 0.2 | 0.291 | 0.24557 | 0.27 | 0.26 | 0.23979 | 0.28 | 0.26 | 0.27 | 0.28 |  |  |  | 0.26428
| 1 | 5 | 0.28 | 0.3 | 0.2642 | 0.27095 | 0.25 | 0.314 | 0.267 | 0.28 | 0.27 | 0.24 | 0.26 | 0.25 | 0.26 | 0.26 | 0.25 | 0.272 | 0.2 | 0.29 | 0.24126 | 0.27 | 0.26 | 0.2424 | 0.29 | 0.27 | 0.27 | 0.26 |  |  |  | 53
| 1 | 6 | 0.28 | 0.28 | 0.2697 | 0.27296 | 0.27 | 0.302 | 0.261 | 0.25 | 0.27 | 0.25 | 0.27 | 0.27 | 0.25 | 0.25 | 0.25 | 0.269 | 0.2 | 0.273 | 0.23586 | 0.275 | 0.26 | 0.24891 | 0.28 | 0.27 | 0.27 | 0.27
| Mean |  | 0.28167 | 0.28667 | 0.26513 | 0.27096 | 0.24833 | 0.30283 | 0.26433 | 0.26 | 0.27167 | 0.24667 | 0.27 | 0.26667 | 0.25667 | 0.25833 | 0.25 | 0.26983 | 0.2 | 0.27833 | 0.24507 | 0.27167 | 0.25833 | 0.23747 | 0.28333 | 0.26333 | 0.26667 | 0.27167
| Median |  | 0.28 | 0.285 | 0.26438 | 0.27102 | 0.25 | 0.302 | 0.264 | 0.255 | 0.27 | 0.25 | 0.27 | 0.27 | 0.26 | 0.255 | 0.25 | 0.269 | 0.2 | 0.2735 | 0.24433 | 0.27 | 0.26 | 0.23946 | 0.28 | 0.265 | 0.27 | 0.27
| Std Dev. |  | 0.0040825 | 0.008165 | 0.0030299 | 0.0013943 | 0.013292 | 0.0069113 | 0.0027325 | 0.012649 | 0.0040825 | 0.005164 | 0.0063246 | 0.010328 | 0.005164 | 0.01169 | 0 | 0.0017224 | 3.0405e-17 | 0.0095219 | 0.0064534 | 0.002582 | 0.0040825 | 0.0087141 | 0.005164 | 0.008165 | 0.005164 | 0.0075277
| Rel.Std.Dev. |  | 1.449% | 2.848% | 1.143% | 0.5146% | 5.352% | 2.282% | 1.034% | 4.865% | 1.503% | 2.094% | 2.342% | 3.873% | 2.012% | 4.525% | 0% | 0.6383% | 1.52e-14% | 3.421% | 2.633% | 0.9504% | 1.58% | 3.67% | 1.823% | 3.101% | 1.936% | 2.771%
| PDM3 |  | 6.578% | 8.469% | 0.3214% | 2.525% | -6.035% | 14.59% | 0.01897% | -1.621% | 2.794% | -6.666% | 2.163% | 0.9019% | -2.882% | -2.251% | -5.405% | 2.1% | -24.32% | 5.316% | -7.271% | 2.794% | -2.251% | -10.14% | 7.208% | -0.3594% | 0.9019% | 2.794%
| Z-Score (Absolute) |  | 1.21 | 1.62 | 0.12 | 0.35 | 1.48 | 2.93 | 0.19 | 0.54 | 0.4 | 1.62 | 0.27 | 0 | 0.81 | 0.67 | 1.35 | 0.26 | NA | 0.94 | 1.75 | 0.4 | 0.67 | 2.36 | 1.35 | 0.27 | 0 | 0.4
Results from laboratory 1.23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 231b (Certified Value 0.82 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.86 | 0.82 | 0.86053 | 0.2 | 0.9 | 0.88 | < 5 | 1 | 0.82 | 0.87 | 0.82 | 0.82 | < 5 | 0.7 | 0.7 | 0.83 | 3.38 | 1 | 0.6965 | 0.81 | 1.0389 | 0.78 | 0.9 | 1.2 | 0.8 |  |  |  |  | 1
| 1 | 2 | 0.85 | 0.85 | 0.83467 | 0.6 | 0.8 | 0.81 | < 5 | 0.9 | 0.84 | 0.95 | 0.76 | 0.89 | < 5 | 0.8 | 0.7 | 0.74 | 6.53 | 1 | 0.68162 | 0.81 | 0.8663 | 0.79 | 0.86 | 1 | 0.76 |  |  |  |  | 33
| 1 | 3 | 1.47 | 0.89 | 0.83588 | 0.3 | 0.9 | 0.79 | < 5 | 1 | 0.84 | 0.85 | 0.78 | 0.83 | < 5 | 0.8 | 0.7 | 0.77 | 6.24 | 1 | 0.66999 | 0.79 | 0.8599 | 0.76 | 0.82 | 1 | 0.77 |  |  |  |  | 16
| 1 | 4 | 1.07 | 0.83 | 0.83994 | 0.2 | 0.8 | 0.85 | < 5 | 1.2 | 0.81 | 0.86 | 0.8 | 0.83 | < 5 | 0.8 | 0.7 | 0.78 | 5.07 | 1 | 0.6973 | 0.85 | 0.8868 | 0.73 | 0.84 | 0.9 | 0.74 |  |  |  |  | 0.82095
| 1 | 5 | 0.98 | 0.86 | 0.85915 | 0.3 | 0.9 | 0.79 | < 5 | 1.3 | 0.81 | 0.84 | 0.76 | 0.88 | < 5 | 0.7 | 0.7 | 0.7 | 5.91 | 1 | 0.69436 | 0.72 | 0.8849 | 0.79 | 0.87 | 0.9 | 0.8 |  |  |  |  | 54
| 1 | 6 | 0.92 | 0.89 | 0.86065 | 0.6 | 0.7 | 0.82 | < 5 | 1.2 | 0.85 | 0.91 | 0.8 | 0.86 | < 5 | 0.7 | 0.7 | 0.8 | 7.11 | 0.5 | 0.68436 | 0.77 | 0.9404 | 0.76 | 0.84 | 0.9 | 0.79
| Mean |  | 1.025 | 0.85667 | 0.84847 | 0.36667 | 0.83333 | 0.82333 |  | 1.1 | 0.82833 | 0.88 | 0.78667 | 0.85167 |  | 0.75 | 0.7 | 0.77 | 5.7067 | 0.91667 | 0.68736 | 0.79167 | 0.91287 | 0.76833 | 0.855 | 0.98333 | 0.77667
| Median |  | 0.95 | 0.855 | 0.84954 | 0.3 | 0.85 | 0.815 |  | 1.1 | 0.83 | 0.865 | 0.79 | 0.845 |  | 0.75 | 0.7 | 0.775 | 6.075 | 1 | 0.68936 | 0.8 | 0.88585 | 0.77 | 0.85 | 0.95 | 0.78
| Std Dev. |  | 0.23279 | 0.029439 | 0.012881 | 0.18619 | 0.08165 | 0.03559 |  | 0.15492 | 0.017224 | 0.041952 | 0.024221 | 0.029269 |  | 0.054772 | 1.2162e-16 | 0.045607 | 1.326 | 0.20412 | 0.010724 | 0.044008 | 0.067938 | 0.023166 | 0.028107 | 0.1169 | 0.024221
| Rel.Std.Dev. |  | 22.71% | 3.436% | 1.518% | 50.78% | 9.798% | 4.323% |  | 14.08% | 2.079% | 4.767% | 3.079% | 3.437% |  | 7.303% | 1.737e-14% | 5.923% | 23.24% | 22.27% | 1.56% | 5.559% | 7.442% | 3.015% | 3.287% | 11.89% | 3.119%
| PDM3 |  | 24.85% | 4.35% | 3.352% | -55.34% | 1.508% | 0.2898% |  | 33.99% | 0.8988% | 7.192% | -4.177% | 3.741% |  | -8.643% | -14.73% | -6.207% | 595.1% | 11.66% | -16.27% | -3.568% | 11.2% | -6.41% | 4.147% | 19.78% | -5.395%
| Z-Score (Absolute) |  | 1.87 | 0.07 | 0.02 | NA | NA | 0.29 | 17.63 | NA | 0.23 | 0.32 | 0.68 | 0.02 | 17.63 | NA | NA | 0.86 | 51.9 | NA | 1.74 | 0.62 | 0.67 | 0.87 | 0.05 | NA | 0.78
Results from laboratories 1.07, 1.10, 1.13, 1.22, 1.23, 1.27 and 1.33 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 231b (Certified Value 39.4 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 42.1 | 39.6 | 33.16 | 39.456 | 39.0484 | 38 | 44.6 | 38 | 41.4 | 38 | 37.3 | 42.8 | 38.8 | 38 | 37.3 | 38 | 39 | 38 | 38.9 | 41 | 40.8312 | 42.1 | 35.576 | 40.7 | 39.7 | 43.5 | 39.8 |  |  | 1
| 1 | 2 | 40.8 | 39.9 | 35.1 | 38.8992 | 39.5228 | 39 | 43.9 | 38 | 40.3 | 39 | 36.2 | 43.6 | 40.2 | 41.9 | 39.5 | 37 | 39 | 38 | 36.7 | 42 | 41.965 | 40.4 | 35.8846 | 40.3 | 41.3 | 41.4 | 37.9 |  |  | 34
| 1 | 3 | 42.6 | 40.8 | 35.74 | 38.984 | 38.9572 | 39 | 43.8 | 38 | 40.8 | 39 | 35.6 | 42.6 | 37.3 | 41.9 | 37.5 | 38 | 39 | 38 | 38.8 | 42 | 40.9276 | 42 | 36.4644 | 41.7 | 38.4 | 41.7 | 38.7 |  |  | 16
| 1 | 4 | 40.2 | 39.5 | 33.44 | 38.6864 | 39.1277 | 39 | 44.6 | 39 | 41.3 | 38 | 35.1 | 43.5 | 38.3 | 40.6 | 37.4 | 38 | 39 | 38 | 39.8 | 42 | 40.7482 | 40.6 | 36.8815 | 41.1 | 39.6 | 40.4 | 37.4 |  |  | 39.4319
| 1 | 5 | 40.4 | 40.5 | 34.04 | 39.2448 | 39.1576 | 39 | 42.6 | 39 | 39.4 | 39 | 35 | 41.8 | 36.6 | 37.3 | 39.1 | 38 | 39 | 38 | 38.6 | 41 | 40.852 | 39.2 | 37.6006 | 41.4 | 39.8 | 42.1 | 39.6 |  |  | 55
| 1 | 6 | 42.1 | 40.6 | 33.49 | 38.1376 | 39.3386 | 40 | 42.2 | 37 | 40.9 | 40 | 34.6 | 43.4 | 38.8 | 43.1 | 37.6 | 37 | 39 | 38 | 40.3 | 42 | 41.0519 | 41.9 | 37.2617 | 39.8 | 41.3 | 41.6 | 38.3
| Mean |  | 41.3667 | 40.15 | 34.1617 | 38.9013 | 39.192 | 39 | 43.6167 | 38.1667 | 40.6833 | 38.8333 | 35.6333 | 42.95 | 38.3333 | 40.4667 | 38.0667 | 37.6667 | 39 | 38 | 38.85 | 41.6667 | 41.0627 | 41.0333 | 36.6115 | 40.8333 | 40.0167 | 41.7833 | 38.6167
| Median |  | 41.45 | 40.2 | 33.765 | 38.9416 | 39.1426 | 39 | 43.85 | 38 | 40.85 | 39 | 35.35 | 43.1 | 38.55 | 41.25 | 37.55 | 38 | 39 | 38 | 38.85 | 42 | 40.8898 | 41.25 | 36.6729 | 40.9 | 39.75 | 41.65 | 38.5
| Std Dev. |  | 1.0211 | 0.55408 | 1.0355 | 0.46086 | 0.20593 | 0.63246 | 1.0088 | 0.75277 | 0.7414 | 0.75277 | 0.98522 | 0.6921 | 1.2644 | 2.3312 | 0.96885 | 0.5164 | 0 | 0 | 1.2406 | 0.5164 | 0.45374 | 1.1639 | 0.78692 | 0.70899 | 1.1161 | 1.0147 | 0.94534
| Rel.Std.Dev. |  | 2.468% | 1.38% | 3.031% | 1.185% | 0.5254% | 1.622% | 2.313% | 1.972% | 1.822% | 1.938% | 2.765% | 1.611% | 3.298% | 5.761% | 2.545% | 1.371% | 0% | 0% | 3.193% | 1.239% | 1.105% | 2.836% | 2.149% | 1.736% | 2.789% | 2.429% | 2.448%
| PDM3 |  | 4.906% | 1.821% | -13.37% | -1.346% | -0.6084% | -1.095% | 10.61% | -3.209% | 3.174% | -1.518% | -9.633% | 8.922% | -2.786% | 2.624% | -3.462% | -4.477% | -1.095% | -3.631% | -1.476% | 5.667% | 4.136% | 4.061% | -7.153% | 3.554% | 1.483% | 5.963% | -2.068%
| Z-Score (Absolute) |  | 1.2 | 0.58 | 2.45 | 0.05 | 0.1 | 0 | 2.33 | 0.42 | 0.85 | 0.08 | 1.7 | 2 | 0.34 | 0.74 | 0.47 | 0.67 | 0 | 0.51 | 0.08 | 1.35 | 1.04 | 1.03 | 1.21 | 0.93 | 0.51 | 1.41 | 0.19
| Analytical results for Se in OREAS 231b (Indicative Value 1 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 2 | < 0.5 | < 1.5 | 0.3 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.36 | < 5 | 1 | 0.5531 | 2 | < 2 | < 1 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 2 | 0.7 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 2.7 | < 5 | 1 | 0.519 | 1.8 | < 2 | < 1 | < 0.5 | [6 empty] | 4
| 1 | 3 | < 2 | < 0.5 | < 1.5 | 0.6 | < 1 | 0.6 | < 5 | 1 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 4.68 | < 5 | 1 | 0.5198 | 2.4 | < 2 | < 1 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 2 | 0.6 | < 1.5 | 0.3 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 5.09 | < 5 | 1 | 0.5187 | 2.5 | < 2 | < 1 | < 0.5 | [6 empty] | 0.99918
| 1 | 5 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | 0.5 | < 5 | 1 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 4.12 | < 5 | 1 | 0.5278 | 2.9 | < 2 | < 1 | < 0.5 | [6 empty] | 10
| 1 | 6 | < 2 | < 0.5 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 1 | < 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 3.9 | < 5 | 1 | 0.7218 | 2.2 | < 2 | < 1 | < 0.5
| Mean |  |  | 0.63333 |  | 0.43333 |  | 0.55 |  | 1 | 1 | 1 | 1.6667 |  |  |  | 1 | 4.475 |  | 1 | 0.56003 | 2.3
| Median |  |  | 0.6 |  | 0.45 |  | 0.55 |  | 1 | 1 | 1 | 2 |  |  |  | 1 | 4.4 |  | 1 | 0.5238 | 2.3
| Std Dev. |  |  | 0.057735 |  | 0.12111 |  | 0.070711 |  | 0 | 0 | 0 | 0.5164 |  |  |  | 0 | 1.2313 |  | 0 | 0.080332 | 0.38987
| Rel.Std.Dev. |  |  | 9.116% |  | 27.95% |  | 12.86% |  | 0% | 0% | 0% | 30.98% |  |  |  | 0% | 27.52% |  | 0% | 14.34% | 16.95%
| PDM3 |  |  | -36.61% |  | -56.63% |  | -44.96% |  | 0.08176% | 0.08176% | 0.08176% | 66.8% |  |  |  | 0.08176% | 347.9% |  | 0.08176% | -43.95% | 130.2%
| Z-Score (Absolute) |  | 0 | 0.75 | 0.34 | 0.76 | 0.67 | 0.88 | 2.02 | 0 | 0.11 | 0 | 0.9 | 2.02 | 0 | 0.67 | 0 | 4.69 | 2.02 | 0 | 0.59 | 1.75 | 0 | 0.67 | 1.01
| Analytical results for Sm in OREAS 231b (Certified Value 2.44 ppm) | [30 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.49 | 2.07 | 2.2585 | 2.7 | 2.5 | 2.69 | 2.4 | 2.4641 | [21 empty] | 1
| 1 | 2 | 2.51 | 2.17 | 2.2924 | 2.7 | 2.5 | 2.58 | 2.4 | 2.689 | [21 empty] | 19
| 1 | 3 | 2.47 | 2.27 | 2.3052 | 2.5 | 2.4 | 2.63 | 2.6 | 2.5557 | [21 empty] | 16
| 1 | 4 | 2.47 | 2.02 | 2.2384 | 2.5 | 2.4 | 2.64 | 2.4 | 2.552 | [21 empty] | 2.44
| 1 | 5 | 2.42 | 2.08 | 2.2399 | 2.8 | 2.4 | 2.52 | 2.4 | 2.497 | [21 empty] | 56
| 1 | 6 | 2.43 | 2.06 | 2.2956 | 2.8 | 2.3 | 2.6 | 2.6 | 2.5027
| Mean |  | 2.465 | 2.1117 | 2.2717 | 2.6667 | 2.4167 | 2.61 | 2.4667 | 2.5434
| Median |  | 2.47 | 2.075 | 2.2755 | 2.7 | 2.4 | 2.615 | 2.4 | 2.5273
| Std Dev. |  | 0.034496 | 0.09196 | 0.029752 | 0.13663 | 0.075277 | 0.057966 | 0.10328 | 0.079404
| Rel.Std.Dev. |  | 1.399% | 4.355% | 1.31% | 5.123% | 3.115% | 2.221% | 4.187% | 3.122%
| PDM3 |  | 1.024% | -13.46% | -6.899% | 9.289% | -0.9567% | 6.967% | 1.092% | 4.238%
| Z-Score (Absolute) |  | 0.01 | 2.15 | 1.18 | 1.22 | 0.3 | 0.88 | 0.01 | 0.47
| Analytical results for Sn in OREAS 231b (Certified Value 0.94 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.1 | 0.9 | 5.26 | 0.92548 | < 1 | 1.06 | 1.6 | 0.9 | < 10 | 1.1 | 0.9 | 1 | 0.8 | 0.8 | < 10 | 0.8 | 0.9 | 0.9 | 2.61 | 1 | 0.79265 | 0.9 | 0.6319 | 1 | 1.1 | 1.3 | 1.6 |  |  | 1
| 1 | 2 | 1.2 | 0.9 | 3.15 | 0.91559 | < 1 | 1.05 | 1.6 | 0.9 | < 10 | 0.9 | 0.9 | 1.1 | 0.9 | 0.9 | < 10 | 0.8 | 0.8 | 0.8 | 1.45 | 1 | 0.79332 | 1 | 0.6451 | 1 | 1.2 | 1.2 | 1.5 |  |  | 20
| 1 | 3 | 1.3 | 0.9 | 2.38 | 0.901 | < 1 | 1.06 | 1.5 | 0.9 | < 10 | 1 | 0.9 | 1 | 0.9 | 0.9 | < 10 | 0.9 | 0.8 | 0.9 | 11.98 | 17.5 | 0.7792 | 0.9 | 0.743 | 1 | 1.1 | 1.1 | 1.5 |  |  | 16
| 1 | 4 | 1.4 | 0.9 | 4.26 | 0.87656 | < 1 | 1.05 | 1.5 | 0.9 | < 10 | 1.1 | 0.9 | 1 | 0.9 | 0.8 | < 10 | 0.8 | 0.8 | 1 | 4.16 | 1 | 0.77642 | 0.8 | 0.7009 | 1 | 1 | 1.1 | 1.5 |  |  | 0.93501
| 1 | 5 | 1.2 | 1 | 4.22 | 0.87443 | < 1 | 1.02 | 1.3 | 0.9 | < 10 | 1 | 0.9 | 1 | 0.8 | 1 | < 10 | 0.7 | 0.9 | 0.9 | 2.7 | 1 | 0.78042 | 0.9 | 0.7499 | 1 | 1 | 1.1 | 1.5 |  |  | 57
| 1 | 6 | 1.2 | 0.9 | 3.62 | 0.9187 | < 1 | 1.05 | 1.4 | 1 | < 10 | 1.1 | 0.9 | 1 | 0.9 | 0.9 | < 10 | 0.7 | 0.8 | 0.9 | 1.88 | 1 | 0.80006 | 0.9 | 0.5966 | 1 | 0.9 | 1.1 | 1.5
| Mean |  | 1.2333 | 0.91667 | 3.815 | 0.90196 |  | 1.0483 | 1.4833 | 0.91667 |  | 1.0333 | 0.9 | 1.0167 | 0.86667 | 0.88333 |  | 0.78333 | 0.83333 | 0.9 | 4.13 | 3.75 | 0.78701 | 0.9 | 0.6779 | 1 | 1.05 | 1.15 | 1.5167
| Median |  | 1.2 | 0.9 | 3.92 | 0.9083 |  | 1.05 | 1.5 | 0.9 |  | 1.05 | 0.9 | 1 | 0.9 | 0.9 |  | 0.8 | 0.8 | 0.9 | 2.655 | 1 | 0.78653 | 0.9 | 0.673 | 1 | 1.05 | 1.1 | 1.5
| Std Dev. |  | 0.10328 | 0.040825 | 0.99895 | 0.022017 |  | 0.01472 | 0.1169 | 0.040825 |  | 0.08165 | 0 | 0.040825 | 0.05164 | 0.075277 |  | 0.075277 | 0.05164 | 0.063246 | 3.9553 | 6.7361 | 0.009576 | 0.063246 | 0.062853 | 0 | 0.10488 | 0.083666 | 0.040825
| Rel.Std.Dev. |  | 8.374% | 4.454% | 26.18% | 2.441% |  | 1.404% | 7.881% | 4.454% |  | 7.902% | 0% | 4.016% | 5.958% | 8.522% |  | 9.61% | 6.197% | 7.027% | 95.77% | 179.6% | 1.217% | 7.027% | 9.272% | 0% | 9.989% | 7.275% | 2.692%
| PDM3 |  | 31.91% | -1.962% | 308% | -3.535% |  | 12.12% | 58.64% | -1.962% |  | 10.52% | -3.744% | 8.733% | -7.309% | -5.527% |  | -16.22% | -10.87% | -3.744% | 341.7% | 301.1% | -15.83% | -3.744% | -27.5% | 6.951% | 12.3% | 22.99% | 62.21%
| Z-Score (Absolute) |  | 1.15 | 0.22 | 12.27 | 0.28 | 2.01 | 0.35 | 2.23 | 0.22 | 17.38 | 0.29 | 0.29 | 0.22 | 0.43 | 0.36 | 17.38 | 0.79 | 0.57 | 0.29 | 13.63 | 11.99 | 0.77 | 0.29 | 1.24 | NA | 0.36 | 0.79 | 2.37
Results from laboratory 1.31 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sr in OREAS 231b (Certified Value 203 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 192.8 | 196.66 | 210 | 204.5946 | 204.7718 | 176 | 210 | 218 | 204.38 | 195 | 195 | 225 | 203 | 207 | 206 | 213 | 197.7 | 194 | 200 | 110.24 | 188.7555 | 198 | 202.5956 | 216 | 240.3563 | 204 | 203 | 204.7 | 198 | 1
| 1 | 2 | 187.7 | 197.51 | 210 | 200.4198 | 206.443 | 184 | 207 | 208 | 199.13 | 192 | 191 | 226 | 210 | 205 | 226 | 211 | 200.1 | 190 | 199 | 110.02 | 196.802 | 197 | 206.0557 | 207 | 238.9431 | 202 | 214.7 | 206.6 | 191 | 21
| 1 | 3 | 195.7 | 204.27 | 216 | 203.9944 | 204.7065 | 196 | 209 | 209 | 205.56 | 197 | 194 | 225 | 202 | 213 | 207 | 215 | 197.3 | 191 | 200 | 96.55 | 196.232 | 198 | 203.1942 | 217 | 243.0346 | 207 | 209.9 | 206.7 | 192 | 16
| 1 | 4 | 186.2 | 197.25 | 202 | 205.6726 | 205.6645 | 204 | 210 | 212 | 201.47 | 196 | 194 | 225 | 203 | 206 | 209 | 210 | 199 | 192 | 211 | 115.38 | 192.9355 | 197 | 203.2479 | 212 | 246.6256 | 205 | 202.5 | 207.3 | 187 | 203.2353
| 1 | 5 | 189.7 | 206.17 | 207 | 205.5942 | 205.449 | 186 | 212 | 212 | 198.3 | 193 | 192 | 209 | 198 | 204 | 216 | 213 | 203.3 | 195 | 213 | 109.17 | 188.1 | 199 | 202.4723 | 216 | 243.5225 | 204 | 211.9 | 208.6 | 199 | 58
| 1 | 6 | 192.3 | 203.57 | 202 | 209.72 | 207.8659 | 200 | 211 | 205 | 204.99 | 194 | 193 | 224 | 207 | 207 | 210 | 209 | 204.4 | 192 | 200 | 113.02 | 187.5205 | 199 | 203.8578 | 220 | 234.5631 | 206 | 214.6 | 210.1 | 188
| Mean |  | 190.7333 | 200.905 | 207.8333 | 204.9993 | 205.8168 | 191 | 209.8333 | 210.6667 | 202.305 | 194.5 | 193.1667 | 222.3333 | 203.8333 | 207 | 212.3333 | 211.8333 | 200.3 | 192.3333 | 203.8333 | 109.0633 | 191.7242 | 198 | 203.5706 | 214.6667 | 241.1742 | 204.6667 | 209.4333 | 207.3333 | 192.5
| Median |  | 191 | 200.54 | 208.5 | 205.0944 | 205.5568 | 191 | 210 | 210.5 | 202.925 | 194.5 | 193.5 | 225 | 203 | 206.5 | 209.5 | 212 | 199.55 | 192 | 200 | 110.13 | 190.8455 | 198 | 203.221 | 216 | 241.6954 | 204.5 | 210.9 | 207 | 191.5
| Std Dev. |  | 3.5274 | 4.2202 | 5.3821 | 3.0079 | 1.1895 | 10.7145 | 1.7224 | 4.4572 | 3.1282 | 1.8708 | 1.472 | 6.5625 | 4.1673 | 3.1623 | 7.5542 | 2.2286 | 2.9428 | 1.8619 | 6.3692 | 6.549 | 4.1759 | 0.89443 | 1.3161 | 4.5461 | 4.1964 | 1.7512 | 5.4808 | 1.8511 | 5.01
| Rel.Std.Dev. |  | 1.849% | 2.101% | 2.59% | 1.467% | 0.5779% | 5.61% | 0.8208% | 2.116% | 1.546% | 0.9619% | 0.762% | 2.952% | 2.044% | 1.528% | 3.558% | 1.052% | 1.469% | 0.9681% | 3.125% | 6.005% | 2.178% | 0.4517% | 0.6465% | 2.118% | 1.74% | 0.8556% | 2.617% | 0.8928% | 2.603%
| PDM3 |  | -6.151% | -1.147% | 2.262% | 0.8679% | 1.27% | -6.02% | 3.246% | 3.657% | -0.4578% | -4.298% | -4.954% | 9.397% | 0.2942% | 1.852% | 4.477% | 4.231% | -1.444% | -5.364% | 0.2942% | -46.34% | -5.664% | -2.576% | 0.165% | 5.625% | 18.67% | 0.7043% | 3.05% | 2.016% | -5.282%
| Z-Score (Absolute) |  | 1.47 | 0.33 | 0.45 | 0.13 | 0.22 | 1.44 | 0.67 | 0.77 | 0.17 | 1.05 | 1.2 | 2.08 | 0 | 0.36 | 0.96 | 0.9 | 0.4 | 1.29 | 0 | 10.65 | 1.36 | 0.66 | 0.03 | 1.22 | 4.2 | 0.09 | 0.63 | 0.39 | 1.27
| Analytical results for Ta in OREAS 231b (Certified Value 0.24 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.91 | 0.26 | 0.55107 | < 0.1 | 0.21 | 0.4 | 0.24 | 13 | 0.5 | 0.23 | 0.25 | 0.2 | 0.27 | 0.2 | 0.2 | 0.24 | < 0.5 | 0.19175 | 0.24 | 0.3569 | 0.55 | 0.61 | 0.33 | 0.26 | [5 empty] | 1
| 1 | 2 | 1.02 | 0.24 | 0.58604 | 0.2 | 0.21 | 0.37 | 0.23 | 13 | 0.5 | 0.23 | 0.25 | 0.2 | 0.27 | 0.21 | 0.2 | 0.23 | < 0.5 | 0.19077 | 0.21 | 0.3647 | 0.49 | 0.45 | 0.3 | 0.25 | [5 empty] | 22
| 1 | 3 | 1.04 | 0.25 | 0.52874 | 0.2 | 0.2 | 0.39 | 0.23 | 13 | 0.4 | 0.23 | 0.24 | 0.22 | 0.28 | 0.23 | 0.2 | 0.23 | < 0.5 | 0.19345 | 0.24 | 0.4319 | 0.56 | 0.5 | 0.29 | 0.24 | [5 empty] | 16
| 1 | 4 | 1.74 | 0.25 | 0.4608 | < 0.1 | 0.21 | 0.42 | 0.23 | 13 | 0.5 | 0.23 | 0.25 | 0.21 | 0.26 | 0.21 | 0.2 | 0.28 | < 0.5 | 0.18854 | 0.24 | 0.4174 | 0.49 | 0.61 | 0.42 | 0.23 | [5 empty] | 0.24188
| 1 | 5 | 0.45 | 0.24 | 0.50668 | < 0.1 | 0.2 | 0.36 | 0.24 | 13 | 0.4 | 0.22 | 0.23 | 0.21 | 0.25 | 0.17 | 0.2 | 0.23 | < 0.5 | 0.18913 | 0.24 | 0.4283 | 0.48 | 0.45 | 0.44 | 0.28 | [5 empty] | 59
| 1 | 6 | 0.42 | 0.26 | 0.45096 | < 0.1 | 0.19 | 0.34 | 0.24 | 13 | 0.4 | 0.23 | 0.25 | 0.21 | 0.26 | 0.19 | 0.2 | 0.24 | < 0.5 | 0.19248 | 0.24 | 0.34 | 0.51 | 0.58 | 0.42 | 0.22
| Mean |  | 0.93 | 0.25 | 0.51405 | 0.2 | 0.20333 | 0.38 | 0.235 | 13 | 0.45 | 0.22833 | 0.245 | 0.20833 | 0.265 | 0.20167 | 0.2 | 0.24167 |  | 0.19102 | 0.235 | 0.38987 | 0.51333 | 0.53333 | 0.36667 | 0.24667
| Median |  | 0.965 | 0.25 | 0.51771 | 0.2 | 0.205 | 0.38 | 0.235 | 13 | 0.45 | 0.23 | 0.25 | 0.21 | 0.265 | 0.205 | 0.2 | 0.235 |  | 0.19126 | 0.24 | 0.39105 | 0.5 | 0.54 | 0.375 | 0.245
| Std Dev. |  | 0.48316 | 0.0089443 | 0.052228 | 0 | 0.008165 | 0.028983 | 0.0054772 | 0 | 0.054772 | 0.0040825 | 0.0083666 | 0.0075277 | 0.010488 | 0.020412 | 3.0405e-17 | 0.019408 |  | 0.001916 | 0.012247 | 0.040519 | 0.033862 | 0.07607 | 0.067429 | 0.021602
| Rel.Std.Dev. |  | 51.95% | 3.578% | 10.16% | 0% | 4.016% | 7.627% | 2.331% | 0% | 12.17% | 1.788% | 3.415% | 3.613% | 3.958% | 10.12% | 1.52e-14% | 8.031% |  | 1.003% | 5.212% | 10.39% | 6.597% | 14.26% | 18.39% | 8.758%
| PDM3 |  | 284.5% | 3.356% | 112.5% | -17.31% | -15.94% | 57.1% | -2.845% | 5275% | 86.04% | -5.601% | 1.289% | -13.87% | 9.558% | -16.63% | -17.31% | -0.0888% |  | -21.03% | -2.845% | 61.18% | 112.2% | 120.5% | 51.59% | 1.978%
| Z-Score (Absolute) |  | 10.24 | 0.05 | 4.01 | 2.2 | 0.65 | 2 | 0.17 | NA | 3.05 | 0.27 | 0.02 | 0.57 | 0.27 | 0.67 | 0.7 | 0.07 | 0.05 | 0.83 | 0.17 | 2.15 | 4 | 4.3 | 1.8 | 0
Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 231b (Certified Value 0.56 ppm) | [30 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 0.62 | 0.57 | 0.52 | 0.55855 | 0.5 | 0.54 | 0.57 | 0.5 | 3.2462 | 0.56 | [19 empty] | 1
| 1 | 2 | 0.61 | 0.56 | 0.56 | 0.53689 | 0.6 | 0.53 | 0.58 | 0.5 | 3.3582 | 0.6 | [19 empty] | 23
| 1 | 3 | 0.64 | 0.58 | 0.58 | 0.52556 | 0.6 | 0.53 | 0.57 | 0.55 | 3.3036 | 0.56 | [19 empty] | 16
| 1 | 4 | 0.63 | 0.56 | 0.52 | 0.52756 | 0.6 | 0.54 | 0.57 | 0.5 | 3.3247 | 0.57 | [19 empty] | 0.55959
| 1 | 5 | 0.62 | 0.59 | 0.53 | 0.53 | 0.6 | 0.53 | 0.55 | 0.55 | 3.2775 | 0.55 | [19 empty] | 60
| 1 | 6 | 0.65 | 0.6 | 0.53 | 0.53645 | 0.6 | 0.51 | 0.57 | 0.45 | 3.2554 | 0.58
| Mean |  | 0.62833 | 0.57667 | 0.54 | 0.53583 | 0.58333 | 0.53 | 0.56833 | 0.50833 | 3.2943 | 0.57
| Median |  | 0.625 | 0.575 | 0.53 | 0.53323 | 0.6 | 0.53 | 0.57 | 0.5 | 3.2906 | 0.565
| Std Dev. |  | 0.01472 | 0.01633 | 0.024495 | 0.012047 | 0.040825 | 0.010954 | 0.0098319 | 0.037639 | 0.042912 | 0.017889
| Rel.Std.Dev. |  | 2.343% | 2.832% | 4.536% | 2.248% | 6.999% | 2.067% | 1.73% | 7.404% | 1.303% | 3.138%
| PDM3 |  | 12.29% | 3.052% | -3.5% | -4.245% | 4.243% | -5.287% | 1.563% | -9.159% | 488.7% | 1.861%
| Z-Score (Absolute) |  | 1.28 | 0.16 | 0.63 | 0.72 | 0.31 | 0.85 | 0.02 | 1.31 | 58.8 | 0.02
| Analytical results for Te in OREAS 231b (Certified Value 0.074 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | < 0.2 | 0.14997 | < 0.1 | < 0.1 | < 0.2 | 16 | < 0.1 | 0.07 | 0.05 | 0.06 | 0.11 | < 5 | < 0.1 | < 0.5 | < 0.05 | < 1 | 0.06347 | 0.07 | 0.0698 | 0.1 | 0.12 | 0.13 | 0.09 | [5 empty] | 1
| 1 | 2 | 0.06 | < 0.2 | 0.14068 | < 0.1 | 0.1 | < 0.2 | 16 | < 0.1 | 0.06 | 0.07 | 0.06 | 0.07 | < 5 | < 0.1 | < 0.5 | 0.07 | < 1 | 0.06953 | 0.06 | 0.0863 | 0.03 | 0.08 | 0.12 | 0.05 | [5 empty] | 24
| 1 | 3 | 0.17 | < 0.2 | 0.12997 | < 0.1 | 0.1 | < 0.2 | 16 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.06 | < 5 | < 0.1 | < 0.5 | 0.05 | < 1 | 0.05526 | 0.07 | 0.0904 | < 0.01 | 0.08 | 0.1 | 0.07 | [5 empty] | 16
| 1 | 4 | 0.1 | < 0.2 | 0.15019 | < 0.1 | < 0.1 | < 0.2 | 16 | < 0.1 | 0.07 | 0.05 | 0.05 | 0.11 | < 5 | < 0.1 | < 0.5 | 0.06 | < 1 | 0.07496 | 0.07 | 0.0781 | 0.07 | 0.09 | 0.11 | 0.07 | [5 empty] | 0.073508
| 1 | 5 | 0.1 | < 0.2 | 0.13678 | < 0.1 | 0.1 | < 0.2 | 16 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.05 | < 5 | < 0.1 | < 0.5 | 0.06 | < 1 | 0.06505 | 0.08 | 0.0904 | < 0.01 | 0.1 | 0.1 | 0.07 | [5 empty] | 61
| 1 | 6 | < 0.05 | < 0.2 | 0.13043 | < 0.1 | < 0.1 | < 0.2 | 16 | < 0.1 | 0.07 | 0.06 | 0.07 | 0.07 | < 5 | < 0.1 | < 0.5 | 0.08 | < 1 | 0.07232 | 0.06 | 0.3369 | 0.12 | 0.12 | 0.12 | 0.06
| Mean |  | 0.102 |  | 0.13967 |  | 0.1 |  | 16 |  | 0.065 | 0.061667 | 0.063333 | 0.078333 |  |  |  | 0.064 |  | 0.066765 | 0.068333 | 0.12532 | 0.08 | 0.098333 | 0.11333 | 0.068333
| Median |  | 0.1 |  | 0.13873 |  | 0.1 |  | 16 |  | 0.065 | 0.065 | 0.065 | 0.07 |  |  |  | 0.06 |  | 0.06729 | 0.07 | 0.08835 | 0.085 | 0.095 | 0.115 | 0.07
| Std Dev. |  | 0.041473 |  | 0.0090083 |  | 1.6997e-17 |  | 0 |  | 0.0054772 | 0.0098319 | 0.008165 | 0.025626 |  |  |  | 0.011402 |  | 0.0070968 | 0.0075277 | 0.10396 | 0.039158 | 0.018348 | 0.012111 | 0.013292
| Rel.Std.Dev. |  | 40.66% |  | 6.45% |  | 1.7e-14% |  | 0% |  | 8.427% | 15.94% | 12.89% | 32.71% |  |  |  | 17.82% |  | 10.63% | 11.02% | 82.96% | 48.95% | 18.66% | 10.69% | 19.45%
| PDM3 |  | 38.76% |  | 90% |  | 36.04% |  | 21666% |  | -11.57% | -16.11% | -13.84% | 6.564% |  |  |  | -12.93% |  | -9.173% | -7.04% | 70.48% | 8.831% | 33.77% | 54.18% | -7.04%
| Z-Score (Absolute) |  | 0.46 | 0.77 | 1.92 | 0.67 | NA | 0.77 | NA | 0.67 | 0.24 | 0.34 | 0.29 | 0.14 | 70.13 | 0.67 | 5.11 | 0.46 | 12.33 | 0.19 | 0.14 | 1.5 | 0.53 | 0.72 | 1.16 | 0.14
Results from laboratory 1.10 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 231b (Certified Value 0.65 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1 | 0.65 | 0.66 | 0.62268 | 0.68502 | 0.7 | 0.77 | 0.96 | 0.65 | 0.7 | 0.63 | 0.63 | 0.61 | 0.58 | 0.59 | 0.6 | 0.74 | 0.7 | 0.7106 | 0.66 | 1.1266 | 1.3 | 0.8 | 0.9 | 0.6 |  |  |  |  | 1
| 1 | 2 | 2.1 | 0.63 | 0.63 | 0.6178 | 0.64446 | 0.7 | 0.76 | 0.87 | 0.61 | 0.7 | 0.66 | 0.67 | 0.62 | 0.62 | 0.63 | 0.6 | 0.75 | 0.7 | 0.71989 | 0.66 | 0.9505 | 1.4 | 0.7 | 0.8 | 0.6 |  |  |  |  | 25
| 1 | 3 | 0.9 | 0.64 | 0.64 | 0.59734 | 0.64631 | 0.7 | 0.7 | 0.84 | 0.65 | 0.6 | 0.67 | 0.63 | 0.65 | 0.62 | 0.69 | 0.6 | 0.73 | 0.7 | 0.70987 | 0.68 | 0.9262 | 1.5 | 0.7 | 0.8 | 0.6 |  |  |  |  | 16
| 1 | 4 | 1.5 | 0.65 | 0.56 | 0.6193 | 0.64812 | 0.7 | 0.7 | 0.78 | 0.65 | 0.6 | 0.64 | 0.65 | 0.67 | 0.63 | 0.65 | 0.6 | 0.72 | 0.7 | 0.70101 | 0.68 | 0.8982 | 1.1 | 0.7 | 0.7 | 0.6 |  |  |  |  | 0.65239
| 1 | 5 | 1 | 0.67 | 0.58 | 0.59051 | 0.65175 | 0.7 | 0.74 | 0.76 | 0.65 | 0.6 | 0.62 | 0.62 | 0.61 | 0.62 | 0.54 | 0.6 | 0.69 | 0.7 | 0.70306 | 0.63 | 0.8821 | 1.2 | 0.7 | 0.8 | 0.6 |  |  |  |  | 62
| 1 | 6 | 2.3 | 0.68 | 0.58 | 0.59634 | 0.65872 | 0.7 | 0.66 | 0.74 | 0.63 | 0.6 | 0.65 | 0.63 | 0.65 | 0.66 | 0.6 | 0.6 | 0.71 | 0.7 | 0.70492 | 0.68 | 0.8807 | 1.2 | 0.7 | 0.7 | 0.6
| Mean |  | 1.4667 | 0.65333 | 0.60833 | 0.60733 | 0.65573 | 0.7 | 0.72167 | 0.825 | 0.64 | 0.63333 | 0.645 | 0.63833 | 0.635 | 0.62167 | 0.61667 | 0.6 | 0.72333 | 0.7 | 0.70822 | 0.665 | 0.94405 | 1.2833 | 0.71667 | 0.78333 | 0.6
| Median |  | 1.25 | 0.65 | 0.605 | 0.60757 | 0.64994 | 0.7 | 0.72 | 0.81 | 0.65 | 0.6 | 0.645 | 0.63 | 0.635 | 0.62 | 0.615 | 0.6 | 0.725 | 0.7 | 0.70739 | 0.67 | 0.9122 | 1.25 | 0.7 | 0.8 | 0.6
| Std Dev. |  | 0.60882 | 0.018619 | 0.040208 | 0.014085 | 0.015208 | 1.2162e-16 | 0.042151 | 0.082401 | 0.016733 | 0.05164 | 0.018708 | 0.018348 | 0.0251 | 0.025626 | 0.052026 | 0 | 0.021602 | 1.2162e-16 | 0.0068422 | 0.019748 | 0.093422 | 0.1472 | 0.040825 | 0.075277 | 0
| Rel.Std.Dev. |  | 41.51% | 2.85% | 6.61% | 2.319% | 2.319% | 1.737e-14% | 5.841% | 9.988% | 2.615% | 8.154% | 2.901% | 2.874% | 3.953% | 4.122% | 8.437% | 0% | 2.987% | 1.737e-14% | 0.9661% | 2.97% | 9.896% | 11.47% | 5.696% | 9.61% | 0%
| PDM3 |  | 124.8% | 0.145% | -6.753% | -6.907% | 0.5126% | 7.298% | 10.62% | 26.46% | -1.899% | -2.921% | -1.132% | -2.154% | -2.665% | -4.709% | -5.475% | -8.03% | 10.87% | 7.298% | 8.559% | 1.933% | 44.71% | 96.71% | 9.853% | 20.07% | -8.03%
| Z-Score (Absolute) |  | NA | 0.09 | 0.92 | 0.94 | 0.15 | NA | 1.63 | 3.95 | 0.21 | NA | 0.09 | 0.24 | 0.32 | 0.62 | 0.73 | NA | 1.67 | NA | 1.33 | 0.36 | 6.63 | NA | NA | NA | NA
Results from laboratories 1.01, 1.07, 1.13, 1.23, 1.27, 1.31, 1.32, 1.33 and 1.34 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 231b (Certified Value 0.638 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.63 | 0.6449 | 0.63579 | 0.64812 | 0.403 | 0.6774 | 0.635 | 0.662 | 0.6 | 0.659 | 0.687 | 0.642 | 0.623 | 0.622 | 0.6 | 0.6578 | 0.651 | 0.638 | 0.68 | 0.59079 | 0.627 | 0.62653 | 0.648 | 0.51809 | 0.63 | 0.63 | 0.63 | 0.7 |  | 1
| 1 | 2 | 0.63 | 0.6408 | 0.63288 | 0.65671 | 0.623 | 0.6774 | 0.636 | 0.6536 | 0.61 | 0.629 | 0.693 | 0.658 | 0.628 | 0.664 | 0.59 | 0.661 | 0.641 | 0.628 | 0.67 | 0.61682 | 0.628 | 0.64268 | 0.623 | 0.54039 | 0.62 | 0.65 | 0.64 | 0.8 |  | 26
| 1 | 3 | 0.64 | 0.6538 | 0.63697 | 0.65311 | 0.607 | 0.6834 | 0.635 | 0.6652 | 0.61 | 0.636 | 0.69 | 0.649 | 0.632 | 0.617 | 0.6 | 0.6549 | 0.637 | 0.646 | 0.677 | 0.61545 | 0.626 | 0.62836 | 0.64 | 0.52859 | 0.64 | 0.6 | 0.64 | 0.7 |  | 16
| 1 | 4 | 0.63 | 0.6413 | 0.63175 | 0.65034 | 0.157 | 0.6834 | 0.645 | 0.6545 | 0.6 | 0.645 | 0.708 | 0.639 | 0.613 | 0.627 | 0.59 | 0.6551 | 0.644 | 0.658 | 0.687 | 0.6039 | 0.624 | 0.62534 | 0.637 | 0.53609 | 0.63 | 0.63 | 0.63 | 0.7 |  | 0.63779
| 1 | 5 | 0.64 | 0.6566 | 0.63789 | 0.65386 | 0.306 | 0.6894 | 0.644 | 0.6421 | 0.61 | 0.638 | 0.649 | 0.628 | 0.617 | 0.647 | 0.6 | 0.6616 | 0.65 | 0.673 | 0.682 | 0.59322 | 0.632 | 0.63687 | 0.637 | 0.55409 | 0.63 | 0.63 | 0.64 | 0.7 |  | 63
| 1 | 6 | 0.64 | 0.6545 | 0.63096 | 0.65354 | 0.546 | 0.6834 | 0.625 | 0.6589 | 0.62 | 0.636 | 0.685 | 0.649 | 0.634 | 0.619 | 0.59 | 0.6635 | 0.641 | 0.639 | 0.66 | 0.58749 | 0.627 | 0.62228 | 0.645 | 0.54519 | 0.63 | 0.65 | 0.64 | 0.7
| Mean |  | 0.635 | 0.64865 | 0.63437 | 0.65261 | 0.44033 | 0.6824 | 0.63667 | 0.65605 | 0.60833 | 0.6405 | 0.68533 | 0.64417 | 0.6245 | 0.63267 | 0.595 | 0.65898 | 0.644 | 0.647 | 0.676 | 0.60128 | 0.62733 | 0.63034 | 0.63833 | 0.53707 | 0.63 | 0.63167 | 0.63667 | 0.71667
| Median |  | 0.635 | 0.64935 | 0.63433 | 0.65332 | 0.4745 | 0.6834 | 0.6355 | 0.6567 | 0.61 | 0.637 | 0.6885 | 0.6455 | 0.6255 | 0.6245 | 0.595 | 0.6594 | 0.6425 | 0.6425 | 0.6785 | 0.59856 | 0.627 | 0.62744 | 0.6385 | 0.53824 | 0.63 | 0.63 | 0.64 | 0.7
| Std Dev. |  | 0.0054772 | 0.0071226 | 0.0028936 | 0.0029936 | 0.18549 | 0.0045166 | 0.0072847 | 0.0081272 | 0.0075277 | 0.010407 | 0.019582 | 0.010304 | 0.0083606 | 0.01879 | 0.0054772 | 0.0035908 | 0.0055136 | 0.016149 | 0.009654 | 0.012761 | 0.0026583 | 0.0077896 | 0.0087101 | 0.012647 | 0.0063246 | 0.018348 | 0.005164 | 0.040825
| Rel.Std.Dev. |  | 0.8626% | 1.098% | 0.4561% | 0.4587% | 42.12% | 0.6619% | 1.144% | 1.239% | 1.237% | 1.625% | 2.857% | 1.6% | 1.339% | 2.97% | 0.9205% | 0.5449% | 0.8562% | 2.496% | 1.428% | 2.122% | 0.4237% | 1.236% | 1.365% | 2.355% | 1.004% | 2.905% | 0.8111% | 5.696%
| PDM3 |  | -0.4378% | 1.702% | -0.536% | 2.324% | -30.96% | 6.994% | -0.1765% | 2.863% | -4.619% | 0.4246% | 7.454% | 0.9995% | -2.084% | -0.8036% | -6.709% | 3.323% | 0.9733% | 1.444% | 5.991% | -5.725% | -1.64% | -1.168% | 0.08487% | -15.79% | -1.222% | -0.9604% | -0.1765% | 12.37%
| Z-Score (Absolute) |  | 0.1 | 0.72 | 0.14 | 0.96 | 11.86 | 2.76 | 0 | 1.17 | 1.71 | 0.23 | 2.94 | 0.45 | 0.74 | 0.24 | 2.52 | 1.35 | 0.44 | 0.62 | 2.38 | 2.14 | 0.56 | 0.38 | 0.1 | 6.02 | 0.4 | 0.3 | 0 | 4.83
| Analytical results for Tl in OREAS 231b (Certified Value 0.12 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.1 | 0.13 | 0.17146 | 0.11 | 0.13 | 0.14 | 0.1 | 0.13 | 0.14 | 0.11 | 0.11 | < 5 | 0.12 | < 0.5 | 0.13 | 0.28 | 0.1 | 0.13354 | 0.13 | 0.1652 | 0.07 | 0.13 | 0.14 | 0.13 | [5 empty] | 1
| 1 | 2 | 0.09 | 0.13 | 0.15801 | 0.11 | 0.13 | 0.13 | < 0.1 | 0.12 | 0.13 | 0.11 | 0.13 | < 5 | 0.13 | < 0.5 | 0.13 | < 0.1 | 0.15 | 0.13242 | 0.14 | 0.1465 | 0.08 | 0.12 | 0.13 | 0.12 | [5 empty] | 27
| 1 | 3 | 0.14 | 0.14 | 0.144 | 0.12 | 0.13 | 0.13 | < 0.1 | 0.12 | 0.13 | 0.13 | 0.12 | < 5 | 0.13 | < 0.5 | 0.12 | 0.86 | 0.1 | 0.13288 | 0.11 | 0.1466 | 0.07 | 0.12 | 0.13 | 0.12 | [5 empty] | 16
| 1 | 4 | 0.11 | 0.12 | 0.14053 | 0.12 | 0.12 | 0.12 | < 0.1 | 0.12 | 0.13 | 0.11 | 0.11 | < 5 | 0.13 | < 0.5 | 0.13 | < 0.1 | 0.1 | 0.12916 | 0.12 | 0.1466 | 0.08 | 0.12 | 0.13 | 0.12 | [5 empty] | 0.12208
| 1 | 5 | 0.09 | 0.13 | 0.16292 | 0.11 | 0.14 | 0.14 | < 0.1 | 0.12 | 0.13 | 0.11 | 0.11 | < 5 | 0.12 | < 0.5 | 0.12 | 1.45 | 0.1 | 0.12774 | 0.12 | 0.1428 | 0.08 | 0.12 | 0.13 | 0.11 | [5 empty] | 64
| 1 | 6 | 0.1 | 0.14 | 0.14294 | 0.12 | 0.13 | 0.13 | < 0.1 | 0.13 | 0.14 | 0.13 | 0.11 | < 5 | 0.12 | < 0.5 | 0.12 | 0.37 | 0.1 | 0.12992 | 0.13 | 0.1643 | 0.08 | 0.12 | 0.13 | 0.12
| Mean |  | 0.105 | 0.13167 | 0.15331 | 0.115 | 0.13 | 0.13167 | 0.1 | 0.12333 | 0.13333 | 0.11667 | 0.115 |  | 0.125 |  | 0.125 | 0.74 | 0.10833 | 0.13094 | 0.125 | 0.152 | 0.076667 | 0.12167 | 0.13167 | 0.12
| Median |  | 0.1 | 0.13 | 0.151 | 0.115 | 0.13 | 0.13 | 0.1 | 0.12 | 0.13 | 0.11 | 0.11 |  | 0.125 |  | 0.125 | 0.615 | 0.1 | 0.13117 | 0.125 | 0.1466 | 0.08 | 0.12 | 0.13 | 0.12
| Std Dev. |  | 0.018708 | 0.0075277 | 0.012662 | 0.0054772 | 0.0063246 | 0.0075277 |  | 0.005164 | 0.005164 | 0.010328 | 0.0083666 |  | 0.0054772 |  | 0.0054772 | 0.53759 | 0.020412 | 0.0023308 | 0.010488 | 0.0099874 | 0.005164 | 0.0040825 | 0.0040825 | 0.0063246
| Rel.Std.Dev. |  | 17.82% | 5.717% | 8.259% | 4.763% | 4.865% | 5.717% |  | 4.187% | 3.873% | 8.853% | 7.275% |  | 4.382% |  | 4.382% | 72.65% | 18.84% | 1.78% | 8.39% | 6.571% | 6.736% | 3.355% | 3.101% | 5.27%
| PDM3 |  | -13.99% | 7.851% | 25.58% | -5.801% | 6.485% | 7.851% | -18.09% | 1.025% | 9.216% | -4.436% | -5.801% |  | 2.39% |  | 2.39% | 506.1% | -11.26% | 7.258% | 2.39% | 24.51% | -37.2% | -0.3407% | 7.851% | -1.706%
| Z-Score (Absolute) |  | 1.62 | 0.54 | 2.29 | 0.81 | 0.4 | 0.54 | 5.39 | 0.13 | 0.67 | 0.67 | 0.81 | 192.18 | 0 | 10.11 | 0 | 31.15 | 1.35 | 0.48 | 0 | 2.18 | 3.91 | 0.27 | 0.54 | 0.4
| Analytical results for Tm in OREAS 231b (Certified Value 0.33 ppm) | [30 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | [22 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [22 empty] | 2
3
| 1 | 1 | 0.34 | 0.29 | 0.32176 | 0.3 | 0.32 | 0.35 | 0.35476 | [22 empty] | 1
| 1 | 2 | 0.34 | 0.33 | 0.31876 | 0.3 | 0.32 | 0.35 | 0.35486 | [22 empty] | 28
| 1 | 3 | 0.34 | 0.34 | 0.3103 | 0.3 | 0.33 | 0.4 | 0.36015 | [22 empty] | 16
| 1 | 4 | 0.33 | 0.3 | 0.30976 | 0.3 | 0.32 | 0.35 | 0.36469 | [22 empty] | 0.33216
| 1 | 5 | 0.34 | 0.3 | 0.31508 | 0.3 | 0.31 | 0.3 | 0.34959 | [22 empty] | 65
| 1 | 6 | 0.35 | 0.3 | 0.30845 | 0.3 | 0.33 | 0.35 | 0.3597
| Mean |  | 0.34 | 0.31 | 0.31402 | 0.3 | 0.32167 | 0.35 | 0.35729
| Median |  | 0.34 | 0.3 | 0.31269 | 0.3 | 0.32 | 0.35 | 0.35728
| Std Dev. |  | 0.0063246 | 0.02 | 0.0054112 | 0 | 0.0075277 | 0.031623 | 0.005298
| Rel.Std.Dev. |  | 1.86% | 6.452% | 1.723% | 0% | 2.34% | 9.035% | 1.483%
| PDM3 |  | 2.359% | -6.672% | -5.463% | -9.683% | -3.16% | 5.37% | 7.565%
| Z-Score (Absolute) |  | 0.34 | 0.78 | 0.63 | NA | 0.34 | 0.72 | 0.99
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 231b (Certified Value 0.2 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.3 | 0.21 | 0.21 | 0.16233 | 0.2292 | 0.2 | 0.27 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.17 | 0.2 | 0.22 | 0.2 | 0.20254 | 0.2 | 0.2173 | 0.225 | 0.3 | 0.2 | 0.192 |  |  |  | 1
| 1 | 2 | 0.4 | 0.2 | 0.22 | 0.15985 | 0.25386 | 0.2 | 0.29 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.18 | 0.2 | 0.22 | 0.2 | 0.21374 | 0.2 | 0.2083 | 0.202 | 0.2 | 0.2 | 0.185 |  |  |  | 29
| 1 | 3 | 0.2 | 0.21 | 0.34 | 0.17015 | 0.21144 | 0.2 | 0.4 | 0.22 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.18 | 0.2 | 0.21 | 0.2 | 0.20592 | 0.2 | 0.217 | 0.203 | 0.2 | 0.2 | 0.186 |  |  |  | 16
| 1 | 4 | 0.3 | 0.2 | 0.21 | 0.16721 | 0.24139 | 0.2 | 0.4 | 0.2 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.18 | 0.2 | 0.21 | 0.2 | 0.20022 | 0.2 | 0.22 | 0.201 | 0.2 | 0.2 | 0.18 |  |  |  | 0.20272
| 1 | 5 | 0.3 | 0.21 | 0.21 | 0.15872 | 0.22225 | 0.2 | 0.27 | 0.2 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.16 | 0.2 | 0.21 | 0.2 | 0.19629 | 0.2 | 0.2204 | 0.201 | 0.2 | 0.2 | 0.187 |  |  |  | 66
| 1 | 6 | 0.4 | 0.22 | 0.2 | 0.15935 | 0.2104 | 0.2 | 0.24 | 0.22 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 10 | 0.16 | 0.2 | 0.22 | 0.2 | 0.19824 | 0.2 | 0.2884 | 0.201 | 0.2 | 0.2 | 0.19
| Mean |  | 0.31667 | 0.20833 | 0.23167 | 0.16293 | 0.22809 | 0.2 | 0.31167 | 0.20833 | 0.205 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 |  | 0.17167 | 0.2 | 0.215 | 0.2 | 0.20283 | 0.2 | 0.22857 | 0.2055 | 0.21667 | 0.2 | 0.18667
| Median |  | 0.3 | 0.21 | 0.21 | 0.16109 | 0.22573 | 0.2 | 0.28 | 0.205 | 0.205 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 |  | 0.175 | 0.2 | 0.215 | 0.2 | 0.20138 | 0.2 | 0.21865 | 0.2015 | 0.2 | 0.2 | 0.1865
| Std Dev. |  | 0.075277 | 0.0075277 | 0.053448 | 0.0047059 | 0.017126 | 3.0405e-17 | 0.070261 | 0.0098319 | 0.0054772 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 |  | 0.0098319 | 3.0405e-17 | 0.0054772 | 3.0405e-17 | 0.0063137 | 3.0405e-17 | 0.029636 | 0.0095864 | 0.040825 | 3.0405e-17 | 0.0041793
| Rel.Std.Dev. |  | 23.77% | 3.613% | 23.07% | 2.888% | 7.508% | 1.52e-14% | 22.54% | 4.719% | 2.672% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% |  | 5.727% | 1.52e-14% | 2.548% | 1.52e-14% | 3.113% | 1.52e-14% | 12.97% | 4.665% | 18.84% | 1.52e-14% | 2.239%
| PDM3 |  | 56.21% | 2.77% | 14.28% | -19.62% | 12.52% | -1.341% | 53.74% | 2.77% | 1.126% | -1.341% | -1.341% | -1.341% | -1.341% | -1.341% |  | -15.32% | -1.341% | 6.059% | -1.341% | 0.05282% | -1.341% | 12.75% | 1.372% | 6.881% | -1.341% | -7.918%
| Z-Score (Absolute) |  | 14.12 | 0.85 | 3.71 | 4.71 | 3.27 | 0.17 | 13.51 | 0.85 | 0.44 | 0.17 | 0.17 | 0.17 | 0.17 | 0.17 | 587.87 | 3.64 | 0.17 | 1.66 | 0.17 | 0.17 | 0.17 | 3.33 | 0.5 | 1.87 | 0.17 | 1.81
| Analytical results for V in OREAS 231b (Certified Value 283 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 296 | 290 | 252.1 | 284.6957 | 279.7415 | 106 | 350 | 264 | 274 | 279 | 298 | 304 | 281 | 279 | 273 | 287 | 285 | 289 | 292 | 230.86 | 274.5004 | 280 | 292.6265 | 279 | 247.6557 | 286 | 276 | 298 | 274 | 1
| 1 | 2 | 262 | 284 | 256.04 | 280.9411 | 279.4298 | 274 | 348 | 255 | 282 | 282 | 279 | 303 | 289 | 280 | 296 | 281 | 285 | 285 | 293 | 231.28 | 285.108 | 275 | 302.6284 | 275 | 251.7548 | 286 | 285 | 300 | 268 | 30
| 1 | 3 | 310 | 293 | 254.95 | 289.1279 | 278.5052 | 292 | 352 | 263 | 292 | 280 | 274 | 306 | 281 | 282 | 275 | 289 | 296 | 287 | 286 | 225.06 | 284.5928 | 280 | 291.2775 | 277 | 251.7518 | 296 | 269 | 299 | 268 | 16
| 1 | 4 | 279 | 288 | 253.99 | 280.9806 | 277.5286 | 143 | 354 | 266 | 292 | 279 | 278 | 307 | 280 | 277 | 275 | 282 | 301 | 287 | 303 | 231.7 | 281.3636 | 280 | 293.2931 | 282 | 259.6563 | 292 | 277 | 300 | 257 | 282.8955
| 1 | 5 | 290 | 288 | 253.34 | 278.8175 | 279.762 | 168 | 355 | 272 | 282 | 281 | 273 | 287 | 274 | 278 | 285 | 283 | 294 | 284 | 307 | 218.9 | 277.5824 | 280 | 297.1989 | 287 | 261.8501 | 293 | 273 | 302 | 265 | 67
| 1 | 6 | 299 | 279 | 255.94 | 277.9743 | 281.8737 | 263 | 348 | 260 | 237 | 280 | 278 | 303 | 282 | 282 | 282 | 281 | 297 | 284 | 288 | 219.57 | 272.458 | 280 | 290.1012 | 283 | 258.4981 | 282 | 283 | 303 | 249
| Mean |  | 289.3333 | 287 | 254.3933 | 282.0895 | 279.4735 | 207.6667 | 351.1667 | 263.3333 | 276.5 | 280.1667 | 280 | 301.6667 | 281.1667 | 279.6667 | 281 | 283.8333 | 293 | 286 | 294.8333 | 226.2283 | 279.2675 | 279.1667 | 294.5209 | 280.5 | 255.1945 | 289.1667 | 277.1667 | 300.3333 | 263.5
| Median |  | 293 | 288 | 254.47 | 280.9608 | 279.5857 | 215.5 | 351 | 263.5 | 282 | 280 | 278 | 303.5 | 281 | 279.5 | 278.5 | 282.5 | 295 | 286 | 292.5 | 227.96 | 279.473 | 280 | 292.9598 | 280.5 | 255.1265 | 289 | 276.5 | 300 | 266.5
| Std Dev. |  | 16.8483 | 4.899 | 1.5456 | 4.1598 | 1.4586 | 78.3139 | 2.9944 | 5.7155 | 20.5305 | 1.169 | 9.1433 | 7.3666 | 4.7924 | 2.0656 | 8.6948 | 3.3714 | 6.603 | 2 | 8.3766 | 5.9382 | 5.2687 | 2.0412 | 4.6487 | 4.3704 | 5.5792 | 5.3072 | 6.0139 | 1.8619 | 9.0056
| Rel.Std.Dev. |  | 5.823% | 1.707% | 0.6076% | 1.475% | 0.5219% | 37.71% | 0.8527% | 2.17% | 7.425% | 0.4173% | 3.265% | 2.442% | 1.704% | 0.7386% | 3.094% | 1.188% | 2.254% | 0.6993% | 2.841% | 2.625% | 1.887% | 0.7312% | 1.578% | 1.558% | 2.186% | 1.835% | 2.17% | 0.6199% | 3.418%
| PDM3 |  | 2.276% | 1.451% | -10.08% | -0.2849% | -1.21% | -26.59% | 24.13% | -6.915% | -2.261% | -0.9646% | -1.024% | 6.635% | -0.6111% | -1.141% | -0.67% | 0.3315% | 3.572% | 1.097% | 4.22% | -20.03% | -1.282% | -1.318% | 4.109% | -0.8468% | -9.792% | 2.217% | -2.025% | 6.164% | -6.856%
| Z-Score (Absolute) |  | 0.92 | 0.67 | 2.71 | 0.16 | 0.11 | 7.56 | 7.33 | 1.78 | 0.41 | 0.03 | 0.05 | 2.2 | 0.07 | 0.09 | 0.05 | 0.35 | 1.3 | 0.57 | 1.49 | 5.63 | 0.13 | 0.14 | 1.45 | 0 | 2.63 | 0.9 | 0.35 | 2.06 | 1.76
| Analytical results for W in OREAS 231b (Certified Value 17 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 18 | 16.7 | 18.54 | 11.7006 | 16.6269 | < 0.1 | 16.6 | 16.4 | 17.1 | 17 | 3.4 | 16.9 | 16.2 | 14.4 | 16.9 | 14 | 14.1 | 14.8 | 17 | 19 | 16.7745 | 18.2 | 17.1827 | 18.1 | 16.5 | 19.5 | 17 |  |  | 1
| 1 | 2 | 18.1 | 16.7 | 19.14 | 11.7675 | 16.8887 | 12.2 | 16.6 | 16.3 | 16.4 | 17 | 4.6 | 17.1 | 16.9 | 15.5 | 17.6 | 14 | 15.6 | 15 | 17 | 19 | 16.528 | 16.5 | 17.6154 | 18 | 17.3 | 19 | 15.9 |  |  | 31
| 1 | 3 | 19.4 | 17 | 19.48 | 11.8572 | 16.9144 | 11.4 | 16.5 | 16.2 | 16.8 | 17 | 2.8 | 17.1 | 15.6 | 15.6 | 17.4 | 13 | 15.3 | 14.8 | 17 | 19 | 16.8216 | 17.9 | 17.7811 | 17.5 | 16.2 | 19.1 | 16.1 |  |  | 16
| 1 | 4 | 19.2 | 16.5 | 19.7 | 11.6674 | 17.0675 | 0.4 | 16.8 | 16.2 | 16.8 | 17 | 4 | 17.1 | 15.6 | 15.9 | 16.6 | 14 | 16 | 14.9 | 18 | 19 | 16.8577 | 17.7 | 18.1301 | 18 | 16.4 | 18.2 | 16 |  |  | 16.9595
| 1 | 5 | 18.2 | 17.1 | 19.51 | 11.8904 | 17.15 | 0.1 | 16.6 | 16.1 | 16.4 | 17 | 2.4 | 16.5 | 15.4 | 15 | 16.6 | 12 | 14.3 | 15 | 18 | 19 | 16.6304 | 18.2 | 18.1004 | 18.3 | 16.2 | 19.1 | 16.5 |  |  | 68
| 1 | 6 | 18.8 | 16.6 | 20.69 | 11.6897 | 17.3366 | 5.1 | 16.1 | 16.1 | 16.8 | 17 | 2.7 | 17.4 | 16.2 | 15.7 | 17.1 | 14 | 14.3 | 15 | 17 | 19 | 16.8145 | 18.1 | 17.9998 | 18.3 | 17 | 18.7 | 16.2
| Mean |  | 18.6167 | 16.7667 | 19.51 | 11.7621 | 16.9973 | 5.84 | 16.5333 | 16.2167 | 16.7167 | 17 | 3.3167 | 17.0167 | 15.9833 | 15.35 | 17.0333 | 13.5 | 14.9333 | 14.9167 | 17.3333 | 19 | 16.7378 | 17.7667 | 17.8016 | 18.0333 | 16.6 | 18.9333 | 16.2833
| Median |  | 18.5 | 16.7 | 19.495 | 11.734 | 16.9909 | 5.1 | 16.6 | 16.2 | 16.8 | 17 | 3.1 | 17.1 | 15.9 | 15.55 | 17 | 14 | 14.8 | 14.95 | 17 | 19 | 16.7945 | 18 | 17.8905 | 18.05 | 16.45 | 19.05 | 16.15
| Std Dev. |  | 0.60139 | 0.23381 | 0.70812 | 0.093298 | 0.24466 | 5.7977 | 0.23381 | 0.1169 | 0.27142 | 0 | 0.84951 | 0.29944 | 0.56006 | 0.55408 | 0.41312 | 0.83666 | 0.80166 | 0.098319 | 0.5164 | 0 | 0.12977 | 0.65013 | 0.36183 | 0.29439 | 0.45166 | 0.44121 | 0.40702
| Rel.Std.Dev. |  | 3.23% | 1.394% | 3.63% | 0.7932% | 1.439% | 99.28% | 1.414% | 0.7209% | 1.624% | 0% | 25.61% | 1.76% | 3.504% | 3.61% | 2.425% | 6.197% | 5.368% | 0.6591% | 2.979% | 0% | 0.7753% | 3.659% | 2.033% | 1.632% | 2.721% | 2.33% | 2.5%
| PDM3 |  | 9.771% | -1.137% | 15.04% | -30.65% | 0.2232% | -65.57% | -2.513% | -4.38% | -1.432% | 0.2388% | -80.44% | 0.3371% | -5.756% | -9.49% | 0.4353% | -20.4% | -11.95% | -12.05% | 2.204% | 12.03% | -1.307% | 4.759% | 4.965% | 6.332% | -2.12% | 11.64% | -3.987%
| Z-Score (Absolute) |  | 1.22 | 0.03 | 1.79 | 3.18 | 0.18 | 7.6 | 0.12 | 0.32 | 0 | NA | 8.61 | 0.19 | 0.47 | 0.88 | 0.2 | NA | 1.15 | 1.16 | NA | NA | 0.01 | 0.67 | 0.7 | 0.85 | 0.07 | 1.42 | 0.28
Results from laboratories 1.12, 1.19, 1.24 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 231b (Certified Value 20.5 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 18.9 | 20.07 | 17.36 | 19.0609 | 20.0678 | 18.5 | 20.3 | 22.2 | 20.44 | 20 | 21.4 | 22.7 | 21.2 | 19.6 | 21.1 | 19 | 18.2 | 19.4 | 21 | 20.9 | 21.0114 | 21.8 | 20.6194 | 21.4 | 20 | 23.9 | 18.8 |  |  | 1
| 1 | 2 | 18.7 | 20.57 | 18.69 | 18.7109 | 20.0941 | 19 | 20.2 | 21.7 | 20.12 | 20 | 21.1 | 23.2 | 21.4 | 24.8 | 22.9 | 19 | 20.1 | 19.5 | 19.5 | 21.2 | 21.3015 | 20.2 | 21.0025 | 21 | 21.1 | 23.6 | 17.8 |  |  | 32
| 1 | 3 | 19.2 | 20.83 | 19.29 | 18.9097 | 19.8499 | 19.9 | 19.9 | 21.6 | 20.47 | 20 | 21.3 | 23 | 20.8 | 22.2 | 22 | 19 | 21.2 | 20.2 | 20.2 | 21.9 | 20.9968 | 22.2 | 21.4114 | 20.8 | 19.5 | 24 | 18.2 |  |  | 16
| 1 | 4 | 18.7 | 19.98 | 18.17 | 19.0779 | 19.9336 | 20.1 | 20.8 | 22.4 | 20.5 | 20 | 21.2 | 22.3 | 20.6 | 20.8 | 21.3 | 19 | 20 | 19.4 | 21 | 20.9 | 20.7837 | 21.1 | 21.6989 | 21.3 | 20.4 | 22.5 | 17.5 |  |  | 20.4727
| 1 | 5 | 19.1 | 20.69 | 18.32 | 19.0062 | 20.0367 | 19 | 20 | 22.4 | 19.74 | 20 | 20.3 | 22 | 21.6 | 19.9 | 22.3 | 19 | 20.1 | 19.4 | 20.5 | 20.8 | 20.5539 | 20.9 | 21.4598 | 21.8 | 19.8 | 23.6 | 18.6 |  |  | 69
| 1 | 6 | 19.7 | 20.91 | 18.29 | 19.2834 | 20.3075 | 19.9 | 19.4 | 21 | 20.34 | 20 | 21.3 | 22.7 | 20.9 | 21.5 | 21.8 | 19 | 18.5 | 19.2 | 21.3 | 21.7 | 20.7941 | 22.5 | 20.9578 | 20.7 | 20.3 | 22.9 | 18
| Mean |  | 19.05 | 20.5083 | 18.3533 | 19.0082 | 20.0483 | 19.4 | 20.1 | 21.8833 | 20.2683 | 20 | 21.1 | 22.65 | 21.0833 | 21.4667 | 21.9 | 19 | 19.6833 | 19.5167 | 20.5833 | 21.2333 | 20.9069 | 21.45 | 21.1916 | 21.1667 | 20.1833 | 23.4167 | 18.15
| Median |  | 19 | 20.63 | 18.305 | 19.0335 | 20.0523 | 19.45 | 20.1 | 21.95 | 20.39 | 20 | 21.25 | 22.7 | 21.05 | 21.15 | 21.9 | 19 | 20.05 | 19.4 | 20.75 | 21.05 | 20.8955 | 21.45 | 21.2069 | 21.15 | 20.15 | 23.6 | 18.1
| Std Dev. |  | 0.37815 | 0.39316 | 0.635 | 0.19049 | 0.1564 | 0.65115 | 0.46476 | 0.55287 | 0.29329 | 0 | 0.40497 | 0.44159 | 0.38166 | 1.8991 | 0.6603 | 0 | 1.1268 | 0.34881 | 0.66156 | 0.46332 | 0.25585 | 0.86891 | 0.39886 | 0.41312 | 0.55648 | 0.59133 | 0.48888
| Rel.Std.Dev. |  | 1.985% | 1.917% | 3.46% | 1.002% | 0.7801% | 3.356% | 2.312% | 2.526% | 1.447% | 0% | 1.919% | 1.95% | 1.81% | 8.847% | 3.015% | 0% | 5.725% | 1.787% | 3.214% | 2.182% | 1.224% | 4.051% | 1.882% | 1.952% | 2.757% | 2.525% | 2.694%
| PDM3 |  | -6.949% | 0.1739% | -10.35% | -7.154% | -2.073% | -5.24% | -1.821% | 6.89% | -0.9984% | -2.309% | 3.064% | 10.63% | 2.982% | 4.855% | 6.972% | -7.194% | -3.856% | -4.67% | 0.5402% | 3.715% | 2.121% | 4.773% | 3.511% | 3.39% | -1.414% | 14.38% | -11.35%
| Z-Score (Absolute) |  | 1.19 | 0 | 1.76 | 1.23 | 0.38 | 0.91 | 0.33 | 1.12 | 0.2 | 0.42 | 0.48 | 1.75 | 0.47 | 0.78 | 1.14 | 1.23 | 0.67 | 0.81 | 0.06 | 0.59 | 0.33 | 0.77 | 0.56 | 0.54 | 0.27 | 2.38 | 1.93
| Analytical results for Yb in OREAS 231b (Certified Value 2.13 ppm) | [30 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 2.4 | 2.15 | 1.83 | 2.0079 | 2 | 2.04 | 2.2 | 2 | 2.3705 | 2 | [19 empty] | 1
| 1 | 2 | 2.4 | 2.1 | 1.95 | 2.01 | 2.1 | 2.07 | 2.2 | 2.1 | 2.3392 | 2.2 | [19 empty] | 33
| 1 | 3 | 2.4 | 2.23 | 2.03 | 1.9948 | 2.2 | 2 | 2.4 | 2 | 2.321 | 2 | [19 empty] | 16
| 1 | 4 | 2.4 | 2.1 | 1.82 | 2.0586 | 2.3 | 2.09 | 2.3 | 2 | 2.3141 | 2 | [19 empty] | 2.1304
| 1 | 5 | 2.4 | 2.14 | 1.85 | 2.0701 | 2.1 | 2.01 | 2.3 | 2 | 2.2698 | 2 | [19 empty] | 70
| 1 | 6 | 2.5 | 2.25 | 1.86 | 2.0055 | 2.2 | 1.96 | 2.3 | 1.9 | 2.3494 | 2.1
| Mean |  | 2.4167 | 2.1617 | 1.89 | 2.0245 | 2.15 | 2.0283 | 2.2833 | 2 | 2.3273 | 2.05
| Median |  | 2.4 | 2.145 | 1.855 | 2.009 | 2.15 | 2.025 | 2.3 | 2 | 2.3301 | 2
| Std Dev. |  | 0.040825 | 0.064317 | 0.082704 | 0.031524 | 0.10488 | 0.047924 | 0.075277 | 0.063246 | 0.034673 | 0.083666
| Rel.Std.Dev. |  | 1.689% | 2.975% | 4.376% | 1.557% | 4.878% | 2.363% | 3.297% | 3.162% | 1.49% | 4.081%
| PDM3 |  | 13.44% | 1.468% | -11.28% | -4.97% | 0.9208% | -4.79% | 7.179% | -6.12% | 9.245% | -3.773%
| Z-Score (Absolute) |  | 2.43 | 0.47 | 1.61 | 0.58 | 0.38 | 0.55 | 1.41 | 0.77 | 1.75 | 0.38
| Analytical results for Zn in OREAS 231b (Certified Value 118 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 120 | 123 | 116.94 | 118.48 | 118.4832 | 108 | 127 | 127 | 122 | 111 | 112 | 130 | 122 | 122 | 123 | 114 | 125 | 119 | 113 | 66.63 | 105.0493 | 120 | 118.079 | 126 | 143.0962 | 110 | 116 | 108 | 115 | 1
| 1 | 2 | 129 | 123 | 117.27 | 115.79 | 118.8012 | 110 | 125 | 118 | 119 | 111 | 110 | 131 | 125 | 119 | 132 | 112 | 127 | 113 | 111 | 67.84 | 108.7623 | 120 | 121.2986 | 125 | 145.6891 | 109 | 118 | 108 | 110 | 34
| 1 | 3 | 145 | 127 | 116.66 | 120.09 | 118.905 | 111 | 125 | 118 | 121 | 114 | 110 | 132 | 121 | 124 | 121 | 114 | 127 | 115 | 111 | 67.57 | 108.3064 | 110 | 117.0381 | 125 | 150.111 | 110 | 111 | 109 | 112 | 16
| 1 | 4 | 137 | 124 | 116.7 | 121.05 | 118.9524 | 111 | 127 | 120 | 121 | 110 | 112 | 134 | 122 | 121 | 123 | 113 | 125 | 115 | 117 | 71.23 | 103.7377 | 110 | 117.4317 | 124 | 152.824 | 111 | 115 | 109 | 110 | 117.7622
| 1 | 5 | 120 | 127 | 116.06 | 118.77 | 120.9222 | 112 | 123 | 121 | 120 | 110 | 112 | 124 | 119 | 121 | 126 | 111 | 126 | 117 | 119 | 73.6 | 101.8947 | 110 | 119.5628 | 127 | 152.5975 | 111 | 112 | 110 | 116 | 71
| 1 | 6 | 122 | 127 | 117.77 | 121.66 | 119.7549 | 113 | 120 | 115 | 120 | 114 | 112 | 130 | 124 | 121 | 123 | 111 | 126 | 115 | 112 | 61.91 | 100.1681 | 120 | 117.2855 | 127 | 154.8096 | 109 | 115 | 110 | 112
| Mean |  | 128.8333 | 125.1667 | 116.9 | 119.3067 | 119.3032 | 110.8333 | 124.5 | 119.8333 | 120.5 | 111.6667 | 111.3333 | 130.1667 | 122.1667 | 121.3333 | 124.6667 | 112.5 | 126 | 115.6667 | 113.8333 | 68.13 | 104.6531 | 115 | 118.4493 | 125.6667 | 149.8546 | 110 | 114.5 | 109 | 112.5
| Median |  | 125.5 | 125.5 | 116.82 | 119.43 | 118.9287 | 111 | 125 | 119 | 120.5 | 111 | 112 | 130.5 | 122 | 121 | 123 | 112.5 | 126 | 115 | 112.5 | 67.705 | 104.3935 | 115 | 117.7553 | 125.5 | 151.3542 | 110 | 115 | 109 | 112
| Std Dev. |  | 10.3037 | 2.0412 | 0.58286 | 2.1227 | 0.89792 | 1.7224 | 2.6646 | 4.0702 | 1.0488 | 1.8619 | 1.0328 | 3.3714 | 2.137 | 1.633 | 3.9328 | 1.3784 | 0.89443 | 2.0656 | 3.3714 | 4.0216 | 3.4337 | 5.4772 | 1.6659 | 1.2111 | 4.5604 | 0.89443 | 2.5884 | 0.89443 | 2.51
| Rel.Std.Dev. |  | 7.998% | 1.631% | 0.4986% | 1.779% | 0.7526% | 1.554% | 2.14% | 3.397% | 0.8704% | 1.667% | 0.9277% | 2.59% | 1.749% | 1.346% | 3.155% | 1.225% | 0.7099% | 1.786% | 2.962% | 5.903% | 3.281% | 4.763% | 1.406% | 0.9637% | 3.043% | 0.8131% | 2.261% | 0.8206% | 2.231%
| PDM3 |  | 9.401% | 6.288% | -0.7321% | 1.312% | 1.309% | -5.884% | 5.722% | 1.759% | 2.325% | -5.176% | -5.459% | 10.53% | 3.74% | 3.033% | 5.863% | -4.468% | 6.995% | -1.779% | -3.336% | -42.15% | -11.13% | -2.346% | 0.5835% | 6.712% | 27.25% | -6.591% | -2.77% | -7.441% | -4.468%
| Z-Score (Absolute) |  | 1.16 | 0.75 | 0.17 | 0.1 | 0.1 | 0.85 | 0.67 | 0.15 | 0.23 | 0.76 | 0.79 | 1.31 | 0.41 | 0.32 | 0.69 | 0.66 | 0.84 | 0.31 | 0.51 | 5.61 | 1.54 | 0.38 | 0 | 0.8 | 3.5 | 0.94 | 0.44 | 1.05 | 0.66
| Analytical results for Zr in OREAS 231b (Certified Value 45.4 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
1
| 1 | 1 | 47.8 | 43.6 | 68.44 | 53.7935 | 46.8453 | 25 | 55 | 49 | 41.8 | 51 | 56.5 | 46.8 | 47.6 | 44.6 | 45.3 | 39 | 37.4 | 96.9 | 42.3 | 45.1444 | 45 | 67.2837 | 50.2 | 29.7045 | 43.3 | 84.9 | 50.5 | 37.4 |  | 1
| 1 | 2 | 67.6 | 43.5 | 70.41 | 52.5546 | 46.544 | 46 | 54 | 46.4 | 44.7 | 52 | 54.7 | 46.3 | 47.4 | 39.7 | 48.7 | 37 | 42.6 | 39.5 | 42.6 | 45.7148 | 45 | 68.8732 | 43.1 | 29.8453 | 42.9 | 79.5 | 53.9 | 35.6 |  | 35
| 1 | 3 | 48.8 | 48.8 | 64.39 | 53.4258 | 45.1093 | 47 | 54 | 44.5 | 43.5 | 52 | 55.7 | 47.5 | 45.5 | 50 | 49.4 | 38 | 40.9 | 43.5 | 41.9 | 46.69 | 45 | 66.7952 | 64.8 | 30.144 | 43.6 | 76.9 | 48.7 | 36.1 |  | 16
| 1 | 4 | 55.2 | 45.3 | 71.59 | 53.2336 | 45.9449 | 14 | 55 | 45.3 | 40.8 | 50 | 53.4 | 45.5 | 44.2 | 47.9 | 52.3 | 36 | 41.9 | 42.9 | 44 | 45.1352 | 45 | 67.3014 | 52.9 | 31.6875 | 44.3 | 73.6 | 54.8 | 35.2 |  | 45.4343
| 1 | 5 | 51.8 | 45.8 | 64.13 | 53.4973 | 46.1571 | 27 | 55 | 48.5 | 41.2 | 52 | 54.6 | 52 | 41.3 | 47.1 | 43.1 | 40 | 38.3 | 41.9 | 45.7 | 44.6292 | 45 | 65.4503 | 52 | 30.2762 | 44.5 | 78.6 | 53.9 | 38 |  | 72
| 1 | 6 | 59.1 | 50.8 | 64.25 | 53.9902 | 46.9207 | 43 | 54 | 49.6 | 41.6 | 52 | 53.9 | 46.6 | 44 | 43.6 | 41.1 | 38 | 38.2 | 41.8 | 43 | 44.2888 | 45 | 66.7419 | 52.1 | 29.3289 | 42.9 | 79.8 | 53.4 | 36.3
| Mean |  | 55.05 | 46.3 | 67.2017 | 53.4158 | 46.2536 | 33.6667 | 54.5 | 47.2167 | 42.2667 | 51.5 | 54.8 | 47.45 | 45 | 45.4833 | 46.65 | 38 | 39.8833 | 51.0833 | 43.25 | 45.2671 | 45 | 67.0743 | 52.5167 | 30.1644 | 43.5833 | 78.8833 | 52.5333 | 36.4333
| Median |  | 53.5 | 45.55 | 66.415 | 53.4616 | 46.3505 | 35 | 54.5 | 47.45 | 41.7 | 52 | 54.65 | 46.7 | 44.85 | 45.85 | 47 | 38 | 39.6 | 42.4 | 42.8 | 45.1398 | 45 | 67.0394 | 52.05 | 29.9947 | 43.45 | 79.05 | 53.65 | 36.2
| Std Dev. |  | 7.4312 | 2.9285 | 3.3805 | 0.50084 | 0.67726 | 13.5892 | 0.54772 | 2.1085 | 1.5095 | 0.83666 | 1.1419 | 2.3227 | 2.3707 | 3.6504 | 4.217 | 1.4142 | 2.1904 | 22.487 | 1.3982 | 0.85106 | 0 | 1.1097 | 7.0073 | 0.81782 | 0.68823 | 3.7242 | 2.3855 | 1.0708
| Rel.Std.Dev. |  | 13.5% | 6.325% | 5.03% | 0.9376% | 1.464% | 40.36% | 1.005% | 4.466% | 3.571% | 1.625% | 2.084% | 4.895% | 5.268% | 8.026% | 9.04% | 3.722% | 5.492% | 44.02% | 3.233% | 1.88% | 0% | 1.654% | 13.34% | 2.711% | 1.579% | 4.721% | 4.541% | 2.939%
| PDM3 |  | 21.16% | 1.906% | 47.91% | 17.57% | 1.803% | -25.9% | 19.95% | 3.923% | -6.972% | 13.35% | 20.61% | 4.437% | -0.9558% | 0.108% | 2.676% | -16.36% | -12.22% | 12.43% | -4.807% | -0.368% | -0.9558% | 47.63% | 15.59% | -33.61% | -4.074% | 73.62% | 15.62% | -19.81%
| Z-Score (Absolute) |  | 1.04 | 0.02 | 2.53 | 0.85 | 0.03 | 1.56 | 0.98 | 0.09 | 0.51 | 0.61 | 1.01 | 0.12 | 0.18 | 0.12 | 0.02 | 1.03 | 0.8 | 0.56 | 0.39 | 0.15 | 0.18 | 2.51 | 0.74 | 1.99 | 0.35 | 3.95 | 0.74 | 1.22
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C312: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C533: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C552: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C644: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C680: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C698: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C717: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C735: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C753: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C771: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C790: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C882: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C919: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C937: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C975: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C993: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1011: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 231b (Certified Value 0.182 ppm) | [32 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.22 | 0.2 | 0.41 | 0.20646 | 0.18 | 0.21 | 0.15 | 0.16 | < 0.5 | 0.3 | 0.18 | 0.19 | 0.17 | < 10 | 0.18 | < 0.2 | < 1 | 0.16 | 0.168 | 0.17 | < 0.2 | 0.2 | < 0.2 | 0.18 | 0.1452 | 0.16 | 0.2 | 0.17 | 0.17 |  |  | 1
| 1 | 2 | 0.23 | 0.19 | 0.35 | 0.1941 | 0.176 | 0.22 | 0.18 | 0.17 | < 0.5 | 0.2 | 0.18 | 0.19 | 0.17 | < 10 | 0.17 | < 0.2 | < 1 | 0.15 | 0.168 | 0.19 | < 0.2 | 0.2 | < 0.2 | 0.18 | 0.175 | 0.16 | 0.19 | 0.18 | 0.19 |  |  | 19
| 1 | 3 | 0.26 | 0.19 | 0.24 | 0.19183 | 0.18 | 0.2 | 0.18 | 0.17 | < 0.5 | 0.2 | 0.18 | 0.18 | 0.18 | < 10 | 0.18 | < 0.2 | < 1 | 0.16 | 0.162 | 0.19 | < 0.2 | 0.2 | < 0.2 | 0.18 | 0.176 | 0.17 | 0.19 | 0.18 | 0.19 |  |  | 16
| 1 | 4 | 0.24 | 0.18 | 0.29 | 0.19911 | 0.171 | 0.21 | 0.17 | 0.17 | < 0.5 | 0.2 | 0.18 | 0.18 | 0.18 | < 10 | 0.17 | < 0.2 | < 1 | 0.14 | 0.16 | 0.19 | < 0.2 | 0.2 | < 0.2 | 0.19 | 0.233 | 0.15 | 0.19 | 0.17 | 0.18 |  |  | 0.18226
| 1 | 5 | 0.23 | 0.18 | 0.4 | 0.21063 | 0.181 | 0.22 | 0.17 | 0.16 | < 0.5 | 0.2 | 0.17 | 0.18 | 0.18 | < 10 | 0.18 | < 0.2 | < 1 | 0.14 | 0.168 | 0.18 | 0.21 | 0.2 | < 0.2 | 0.18 | 0.193 | 0.16 | 0.19 | 0.18 | 0.19 |  |  | 74
| 1 | 6 | 0.23 | 0.18 | 0.37 | 0.19794 | 0.182 | 0.2 | 0.19 | 0.16 | < 0.5 | 0.2 | 0.17 | 0.19 | 0.18 | < 10 | 0.17 | < 0.2 | < 1 | 0.16 | 0.16 | 0.19 | < 0.2 | 0.2 | < 0.2 | 0.18 | 0.221 | 0.16 | 0.18 | 0.18 | 0.19
| Mean |  | 0.235 | 0.18667 | 0.34333 | 0.20001 | 0.17833 | 0.21 | 0.17333 | 0.165 |  | 0.21667 | 0.17667 | 0.185 | 0.17667 |  | 0.175 |  |  | 0.15167 | 0.16433 | 0.185 | 0.21 | 0.2 |  | 0.18167 | 0.19053 | 0.16 | 0.19 | 0.17667 | 0.185
| Median |  | 0.23 | 0.185 | 0.36 | 0.19853 | 0.18 | 0.21 | 0.175 | 0.165 |  | 0.2 | 0.18 | 0.185 | 0.18 |  | 0.175 |  |  | 0.155 | 0.165 | 0.19 | 0.21 | 0.2 |  | 0.18 | 0.1845 | 0.16 | 0.19 | 0.18 | 0.19
| Std Dev. |  | 0.013784 | 0.008165 | 0.066232 | 0.0072298 | 0.0041312 | 0.0089443 | 0.013663 | 0.0054772 |  | 0.040825 | 0.005164 | 0.0054772 | 0.005164 |  | 0.0054772 |  |  | 0.0098319 | 0.0040825 | 0.0083666 |  | 3.0405e-17 |  | 0.0040825 | 0.032389 | 0.0063246 | 0.0063246 | 0.005164 | 0.0083666
| Rel.Std.Dev. |  | 5.866% | 4.374% | 19.29% | 3.615% | 2.317% | 4.259% | 7.882% | 3.32% |  | 18.84% | 2.923% | 2.961% | 2.923% |  | 3.13% |  |  | 6.483% | 2.484% | 4.522% |  | 1.52e-14% |  | 2.247% | 17% | 3.953% | 3.329% | 2.923% | 4.522%
| PDM3 |  | 28.94% | 2.42% | 88.38% | 9.743% | -2.152% | 15.22% | -4.895% | -9.468% |  | 18.88% | -3.066% | 1.506% | -3.066% |  | -3.981% |  |  | -16.78% | -9.833% | 1.506% | 15.22% | 9.736% |  | -0.323% | 4.542% | -12.21% | 4.249% | -3.066% | 1.506%
| Z-Score (Absolute) |  | 2.16 | 0.2 | 6.54 | 0.74 | 0.13 | 1.15 | 0.34 | 0.67 | 2.76 | NA | 0.2 | 0.13 | 0.2 | 194.94 | 0.27 | 3.3 | 12.88 | 1.21 | 0.7 | 0.13 | 2.56 | NA | 3.3 | 0 | 0.36 | 0.88 | 0.34 | 0.2 | 0.13
Results from laboratories 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 231b (Certified Value 3.13 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.25 | 2.949 | 2.8448 | 3.528 | 3.2171 | 3.02 | 3.371 | 3.58 | 2.9819 | 2.93 | 3.25 | 3.21 | 3.22 | 3.19 | 3.2 | 3.24 | 3.3108 | 3.371 | 3 | 3.47 | 3.4204 | 2.9036 | 2.799 | 3.04 | 2.9 | 3.45 | 2.99 | 2.85 |  |  |  | 1
| 1 | 2 | 3.16 | 3.006 | 2.6157 | 3.5181 | 3.2348 | 2.9 | 3.461 | 3.47 | 2.9812 | 2.93 | 3.3 | 3.2 | 3.22 | 3.14 | 3.21 | 3.15 | 3.204 | 3.551 | 3.1 | 3.39 | 3.1789 | 2.7098 | 2.802 | 3.04 | 2.85 | 3.36 | 3.09 | 2.86 |  |  |  | #N/A
| 1 | 3 | 3.12 | 2.938 | 2.6469 | 3.5677 | 3.2154 | 3.09 | 3.477 | 3.48 | 2.9716 | 2.95 | 3.33 | 3.17 | 3.21 | 3.17 | 3.26 | 3.05 | 3.3395 | 3.4 | 3.01 | 3.41 | 3.5486 | 2.7342 | 2.7775 | 3.02 | 2.86 | 3.35 | 3.04 | 2.83 |  |  |  | 16
| 1 | 4 | 3.1 | 2.914 | 2.8235 | 3.5642 | 3.0737 | 2.85 | 3.414 | 3.51 | 2.8362 | 2.93 | 3.37 | 3.2 | 3.25 | 3.18 | 3.21 | 3.06 | 3.2791 | 3.546 | 2.88 | 3.4 | 2.9955 | 2.7526 | 2.842 | 3.03 | 2.77 | 3.37 | 2.97 | 2.8 |  |  |  | 3.1318
| 1 | 5 | 3.06 | 2.985 | 2.6272 | 3.5086 | 3.3637 | 2.92 | 3.482 | 3.49 | 2.8549 | 2.92 | 3.36 | 3.18 | 3.22 | 3.2 | 3.22 | 3.23 | 3.4216 | 3.504 | 2.88 | 3.49 | 3.5099 | 2.7126 | 2.745 | 3.05 | 2.77 | 3.35 | 2.98 | 2.84 |  |  |  | 75
| 1 | 6 | 3.05 | 2.964 | 2.8945 | 3.5545 | 3.0355 | 2.94 | 3.451 | 3.46 | 2.9904 | 2.94 | 3.34 | 3.17 | 3.21 | 3.19 | 3.23 | 3.21 | 3.4652 | 3.435 | 2.9 | 3.41 | 2.9769 | 2.683 | 2.7325 | 3.05 | 2.85 | 3.39 | 3.01 | 2.91
| Mean |  | 3.1233 | 2.9593 | 2.7421 | 3.5402 | 3.19 | 2.9533 | 3.4427 | 3.4983 | 2.936 | 2.9333 | 3.325 | 3.1883 | 3.2217 | 3.1783 | 3.2217 | 3.1567 | 3.3367 | 3.4678 | 2.9617 | 3.4283 | 3.2717 | 2.7493 | 2.783 | 3.0383 | 2.8333 | 3.3783 | 3.0133 | 2.8483
| Median |  | 3.11 | 2.9565 | 2.7352 | 3.5412 | 3.2163 | 2.93 | 3.456 | 3.485 | 2.9764 | 2.93 | 3.335 | 3.19 | 3.22 | 3.185 | 3.215 | 3.18 | 3.3252 | 3.4695 | 2.95 | 3.41 | 3.2996 | 2.7234 | 2.7883 | 3.04 | 2.85 | 3.365 | 3 | 2.845
| Std Dev. |  | 0.073937 | 0.033092 | 0.12542 | 0.025189 | 0.11912 | 0.087101 | 0.042636 | 0.043551 | 0.070589 | 0.010328 | 0.044159 | 0.017224 | 0.01472 | 0.02137 | 0.02137 | 0.084774 | 0.095207 | 0.076669 | 0.089536 | 0.041191 | 0.25586 | 0.0792 | 0.040305 | 0.01169 | 0.052409 | 0.038166 | 0.045019 | 0.03656
| Rel.Std.Dev. |  | 2.367% | 1.118% | 4.574% | 0.7115% | 3.734% | 2.949% | 1.238% | 1.245% | 2.404% | 0.3521% | 1.328% | 0.5402% | 0.4569% | 0.6724% | 0.6633% | 2.686% | 2.853% | 2.211% | 3.023% | 1.201% | 7.821% | 2.881% | 1.448% | 0.3848% | 1.85% | 1.13% | 1.494% | 1.284%
| PDM3 |  | -0.2696% | -5.506% | -12.44% | 13.04% | 1.86% | -5.698% | 9.927% | 11.7% | -6.25% | -6.336% | 6.17% | 1.806% | 2.87% | 1.487% | 2.87% | 0.7948% | 6.544% | 10.73% | -5.432% | 9.469% | 4.467% | -12.21% | -11.14% | -2.984% | -9.53% | 7.873% | -3.782% | -9.051%
| Z-Score (Absolute) |  | 0.14 | 0.67 | 1.37 | 1.2 | 0.07 | 0.69 | 0.89 | 1.06 | 0.75 | 0.75 | 0.51 | 0.07 | 0.17 | 0.03 | 0.17 | 0.03 | 0.54 | 0.97 | 0.66 | 0.84 | 0.34 | 1.35 | 1.24 | 0.42 | 1.08 | 0.68 | 0.5 | 1.03
| Analytical results for As in OREAS 231b (Certified Value 33.3 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 2
| 1 | 1 | 35 | 33.54 | 31.09 | 32.6724 | 34.078 | 35.1 | 39.2 | 36 | 34 | 31 | 33 | 34.8 | 32.4 | 31.8 | < 100 | 34.5 | 32 | 22.8027 | 30 | 32 | 34 | 24.85 | 35.42 | 35 | 35.4519 | 35.7 | 31.312 | 34.2 | 34 | 34.4 | 33.2 | 1
| 1 | 2 | 35 | 34.2 | 28.09 | 32.0361 | 33.808 | 33.7 | 36.1 | 35 | 34 | 31 | 33.1 | 34.1 | 32.4 | 32.1 | < 100 | 33.4 | 32 | 21.9478 | 31 | 31.9 | 33 | 28.21 | 31.8504 | 35 | 35.697 | 34.8 | 31.377 | 33.7 | 32 | 34.4 | 33.3 | 20
| 1 | 3 | 35 | 33.9 | 27.94 | 31.7955 | 34.2278 | 34.6 | 35 | 35 | 33 | 31 | 32.2 | 34.5 | 34.5 | 34.5 | < 100 | 35.8 | 30 | 19.8397 | 30 | 30.8 | 33 | 27.25 | 31.5836 | 35 | 35.5882 | 35.4 | 29.898 | 34 | 33 | 34.5 | 33.2 | 16
| 1 | 4 | 35 | 33.65 | 30.68 | 32.5093 | 33.457 | 34 | 35.4 | 35 | 32 | 30 | 33.6 | 33.7 | 33.5 | 35 | < 100 | 35.1 | 31 | 21.8082 | 31 | 31.5 | 33 | 27.22 | 31.7492 | 30 | 34.4537 | 35.2 | 31.171 | 35 | 34 | 33.8 | 33.3 | 33.3234
| 1 | 5 | 35 | 33.79 | 27.45 | 32.9001 | 36.8316 | 33.3 | 35.9 | 35 | 33 | 30 | 32.8 | 33.7 | 33.5 | 35.3 | < 100 | 33.6 | 32 | 20.1053 | 31 | 31.3 | 33 | 27.22 | 33.028 | 35 | 37.2218 | 35.6 | 31.194 | 35 | 32 | 34.3 | 32.8 | 76
| 1 | 6 | 35 | 33.63 | 31.37 | 32.2521 | 32.9521 | 33.2 | 35 | 35 | 34 | 30 | 33.2 | 33.7 | 33.8 | 33.8 | < 100 | 35.3 | 32 | 21.733 | 31 | 30.7 | 33 | 27.84 | 32.2092 | 35 | 34.4351 | 35.4 | 30.41 | 34.3 | 33 | 34.7 | 33.6
| Mean |  | 35 | 33.785 | 29.4367 | 32.3609 | 34.2257 | 33.9833 | 36.1 | 35.1667 | 33.3333 | 30.5 | 32.9833 | 34.0833 | 33.35 | 33.75 |  | 34.6167 | 31.5 | 21.3728 | 30.6667 | 31.3667 | 33.1667 | 27.0983 | 32.6401 | 34.1667 | 35.4746 | 35.35 | 30.8937 | 34.3667 | 33 | 34.35 | 33.2333
| Median |  | 35 | 33.72 | 29.385 | 32.3807 | 33.943 | 33.85 | 35.65 | 35 | 33.5 | 30.5 | 33.05 | 33.9 | 33.5 | 34.15 |  | 34.8 | 32 | 21.7706 | 31 | 31.4 | 33 | 27.235 | 32.0298 | 35 | 35.52 | 35.4 | 31.1825 | 34.25 | 33 | 34.4 | 33.25
| Std Dev. |  | 0 | 0.2399 | 1.7898 | 0.41156 | 1.3565 | 0.74677 | 1.5849 | 0.40825 | 0.8165 | 0.54772 | 0.46655 | 0.47504 | 0.82158 | 1.4869 |  | 0.96212 | 0.83666 | 1.1534 | 0.5164 | 0.54283 | 0.40825 | 1.1742 | 1.4559 | 2.0412 | 1.0236 | 0.32094 | 0.60017 | 0.53166 | 0.89443 | 0.30166 | 0.2582
| Rel.Std.Dev. |  | 0% | 0.7101% | 6.08% | 1.272% | 3.963% | 2.197% | 4.39% | 1.161% | 2.449% | 1.796% | 1.414% | 1.394% | 2.464% | 4.406% |  | 2.779% | 2.656% | 5.396% | 1.684% | 1.731% | 1.231% | 4.333% | 4.46% | 5.974% | 2.885% | 0.9079% | 1.943% | 1.547% | 2.71% | 0.8782% | 0.7769%
| PDM3 |  | 5.031% | 1.385% | -11.66% | -2.888% | 2.708% | 1.98% | 8.332% | 5.531% | 0.02971% | -8.473% | -1.021% | 2.28% | 0.07972% | 1.28% |  | 3.881% | -5.472% | -35.86% | -7.973% | -5.872% | -0.4704% | -18.68% | -2.051% | 2.53% | 6.455% | 6.082% | -7.291% | 3.131% | -0.9706% | 3.081% | -0.2704%
| Z-Score (Absolute) |  | 1.09 | 0.29 | 2.6 | 0.66 | 0.58 | 0.42 | 1.82 | 1.2 | 0.01 | 1.89 | 0.24 | 0.49 | 0 | 0.27 | 11.04 | 0.84 | 1.23 | 7.94 | 1.78 | 1.32 | 0.12 | 4.15 | 0.47 | 0.54 | 1.41 | 1.33 | 1.63 | 0.67 | 0.23 | 0.66 | 0.08
| Analytical results for B in OREAS 231b (Certified Value 28.6 ppm) | [32 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.11 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 28 | 28 | 44.71 | 34.6115 | 43 | 32 | 30 | 20 | 30 | 20 | 26 | 21 | 69.83 | 20 | 20 | 19 | < 10 | 32 | 18.5 | [12 empty] | 1
| 1 | 2 | 25 | 25 | 40.42 | 37.325 | 43 | 32 | 30 | 20 | 20 | 20 | 26 | 22 | 68.05 | 20 | 20 | 17 | < 10 | 34 | 18.5 | [12 empty] | 3
| 1 | 3 | 25 | 26 | 36.79 | 36.5108 | 42 | 32 | 30 | 20 | 30 | 30 | 26 | 22 | 87.26 | 20 | 20 | 17 | < 10 | 34 | 19.3 | [12 empty] | 16
| 1 | 4 | 25 | 24 | 38.97 | 35.0307 | 39 | 31 | 30 | 20 | 30 | 20 | 26 | 21 | 71.45 | 20 | 20 | 17 | < 10 | 32 | 18.1 | [12 empty] | 28.5899
| 1 | 5 | 24 | 25 | 34.92 | 37.2352 | 39 | 32 | 30 | 20 | 30 | 20 | 26 | 21 | 70 | 20 | 20 | 16 | < 10 | 32 | 19.2 | [12 empty] | 77
| 1 | 6 | 24 | 25 | 40.28 | 33.9292 | 41 | 34 | 30 | 20 | 20 | 20 | 27 | 22 | 70.58 | 20 | 20 | 16 | < 10 | 33 | 20
| Mean |  | 25.1667 | 25.5 | 39.3483 | 35.7737 | 41.1667 | 32.1667 | 30 | 20 | 26.6667 | 21.6667 | 26.1667 | 21.5 | 72.8617 | 20 | 20 | 17 |  | 32.8333 | 18.9333
| Median |  | 25 | 25 | 39.625 | 35.7708 | 41.5 | 32 | 30 | 20 | 30 | 20 | 26 | 21.5 | 70.29 | 20 | 20 | 17 |  | 32.5 | 18.85
| Std Dev. |  | 1.472 | 1.3784 | 3.3782 | 1.4415 | 1.8348 | 0.98319 | 0 | 0 | 5.164 | 4.0825 | 0.40825 | 0.54772 | 7.142 | 0 | 0 | 1.0954 |  | 0.98319 | 0.69474
| Rel.Std.Dev. |  | 5.849% | 5.406% | 8.585% | 4.03% | 4.457% | 3.057% | 0% | 0% | 19.36% | 18.84% | 1.56% | 2.548% | 9.802% | 0% | 0% | 6.444% |  | 2.994% | 3.669%
| PDM3 |  | -11.97% | -10.81% | 37.63% | 25.13% | 43.99% | 12.51% | 4.932% | -30.05% | -6.727% | -24.22% | -8.476% | -24.8% | 154.9% | -30.05% | -30.05% | -40.54% |  | 14.84% | -33.78%
| Z-Score (Absolute) |  | 0.09 | 0.06 | 1.23 | 0.9 | 1.4 | 0.56 | NA | NA | NA | NA | 0 | 0.44 | 4.35 | NA | NA | 0.85 | 1.97 | 0.62 | 0.67
Results from laboratories 1.14, 1.15, 1.16, 1.18, 1.27 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Ba in OREAS 231b (Certified Value 35.8 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 37 | 35.75 | 40.78 | 41.58 | 38.461 | 36.6 | 46 | 44 | 36 | 39 | 34 | 30 | 40 | 40 | 48.569 | 40 | 25 | 36.5811 | 35 | 27.2 | 36 | 31.5 | 36.189 | 29 | 30 | 34 | 40 | 33 | 32 |  |  | 1
| 1 | 2 | 35 | 36.81 | 37.83 | 42.83 | 38.619 | 36 | 45 | 44 | 35 | 38 | 35 | 40 | 40 | 30 | 49.948 | 40 | 25 | 34.9538 | 35 | 29.3 | 36 | 31.04 | 33.345 | 29 | 30 | 34 | 38 | 33 | 31 |  |  | 21
| 1 | 3 | 35 | 35.6 | 36.71 | 42.4 | 38.9681 | 36 | 44 | 45 | 36 | 38 | 35 | 30 | 40 | 40 | 48.889 | 40 | 24 | 36.561 | 36 | 27.6 | 35 | 32.66 | 34.236 | 27 | 30 | 34 | 37 | 33 | 32 |  |  | 16
| 1 | 4 | 34 | 35.03 | 39.92 | 41.44 | 37.9403 | 35.7 | 43 | 45 | 34 | 38 | 35 | 30 | 40 | 40 | 52.398 | 40 | 24 | 36.065 | 36 | 27.3 | 36 | 31.61 | 34.362 | 28 | 30 | 33 | 39 | 32 | 32 |  |  | 35.8332
| 1 | 5 | 34 | 36.19 | 36.26 | 41.06 | 41.1233 | 35.7 | 46 | 43 | 35 | 39 | 35 | 30 | 40 | 40 | 49.024 | 40 | 25 | 37.642 | 35 | 27.3 | 37 | 30.59 | 33.588 | 26 | 30 | 33 | 39 | 32 | 31 |  |  | 78
| 1 | 6 | 34 | 35.5 | 40.37 | 42.2 | 37.4551 | 36.9 | 44 | 45 | 37 | 37 | 35 | 30 | 40 | 40 | 50.645 | 40 | 24 | 39.3928 | 35 | 26.8 | 36 | 31.05 | 33.282 | 27 | 30 | 34 | 38 | 32 | 33
| Mean |  | 34.8333 | 35.8133 | 38.645 | 41.9183 | 38.7611 | 36.15 | 44.6667 | 44.3333 | 35.5 | 38.1667 | 34.8333 | 31.6667 | 40 | 38.3333 | 49.9122 | 40 | 24.5 | 36.8659 | 35.3333 | 27.5833 | 36 | 31.4083 | 34.167 | 27.6667 | 30 | 33.6667 | 38.5 | 32.5 | 31.8333
| Median |  | 34.5 | 35.675 | 38.875 | 41.89 | 38.54 | 36 | 44.5 | 44.5 | 35.5 | 38 | 35 | 30 | 40 | 40 | 49.486 | 40 | 24.5 | 36.571 | 35 | 27.3 | 36 | 31.275 | 33.912 | 27.5 | 30 | 34 | 38.5 | 32.5 | 32
| Std Dev. |  | 1.169 | 0.61588 | 1.9625 | 0.66671 | 1.2738 | 0.49295 | 1.2111 | 0.8165 | 1.0488 | 0.75277 | 0.40825 | 4.0825 | 0 | 4.0825 | 1.4393 | 0 | 0.54772 | 1.5139 | 0.5164 | 0.87958 | 0.63246 | 0.71407 | 1.0886 | 1.2111 | 0 | 0.5164 | 1.0488 | 0.54772 | 0.75277
| Rel.Std.Dev. |  | 3.356% | 1.72% | 5.078% | 1.59% | 3.286% | 1.364% | 2.711% | 1.842% | 2.954% | 1.972% | 1.172% | 12.89% | 0% | 10.65% | 2.884% | 0% | 2.236% | 4.107% | 1.462% | 3.189% | 1.757% | 2.274% | 3.186% | 4.377% | 0% | 1.534% | 2.724% | 1.685% | 2.365%
| PDM3 |  | -2.79% | -0.05539% | 7.847% | 16.98% | 8.171% | 0.8842% | 24.65% | 23.72% | -0.9298% | 6.512% | -2.79% | -11.63% | 11.63% | 6.977% | 39.29% | 11.63% | -31.63% | 2.882% | -1.395% | -23.02% | 0.4655% | -12.35% | -4.65% | -22.79% | -16.28% | -6.046% | 7.442% | -9.302% | -11.16%
| Z-Score (Absolute) |  | 0.19 | 0.04 | 0.69 | 1.45 | 0.72 | 0.11 | 2.08 | 2.01 | 0.04 | 0.58 | 0.19 | NA | NA | NA | 3.3 | NA | 2.58 | 0.28 | 0.07 | 1.87 | 0.08 | 0.98 | 0.34 | 1.85 | NA | 0.46 | 0.66 | 0.73 | 0.88
Results from laboratories 1.14, 1.15, 1.16, 1.18 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 231b (Certified Value 0.17 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.2 | 0.21 | < 1 | 0.15165 | 0.19455 | 0.2 | 0.27 | < 0.5 | 0.18 | < 5 | < 1 | 0.17 | 0.17 | 0.16 | < 10 | 0.16 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.15 | < 0.05 | 0.2 | 0.21 | 0.14 | [5 empty] | 1
| 1 | 2 | 0.1 | 0.19 | < 1 | 0.14742 | 0.18713 | 0.2 | 0.25 | < 0.5 | 0.19 | < 5 | < 1 | 0.17 | 0.17 | 0.16 | < 10 | 0.16 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.08 | 0.2 | 0.21 | 0.08 | [5 empty] | 22
| 1 | 3 | 0.2 | 0.2 | < 1 | 0.14481 | 0.1938 | 0.2 | 0.13 | < 0.5 | 0.17 | < 5 | < 1 | 0.17 | 0.17 | 0.17 | < 10 | 0.17 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.15 | < 0.05 | 0.2 | 0.2 | 0.16 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.19 | < 1 | 0.14872 | 0.18945 | 0.2 | 0.18 | < 0.5 | 0.15 | < 5 | < 1 | 0.17 | 0.17 | 0.16 | < 10 | 0.16 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.07 | 0.2 | 0.21 | 0.16 | [5 empty] | 0.17019
| 1 | 5 | 0.1 | 0.17 | < 1 | 0.14143 | 0.20806 | 0.2 | 0.19 | < 0.5 | 0.17 | < 5 | < 1 | 0.16 | 0.18 | 0.17 | < 10 | 0.16 | < 0.5 | 1.0864 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.07 | 0.2 | 0.2 | 0.11 | [5 empty] | 79
| 1 | 6 | 0.1 | 0.18 | < 1 | 0.14045 | 0.19459 | 0.2 | 0.15 | < 0.5 | 0.18 | < 5 | < 1 | 0.17 | 0.17 | 0.17 | < 10 | 0.17 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.14 | < 0.05 | 0.2 | 0.21 | 0.12
| Mean |  | 0.15 | 0.19 |  | 0.14575 | 0.1946 | 0.2 | 0.195 |  | 0.17333 |  |  | 0.16833 | 0.17167 | 0.165 |  | 0.16333 |  | 1.0864 |  |  | 0.2 | 0.14833 | 0.073333 | 0.2 | 0.20667 | 0.12833
| Median |  | 0.15 | 0.19 |  | 0.14611 | 0.19418 | 0.2 | 0.185 |  | 0.175 |  |  | 0.17 | 0.17 | 0.165 |  | 0.16 |  | 1.0864 |  |  | 0.2 | 0.15 | 0.07 | 0.2 | 0.21 | 0.13
| Std Dev. |  | 0.054772 | 0.014142 |  | 0.0043359 | 0.0072695 | 3.0405e-17 | 0.055045 |  | 0.013663 |  |  | 0.0040825 | 0.0040825 | 0.0054772 |  | 0.005164 |  |  |  |  | 3.0405e-17 | 0.0040825 | 0.0057735 | 3.0405e-17 | 0.005164 | 0.031252
| Rel.Std.Dev. |  | 36.51% | 7.443% |  | 2.975% | 3.736% | 1.52e-14% | 28.23% |  | 7.882% |  |  | 2.425% | 2.378% | 3.32% |  | 3.162% |  |  |  |  | 1.52e-14% | 2.752% | 7.873% | 1.52e-14% | 2.499% | 24.35%
| PDM3 |  | -11.86% | 11.64% |  | -14.36% | 14.34% | 17.51% | 14.58% |  | 1.845% |  |  | -1.093% | 0.8657% | -3.051% |  | -4.031% |  | 538.3% |  |  | 17.51% | -12.84% | -56.91% | 17.51% | 21.43% | -24.6%
| Z-Score (Absolute) |  | NA | 0.06 | 3.95 | 0.63 | 0 | NA | 0 | 0.71 | 0.28 | 29.82 | 3.95 | 0.34 | 0.3 | 0.39 | 62.16 | 0.41 | 0.71 | 5.21 | 0.71 | 0.71 | NA | 0.6 | 1.88 | NA | 0.15 | 0.86
| Analytical results for Bi in OREAS 231b (Certified Value 0.027 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.02 | 0.03 | < 1 | 0.020335 | 0.03 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.03 | 0.03 | 0.03 | < 10 | 0.03 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.02837 | 0.03 | 0.03 | 0.11 | 0.03 | 0.04 | 0.02 |  |  |  |  | 1
| 1 | 2 | < 0.02 | 0.02 | < 1 | 0.015136 | 0.02 | 0.06 | < 0.05 | 0.03 | < 2 | < 0.1 | 0.03 | 0.03 | 0.02 | < 10 | 0.02 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.02453 | 0.02 | < 0.02 | 0.09 | 0.03 | 0.03 | 0.02 |  |  |  |  | 23
| 1 | 3 | < 0.02 | 0.03 | < 1 | 0.010693 | 0.02 | 0.05 | < 0.05 | 0.03 | < 2 | < 0.1 | 0.02 | 0.03 | 0.03 | < 10 | 0.03 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.02861 | 0.02 | < 0.02 | 0.08 | 0.03 | 0.03 | 0.02 |  |  |  |  | 16
| 1 | 4 | < 0.02 | 0.03 | < 1 | 0.013078 | 0.03 | 0.04 | < 0.05 | 0.03 | < 2 | < 0.1 | 0.03 | 0.02 | 0.03 | < 10 | 0.03 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.02832 | 0.03 | < 0.02 | 0.07 | 0.03 | 0.03 | 0.02 |  |  |  |  | 0.027289
| 1 | 5 | < 0.02 | 0.03 | < 1 | 0.014053 | 0.02 | 0.04 | < 0.05 | 0.03 | < 2 | < 0.1 | 0.02 | 0.02 | 0.03 | < 10 | 0.03 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.0278 | 0.03 | < 0.02 | 0.06 | 0.03 | 0.03 | 0.01 |  |  |  |  | 80
| 1 | 6 | < 0.02 | 0.03 | < 1 | 0.011208 | 0.02 | 0.05 | < 0.05 | 0.02 | < 2 | < 0.1 | 0.03 | 0.02 | 0.03 | < 10 | 0.03 | < 5 | < 0.05 | < 0.02 | 0.03 | < 0.1 | 0.02735 | 0.02 | 0.194 | 0.06 | 0.03 | 0.03 | 0.02
| Mean |  |  | 0.028333 |  | 0.014084 | 0.023333 | 0.05 |  | 0.03 |  |  | 0.026667 | 0.025 | 0.028333 |  | 0.028333 |  |  |  | 0.03 |  | 0.027497 | 0.025 | 0.112 | 0.078333 | 0.03 | 0.031667 | 0.018333
| Median |  |  | 0.03 |  | 0.013566 | 0.02 | 0.05 |  | 0.03 |  |  | 0.03 | 0.025 | 0.03 |  | 0.03 |  |  |  | 0.03 |  | 0.02806 | 0.025 | 0.112 | 0.075 | 0.03 | 0.03 | 0.02
| Std Dev. |  |  | 0.0040825 |  | 0.0034917 | 0.005164 | 0.0089443 |  | 0.0063246 |  |  | 0.005164 | 0.0054772 | 0.0040825 |  | 0.0040825 |  |  |  | 0 |  | 0.0015228 | 0.0054772 | 0.11597 | 0.019408 | 0 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  |  | 14.41% |  | 24.79% | 22.13% | 17.89% |  | 21.08% |  |  | 19.36% | 21.91% | 14.41% |  | 14.41% |  |  |  | 0% |  | 5.538% | 21.91% | 103.5% | 24.78% | 0% | 12.89% | 22.27%
| PDM3 |  |  | 3.827% |  | -48.39% | -14.5% | 83.22% |  | 9.935% |  |  | -2.28% | -8.388% | 3.827% |  | 3.827% |  |  |  | 9.935% |  | 0.7611% | -8.388% | 310.4% | 187.1% | 9.935% | 16.04% | -32.82%
| Z-Score (Absolute) |  | 2.47 | 0 | 63.61 | 1.92 | 0.67 | 2.92 | 0.45 | 0.22 | 131.04 | 2.92 | 0.22 | 0.45 | 0 | 670.49 | 0 | 333.33 | 0.45 | 2.47 | 0.22 | 2.92 | 0.11 | 0.45 | 2.11 | 6.74 | 0.22 | 0.45 | 1.35
| Analytical results for Ca in OREAS 231b (Certified Value 2.11 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.22 | 2.069 | 2.324 | 2.5295 | 2.2822 | 2.77 | 2.409 | 2.84 | 2.43 | 2.28 | 1.79 | 2.13 | 2.15 | 2.03 | 1.3932 | 2.12 | 1.97 | 2.401 | 2.3 | 1.85 | 2.08 | 1.7351 | 2.2762 | 1.4265 | 1.99 | 1.78 | 1.91 | 1.96 | 1.54 |  |  | 1
| 1 | 2 | 2.04 | 2.138 | 2.172 | 2.5173 | 2.2944 | 2.72 | 2.537 | 2.76 | 2.33 | 2.28 | 1.83 | 2.14 | 2.15 | 2.02 | 1.4461 | 2.1 | 1.99 | 2.2143 | 2.33 | 1.92 | 2.14 | 1.5864 | 2.1295 | 1.462 | 1.99 | 1.74 | 1.87 | 2 | 1.56 |  |  | #N/A
| 1 | 3 | 2 | 2.099 | 2.1024 | 2.5377 | 2.3099 | 2.81 | 2.523 | 2.77 | 2.35 | 2.25 | 1.85 | 2.1 | 2.13 | 2.03 | 1.4356 | 2.2 | 1.95 | 2.4273 | 2.29 | 1.86 | 2.16 | 1.776 | 2.1922 | 1.409 | 1.96 | 1.7 | 1.87 | 1.97 | 1.5 |  |  | 16
| 1 | 4 | 1.95 | 2.077 | 2.2978 | 2.5519 | 2.2792 | 2.6 | 2.509 | 2.8 | 2.24 | 2.17 | 1.89 | 2.13 | 2.16 | 2.01 | 1.5381 | 2.11 | 1.93 | 2.311 | 2.3 | 1.78 | 2.14 | 1.5697 | 2.2224 | 1.4365 | 2 | 1.6 | 1.89 | 1.93 | 1.44 |  |  | 2.1067
| 1 | 5 | 1.9 | 2.117 | 2.0917 | 2.5622 | 2.4557 | 2.59 | 2.466 | 2.77 | 2.31 | 2.21 | 1.91 | 2.1 | 2.15 | 2.05 | 1.4614 | 2.13 | 2.03 | 2.4966 | 2.33 | 1.78 | 2.14 | 1.7336 | 2.1665 | 1.375 | 2.02 | 1.54 | 1.89 | 1.93 | 1.56 |  |  | 81
| 1 | 6 | 1.89 | 2.074 | 2.4065 | 2.5666 | 2.2233 | 2.66 | 2.494 | 2.75 | 2.42 | 2.2 | 1.91 | 2.1 | 2.15 | 2.06 | 1.5139 | 2.14 | 1.96 | 2.544 | 2.29 | 1.83 | 2.17 | 1.5741 | 2.1796 | 1.3625 | 2 | 1.63 | 1.9 | 1.96 | 1.62
| Mean |  | 2 | 2.0957 | 2.2324 | 2.5442 | 2.3075 | 2.6917 | 2.4897 | 2.7817 | 2.3467 | 2.2317 | 1.8633 | 2.1167 | 2.1483 | 2.0333 | 1.4647 | 2.1333 | 1.9717 | 2.399 | 2.3067 | 1.8367 | 2.1383 | 1.6625 | 2.1944 | 1.4119 | 1.9933 | 1.665 | 1.8883 | 1.9583 | 1.5367
| Median |  | 1.975 | 2.088 | 2.2349 | 2.5448 | 2.2883 | 2.69 | 2.5015 | 2.77 | 2.34 | 2.23 | 1.87 | 2.115 | 2.15 | 2.03 | 1.4537 | 2.125 | 1.965 | 2.4142 | 2.3 | 1.84 | 2.14 | 1.66 | 2.1859 | 1.4177 | 1.995 | 1.665 | 1.89 | 1.96 | 1.55
| Std Dev. |  | 0.12215 | 0.027522 | 0.1291 | 0.019324 | 0.078329 | 0.090203 | 0.0465 | 0.033116 | 0.071181 | 0.04535 | 0.048442 | 0.018619 | 0.0098319 | 0.018619 | 0.053158 | 0.03559 | 0.034881 | 0.12099 | 0.018619 | 0.053166 | 0.031252 | 0.095316 | 0.050378 | 0.037785 | 0.019664 | 0.090719 | 0.016021 | 0.026394 | 0.06121
| Rel.Std.Dev. |  | 6.107% | 1.313% | 5.783% | 0.7595% | 3.395% | 3.351% | 1.868% | 1.191% | 3.033% | 2.032% | 2.6% | 0.8796% | 0.4577% | 0.9157% | 3.629% | 1.668% | 1.769% | 5.043% | 0.8072% | 2.895% | 1.461% | 5.733% | 2.296% | 2.676% | 0.9865% | 5.449% | 0.8484% | 1.348% | 3.983%
| PDM3 |  | -5.066% | -0.5253% | 5.965% | 20.77% | 9.528% | 27.76% | 18.18% | 32.04% | 11.39% | 5.93% | -11.55% | 0.4715% | 1.975% | -3.484% | -30.47% | 1.263% | -6.411% | 13.87% | 9.49% | -12.82% | 1.5% | -21.09% | 4.161% | -32.98% | -5.383% | -20.97% | -10.37% | -7.044% | -27.06%
| Z-Score (Absolute) |  | 0.41 | 0.07 | 0.41 | 1.51 | 0.67 | 2.03 | 1.32 | 2.35 | 0.81 | 0.41 | 0.9 | 0 | 0.11 | 0.29 | 2.3 | 0.06 | 0.51 | 1 | 0.67 | 0.99 | 0.08 | 1.61 | 0.27 | 2.49 | 0.44 | 1.6 | 0.81 | 0.56 | 2.05
| Analytical results for Cd in OREAS 231b (Certified Value 0.34 ppm) | [32 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.37 | 0.353 | < 1 | 0.2372 | 0.37038 | 0.41 | 0.32 | 0.36 | 0.33 | 0.5 | 0.3 | 0.35 | 0.35 | 0.3 | < 10 | 0.38 | < 1 | < 1 | 0.31 | 0.32 | 0.35 | 1.32 | < 0.5 | 0.28479 | 0.33 | 0.3 | 0.38 | 0.32 | 0.36 |  |  | 1
| 1 | 2 | 0.42 | 0.344 | < 1 | 0.23187 | 0.36931 | 0.37 | 0.36 | 0.36 | 0.39 | 0.5 | 0.3 | 0.34 | 0.35 | 0.32 | < 10 | 0.37 | < 1 | < 1 | 0.32 | 0.31 | 0.36 | 1.45 | < 0.5 | 0.33557 | 0.34 | 0.29 | 0.34 | 0.33 | 0.33 |  |  | 24
| 1 | 3 | 0.39 | 0.348 | < 1 | 0.24295 | 0.36678 | 0.3 | 0.37 | 0.35 | 0.34 | 0.5 | 0.4 | 0.32 | 0.34 | 0.34 | < 10 | 0.38 | < 1 | < 1 | 0.32 | 0.31 | 0.36 | 1.46 | < 0.5 | 0.35811 | 0.34 | 0.31 | 0.36 | 0.34 | 0.34 |  |  | 16
| 1 | 4 | 0.37 | 0.335 | < 1 | 0.2647 | 0.3738 | 0.33 | 0.42 | 0.34 | 0.34 | 0.5 | 0.3 | 0.35 | 0.35 | 0.35 | < 10 | 0.38 | < 1 | < 1 | 0.3 | 0.31 | 0.34 | 1.33 | < 0.5 | 0.34503 | 0.35 | 0.29 | 0.38 | 0.32 | 0.35 |  |  | 0.34391
| 1 | 5 | 0.37 | 0.348 | < 1 | 0.24511 | 0.40512 | 0.31 | 0.4 | 0.36 | 0.31 | 0.5 | 0.4 | 0.33 | 0.34 | 0.33 | < 10 | 0.39 | < 1 | < 1 | 0.31 | 0.3 | 0.35 | 1.19 | < 0.5 | 0.37268 | 0.34 | 0.3 | 0.36 | 0.33 | 0.36 |  |  | 82
| 1 | 6 | 0.35 | 0.343 | < 1 | 0.24622 | 0.35971 | 0.32 | 0.38 | 0.39 | 0.36 | 0.5 | 0.4 | 0.33 | 0.35 | 0.35 | < 10 | 0.39 | < 1 | < 1 | 0.32 | 0.29 | 0.35 | 1.41 | < 0.5 | 0.27347 | 0.35 | 0.31 | 0.35 | 0.32 | 0.35
| Mean |  | 0.37833 | 0.34517 |  | 0.24467 | 0.37419 | 0.34 | 0.375 | 0.36 | 0.345 | 0.5 | 0.35 | 0.33667 | 0.34667 | 0.33167 |  | 0.38167 |  |  | 0.31333 | 0.30667 | 0.35167 | 1.36 |  | 0.32828 | 0.34167 | 0.3 | 0.36167 | 0.32667 | 0.34833
| Median |  | 0.37 | 0.346 |  | 0.24403 | 0.36985 | 0.325 | 0.375 | 0.36 | 0.34 | 0.5 | 0.35 | 0.335 | 0.35 | 0.335 |  | 0.38 |  |  | 0.315 | 0.31 | 0.35 | 1.37 |  | 0.3403 | 0.34 | 0.3 | 0.36 | 0.325 | 0.35
| Std Dev. |  | 0.024014 | 0.0061128 |  | 0.011194 | 0.015873 | 0.041952 | 0.034496 | 0.016733 | 0.027386 | 0 | 0.054772 | 0.012111 | 0.005164 | 0.019408 |  | 0.0075277 |  |  | 0.008165 | 0.010328 | 0.0075277 | 0.10198 |  | 0.040225 | 0.0075277 | 0.0089443 | 0.016021 | 0.008165 | 0.01169
| Rel.Std.Dev. |  | 6.347% | 1.771% |  | 4.575% | 4.242% | 12.34% | 9.199% | 4.648% | 7.938% | 0% | 15.65% | 3.597% | 1.49% | 5.852% |  | 1.972% |  |  | 2.606% | 3.368% | 2.141% | 7.499% |  | 12.25% | 2.203% | 2.981% | 4.43% | 2.499% | 3.356%
| PDM3 |  | 10.01% | 0.3668% |  | -28.85% | 8.805% | -1.136% | 9.042% | 4.68% | 0.3183% | 45.39% | 1.772% | -2.105% | 0.8029% | -3.559% |  | 10.98% |  |  | -8.89% | -10.83% | 2.257% | 295.5% |  | -4.545% | -0.651% | -12.77% | 5.165% | -5.013% | 1.288%
| Z-Score (Absolute) |  | 0.78 | 0.04 | 3.76 | 2.5 | 0.67 | 0.16 | 0.69 | 0.33 | 0.04 | NA | NA | 0.25 | 0 | 0.37 | 114.02 | 0.86 | 3.76 | 3.76 | 0.82 | 0.98 | 0.12 | 24.83 | 2.37 | 0.45 | 0.12 | 1.14 | 0.37 | 0.49 | 0.04
Results from laboratories 1.12 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 231b (Certified Value 7.94 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 7.44 | 7.962 | 7.636 | 7.1253 | 8.3036 | 8.65 | < 1 | 9.74 | 8.097 | 8 | 7.45 | 8.03 | 7.32 | 7.94 | < 10 | 8.12 | 8.63 | 7.5 | 8.19 | 8.45 | 8.2 | 8.17 | 8.03 | [8 empty] | 1
| 1 | 2 | 7.34 | 8.026 | 7.428 | 7.2715 | 7.9722 | 8.49 | < 1 | 9.46 | 7.894 | 8 | 7.35 | 8.07 | 7.18 | 7.78 | < 10 | 8.39 | 8.53 | 8 | 8.07 | 8.44 | 7.78 | 8.22 | 8.02 | [8 empty] | 25
| 1 | 3 | 7.38 | 8.17 | 7.37 | 7.2586 | 8.0483 | 8.68 | < 1 | 9.6 | 7.888 | 8 | 7.28 | 8.07 | 7.75 | 8.03 | < 10 | 8.26 | 8.67 | 8 | 8.24 | 8.48 | 7.89 | 8.34 | 8.01 | [8 empty] | 16
| 1 | 4 | 7.31 | 7.92 | 7.66 | 7.2683 | 8.0391 | 8.56 | < 1 | 9.58 | 7.7 | 8 | 7.14 | 8.08 | 7.8 | 7.91 | < 10 | 8.29 | 8.72 | 7.5 | 8.14 | 8.46 | 8.07 | 8.01 | 8 | [8 empty] | 7.936
| 1 | 5 | 7.26 | 8.053 | 7.846 | 7.3122 | 8.7951 | 8.5 | < 1 | 9.49 | 7.686 | 9 | 7.17 | 7.9 | 7.99 | 7.87 | < 10 | 8.37 | 8.44 | 7.5 | 8.24 | 8.46 | 7.92 | 8.08 | 8.06 | [8 empty] | 83
| 1 | 6 | 7.21 | 7.926 | 7.256 | 7.017 | 8.1011 | 8.4 | < 1 | 9.63 | 7.882 | 8 | 7.08 | 8.17 | 7.61 | 8.02 | < 10 | 8.15 | 8.39 | 7.5 | 8.31 | 8.49 | 7.76 | 8.31 | 8.25
| Mean |  | 7.3233 | 8.0095 | 7.5327 | 7.2088 | 8.2099 | 8.5467 |  | 9.5833 | 7.8578 | 8.1667 | 7.245 | 8.0533 | 7.6083 | 7.925 |  | 8.2633 | 8.5633 | 7.6667 | 8.1983 | 8.4633 | 7.9367 | 8.1883 | 8.0617
| Median |  | 7.325 | 7.994 | 7.532 | 7.2635 | 8.0747 | 8.53 |  | 9.59 | 7.885 | 8 | 7.225 | 8.07 | 7.68 | 7.925 |  | 8.275 | 8.58 | 7.5 | 8.215 | 8.46 | 7.905 | 8.195 | 8.025
| Std Dev. |  | 0.082624 | 0.095041 | 0.21865 | 0.11352 | 0.30817 | 0.10539 |  | 0.10093 | 0.15129 | 0.40825 | 0.14011 | 0.088242 | 0.3063 | 0.094393 |  | 0.11094 | 0.13171 | 0.2582 | 0.084715 | 0.018619 | 0.17049 | 0.12859 | 0.094534
| Rel.Std.Dev. |  | 1.128% | 1.187% | 2.903% | 1.575% | 3.754% | 1.233% |  | 1.053% | 1.925% | 4.999% | 1.934% | 1.096% | 4.026% | 1.191% |  | 1.343% | 1.538% | 3.368% | 1.033% | 0.22% | 2.148% | 1.57% | 1.173%
| PDM3 |  | -7.72% | 0.9266% | -5.082% | -9.162% | 3.452% | 7.695% |  | 20.76% | -0.9846% | 2.907% | -8.707% | 1.479% | -4.128% | -0.1382% |  | 4.125% | 7.905% | -3.393% | 3.306% | 6.645% | 0.008803% | 3.18% | 1.584%
| Z-Score (Absolute) |  | 1.31 | 0.07 | 0.88 | 1.54 | 0.48 | 1.15 | 15.02 | 3.24 | 0.23 | NA | 1.46 | 0.16 | 0.73 | 0.1 | 5.97 | 0.58 | 1.19 | 0.62 | 0.45 | 0.99 | 0.07 | 0.43 | 0.18
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 231b (Certified Value 33.4 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 32.3 | 33.21 | 30.83 | 26.6291 | 34.9364 | 34.7 | 32.7 | 37.8 | 34.5 | 30 | 34.5 | 35.6 | 34.1 | 32.2 | 26.577 | 32.3 | 32 | 34.7936 | 33.3 | 32.4 | 35 | 30.21 | 30 | 34.9257 | 32.7 | 37 | 35.4 | 33.4 | 33.6 |  |  | 1
| 1 | 2 | 32.6 | 33.54 | 27.44 | 26.906 | 35.0057 | 34.5 | 32.1 | 36.4 | 33.7 | 30 | 34.4 | 35.6 | 34.1 | 31.7 | 24.722 | 31.8 | 31 | 35.4003 | 33.5 | 32.7 | 34 | 29.91 | 31 | 35.1931 | 32.3 | 36.3 | 34.5 | 33.7 | 33.1 |  |  | 26
| 1 | 3 | 32.5 | 33.28 | 28.1 | 26.5915 | 35.3361 | 34.8 | 34.5 | 37.2 | 33.7 | 30 | 34.7 | 35.9 | 34.5 | 33.8 | 25.712 | 33.5 | 30 | 34.3962 | 34.3 | 31.9 | 34 | 29.52 | 29 | 34.8813 | 32.7 | 37 | 34.5 | 34 | 33.1 |  |  | 16
| 1 | 4 | 32.2 | 33.28 | 30.3 | 26.1272 | 34.4508 | 34 | 33.2 | 37 | 33 | 30 | 35.1 | 34.2 | 34.6 | 33.8 | 29.232 | 32.8 | 30 | 35.3322 | 34.5 | 31.9 | 34 | 30.93 | 30 | 36.1042 | 32.4 | 37 | 35.9 | 32.6 | 33 |  |  | 33.3786
| 1 | 5 | 32.1 | 32.83 | 27.91 | 26.443 | 37.7922 | 34 | 33.7 | 37.3 | 33.7 | 30 | 35.3 | 34.9 | 33.9 | 34.1 | 25.651 | 32.4 | 31 | 36.0886 | 34.2 | 31.1 | 34 | 29.67 | 29 | 36.4414 | 32.9 | 37.5 | 35.3 | 33.1 | 33.3 |  |  | 84
| 1 | 6 | 31.7 | 33.09 | 31.32 | 26.6278 | 34.0326 | 34.7 | 33.1 | 37.6 | 34.7 | 30 | 34.1 | 35.2 | 34.7 | 33.4 | 27.887 | 33.3 | 32 | 36.2556 | 34.5 | 32.3 | 34 | 28.23 | 29 | 34.8142 | 32.6 | 37.2 | 35 | 33.8 | 33.6
| Mean |  | 32.2333 | 33.205 | 29.3167 | 26.5541 | 35.259 | 34.45 | 33.2167 | 37.2167 | 33.8833 | 30 | 34.6833 | 35.2333 | 34.3167 | 33.1667 | 26.6302 | 32.6833 | 31 | 35.3778 | 34.05 | 32.05 | 34.1667 | 29.745 | 29.6667 | 35.3933 | 32.6 | 37 | 35.1 | 33.4333 | 33.2833
| Median |  | 32.25 | 33.245 | 29.2 | 26.6097 | 34.971 | 34.6 | 33.15 | 37.25 | 33.7 | 30 | 34.6 | 35.4 | 34.3 | 33.6 | 26.1445 | 32.6 | 31 | 35.3663 | 34.25 | 32.1 | 34 | 29.79 | 29.5 | 35.0594 | 32.65 | 37 | 35.15 | 33.55 | 33.2
| Std Dev. |  | 0.32042 | 0.23552 | 1.6883 | 0.25724 | 1.3228 | 0.36194 | 0.82563 | 0.4916 | 0.62102 | 0 | 0.44907 | 0.61536 | 0.32506 | 0.98116 | 1.661 | 0.64317 | 0.89443 | 0.71876 | 0.52058 | 0.55767 | 0.40825 | 0.89433 | 0.8165 | 0.70146 | 0.21909 | 0.39497 | 0.54772 | 0.5164 | 0.26394
| Rel.Std.Dev. |  | 0.9941% | 0.7093% | 5.759% | 0.9687% | 3.752% | 1.051% | 2.486% | 1.321% | 1.833% | 0% | 1.295% | 1.747% | 0.9472% | 2.958% | 6.237% | 1.968% | 2.885% | 2.032% | 1.529% | 1.74% | 1.195% | 3.007% | 2.752% | 1.982% | 0.6721% | 1.067% | 1.56% | 1.545% | 0.793%
| PDM3 |  | -3.431% | -0.5202% | -12.17% | -20.45% | 5.633% | 3.21% | -0.4853% | 11.5% | 1.512% | -10.12% | 3.909% | 5.557% | 2.81% | -0.6351% | -20.22% | -2.083% | -7.126% | 5.989% | 2.011% | -3.981% | 2.361% | -10.89% | -11.12% | 6.036% | -2.333% | 10.85% | 5.157% | 0.1638% | -0.2855%
| Z-Score (Absolute) |  | 0.51 | 0.04 | 1.91 | 3.24 | 0.95 | 0.56 | 0.03 | 1.89 | 0.29 | 1.58 | 0.67 | 0.94 | 0.5 | 0.06 | 3.2 | 0.29 | 1.1 | 1.01 | 0.37 | 0.59 | 0.43 | 1.7 | 1.74 | 1.02 | 0.33 | 1.79 | 0.87 | 0.07 | 0
| Analytical results for Cr in OREAS 231b (Certified Value 23.9 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 27 | 24.7 | 26.65 | 25.6105 | 32 | 28 | 33 | 30 | 25 | 28 | 24 | 24 | 22 | 18.766 | 23 | 21 | 23.6087 | 24 | 23.6 | 24 | 23.95 | 25.2123 | 20 | 23 | 22.2 | 24 | 22 | 21.9 |  |  |  | 1
| 1 | 2 | 26 | 25.1 | 25.49 | 25.521 | 31 | 29 | 32 | 27 | 25 | 27 | 23 | 24 | 23 | 19.038 | 23 | 21 | 22.2397 | 26 | 24.3 | 24 | 53.93 | 23.715 | 20 | 23 | 21.5 | 24 | 22 | 22 |  |  |  | 27
| 1 | 3 | 26 | 24.9 | 26.12 | 25.7854 | 31 | 28 | 32 | 27 | 25 | 27 | 23 | 24 | 23 | 19.079 | 23 | 20 | 23.6327 | 25 | 23.4 | 24 | 25.28 | 24.0963 | 20 | 23 | 21.5 | 23 | 22 | 22 |  |  |  | 16
| 1 | 4 | 25 | 24.8 | 26.92 | 25.4423 | 31 | 29 | 32 | 27 | 24 | 28 | 23 | 24 | 23 | 20.51 | 23 | 20 | 22.9411 | 27 | 23.3 | 25 | 25.01 | 24.1149 | 20 | 23 | 20.9 | 24 | 21 | 22.2 |  |  |  | 23.8965
| 1 | 5 | 25 | 24.8 | 27.82 | 27.5797 | 30 | 30 | 33 | 27 | 24 | 27 | 23 | 24 | 23 | 18.891 | 23 | 21 | 24.5543 | 27 | 23.6 | 25 | 25.85 | 23.3151 | 20 | 23 | 21.2 | 25 | 21 | 22 |  |  |  | 85
| 1 | 6 | 25 | 24.4 | 27.3 | 24.708 | 31 | 28 | 33 | 27 | 24 | 27 | 23 | 24 | 23 | 19.579 | 24 | 21 | 26.1142 | 25 | 23.1 | 25 | 22.96 | 23.529 | 20 | 23 | 21.5 | 24 | 22 | 23.4
| Mean |  | 25.6667 | 24.7833 | 26.7167 | 25.7745 | 31 | 28.6667 | 32.5 | 27.5 | 24.5 | 27.3333 | 23.1667 | 24 | 22.8333 | 19.3105 | 23.1667 | 20.6667 | 23.8485 | 25.6667 | 23.55 | 24.5 | 29.4967 | 23.9971 | 20 | 23 | 21.4667 | 24 | 21.6667 | 22.25
| Median |  | 25.5 | 24.8 | 26.785 | 25.5658 | 31 | 28.5 | 32.5 | 27 | 24.5 | 27 | 23 | 24 | 23 | 19.0585 | 23 | 21 | 23.6207 | 25.5 | 23.5 | 24.5 | 25.145 | 23.9057 | 20 | 23 | 21.5 | 24 | 22 | 22
| Std Dev. |  | 0.8165 | 0.23166 | 0.83236 | 0.95893 | 0.63246 | 0.8165 | 0.54772 | 1.2247 | 0.54772 | 0.5164 | 0.40825 | 0 | 0.40825 | 0.6498 | 0.40825 | 0.5164 | 1.3526 | 1.2111 | 0.41352 | 0.54772 | 12.0141 | 0.67294 | 0 | 0 | 0.43205 | 0.63246 | 0.5164 | 0.57184
| Rel.Std.Dev. |  | 3.181% | 0.9347% | 3.116% | 3.72% | 2.04% | 2.848% | 1.685% | 4.454% | 2.236% | 1.889% | 1.762% | 0% | 1.788% | 3.365% | 1.762% | 2.499% | 5.672% | 4.718% | 1.756% | 2.236% | 40.73% | 2.804% | 0% | 0% | 2.013% | 2.635% | 2.383% | 2.57%
| PDM3 |  | 7.408% | 3.711% | 11.8% | 7.859% | 29.73% | 19.96% | 36% | 15.08% | 2.525% | 14.38% | -3.054% | 0.433% | -4.449% | -19.19% | -3.054% | -13.52% | -0.2012% | 7.408% | -1.45% | 2.525% | 23.43% | 0.4209% | -16.31% | -3.752% | -10.17% | 0.433% | -9.331% | -6.89%
| Z-Score (Absolute) |  | 0.66 | 0.31 | 1.07 | 0.7 | 2.76 | 1.84 | 3.36 | 1.38 | 0.2 | 1.32 | 0.33 | 0 | 0.46 | 1.85 | 0.33 | 1.32 | 0.06 | 0.66 | 0.18 | 0.2 | 2.17 | 0 | 1.58 | 0.39 | 1 | 0 | 0.92 | 0.69
| Analytical results for Cs in OREAS 231b (Certified Value 0.44 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.48 | 0.44 | 0.38312 | 0.4 | 0.43 | 0.45 | 0.44 | 0.5 | 0.43 | 0.44 | 0.43 | 0.45 | 0.44 | < 0.5 | 0.43 | 0.41 | 0.48 | 0.47 | 0.47 | [12 empty] | 1
| 1 | 2 | 0.49 | 0.44 | 0.37694 | 0.39 | 0.43 | 0.46 | 0.44 | 0.5 | 0.44 | 0.44 | 0.42 | 0.44 | 0.44 | < 0.5 | 0.43 | 0.38 | 0.47 | 0.48 | 0.45 | [12 empty] | 28
| 1 | 3 | 0.5 | 0.43 | 0.37292 | 0.4 | 0.43 | 0.46 | 0.45 | 0.5 | 0.42 | 0.44 | 0.46 | 0.45 | 0.44 | < 0.5 | 0.44 | 0.37 | 0.47 | 0.46 | 0.48 | [12 empty] | 16
| 1 | 4 | 0.5 | 0.44 | 0.38454 | 0.41 | 0.42 | 0.47 | 0.43 | 0.5 | 0.4 | 0.44 | 0.46 | 0.45 | 0.46 | < 0.5 | 0.44 | 0.41 | 0.49 | 0.45 | 0.46 | [12 empty] | 0.4401
| 1 | 5 | 0.48 | 0.45 | 0.38828 | 0.39 | 0.44 | 0.47 | 0.44 | 0.5 | 0.42 | 0.43 | 0.45 | 0.45 | 0.46 | < 0.5 | 0.44 | 0.43 | 0.48 | 0.44 | 0.45 | [12 empty] | 86
| 1 | 6 | 0.48 | 0.44 | 0.36423 | 0.39 | 0.42 | 0.48 | 0.46 | 0.5 | 0.42 | 0.45 | 0.45 | 0.45 | 0.45 | < 0.5 | 0.44 | 0.4 | 0.46 | 0.47 | 0.48
| Mean |  | 0.48833 | 0.44 | 0.37834 | 0.39667 | 0.42833 | 0.465 | 0.44333 | 0.5 | 0.42167 | 0.44 | 0.445 | 0.44833 | 0.44833 |  | 0.43667 | 0.4 | 0.475 | 0.46167 | 0.465
| Median |  | 0.485 | 0.44 | 0.38003 | 0.395 | 0.43 | 0.465 | 0.44 | 0.5 | 0.42 | 0.44 | 0.45 | 0.45 | 0.445 |  | 0.44 | 0.405 | 0.475 | 0.465 | 0.465
| Std Dev. |  | 0.0098319 | 0.0063246 | 0.0088382 | 0.008165 | 0.0075277 | 0.010488 | 0.010328 | 0 | 0.013292 | 0.0063246 | 0.016432 | 0.0040825 | 0.0098319 |  | 0.005164 | 0.021909 | 0.010488 | 0.01472 | 0.013784
| Rel.Std.Dev. |  | 2.013% | 1.437% | 2.336% | 2.058% | 1.757% | 2.256% | 2.33% | 0% | 3.152% | 1.437% | 3.693% | 0.9106% | 2.193% |  | 1.183% | 5.477% | 2.208% | 3.188% | 2.964%
| PDM3 |  | 10.96% | -0.02235% | -14.03% | -9.869% | -2.673% | 5.658% | 0.7351% | 13.61% | -4.188% | -0.02235% | 1.114% | 1.871% | 1.871% |  | -0.7798% | -9.111% | 7.93% | 4.901% | 5.658%
| Z-Score (Absolute) |  | 1.57 | 0.06 | 2.14 | 1.52 | 0.45 | 0.79 | 0.06 | NA | 0.67 | 0.06 | 0.11 | 0.22 | 0.22 | 6.46 | 0.17 | 1.4 | 1.12 | 0.67 | 0.79
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 231b (Certified Value 155 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 171.9 | 160.69 | 153.43 | 168.05 | 164.7288 | 151 | 151 | 152 | 157.7 | 153 | 160.3 | 158.5 | 154.5 | 154 | 148.927 | 154 | 169 | 153.4733 | 157 | 147.06 | 155 | 156.47 | 159.2336 | 151 | 162.7139 | 157.5 | 133.0067 | 159 | 159.4 | 147.5 | 155 | 1
| 1 | 2 | 175.1 | 163.25 | 165.58 | 166.97 | 164.7214 | 151 | 141 | 152 | 154.9 | 153 | 152.4 | 159 | 153 | 151 | 153.167 | 153.5 | 163 | 155.9311 | 155 | 148.84 | 154 | 155.43 | 149.4448 | 153 | 160.8182 | 156.5 | 138.351 | 158 | 156.9 | 152.5 | 157 | 29
| 1 | 3 | 171 | 162.5 | 159.85 | 168.15 | 166.4573 | 155 | 142 | 154 | 154.6 | 152 | 152.9 | 158.5 | 156.5 | 158 | 147.797 | 156 | 160 | 155.1734 | 155 | 142.38 | 151 | 158.07 | 149.1504 | 152 | 159.9076 | 157 | 132.573 | 159 | 156.9 | 151.7 | 154 | 16
| 1 | 4 | 172.4 | 162.3 | 155.26 | 171.78 | 160.993 | 149 | 149 | 152 | 152.8 | 152 | 153.5 | 157 | 156.5 | 158 | 161.566 | 153 | 159 | 149.7374 | 156 | 147.49 | 151 | 157.23 | 148.2764 | 150 | 157.4118 | 158 | 134.883 | 157 | 161.4 | 146.2 | 157 | 154.8415
| 1 | 5 | 173.3 | 162.49 | 156.03 | 172.12 | 175.8562 | 143 | 148 | 152 | 153.3 | 150 | 152.8 | 156 | 154 | 159.5 | 150.395 | 153.5 | 165 | 158.2401 | 156 | 143.41 | 157 | 157.18 | 146.9792 | 151 | 160.6156 | 157.5 | 131.728 | 157 | 159.3 | 145.7 | 155 | 87
| 1 | 6 | 170.3 | 160.67 | 148.33 | 161.55 | 159.7292 | 150 | 148 | 154 | 158.1 | 150 | 150.1 | 158 | 155 | 157.5 | 153.994 | 154.5 | 167 | 160.4217 | 154 | 142.73 | 152 | 159.74 | 146.5008 | 153 | 156.5513 | 157 | 251.274 | 160 | 155.9 | 148 | 158
| Mean |  | 172.3333 | 161.9833 | 156.4133 | 168.1033 | 165.4143 | 149.8333 | 146.5 | 152.6667 | 155.2333 | 151.6667 | 153.6667 | 157.8333 | 154.9167 | 156.3333 | 152.641 | 154.0833 | 163.8333 | 155.4962 | 155.5 | 145.3183 | 153.3333 | 157.3533 | 149.9309 | 151.6667 | 159.6697 | 157.25 | 153.636 | 158.3333 | 158.3 | 148.6 | 156
| Median |  | 172.15 | 162.395 | 155.645 | 168.1 | 164.7251 | 150.5 | 148 | 152 | 154.75 | 152 | 152.85 | 158.25 | 154.75 | 157.75 | 151.781 | 153.75 | 164 | 155.5523 | 155.5 | 145.235 | 153 | 157.205 | 148.7134 | 151.5 | 160.2616 | 157.25 | 133.9449 | 158.5 | 158.1 | 147.75 | 156
| Std Dev. |  | 1.7143 | 1.0607 | 5.8576 | 3.8441 | 5.7086 | 3.92 | 4.0373 | 1.0328 | 2.2124 | 1.3663 | 3.4552 | 1.1255 | 1.3934 | 3.1885 | 4.9813 | 1.0685 | 3.92 | 3.7239 | 1.0488 | 2.7974 | 2.4221 | 1.4652 | 4.7032 | 1.2111 | 2.2961 | 0.5244 | 47.8908 | 1.2111 | 2.0736 | 2.8482 | 1.5492
| Rel.Std.Dev. |  | 0.9947% | 0.6549% | 3.745% | 2.287% | 3.451% | 2.616% | 2.756% | 0.6765% | 1.425% | 0.9008% | 2.249% | 0.7131% | 0.8995% | 2.04% | 3.263% | 0.6934% | 2.393% | 2.395% | 0.6745% | 1.925% | 1.58% | 0.9312% | 3.137% | 0.7985% | 1.438% | 0.3335% | 31.17% | 0.7649% | 1.31% | 1.917% | 0.9931%
| PDM3 |  | 11.3% | 4.612% | 1.015% | 8.565% | 6.828% | -3.234% | -5.387% | -1.405% | 0.2531% | -2.05% | -0.7587% | 1.932% | 0.04856% | 0.9635% | -1.421% | -0.4896% | 5.807% | 0.4228% | 0.4253% | -6.15% | -0.974% | 1.622% | -3.171% | -2.05% | 3.118% | 1.555% | -0.7786% | 2.255% | 2.234% | -4.031% | 0.7482%
| Z-Score (Absolute) |  | 4.01 | 1.55 | 0.22 | 3 | 2.36 | 1.35 | 2.14 | 0.67 | 0.06 | 0.91 | 0.44 | 0.56 | 0.14 | 0.2 | 0.68 | 0.34 | 1.99 | 0 | 0 | 2.43 | 0.52 | 0.44 | 1.33 | 0.91 | 0.99 | 0.42 | 0.44 | 0.68 | 0.67 | 1.64 | 0.12
| Analytical results for Dy in OREAS 231b (Certified Value 2.12 ppm) | [32 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 2.117 | 1.998 | 1.8731 | 2.5 | 2.62 | 1.8 | [25 empty] | 1
| 1 | 2 | 2.194 | 1.971 | 1.8681 | 2.4 | 2.53 | 1.8 | [25 empty] | 30
| 1 | 3 | 2.147 | 1.884 | 1.8692 | 2.3 | 2.62 | 1.9 | [25 empty] | 16
| 1 | 4 | 2.094 | 1.947 | 1.9186 | 2.2 | 2.57 | 1.8 | [25 empty] | 2.115
| 1 | 5 | 2.169 | 2.062 | 1.9253 | 2.3 | 2.6 | 1.7 | [25 empty] | 88
| 1 | 6 | 2.073 | 1.884 | 1.8457 | 2.3 | 2.66 | 1.7
| Mean |  | 2.1323 | 1.9577 | 1.8833 | 2.3333 | 2.6 | 1.7833
| Median |  | 2.132 | 1.959 | 1.8711 | 2.3 | 2.61 | 1.8
| Std Dev. |  | 0.046016 | 0.068786 | 0.031474 | 0.10328 | 0.045166 | 0.075277
| Rel.Std.Dev. |  | 2.158% | 3.514% | 1.671% | 4.426% | 1.737% | 4.221%
| PDM3 |  | 0.8195% | -7.439% | -10.95% | 10.32% | 22.93% | -15.68%
| Z-Score (Absolute) |  | 0.28 | 0.28 | 0.52 | 0.92 | 1.77 | 0.83
| Analytical results for Er in OREAS 231b (Certified Value 1.22 ppm) | [32 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.223 | 1.148 | 1.1332 | 1.3 | 1.46 | 1 | [25 empty] | 1
| 1 | 2 | 1.24 | 1.135 | 1.1296 | 1.5 | 1.41 | 1 | [25 empty] | 31
| 1 | 3 | 1.261 | 1.074 | 1.1215 | 1.5 | 1.46 | 0.95 | [25 empty] | 16
| 1 | 4 | 1.243 | 1.101 | 1.1443 | 1.4 | 1.45 | 0.95 | [25 empty] | 1.2177
| 1 | 5 | 1.256 | 1.173 | 1.1512 | 1.3 | 1.48 | 0.9 | [25 empty] | 89
| 1 | 6 | 1.223 | 1.074 | 1.1557 | 1.4 | 1.44 | 0.95
| Mean |  | 1.241 | 1.1175 | 1.1393 | 1.4 | 1.45 | 0.95833
| Median |  | 1.2415 | 1.118 | 1.1387 | 1.4 | 1.455 | 0.95
| Std Dev. |  | 0.015987 | 0.040923 | 0.01332 | 0.089443 | 0.023664 | 0.037639
| Rel.Std.Dev. |  | 1.288% | 3.662% | 1.169% | 6.389% | 1.632% | 3.928%
| PDM3 |  | 1.915% | -8.227% | -6.441% | 14.97% | 19.08% | -21.3%
| Z-Score (Absolute) |  | 0.24 | 0.35 | 0.24 | 1 | 1.24 | 1.11
| Analytical results for Eu in OREAS 231b (Certified Value 0.44 ppm) | [32 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 0.442 | 0.378 | 0.49313 | 0.6 | 0.54 | 0.3 | [25 empty] | 1
| 1 | 2 | 0.44 | 0.367 | 0.50062 | 0.5 | 0.5 | 0.3 | [25 empty] | 32
| 1 | 3 | 0.438 | 0.342 | 0.48405 | 0.6 | 0.53 | 0.3 | [25 empty] | 16
| 1 | 4 | 0.437 | 0.364 | 0.50161 | 0.5 | 0.53 | 0.3 | [25 empty] | 0.44405
| 1 | 5 | 0.441 | 0.377 | 0.49712 | 0.5 | 0.53 | 0.3 | [25 empty] | 90
| 1 | 6 | 0.446 | 0.349 | 0.48421 | 0.5 | 0.54 | 0.3
| Mean |  | 0.44067 | 0.36283 | 0.49346 | 0.53333 | 0.52833 | 0.3
| Median |  | 0.4405 | 0.3655 | 0.49513 | 0.5 | 0.53 | 0.3
| Std Dev. |  | 0.0032042 | 0.014662 | 0.0078145 | 0.05164 | 0.01472 | 0
| Rel.Std.Dev. |  | 0.7271% | 4.041% | 1.584% | 9.682% | 2.786% | 0%
| PDM3 |  | -0.7616% | -18.29% | 11.13% | 20.11% | 18.98% | -32.44%
| Z-Score (Absolute) |  | 0.28 | 1.1 | 0.28 | 0.7 | 0.65 | 1.77
| Analytical results for Fe in OREAS 231b (Certified Value 5.91 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 6.15 | 5.967 | 4.1008 | 5.6887 | 6.157 | 6.17 | 6.001 | 6.57 | 6 | 5.97 | 5.66 | 5.87 | 5.86 | 5.9 | 5.6968 | 5.92 | 6.12 | > 5 | 5.96 | 5.37 | 5.94 | 7.75 | 6.035 | 5.342 | 5.73 | 5.72 | 6.06 | 5.75 | 5.74 |  |  | 1
| 1 | 2 | 6.13 | 6.06 | 3.7779 | 5.6568 | 6.1671 | 6.22 | 6.106 | 6.33 | 5.85 | 5.95 | 5.74 | 5.82 | 5.89 | 5.81 | 5.8638 | 5.99 | 5.99 | > 5 | 6.32 | 5.42 | 5.82 | 7.65 | 5.8498 | 5.442 | 5.71 | 5.62 | 5.88 | 5.95 | 5.76 |  |  | #N/A
| 1 | 3 | 6.05 | 5.959 | 3.858 | 5.7141 | 6.2399 | 6.3 | 6.092 | 6.36 | 5.84 | 6 | 5.77 | 5.78 | 5.89 | 5.88 | 5.7519 | 6.02 | 5.79 | > 5 | 6.01 | 5.19 | 5.8 | 6.45 | 5.8339 | 5.288 | 5.69 | 5.71 | 5.88 | 5.85 | 5.77 |  |  | 16
| 1 | 4 | 6.05 | 5.959 | 4.0949 | 5.7478 | 6.1418 | 6.12 | 6.113 | 6.4 | 5.66 | 5.99 | 5.8 | 5.84 | 5.91 | 5.86 | 6.1006 | 5.97 | 5.78 | > 5 | 6.22 | 5.27 | 5.8 | 5.99 | 5.8534 | 5.464 | 5.69 | 5.56 | 5.89 | 5.74 | 5.82 |  |  | 5.9057
| 1 | 5 | 6.03 | 6.079 | 3.848 | 5.7259 | 6.5973 | 6 | 6.141 | 6.43 | 5.83 | 6.01 | 5.76 | 5.76 | 5.88 | 5.89 | 5.7436 | 5.99 | 6.05 | > 5 | 6.23 | 5.27 | 5.94 | 5.48 | 5.8378 | 5.275 | 5.68 | 5.62 | 5.83 | 5.76 | 5.79 |  |  | 91
| 1 | 6 | 6 | 5.929 | 4.1262 | 5.6805 | 6.0545 | 6.2 | 6.169 | 6.39 | 6.03 | 5.96 | 5.73 | 5.82 | 5.85 | 5.89 | 5.9394 | 5.99 | 6.03 | > 5 | 6.11 | 5.18 | 5.82 | 4.44 | 5.8174 | 5.2525 | 5.7 | 5.78 | 5.86 | 5.76 | 5.84
| Mean |  | 6.0683 | 5.9922 | 3.9677 | 5.7023 | 6.2263 | 6.1683 | 6.1037 | 6.4133 | 5.8683 | 5.98 | 5.7433 | 5.815 | 5.88 | 5.8717 | 5.8494 | 5.98 | 5.96 |  | 6.1417 | 5.2833 | 5.8533 | 6.2933 | 5.8712 | 5.3439 | 5.7 | 5.6683 | 5.9 | 5.8017 | 5.7867
| Median |  | 6.05 | 5.963 | 3.9765 | 5.7014 | 6.162 | 6.185 | 6.1095 | 6.395 | 5.845 | 5.98 | 5.75 | 5.82 | 5.885 | 5.885 | 5.8079 | 5.99 | 6.01 |  | 6.165 | 5.27 | 5.82 | 6.22 | 5.8438 | 5.315 | 5.695 | 5.665 | 5.88 | 5.76 | 5.78
| Std Dev. |  | 0.058793 | 0.061587 | 0.15583 | 0.033157 | 0.19122 | 0.10167 | 0.057326 | 0.084063 | 0.13378 | 0.023664 | 0.04761 | 0.039875 | 0.021909 | 0.033116 | 0.15193 | 0.033466 | 0.14199 |  | 0.13934 | 0.095847 | 0.067725 | 1.2787 | 0.081264 | 0.089752 | 0.017889 | 0.081588 | 0.08124 | 0.082805 | 0.037771
| Rel.Std.Dev. |  | 0.9689% | 1.028% | 3.928% | 0.5815% | 3.071% | 1.648% | 0.9392% | 1.311% | 2.28% | 0.3957% | 0.829% | 0.6857% | 0.3726% | 0.564% | 2.597% | 0.5596% | 2.382% |  | 2.269% | 1.814% | 1.157% | 20.32% | 1.384% | 1.68% | 0.3138% | 1.439% | 1.377% | 1.427% | 0.6527%
| PDM3 |  | 2.754% | 1.464% | -32.82% | -3.444% | 5.428% | 4.447% | 3.352% | 8.596% | -0.6325% | 1.258% | -2.749% | -1.536% | -0.435% | -0.5761% | -0.9538% | 1.258% | 0.9197% |  | 3.996% | -10.54% | -0.8865% | 6.564% | -0.5838% | -9.512% | -3.483% | -4.019% | -0.0963% | -1.761% | -2.015%
| Z-Score (Absolute) |  | 1.07 | 0.65 | 10.32 | 0.92 | 1.92 | 1.61 | 1.26 | 2.94 | 0.02 | 0.59 | 0.69 | 0.31 | 0.05 | 0 | 0.12 | 0.59 | 0.48 | NA | 1.46 | 3.19 | 0.1 | 2.29 | 0 | 2.86 | 0.93 | 1.1 | 0.15 | 0.38 | 0.46
| Analytical results for Ga in OREAS 231b (Certified Value 10.2 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 9.8 | 10.7 | 9.5296 | 9.3059 | 10.5 | 11.8 | 11 | 10.3 | 11 | 10 | 10.3 | 10 | 9.28 | 10.4 | 7 | 10.8 | 9.7 | 12.2 | 13.19 | 10 | 10.05 | 11.6 | 10.5 | 8.28 | 9.1 | [6 empty] | 1
| 1 | 2 | 9.5 | 11 | 9.5976 | 9.2866 | 10.3 | 11.3 | 10.7 | 10.09 | 10 | 10 | 10.25 | 10 | 9.32 | 10.15 | 7 | 11.2 | 9.9 | 12.2 | 12.89 | 10 | 10.15 | 11 | 10.2 | 8.31 | 8.95 | [6 empty] | 33
| 1 | 3 | 9.4 | 11.1 | 9.5958 | 9.0298 | 10.6 | 11.6 | 11 | 9.88 | 10 | 10 | 10.25 | 10.45 | 10.3 | 10.8 | 7 | 10.8 | 9.5 | 12 | 13.27 | 10 | 10.05 | 11.3 | 10.1 | 8.35 | 8.96 | [6 empty] | 16
| 1 | 4 | 9.2 | 10.8 | 9.6733 | 8.8632 | 9.92 | 11.9 | 10.8 | 9.65 | 10 | 10 | 9.87 | 10.55 | 10.05 | 10.5 | 6 | 10.9 | 9.4 | 12.1 | 2.49 | 10 | 10.1 | 11.1 | 10.6 | 8.03 | 8.83 | [6 empty] | 10.2022
| 1 | 5 | 9.2 | 10.9 | 9.6859 | 9.6536 | 9.73 | 11.4 | 10.8 | 10.11 | 10 | 10 | 9.96 | 10.3 | 10.25 | 10.25 | 7 | 11.4 | 9.3 | 12 | 12.63 | 9 | 10.15 | 11 | 10.4 | 8.27 | 9.06 | [6 empty] | 92
| 1 | 6 | 9 | 10.8 | 9.5175 | 9.0864 | 10.1 | 11.7 | 10.9 | 10.52 | 10 | 10 | 9.9 | 10.3 | 9.88 | 10.65 | 7 | 11.1 | 9.3 | 12.2 | 12.72 | 9 | 10.1 | 11.2 | 10.1 | 8.29 | 9.08
| Mean |  | 9.35 | 10.8833 | 9.5999 | 9.2043 | 10.1917 | 11.6167 | 10.8667 | 10.0917 | 10.1667 | 10 | 10.0883 | 10.2667 | 9.8467 | 10.4583 | 6.8333 | 11.0333 | 9.5167 | 12.1167 | 11.1983 | 9.6667 | 10.1 | 11.2 | 10.3167 | 8.255 | 8.9967
| Median |  | 9.3 | 10.85 | 9.5967 | 9.1865 | 10.2 | 11.65 | 10.85 | 10.1 | 10 | 10 | 10.105 | 10.3 | 9.965 | 10.45 | 7 | 11 | 9.45 | 12.15 | 12.805 | 10 | 10.1 | 11.15 | 10.3 | 8.285 | 9.01
| Std Dev. |  | 0.28107 | 0.1472 | 0.070065 | 0.27556 | 0.33731 | 0.23166 | 0.12111 | 0.30564 | 0.40825 | 0 | 0.19833 | 0.2273 | 0.44925 | 0.24376 | 0.40825 | 0.24221 | 0.24014 | 0.098319 | 4.2737 | 0.5164 | 0.044721 | 0.22804 | 0.2137 | 0.1138 | 0.10289
| Rel.Std.Dev. |  | 3.006% | 1.352% | 0.7298% | 2.994% | 3.31% | 1.994% | 1.114% | 3.029% | 4.016% | 0% | 1.966% | 2.214% | 4.562% | 2.331% | 5.974% | 2.195% | 2.523% | 0.8114% | 38.16% | 5.342% | 0.4428% | 2.036% | 2.071% | 1.379% | 1.144%
| PDM3 |  | -8.353% | 6.676% | -5.903% | -9.782% | -0.1033% | 13.86% | 6.513% | -1.084% | -0.3484% | -1.982% | -1.116% | 0.6318% | -3.485% | 2.51% | -33.02% | 8.147% | -6.72% | 18.77% | 9.764% | -5.249% | -1.002% | 9.78% | 1.122% | -19.09% | -11.82%
| Z-Score (Absolute) |  | 0.84 | 0.69 | 0.59 | 0.99 | 0 | 1.42 | 0.67 | 0.1 | NA | NA | 0.1 | 0.07 | 0.34 | 0.27 | NA | 0.84 | 0.67 | 1.92 | 1.01 | NA | 0.09 | 1.01 | 0.12 | 1.93 | 1.19
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 231b (Certified Value 1.82 ppm) | [32 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.916 | 1.553 | 1.8372 | 2 | 1.97 | 1.6 | [25 empty] | 1
| 1 | 2 | 1.955 | 1.513 | 1.8941 | 2.1 | 1.82 | 1.7 | [25 empty] | 34
| 1 | 3 | 2.029 | 1.462 | 1.8875 | 2 | 1.91 | 1.6 | [25 empty] | 16
| 1 | 4 | 1.916 | 1.521 | 1.8951 | 2 | 1.87 | 1.6 | [25 empty] | 1.8169
| 1 | 5 | 1.982 | 1.564 | 1.9206 | 2.1 | 1.91 | 1.6 | [25 empty] | 93
| 1 | 6 | 1.949 | 1.451 | 1.8144 | 2.1 | 1.87 | 1.6
| Mean |  | 1.9578 | 1.5107 | 1.8748 | 2.05 | 1.8917 | 1.6167
| Median |  | 1.952 | 1.517 | 1.8908 | 2.05 | 1.89 | 1.6
| Std Dev. |  | 0.042976 | 0.046211 | 0.040251 | 0.054772 | 0.050761 | 0.040825
| Rel.Std.Dev. |  | 2.195% | 3.059% | 2.147% | 2.672% | 2.683% | 2.525%
| PDM3 |  | 7.754% | -16.86% | 3.185% | 12.83% | 4.113% | -11.02%
| Z-Score (Absolute) |  | 0.42 | 2.08 | 0.05 | 0.93 | 0.05 | 1.49
| Analytical results for Ge in OREAS 231b (Certified Value 0.13 ppm) | [32 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.1 | 0.1 | 0.25219 | < 0.2 | < 0.1 | 0.15 | 2.1 | 0.12 | 0.14 | 0.1 | 0.15 | < 10 | 0.1 | 0.17 | 0.17 | 0.2 | 0.11 | 0.12 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.1 | 0.26454 | < 0.2 | 0.1 | 0.16 | 2.1 | 0.14 | 0.16 | 0.1 | 0.14 | < 10 | 0.1 | 0.17 | 0.12 | 0.2 | 0.11 | 0.1 | [13 empty] | 4
| 1 | 3 | < 0.1 | 0.09 | 0.25005 | < 0.2 | 0.2 | 0.15 | 2 | 0.13 | 0.16 | 0.11 | 0.16 | < 10 | 0.1 | 0.17 | 0.13 | 0.2 | 0.12 | 0.1 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.09 | 0.26016 | < 0.2 | 0.1 | 0.16 | 2.1 | 0.13 | 0.16 | 0.1 | 0.14 | < 10 | 0.1 | 0.16 | 0.08 | 0.2 | 0.11 | 0.1 | [13 empty] | 0.13
| 1 | 5 | < 0.1 | 0.09 | 0.2369 | < 0.2 | 0.1 | 0.17 | 2.1 | 0.14 | 0.17 | 0.11 | 0.14 | < 10 | 0.1 | 0.16 | 0.1 | 0.2 | 0.12 | 0.09 | [13 empty] | 94
| 1 | 6 | < 0.1 | 0.09 | 0.25108 | < 0.2 | < 0.1 | 0.15 | 2.1 | 0.13 | 0.16 | 0.12 | 0.14 | < 10 | 0.1 | 0.18 | 0.15 | 0.2 | 0.11 | 0.1
| Mean |  |  | 0.093333 | 0.25249 |  | 0.125 | 0.15667 | 2.0833 | 0.13167 | 0.15833 | 0.10667 | 0.145 |  | 0.1 | 0.16833 | 0.125 | 0.2 | 0.11333 | 0.10167
| Median |  |  | 0.09 | 0.25163 |  | 0.1 | 0.155 | 2.1 | 0.13 | 0.16 | 0.105 | 0.14 |  | 0.1 | 0.17 | 0.125 | 0.2 | 0.11 | 0.1
| Std Dev. |  |  | 0.005164 | 0.0095417 |  | 0.05 | 0.008165 | 0.040825 | 0.0075277 | 0.0098319 | 0.008165 | 0.0083666 |  | 1.5202e-17 | 0.0075277 | 0.032711 | 3.0405e-17 | 0.005164 | 0.0098319
| Rel.Std.Dev. |  |  | 5.533% | 3.779% |  | 40% | 5.212% | 1.96% | 5.717% | 6.21% | 7.655% | 5.77% |  | 1.52e-14% | 4.472% | 26.17% | 1.52e-14% | 4.556% | 9.671%
| PDM3 |  |  | -28.21% | 94.22% |  | -3.846% | 20.51% | 1503% | 1.282% | 21.79% | -17.95% | 11.54% |  | -23.08% | 29.49% | -3.846% | 53.85% | -12.82% | -21.79%
| Z-Score (Absolute) |  | 1.86 | 0.83 | 2.95 | 0.67 | NA | 0.67 | NA | 0.08 | 0.71 | 0.52 | 0.4 | 115.94 | NA | 0.95 | 0.08 | NA | 0.36 | 0.63
Results from laboratories 1.07, 1.13, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 231b (Certified Value 0.4 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.34 | 0.42 | 0.31644 | 0.4 | 0.43 | 0.38 | 0.53 | 0.3 | 0.43 | 0.43 | 0.37 | 0.41 | 0.44 | < 0.5 | 0.41 | 0.99 | 0.4 | 0.46 | 0.34 | [12 empty] | 1
| 1 | 2 | 0.28 | 0.43 | 0.34383 | 0.4 | 0.36 | 0.39 | 0.53 | 0.3 | 0.42 | 0.42 | 0.39 | 0.36 | 0.43 | < 0.5 | 0.41 | 0.89 | 0.39 | 0.44 | 0.36 | [12 empty] | 19
| 1 | 3 | 0.29 | 0.43 | 0.3251 | 0.4 | 0.4 | 0.42 | 0.53 | 0.3 | 0.37 | 0.43 | 0.42 | 0.39 | 0.45 | < 0.5 | 0.39 | 0.8 | 0.39 | 0.46 | 0.34 | [12 empty] | 16
| 1 | 4 | 0.29 | 0.4 | 0.32061 | 0.4 | 0.37 | 0.41 | 0.53 | 0.3 | 0.41 | 0.42 | 0.4 | 0.38 | 0.45 | < 0.5 | 0.42 | 0.74 | 0.4 | 0.44 | 0.32 | [12 empty] | 0.39831
| 1 | 5 | 0.26 | 0.41 | 0.32048 | 0.4 | 0.34 | 0.39 | 0.53 | 0.3 | 0.36 | 0.43 | 0.45 | 0.38 | 0.45 | < 0.5 | 0.42 | 0.71 | 0.41 | 0.45 | 0.37 | [12 empty] | 95
| 1 | 6 | 0.27 | 0.41 | 0.32443 | 0.3 | 0.37 | 0.4 | 0.52 | 0.3 | 0.35 | 0.44 | 0.42 | 0.39 | 0.43 | < 0.5 | 0.42 | 0.67 | 0.38 | 0.47 | 0.37
| Mean |  | 0.28833 | 0.41667 | 0.32515 | 0.38333 | 0.37833 | 0.39833 | 0.52833 | 0.3 | 0.39 | 0.42833 | 0.40833 | 0.385 | 0.44167 |  | 0.41167 | 0.8 | 0.395 | 0.45333 | 0.35
| Median |  | 0.285 | 0.415 | 0.32252 | 0.4 | 0.37 | 0.395 | 0.53 | 0.3 | 0.39 | 0.43 | 0.41 | 0.385 | 0.445 |  | 0.415 | 0.77 | 0.395 | 0.455 | 0.35
| Std Dev. |  | 0.027869 | 0.012111 | 0.0096728 | 0.040825 | 0.031885 | 0.01472 | 0.0040825 | 0 | 0.034059 | 0.0075277 | 0.027869 | 0.016432 | 0.0098319 |  | 0.01169 | 0.12066 | 0.010488 | 0.012111 | 0.02
| Rel.Std.Dev. |  | 9.665% | 2.907% | 2.975% | 10.65% | 8.428% | 3.695% | 0.7727% | 0% | 8.733% | 1.757% | 6.825% | 4.268% | 2.226% |  | 2.84% | 15.08% | 2.655% | 2.671% | 5.714%
| PDM3 |  | -27.61% | 4.608% | -18.37% | -3.761% | -5.016% | 0.004942% | 32.64% | -24.68% | -2.087% | 7.537% | 2.516% | -3.343% | 10.88% |  | 3.352% | 100.8% | -0.8319% | 13.81% | -12.13%
| Z-Score (Absolute) |  | 2.47 | 0.41 | 1.64 | NA | 0.45 | 0 | 2.92 | NA | 0.19 | 0.67 | 0.22 | 0.3 | 0.97 | 3.33 | 0.3 | 9.03 | 0.07 | 1.24 | 1.09
Results from laboratories 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 231b (Indicative Value 0.036 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.03 | < 1 | 0.05 | 0.43 | < 0.1 | < 5 | 0.48 | 0.02 | 0.02 | < 0.01 | 0.05 | < 1 | 0.05 | 0.017 | 0.05 | 0.42 | 0.04 | < 0.08 | 0.03 | 0.079 | 0.14 | < 0.01 | 0.018 | 0.21 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.02 | < 1 | 0.05 | 0.44 | < 0.1 | < 5 | 0.567 | 0.02 | 0.02 | < 0.01 | 0.04 | < 1 | 0.06 | 0.011 | 0.05 | < 0.1 | 0.04 | < 0.08 | 0.03 | 0.068 | 0.15 | < 0.01 | 0.02 | 0.29 | [6 empty] | 6
| 1 | 3 | < 0.01 | 0.02 | < 1 | 0.07 | 0.44 | < 0.1 | < 5 | 0.586 | 0.01 | 0.02 | < 0.01 | 0.04 | < 1 | 0.06 | 0.011 | 0.04 | 0.73 | 0.04 | < 0.08 | 0.03 | 0.067 | 0.16 | < 0.01 | 0.019 | 0.26 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.02 | < 1 | 0.08 | 0.44 | < 0.1 | < 5 | 0.666 | 0.02 | 0.02 | < 0.01 | 0.03 | < 1 | 0.06 | 0.015 | 0.05 | 0.37 | 0.04 | < 0.08 | 0.03 | 0.067 | 0.14 | < 0.01 | 0.022 | 0.25 | [6 empty] | 0.035994
| 1 | 5 | < 0.01 | 0.02 | < 1 | 0.07 | 0.44 | < 0.1 | < 5 | 0.671 | < 0.01 | 0.02 | < 0.01 | 0.02 | < 1 | 0.05 | 0.015 | 0.04 | < 0.1 | 0.04 | < 0.08 | 0.03 | 0.057 | 0.15 | < 0.01 | 0.02 | 0.22 | [6 empty] | 12
| 1 | 6 | < 0.01 | 0.02 | < 1 | 0.06 | 0.43 | < 0.1 | < 5 | 0.559 | 0.01 | 0.02 | < 0.01 | 0.06 | < 1 | 0.06 | 0.012 | 0.04 | 0.388 | 0.04 | < 0.08 | 0.02 | 0.042 | 0.16 | < 0.01 | 0.02 | 0.23
| Mean |  |  | 0.021667 |  | 0.063333 | 0.43667 |  |  | 0.58817 | 0.016 | 0.02 |  | 0.04 |  | 0.056667 | 0.0135 | 0.045 | 0.477 | 0.04 |  | 0.028333 | 0.063333 | 0.15 |  | 0.019833 | 0.24333
| Median |  |  | 0.02 |  | 0.065 | 0.44 |  |  | 0.5765 | 0.02 | 0.02 |  | 0.04 |  | 0.06 | 0.0135 | 0.045 | 0.404 | 0.04 |  | 0.03 | 0.067 | 0.15 |  | 0.02 | 0.24
| Std Dev. |  |  | 0.0040825 |  | 0.012111 | 0.005164 |  |  | 0.072004 | 0.0054772 | 0 |  | 0.014142 |  | 0.005164 | 0.00251 | 0.0054772 | 0.16993 | 0 |  | 0.0040825 | 0.012565 | 0.0089443 |  | 0.0013292 | 0.029439
| Rel.Std.Dev. |  |  | 18.84% |  | 19.12% | 1.183% |  |  | 12.24% | 34.23% | 0% |  | 35.36% |  | 9.113% | 18.59% | 12.17% | 35.62% | 0% |  | 14.41% | 19.84% | 5.963% |  | 6.702% | 12.1%
| PDM3 |  |  | -39.81% |  | 75.95% | 1113% |  |  | 1534% | -55.55% | -44.44% |  | 11.13% |  | 57.43% | -62.49% | 25.02% | 1225% | 11.13% |  | -21.28% | 75.95% | 316.7% |  | -44.9% | 576%
| Z-Score (Absolute) |  | 0.87 | 0.51 | 9.95 | 0.4 | 8.57 | 0.11 | 53.69 | 11.88 | 0.67 | 0.55 | 0.87 | 0.11 | 9.95 | 0.26 | 0.69 | 0 | 6.33 | 0.11 | 0.11 | 0.36 | 0.4 | 2.3 | 0.87 | 0.55 | 4.34
| Analytical results for Ho in OREAS 231b (Certified Value 0.42 ppm) | [32 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 0.42 | 0.417 | 0.38147 | 0.5 | 0.56 | 0.4 | [25 empty] | 1
| 1 | 2 | 0.434 | 0.411 | 0.3854 | 0.5 | 0.53 | 0.45 | [25 empty] | 20
| 1 | 3 | 0.431 | 0.39 | 0.38975 | 0.5 | 0.56 | 0.4 | [25 empty] | 16
| 1 | 4 | 0.418 | 0.404 | 0.39291 | 0.5 | 0.54 | 0.4 | [25 empty] | 0.42389
| 1 | 5 | 0.434 | 0.423 | 0.39439 | 0.5 | 0.56 | 0.4 | [25 empty] | 96
| 1 | 6 | 0.427 | 0.393 | 0.37092 | 0.5 | 0.55 | 0.4
| Mean |  | 0.42733 | 0.40633 | 0.38581 | 0.5 | 0.55 | 0.40833
| Median |  | 0.429 | 0.4075 | 0.38758 | 0.5 | 0.555 | 0.4
| Std Dev. |  | 0.0069761 | 0.01314 | 0.0087194 | 0 | 0.012649 | 0.020412
| Rel.Std.Dev. |  | 1.632% | 3.234% | 2.26% | 0% | 2.3% | 4.999%
| PDM3 |  | 0.8112% | -4.143% | -8.985% | 17.95% | 29.75% | -3.671%
| Z-Score (Absolute) |  | 0.29 | 0.36 | 0.99 | 2.55 | 4.1 | 0.29
| Analytical results for In in OREAS 231b (Certified Value 0.024 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.02 | 0.023915 | 0.03 | < 0.05 | 0.03 | < 5 | 0.04 | 0.026 | 0.025 | 0.019 | 0.024 | < 10 | < 0.05 | 0.53 | < 0.1 | 0.02156 | 0.025 | 0.034 | 0.02 | 0.024 | 0.02 | [9 empty] | 1
| 1 | 2 | 0.02 | 0.02 | 0.021762 | 0.03 | < 0.05 | 0.02 | < 5 | 0.04 | 0.025 | 0.024 | 0.021 | 0.026 | < 10 | < 0.05 | 0.9 | < 0.1 | 0.02385 | 0.028 | 0.033 | 0.02 | 0.025 | 0.022 | [9 empty] | 21
| 1 | 3 | 0.02 | 0.02 | 0.023683 | 0.02 | < 0.05 | 0.03 | < 5 | 0.03 | 0.028 | 0.026 | 0.022 | 0.024 | < 10 | < 0.05 | 1.16 | < 0.1 | 0.02715 | 0.023 | 0.024 | 0.02 | 0.024 | 0.019 | [9 empty] | 16
| 1 | 4 | < 0.02 | 0.02 | 0.022789 | 0.03 | < 0.05 | 0.02 | < 5 | 0.03 | 0.026 | 0.023 | 0.024 | 0.025 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.02542 | 0.024 | 0.034 | 0.02 | 0.024 | 0.02 | [9 empty] | 0.024324
| 1 | 5 | 0.02 | 0.02 | 0.023949 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.025 | 0.026 | 0.024 | 0.022 | < 10 | < 0.05 | 0.82 | < 0.1 | 0.02577 | 0.026 | 0.028 | 0.02 | 0.023 | 0.018 | [9 empty] | 97
| 1 | 6 | < 0.02 | 0.02 | 0.022391 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.024 | 0.025 | 0.022 | 0.026 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.02683 | 0.025 | 0.026 | 0.02 | 0.025 | 0.02
| Mean |  | 0.02 | 0.02 | 0.023082 | 0.028333 |  | 0.026667 |  | 0.033333 | 0.025667 | 0.024833 | 0.022 | 0.0245 |  |  | 0.8525 |  | 0.025097 | 0.025167 | 0.029833 | 0.02 | 0.024167 | 0.019833
| Median |  | 0.02 | 0.02 | 0.023236 | 0.03 |  | 0.03 |  | 0.03 | 0.0255 | 0.025 | 0.022 | 0.0245 |  |  | 0.86 |  | 0.025595 | 0.025 | 0.0305 | 0.02 | 0.024 | 0.02
| Std Dev. |  | 0 | 0 | 0.00090686 | 0.0040825 |  | 0.005164 |  | 0.005164 | 0.0013663 | 0.001169 | 0.0018974 | 0.0015166 |  |  | 0.25941 |  | 0.00209 | 0.0017224 | 0.0044008 | 0 | 0.00075277 | 0.0013292
| Rel.Std.Dev. |  | 0% | 0% | 3.929% | 14.41% |  | 19.36% |  | 15.49% | 5.323% | 4.708% | 8.624% | 6.19% |  |  | 30.43% |  | 8.328% | 6.844% | 14.75% | 0% | 3.115% | 6.702%
| PDM3 |  | -17.78% | -17.78% | -5.106% | 16.48% |  | 9.633% |  | 37.04% | 5.521% | 2.095% | -9.553% | 0.7251% |  |  | 3405% |  | 3.178% | 3.466% | 22.65% | -17.78% | -0.6453% | -18.46%
| Z-Score (Absolute) |  | 1.78 | 1.07 | 0.42 | 0.7 | 0.01 | 0.34 | 527.02 | 1.76 | 0.13 | 0.05 | 0.65 | 0.12 | 1059.36 | 0.01 | 119.24 | 5.31 | 0.01 | 0.03 | 1.02 | 1.07 | 0.19 | 1.11
| Analytical results for Ir in OREAS 231b (Indicative Value < 0.003 ppm) | [32 empty] | Indicative
Ir | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.27 | 1.28 | [29 empty] | ppm
|  |  | AR*MS | AR*MS | [29 empty] | 3
|  |  | 25g | 0.5g | [29 empty] | 3
| 1 | 1 | N.A. | < 0.003 | [29 empty] | 1
| 1 | 2 | N.A. | < 0.003 | [29 empty] | 7
| 1 | 3 | N.A. | < 0.003 | [29 empty] | 16
| 1 | 4 | N.A. | < 0.003 | [29 empty] | < 0.003
| 1 | 5 | N.A. | < 0.003 | [29 empty] | 13
| 1 | 6 | N.A. | < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 231b (Certified Value 0.104 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.11 | 0.095 | 0.15631 | 0.11931 | 0.1082 | 0.09 | 0.1104 | 0.127 | 0.1154 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10119 | 0.1 | 0.09 | 0.10889 | 0.112 | 0.1 | 0.11 | 0.10262 | 0.11005 | 0.095 | 0.1 | 0.11 | 0.13 | 0.1 | 0.1 |  |  | 1
| 1 | 2 | 0.1 | 0.098 | 0.15963 | 0.12061 | 0.10728 | 0.09 | 0.1146 | 0.126 | 0.1095 | 0.11 | 0.1 | 0.1 | 0.11 | 0.1 | 0.10297 | 0.1 | 0.09 | 0.10469 | 0.112 | 0.11 | 0.11 | 0.10271 | 0.1002 | 0.093 | 0.1 | 0.11 | 0.12 | 0.11 | 0.1 |  |  | #N/A
| 1 | 3 | 0.11 | 0.096 | 0.15169 | 0.12306 | 0.10803 | 0.09 | 0.1129 | 0.129 | 0.1099 | 0.11 | 0.12 | 0.1 | 0.11 | 0.1 | 0.10084 | 0.1 | 0.09 | 0.10689 | 0.111 | 0.1 | 0.11 | 0.10312 | 0.10235 | 0.09 | 0.1 | 0.11 | 0.12 | 0.1 | 0.1 |  |  | 16
| 1 | 4 | 0.1 | 0.095 | 0.15853 | 0.11668 | 0.10705 | 0.08 | 0.1137 | 0.129 | 0.1057 | 0.1 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10771 | 0.1 | 0.09 | 0.1056 | 0.113 | 0.1 | 0.11 | 0.1001 | 0.10243 | 0.093 | 0.1 | 0.1 | 0.12 | 0.1 | 0.1 |  |  | 0.10389
| 1 | 5 | 0.1 | 0.097 | 0.14715 | 0.11522 | 0.11416 | 0.09 | 0.1162 | 0.129 | 0.1084 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10158 | 0.1 | 0.09 | 0.10646 | 0.113 | 0.1 | 0.11 | 0.099051 | 0.10006 | 0.0875 | 0.1 | 0.1 | 0.12 | 0.1 | 0.1 |  |  | 98
| 1 | 6 | 0.1 | 0.097 | 0.16013 | 0.12149 | 0.10362 | 0.09 | 0.1154 | 0.125 | 0.1116 | 0.1 | 0.1 | 0.1 | 0.11 | 0.1 | 0.10445 | 0.1 | 0.09 | 0.11535 | 0.111 | 0.1 | 0.11 | 0.10016 | 0.10281 | 0.0885 | 0.1 | 0.1 | 0.12 | 0.1 | 0.1
| Mean |  | 0.10333 | 0.096333 | 0.15557 | 0.11939 | 0.10806 | 0.088333 | 0.11387 | 0.1275 | 0.11008 | 0.10667 | 0.10833 | 0.1 | 0.11 | 0.1 | 0.10312 | 0.1 | 0.09 | 0.10798 | 0.112 | 0.10167 | 0.11 | 0.1013 | 0.10298 | 0.091167 | 0.1 | 0.105 | 0.12167 | 0.10167 | 0.1
| Median |  | 0.1 | 0.0965 | 0.15742 | 0.11996 | 0.10766 | 0.09 | 0.11415 | 0.128 | 0.1097 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10228 | 0.1 | 0.09 | 0.10667 | 0.112 | 0.1 | 0.11 | 0.10139 | 0.10239 | 0.0915 | 0.1 | 0.105 | 0.12 | 0.1 | 0.1
| Std Dev. |  | 0.005164 | 0.0012111 | 0.0051487 | 0.0029701 | 0.0034226 | 0.0040825 | 0.0020646 | 0.0017607 | 0.0032554 | 0.005164 | 0.0075277 | 1.5202e-17 | 0 | 1.5202e-17 | 0.0026134 | 1.5202e-17 | 1.5202e-17 | 0.0038756 | 0.00089443 | 0.0040825 | 0 | 0.0017239 | 0.0036597 | 0.0029439 | 1.5202e-17 | 0.0054772 | 0.0040825 | 0.0040825 | 1.5202e-17
| Rel.Std.Dev. |  | 4.997% | 1.257% | 3.309% | 2.488% | 3.167% | 4.622% | 1.813% | 1.381% | 2.957% | 4.841% | 6.949% | 1.52e-14% | 0% | 1.52e-14% | 2.534% | 1.52e-14% | 1.689e-14% | 3.589% | 0.7986% | 4.016% | 0% | 1.702% | 3.554% | 3.229% | 1.52e-14% | 5.216% | 3.355% | 4.016% | 1.52e-14%
| PDM3 |  | -0.535% | -7.273% | 49.75% | 14.93% | 4.011% | -14.97% | 9.604% | 22.73% | 5.962% | 2.674% | 4.278% | -3.744% | 5.882% | -3.744% | -0.7356% | -3.744% | -13.37% | 3.936% | 7.807% | -2.139% | 5.882% | -2.497% | -0.8719% | -12.25% | -3.744% | 1.069% | 17.11% | -2.139% | -3.744%
| Z-Score (Absolute) |  | 0 | 1 | 7.46 | 2.29 | 0.67 | 2.14 | 1.5 | 3.45 | 0.96 | 0.48 | 0.71 | 0.48 | 0.95 | 0.48 | 0.03 | 0.48 | 1.9 | 0.66 | 1.24 | 0.24 | 0.95 | 0.29 | 0.05 | 1.74 | 0.48 | 0.24 | 2.62 | 0.24 | 0.48
| Analytical results for La in OREAS 231b (Certified Value 3.34 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 3.1 | 3.354 | 3.067 | 3.59 | 3.9771 | 3.4 | < 1 | 4 | 3.448 | < 5 | 3 | 3.2 | 3.4 | 3 | 3.3 | 3 | 3.58 | 3 | 3.3 | 3.5 | 3.3 | 3.5 | 3.4 | 3.4 | 3.4 | [6 empty] | 1
| 1 | 2 | 3.1 | 3.391 | 2.961 | 3.67 | 3.9445 | 3.4 | < 1 | 3.88 | 3.344 | < 5 | 4 | 3.2 | 3.4 | 3 | 3.3 | 3 | 3.52 | 3.1 | 3.3 | 3.5 | 3.3 | 3.5 | 3.3 | 3.5 | 3.4 | [6 empty] | 22
| 1 | 3 | 3.1 | 3.421 | 2.918 | 3.71 | 3.9817 | 3.5 | < 1 | 3.97 | 3.38 | < 5 | 4 | 3.2 | 3.4 | 3.3 | 3.4 | 3 | 3.27 | 2.9 | 3.3 | 3.5 | 3.3 | 3.5 | 3.3 | 3.5 | 3.4 | [6 empty] | 16
| 1 | 4 | 3.1 | 3.335 | 3.06 | 3.7 | 3.923 | 3.4 | < 1 | 3.94 | 3.261 | < 5 | 4 | 3.1 | 3.5 | 3.3 | 3.3 | 3 | 3.23 | 3 | 3.2 | 3.5 | 3.4 | 3.5 | 3.4 | 3.4 | 3.4 | [6 empty] | 3.3424
| 1 | 5 | 3 | 3.434 | 3.174 | 3.65 | 4.2242 | 3.4 | < 1 | 3.89 | 3.275 | < 5 | 4 | 3.1 | 3.4 | 3.3 | 3.3 | 3 | 3.23 | 2.9 | 3.3 | 3.5 | 3.3 | 3.5 | 3.3 | 3.4 | 3.4 | [6 empty] | 99
| 1 | 6 | 3 | 3.355 | 2.897 | 3.61 | 3.7741 | 3.4 | < 1 | 3.95 | 3.359 | < 5 | 4 | 3.1 | 3.5 | 3.2 | 3.4 | 3 | 3.38 | 2.9 | 3.2 | 3.5 | 3.4 | 3.6 | 3.3 | 3.5 | 3.5
| Mean |  | 3.0667 | 3.3817 | 3.0128 | 3.655 | 3.9708 | 3.4167 |  | 3.9383 | 3.3445 |  | 3.8333 | 3.15 | 3.4333 | 3.1833 | 3.3333 | 3 | 3.3683 | 2.9667 | 3.2667 | 3.5 | 3.3333 | 3.5167 | 3.3333 | 3.45 | 3.4167
| Median |  | 3.1 | 3.373 | 3.0105 | 3.66 | 3.9608 | 3.4 |  | 3.945 | 3.3515 |  | 4 | 3.15 | 3.4 | 3.25 | 3.3 | 3 | 3.325 | 2.95 | 3.3 | 3.5 | 3.3 | 3.5 | 3.3 | 3.45 | 3.4
| Std Dev. |  | 0.05164 | 0.040068 | 0.10605 | 0.048062 | 0.14561 | 0.040825 |  | 0.046224 | 0.069261 |  | 0.40825 | 0.054772 | 0.05164 | 0.1472 | 0.05164 | 0 | 0.15224 | 0.08165 | 0.05164 | 0 | 0.05164 | 0.040825 | 0.05164 | 0.054772 | 0.040825
| Rel.Std.Dev. |  | 1.684% | 1.185% | 3.52% | 1.315% | 3.667% | 1.195% |  | 1.174% | 2.071% |  | 10.65% | 1.739% | 1.504% | 4.624% | 1.549% | 0% | 4.52% | 2.752% | 1.581% | 0% | 1.549% | 1.161% | 1.549% | 1.588% | 1.195%
| PDM3 |  | -8.248% | 1.176% | -9.859% | 9.354% | 18.8% | 2.223% |  | 17.83% | 0.06427% |  | 14.69% | -5.755% | 2.722% | -4.758% | -0.2698% | -10.24% | 0.7773% | -11.24% | -2.264% | 4.717% | -0.2698% | 5.215% | -0.2698% | 3.221% | 2.223%
| Z-Score (Absolute) |  | 1.2 | 0.16 | 1.44 | 1.35 | 2.72 | 0.31 | 12.33 | 2.58 | 0 | 3.66 | NA | 0.84 | 0.39 | 0.7 | 0.05 | NA | 0.1 | 1.64 | 0.34 | 0.67 | 0.05 | 0.75 | 0.05 | 0.46 | 0.31
Results from laboratories 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 231b (Certified Value 10.6 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 1 | 10.47 | 16.78 | 5.4021 | 11.1797 | 10.4 | 6 | 11.9 | 10.2 | 12 | 16.1 | 11.1 | 10.9 | 11.1 | 10.6 | < 5 | 10.0171 | 9.6 | 10 | 10 | 10.9 | 10 | 12 | 10.9 | 10.1 | [6 empty] | 1
| 1 | 2 | 2 | 10.26 | 18.54 | 5.4157 | 11.2017 | 10.2 | 6 | 11.6 | 10.3 | 11 | 16.2 | 10.9 | 11 | 11 | 10.7 | < 5 | 10.5846 | 10.7 | 10 | 10 | 10.8 | 9.6 | 12 | 11.2 | 10.1 | [6 empty] | #N/A
| 1 | 3 | 2 | 10.4 | 17.16 | 5.135 | 11.2913 | 10.6 | 6 | 11.7 | 10.3 | 11 | 16.2 | 10.9 | 11 | 11.2 | 10.7 | < 5 | 9.4235 | 10.2 | 10 | 10 | 10.8 | 9.7 | 11 | 10.8 | 9.8 | [6 empty] | 16
| 1 | 4 | 2 | 10.16 | 17.49 | 5.0455 | 10.8487 | 9.7 | 6 | 11.8 | 10.3 | 11 | 16.1 | 11 | 11.1 | 11 | 10.7 | < 5 | 9.2233 | 10.4 | 10 | 10 | 10.8 | 9.5 | 11 | 10.9 | 10.1 | [6 empty] | 10.5797
| 1 | 5 | 2 | 10.26 | 15.71 | 5.358 | 11.7228 | 9.8 | 6 | 12.1 | 10.2 | 11 | 16.3 | 10.9 | 11 | 11 | 10.7 | < 5 | 8.7489 | 10.7 | 10 | 9 | 10.9 | 9.4 | 11 | 10.9 | 10.3 | [6 empty] | 100
| 1 | 6 | 2 | 10.06 | 17.46 | 5.4569 | 10.4856 | 9.7 | 6 | 12.2 | 10.4 | 11 | 14.9 | 11 | 11 | 11 | 10.6 | < 5 | 10.336 | 9.7 | 10 | 9 | 10.9 | 9.9 | 12 | 10.8 | 9.6
| Mean |  | 1.8333 | 10.2683 | 17.19 | 5.3022 | 11.1216 | 10.0667 | 6 | 11.8833 | 10.2833 | 11.1667 | 15.9667 | 10.9667 | 11 | 11.05 | 10.6667 |  | 9.7222 | 10.2167 | 10 | 9.6667 | 10.85 | 9.6833 | 11.5 | 10.9167 | 10
| Median |  | 2 | 10.26 | 17.31 | 5.3801 | 11.1907 | 10 | 6 | 11.85 | 10.3 | 11 | 16.15 | 10.95 | 11 | 11 | 10.7 |  | 9.7203 | 10.3 | 10 | 10 | 10.85 | 9.65 | 11.5 | 10.9 | 10.1
| Std Dev. |  | 0.40825 | 0.15052 | 0.93218 | 0.16957 | 0.41931 | 0.38816 | 0 | 0.23166 | 0.075277 | 0.40825 | 0.52789 | 0.08165 | 0.063246 | 0.083666 | 0.05164 |  | 0.70608 | 0.47924 | 0 | 0.5164 | 0.054772 | 0.23166 | 0.54772 | 0.1472 | 0.25298
| Rel.Std.Dev. |  | 22.27% | 1.466% | 5.423% | 3.198% | 3.77% | 3.856% | 0% | 1.949% | 0.732% | 3.656% | 3.306% | 0.7445% | 0.575% | 0.7572% | 0.4841% |  | 7.263% | 4.691% | 0% | 5.342% | 0.5048% | 2.392% | 4.763% | 1.348% | 2.53%
| PDM3 |  | -82.67% | -2.943% | 62.48% | -49.88% | 5.122% | -4.849% | -43.29% | 12.32% | -2.801% | 5.548% | 50.92% | 3.658% | 3.973% | 4.445% | 0.822% |  | -8.105% | -3.431% | -5.479% | -8.63% | 2.555% | -8.473% | 8.699% | 3.185% | -5.479%
| Z-Score (Absolute) |  | NA | 0.59 | 9.67 | 7.95 | 0.67 | 0.89 | NA | 1.8 | 0.57 | NA | 7.86 | 0.44 | 0.49 | 0.57 | 0 | 12.1 | 1.4 | 0.67 | NA | NA | 0.27 | 1.46 | NA | 0.37 | 0.99
Results from laboratories 1.01, 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 231b (Certified Value 0.13 ppm) | [32 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | 1.32 | [24 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [24 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [24 empty] | 2
| 1 | 1 | 0.13 | 0.145 | 0.135 | 0.34865 | 0.2 | 0.1 | 0.13 | [24 empty] | 1
| 1 | 2 | 0.12 | 0.147 | 0.138 | 0.34597 | 0.2 | 0.1 | 0.13 | [24 empty] | 23
| 1 | 3 | 0.12 | 0.15 | 0.124 | 0.34517 | 0.1 | 0.1 | 0.13 | [24 empty] | 16
| 1 | 4 | 0.12 | 0.15 | 0.136 | 0.34381 | 0.2 | 0.1 | 0.12 | [24 empty] | 0.12997
| 1 | 5 | 0.11 | 0.15 | 0.138 | 0.36182 | 0.1 | 0.1 | 0.13 | [24 empty] | 101
| 1 | 6 | 0.12 | 0.146 | 0.13 | 0.33862 | 0.1 | 0.1 | 0.13
| Mean |  | 0.12 | 0.148 | 0.1335 | 0.34734 | 0.15 | 0.1 | 0.12833
| Median |  | 0.12 | 0.1485 | 0.1355 | 0.34557 | 0.15 | 0.1 | 0.13
| Std Dev. |  | 0.0063246 | 0.0022804 | 0.0055045 | 0.0078283 | 0.054772 | 1.5202e-17 | 0.0040825
| Rel.Std.Dev. |  | 5.27% | 1.541% | 4.123% | 2.254% | 36.51% | 1.52e-14% | 3.181%
| PDM3 |  | -7.673% | 13.87% | 2.714% | 167.2% | 15.41% | -23.06% | -1.261%
| Z-Score (Absolute) |  | 0.63 | 0.67 | 0 | 9.94 | 0.77 | 1.56 | 0.24
| Analytical results for Mg in OREAS 231b (Certified Value 1.96 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 2.03 | 1.91 | 1.9017 | 2.0095 | 2.0159 | 2.08 | 2.1107 | 2.17 | 1.96 | 2.05 | 1.84 | 1.95 | 1.93 | 1.9 | 1.8961 | 2 | 1.89 | 1.9932 | 1.96 | 1.92 | 2 | 2.5042 | 1.6659 | 1.8145 | 1.91 | 1.92 | 1.95 | 1.91 | 2 |  |  | 1
| 1 | 2 | 2.03 | 1.955 | 1.7737 | 2.016 | 2.027 | 2.04 | 2.1409 | 2.12 | 1.93 | 2.05 | 1.86 | 1.94 | 1.94 | 1.87 | 1.9466 | 2.01 | 1.85 | 1.9797 | 2.07 | 1.93 | 1.95 | 2.3116 | 1.6423 | 1.862 | 1.9 | 1.89 | 1.91 | 1.97 | 2.01 |  |  | #N/A
| 1 | 3 | 2.01 | 1.917 | 1.7646 | 2.0649 | 2.0402 | 2.09 | 2.1168 | 2.14 | 1.93 | 2.04 | 1.88 | 1.92 | 1.94 | 1.89 | 1.9062 | 2.02 | 1.79 | 2.0045 | 1.99 | 1.88 | 1.95 | 2.5847 | 1.6481 | 1.8025 | 1.9 | 1.92 | 1.91 | 1.95 | 2.02 |  |  | 16
| 1 | 4 | 2.01 | 1.911 | 1.9312 | 2.0872 | 1.9906 | 1.93 | 2.159 | 2.16 | 1.86 | 1.96 | 1.89 | 1.95 | 1.96 | 1.89 | 2.0333 | 2.01 | 1.78 | 1.9925 | 2.03 | 1.87 | 1.95 | 2.1962 | 1.6565 | 1.871 | 1.9 | 1.87 | 1.91 | 1.91 | 2.02 |  |  | 1.9568
| 1 | 5 | 2 | 1.943 | 1.7174 | 2.0621 | 2.0175 | 1.94 | 2.1469 | 2.17 | 1.89 | 2.02 | 1.88 | 1.9 | 1.94 | 1.9 | 1.9094 | 2.02 | 1.85 | 2.0507 | 2.03 | 1.86 | 1.99 | 2.5412 | 1.6389 | 1.7615 | 1.9 | 1.87 | 1.9 | 1.91 | 2.01 |  |  | 102
| 1 | 6 | 1.99 | 1.933 | 1.9479 | 2.0125 | 1.958 | 1.94 | 2.1771 | 2.16 | 1.98 | 1.96 | 1.87 | 1.93 | 1.93 | 1.9 | 1.9742 | 2.02 | 1.86 | 2.0582 | 2 | 1.84 | 1.95 | 2.1953 | 1.6341 | 1.785 | 1.9 | 1.9 | 1.9 | 1.91 | 2.06
| Mean |  | 2.0117 | 1.9282 | 1.8394 | 2.042 | 2.0082 | 2.0033 | 2.1419 | 2.1533 | 1.925 | 2.0133 | 1.87 | 1.9317 | 1.94 | 1.8917 | 1.9443 | 2.0133 | 1.8367 | 2.0131 | 2.0133 | 1.8833 | 1.965 | 2.3889 | 1.6476 | 1.8161 | 1.9017 | 1.895 | 1.9133 | 1.9267 | 2.02
| Median |  | 2.01 | 1.925 | 1.8377 | 2.0391 | 2.0167 | 1.99 | 2.1439 | 2.16 | 1.93 | 2.03 | 1.875 | 1.935 | 1.94 | 1.895 | 1.928 | 2.015 | 1.85 | 1.9988 | 2.015 | 1.875 | 1.95 | 2.4079 | 1.6452 | 1.8085 | 1.9 | 1.895 | 1.91 | 1.91 | 2.015
| Std Dev. |  | 0.016021 | 0.018509 | 0.098854 | 0.033388 | 0.029526 | 0.075011 | 0.025143 | 0.019664 | 0.044159 | 0.04274 | 0.017889 | 0.019408 | 0.010954 | 0.01169 | 0.052574 | 0.008165 | 0.04274 | 0.033044 | 0.038297 | 0.035024 | 0.023452 | 0.17631 | 0.011829 | 0.043027 | 0.0040825 | 0.022583 | 0.018619 | 0.026583 | 0.020976
| Rel.Std.Dev. |  | 0.7964% | 0.9599% | 5.374% | 1.635% | 1.47% | 3.744% | 1.174% | 0.9132% | 2.294% | 2.123% | 0.9566% | 1.005% | 0.5647% | 0.618% | 2.704% | 0.4055% | 2.327% | 1.641% | 1.902% | 1.86% | 1.193% | 7.38% | 0.718% | 2.369% | 0.2147% | 1.192% | 0.9731% | 1.38% | 1.038%
| PDM3 |  | 2.805% | -1.462% | -5.997% | 4.357% | 2.628% | 2.379% | 9.461% | 10.05% | -1.624% | 2.89% | -4.435% | -1.283% | -0.8572% | -3.327% | -0.6366% | 2.89% | -6.138% | 2.879% | 2.89% | -3.753% | 0.4204% | 22.08% | -15.8% | -7.19% | -2.816% | -3.157% | -2.22% | -1.539% | 3.231%
| Z-Score (Absolute) |  | 0.69 | 0.11 | 0.97 | 0.98 | 0.66 | 0.61 | 1.94 | 2.06 | 0.14 | 0.71 | 0.67 | 0.08 | 0 | 0.47 | 0.04 | 0.71 | 1 | 0.7 | 0.71 | 0.55 | 0.24 | 4.32 | 2.82 | 1.19 | 0.37 | 0.43 | 0.26 | 0.13 | 0.77
| Analytical results for Mn in OREAS 231b (Certified Value 0.084 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.0904 | 0.08684 | 0.07416 | 0.089407 | 0.086784 | 0.0918 | 0.0922 | 0.0898 | 0.0872 | 0.0781 | 0.0968 | 0.0809 | 0.0814 | 0.0821 | 0.080281 | 0.086 | 0.0806 | 0.087312 | 0.0847 | 0.0808 | 0.0877 | 0.080111 | 0.083852 | 0.0755 | 0.0792 | 0.0794 | 0.0903 | 0.0775 | 0.0857 |  |  | 1
| 1 | 2 | 0.0896 | 0.08701 | 0.068857 | 0.089795 | 0.086786 | 0.0911 | 0.0929 | 0.0885 | 0.0848 | 0.0784 | 0.0998 | 0.0792 | 0.082 | 0.0808 | 0.082928 | 0.0855 | 0.0794 | 0.08624 | 0.0869 | 0.0822 | 0.0856 | 0.079099 | 0.083365 | 0.077 | 0.079 | 0.0789 | 0.0887 | 0.08 | 0.0856 |  |  | 24
| 1 | 3 | 0.0896 | 0.08805 | 0.070063 | 0.089985 | 0.087897 | 0.0931 | 0.0929 | 0.0893 | 0.085 | 0.0784 | 0.1021 | 0.0796 | 0.0821 | 0.0811 | 0.081011 | 0.0864 | 0.0769 | 0.087843 | 0.0855 | 0.08 | 0.0858 | 0.080084 | 0.082527 | 0.0745 | 0.0784 | 0.0802 | 0.0884 | 0.0785 | 0.0868 |  |  | 16
| 1 | 4 | 0.0887 | 0.08719 | 0.074477 | 0.089951 | 0.085527 | 0.0912 | 0.0929 | 0.09 | 0.0814 | 0.0793 | 0.0991 | 0.0807 | 0.0821 | 0.0806 | 0.086067 | 0.0858 | 0.0769 | 0.087313 | 0.0879 | 0.0794 | 0.0856 | 0.079426 | 0.084285 | 0.077 | 0.0785 | 0.0775 | 0.0923 | 0.0771 | 0.086 |  |  | 0.084178
| 1 | 5 | 0.0882 | 0.08848 | 0.069688 | 0.089126 | 0.093348 | 0.0873 | 0.0945 | 0.0905 | 0.0849 | 0.0782 | 0.0999 | 0.0785 | 0.0816 | 0.0813 | 0.081268 | 0.0849 | 0.0799 | 0.089435 | 0.0879 | 0.0788 | 0.0873 | 0.078711 | 0.08304 | 0.0735 | 0.0791 | 0.0773 | 0.0906 | 0.0772 | 0.0855 |  |  | 103
| 1 | 6 | 0.088 | 0.0872 | 0.075239 | 0.08962 | 0.084609 | 0.0894 | 0.0937 | 0.0897 | 0.0873 | 0.0789 | 0.0996 | 0.08 | 0.0819 | 0.082 | 0.083934 | 0.0857 | 0.0799 | 0.090782 | 0.0864 | 0.0781 | 0.0857 | 0.08011 | 0.083341 | 0.073 | 0.0786 | 0.0787 | 0.0866 | 0.0775 | 0.0877
| Mean |  | 0.089083 | 0.087462 | 0.072081 | 0.089647 | 0.087492 | 0.09065 | 0.093183 | 0.089633 | 0.0851 | 0.07855 | 0.09955 | 0.079817 | 0.08185 | 0.081317 | 0.082581 | 0.085717 | 0.078933 | 0.088154 | 0.08655 | 0.079883 | 0.086283 | 0.07959 | 0.083402 | 0.075083 | 0.0788 | 0.078667 | 0.089483 | 0.077967 | 0.086217
| Median |  | 0.08915 | 0.087195 | 0.072111 | 0.089707 | 0.086785 | 0.09115 | 0.0929 | 0.08975 | 0.08495 | 0.0784 | 0.0997 | 0.0798 | 0.08195 | 0.0812 | 0.082098 | 0.08575 | 0.07965 | 0.087578 | 0.08665 | 0.0797 | 0.08575 | 0.079755 | 0.083353 | 0.075 | 0.0788 | 0.0788 | 0.0895 | 0.0775 | 0.08585
| Std Dev. |  | 0.0009347 | 0.00065052 | 0.0028365 | 0.00033417 | 0.0030866 | 0.0020305 | 0.00080104 | 0.0006802 | 0.0021485 | 0.00045935 | 0.0017003 | 0.00091086 | 0.0002881 | 0.00061779 | 0.0021717 | 0.00050365 | 0.0016207 | 0.0016568 | 0.0012896 | 0.0014703 | 0.00095376 | 0.0006044 | 0.00061387 | 0.0017151 | 0.00034059 | 0.0011112 | 0.0019974 | 0.001113 | 0.00086583
| Rel.Std.Dev. |  | 1.049% | 0.7438% | 3.935% | 0.3728% | 3.528% | 2.24% | 0.8596% | 0.7589% | 2.525% | 0.5848% | 1.708% | 1.141% | 0.352% | 0.7597% | 2.63% | 0.5876% | 2.053% | 1.879% | 1.49% | 1.841% | 1.105% | 0.7594% | 0.736% | 2.284% | 0.4322% | 1.412% | 2.232% | 1.427% | 1.004%
| PDM3 |  | 5.828% | 3.901% | -14.37% | 6.498% | 3.937% | 7.689% | 10.7% | 6.481% | 1.095% | -6.686% | 18.26% | -5.181% | -2.765% | -3.399% | -1.897% | 1.828% | -6.23% | 4.724% | 2.818% | -5.102% | 2.501% | -5.45% | -0.922% | -10.8% | -6.389% | -6.547% | 6.303% | -7.379% | 2.422%
| Z-Score (Absolute) |  | 0.59 | 0.35 | 1.94 | 0.68 | 0.36 | 0.83 | 1.2 | 0.67 | 0 | 0.97 | 2.15 | 0.79 | 0.48 | 0.56 | 0.37 | 0.09 | 0.92 | 0.45 | 0.22 | 0.78 | 0.18 | 0.82 | 0.25 | 1.49 | 0.94 | 0.96 | 0.65 | 1.06 | 0.17
| Analytical results for Mo in OREAS 231b (Certified Value 1.39 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 3
| 1 | 1 | 1.39 | 1.42 | 1.78 | 1.0592 | 1.7656 | 1.47 | 1.49 | 1.5 | 1.9 | < 2 | 1.5 | 1.4 | 1.34 | 1.31 | < 10 | 1.32 | < 2 | < 5 | 1.2 | 1.26 | 1.51 | 1.01 | 1 | 1.3438 | 1.4 | 1.36 | 1.52 | 1.52 | 1.4 | 1.4 |  | 1
| 1 | 2 | 1.43 | 1.45 | 1.68 | 1.1062 | 1.5033 | 1.39 | 1.37 | 1.4 | 1.6 | < 2 | 1.4 | 1.44 | 1.28 | 1.32 | < 10 | 1.28 | < 2 | < 5 | 1.2 | 1.27 | 1.46 | 1.24 | 1 | 1.4467 | 1.38 | 1.296 | 1.48 | 1.47 | 1.43 | 1.4 |  | 25
| 1 | 3 | 1.42 | 1.42 | 1.85 | 1.0539 | 1.5454 | 1.47 | 1.36 | 1.5 | 1.6 | < 2 | 1.4 | 1.4 | 1.4 | 1.45 | < 10 | 1.34 | < 2 | < 5 | 1.2 | 1.2 | 1.48 | 1.78 | 1 | 1.3404 | 1.38 | 1.368 | 1.51 | 1.47 | 1.42 | 1.38 |  | 16
| 1 | 4 | 1.47 | 1.42 | 1.65 | 1.1158 | 1.5453 | 1.41 | 1.34 | 1.5 | 1.6 | < 2 | 1.4 | 1.32 | 1.37 | 1.36 | < 10 | 1.3 | < 2 | < 5 | 1.2 | 1.24 | 1.51 | 1.59 | 1 | 1.3264 | 1.36 | 1.317 | 1.46 | 1.54 | 1.4 | 1.42 |  | 1.3948
| 1 | 5 | 1.41 | 1.38 | 1.56 | 1.0616 | 1.7301 | 1.44 | 1.26 | 1.4 | 1.7 | < 2 | 1.5 | 1.34 | 1.37 | 1.41 | < 10 | 1.3 | < 2 | < 5 | 1.2 | 1.22 | 1.44 | 1.78 | 1 | 1.4038 | 1.41 | 1.291 | 1.53 | 1.58 | 1.41 | 1.37 |  | 104
| 1 | 6 | 1.38 | 1.36 | 1.5 | 1.1137 | 1.5274 | 1.35 | 1.32 | 1.5 | 1.4 | < 2 | 1.4 | 1.38 | 1.35 | 1.35 | < 10 | 1.28 | < 2 | < 5 | 1.2 | 1.21 | 1.49 | 1.31 | 1 | 1.4028 | 1.4 | 1.363 | 1.48 | 1.48 | 1.43 | 1.4
| Mean |  | 1.4167 | 1.4083 | 1.67 | 1.0851 | 1.6029 | 1.4217 | 1.3567 | 1.4667 | 1.6333 |  | 1.4333 | 1.38 | 1.3517 | 1.3667 |  | 1.3033 |  |  | 1.2 | 1.2333 | 1.4817 | 1.4517 | 1 | 1.3773 | 1.3883 | 1.3325 | 1.4967 | 1.51 | 1.415 | 1.395
| Median |  | 1.415 | 1.42 | 1.665 | 1.0839 | 1.5453 | 1.425 | 1.35 | 1.5 | 1.6 |  | 1.4 | 1.39 | 1.36 | 1.355 |  | 1.3 |  |  | 1.2 | 1.23 | 1.485 | 1.45 | 1 | 1.3733 | 1.39 | 1.3385 | 1.495 | 1.5 | 1.415 | 1.4
| Std Dev. |  | 0.032042 | 0.032506 | 0.13115 | 0.029665 | 0.11394 | 0.047504 | 0.07607 | 0.05164 | 0.1633 |  | 0.05164 | 0.043818 | 0.040702 | 0.053914 |  | 0.023381 |  |  | 0 | 0.028048 | 0.027869 | 0.31454 | 0 | 0.047419 | 0.018348 | 0.035331 | 0.027325 | 0.044721 | 0.013784 | 0.017607
| Rel.Std.Dev. |  | 2.262% | 2.308% | 7.853% | 2.734% | 7.109% | 3.341% | 5.607% | 3.521% | 9.998% |  | 3.603% | 3.175% | 3.011% | 3.945% |  | 1.794% |  |  | 0% | 2.274% | 1.881% | 21.67% | 0% | 3.443% | 1.322% | 2.652% | 1.826% | 2.962% | 0.9741% | 1.262%
| PDM3 |  | 1.569% | 0.9718% | 19.73% | -22.21% | 14.92% | 1.928% | -2.732% | 5.154% | 17.1% |  | 2.764% | -1.06% | -3.091% | -2.015% |  | -6.556% |  |  | -13.96% | -11.57% | 6.23% | 4.079% | -28.3% | -1.251% | -0.4621% | -4.465% | 7.305% | 8.261% | 1.45% | 0.01589%
| Z-Score (Absolute) |  | 0.08 | 0 | 2.41 | 2.97 | 1.79 | 0.12 | 0.48 | 0.54 | 2.07 | 3.75 | 0.23 | 0.26 | 0.52 | 0.38 | 33.03 | 0.97 | 3.75 | 10.04 | 1.92 | 1.61 | 0.67 | 0.4 | NA | 0.29 | 0.18 | 0.7 | 0.81 | 0.93 | 0.06 | 0.12
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 231b (Certified Value 0.147 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.16 | 0.148 | 0.9881 | 0.14101 | 0.147 | 0.138 | 0.148 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14 | 0.14 | 0.16 | 0.16888 | 0.15 | 0.14 | 0.13475 | 0.15 | 0.172 | 0.16 | 0.17 | 0.16046 | 0.1405 | 0.14 | 0.13 | 0.2 | 0.15 | 0.15 |  |  | 1
| 1 | 2 | 0.15 | 0.151 | 1.21 | 0.14698 | 0.14668 | 0.136 | 0.148 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14 | 0.16 | 0.17586 | 0.14 | 0.14 | 0.13304 | 0.15 | 0.178 | 0.15 | 0.17 | 0.15414 | 0.139 | 0.14 | 0.13 | 0.19 | 0.16 | 0.15 |  |  | #N/A
| 1 | 3 | 0.15 | 0.148 | 1.0143 | 0.14773 | 0.14648 | 0.135 | 0.148 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.14 | 0.16 | 0.16726 | 0.15 | 0.13 | 0.13895 | 0.15 | 0.172 | 0.15 | 0.18 | 0.15556 | 0.136 | 0.14 | 0.13 | 0.19 | 0.15 | 0.15 |  |  | 16
| 1 | 4 | 0.15 | 0.146 | 1.1088 | 0.14045 | 0.14608 | 0.134 | 0.148 | 0.13 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.16 | 0.18385 | 0.15 | 0.13 | 0.13585 | 0.15 | 0.157 | 0.15 | 0.17 | 0.15393 | 0.139 | 0.13 | 0.13 | 0.19 | 0.15 | 0.15 |  |  | 0.14683
| 1 | 5 | 0.15 | 0.148 | 0.89393 | 0.14844 | 0.15735 | 0.132 | 0.148 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.16 | 0.16899 | 0.15 | 0.14 | 0.13844 | 0.15 | 0.155 | 0.16 | 0.16 | 0.15794 | 0.1355 | 0.14 | 0.13 | 0.2 | 0.15 | 0.15 |  |  | 105
| 1 | 6 | 0.15 | 0.15 | 1.1126 | 0.14368 | 0.14111 | 0.132 | 0.148 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.16 | 0.1771 | 0.15 | 0.14 | 0.14465 | 0.15 | 0.162 | 0.15 | 0.16 | 0.16856 | 0.1375 | 0.14 | 0.13 | 0.2 | 0.15 | 0.15
| Mean |  | 0.15167 | 0.1485 | 1.0546 | 0.14472 | 0.14745 | 0.1345 | 0.148 | 0.13167 | 0.13667 | 0.14333 | 0.14333 | 0.14 | 0.14 | 0.16 | 0.17366 | 0.14833 | 0.13667 | 0.13761 | 0.15 | 0.166 | 0.15333 | 0.16833 | 0.15843 | 0.13792 | 0.13833 | 0.13 | 0.195 | 0.15167 | 0.15
| Median |  | 0.15 | 0.148 | 1.0615 | 0.14533 | 0.14658 | 0.1345 | 0.148 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.16 | 0.17243 | 0.15 | 0.14 | 0.13715 | 0.15 | 0.167 | 0.15 | 0.17 | 0.15675 | 0.13825 | 0.14 | 0.13 | 0.195 | 0.15 | 0.15
| Std Dev. |  | 0.0040825 | 0.0017607 | 0.11164 | 0.0034953 | 0.0053268 | 0.0023452 | 0 | 0.0040825 | 0.005164 | 0.005164 | 0.005164 | 0 | 0 | 0 | 0.0064195 | 0.0040825 | 0.005164 | 0.0041034 | 0 | 0.0093167 | 0.005164 | 0.0075277 | 0.0055472 | 0.0019343 | 0.0040825 | 0 | 0.0054772 | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.692% | 1.186% | 10.59% | 2.415% | 3.613% | 1.744% | 0% | 3.101% | 3.779% | 3.603% | 3.603% | 0% | 0% | 0% | 3.697% | 2.752% | 3.779% | 2.982% | 0% | 5.612% | 3.368% | 4.472% | 3.501% | 1.403% | 2.951% | 0% | 2.809% | 2.692% | 0%
| PDM3 |  | 3.293% | 1.136% | 618.3% | -1.441% | 0.4218% | -8.398% | 0.796% | -10.33% | -6.923% | -2.382% | -2.382% | -4.652% | -4.652% | 8.969% | 18.27% | 1.023% | -6.923% | -6.278% | 2.158% | 13.05% | 4.428% | 14.64% | 7.902% | -6.071% | -5.788% | -11.46% | 32.81% | 3.293% | 2.158%
| Z-Score (Absolute) |  | 0.22 | 0.01 | 61.11 | 0.24 | 0.06 | 0.93 | 0 | 1.12 | 0.79 | 0.34 | 0.34 | 0.56 | 0.56 | 0.79 | 1.71 | 0 | 0.79 | 0.72 | 0.11 | 1.19 | 0.34 | 1.35 | 0.68 | 0.7 | 0.67 | 1.24 | 3.15 | 0.22 | 0.11
| Analytical results for Nb in OREAS 231b (Indicative Value 0.11 ppm) | [32 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.12 | 0.12 | 34.63 | 0.13664 | 0.1 | 0.4 | < 0.2 | 0.11 | < 5 | < 0.1 | 0.09 | 0.15 | 0.08 | 0.07 | 10 | < 0.2 | < 1 | 0.1 | 1.52 | 0.19 | 0.08 | 0.12 | [9 empty] | 1
| 1 | 2 | 0.06 | 0.13 | 35.95 | 0.14401 | 0.1 | 0.4 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.08 | 0.17 | 0.08 | 0.06 | 10 | < 0.2 | < 1 | 0.1 | 1.16 | 0.19 | 0.07 | 0.13 | [9 empty] | 8
| 1 | 3 | 0.05 | 0.12 | 35.6 | 0.12 | < 0.1 | 0.45 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.08 | 0.14 | 0.08 | 0.07 | 10 | < 0.2 | < 1 | 0.09 | 0.97 | 0.18 | 0.08 | 0.11 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.12 | 36.32 | 0.14545 | 0.1 | 0.4 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.16 | 0.07 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.8 | 0.2 | 0.07 | 0.1 | [9 empty] | 0.10883
| 1 | 5 | < 0.05 | 0.13 | 34.22 | 0.14757 | < 0.1 | 0.4 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.08 | 0.17 | 0.08 | 0.07 | < 10 | < 0.2 | < 1 | 0.09 | 0.71 | 0.21 | 0.07 | 0.12 | [9 empty] | 14
| 1 | 6 | < 0.05 | 0.13 | 35.79 | 0.13195 | < 0.1 | 0.41 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.08 | 0.16 | 0.08 | 0.07 | < 10 | < 0.2 | < 1 | 0.1 | 0.66 | 0.19 | 0.07 | 0.13
| Mean |  | 0.076667 | 0.125 | 35.4183 | 0.1376 | 0.1 | 0.41 |  | 0.095 |  |  | 0.083333 | 0.15833 | 0.078333 | 0.068333 | 10 |  |  | 0.098333 | 0.97 | 0.19333 | 0.073333 | 0.11833
| Median |  | 0.06 | 0.125 | 35.695 | 0.14032 | 0.1 | 0.4 |  | 0.09 |  |  | 0.08 | 0.16 | 0.08 | 0.07 | 10 |  |  | 0.1 | 0.885 | 0.19 | 0.07 | 0.12
| Std Dev. |  | 0.037859 | 0.0054772 | 0.81539 | 0.010437 | 1.6997e-17 | 0.02 |  | 0.0083666 |  |  | 0.005164 | 0.01169 | 0.0040825 | 0.0040825 | 0 |  |  | 0.0075277 | 0.32594 | 0.010328 | 0.005164 | 0.01169
| Rel.Std.Dev. |  | 49.38% | 4.382% | 2.302% | 7.585% | 1.7e-14% | 4.878% |  | 8.807% |  |  | 6.197% | 7.383% | 5.212% | 5.974% | 0% |  |  | 7.655% | 33.6% | 5.342% | 7.042% | 9.879%
| PDM3 |  | -29.55% | 14.86% | 32445% | 26.44% | -8.112% | 276.7% |  | -12.71% |  |  | -23.43% | 45.49% | -28.02% | -37.21% | 9089% |  |  | -9.643% | 791.3% | 77.65% | -32.62% | 8.734%
| Z-Score (Absolute) |  | 1.03 | 0.28 | 621.11 | 0.5 | NA | 5.29 | 0.16 | 0.25 | 42.06 | 1.04 | 0.45 | 0.86 | 0.54 | 0.72 | NA | 0.16 | 6.88 | 0.19 | 15.14 | 1.48 | 0.63 | 0.16
| Analytical results for Nd in OREAS 231b (Certified Value 5.31 ppm) | [32 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 5.659 | 5.198 | 4.8236 | 5.73 | 5.86 | 5 | [25 empty] | 1
| 1 | 2 | 5.772 | 5.102 | 4.8232 | 5.57 | 5.5 | 5 | [25 empty] | 27
| 1 | 3 | 5.71 | 4.959 | 4.8179 | 5.71 | 5.53 | 5 | [25 empty] | 16
| 1 | 4 | 5.664 | 5.085 | 4.8671 | 5.59 | 5.66 | 5 | [25 empty] | 5.3144
| 1 | 5 | 5.693 | 5.346 | 4.8332 | 5.52 | 5.65 | 4.5 | [25 empty] | 106
| 1 | 6 | 5.621 | 4.952 | 4.7516 | 5.68 | 5.64 | 5
| Mean |  | 5.6865 | 5.107 | 4.8194 | 5.6333 | 5.64 | 4.9167
| Median |  | 5.6785 | 5.0935 | 4.8234 | 5.635 | 5.645 | 5
| Std Dev. |  | 0.05188 | 0.14956 | 0.037649 | 0.08501 | 0.12696 | 0.20412
| Rel.Std.Dev. |  | 0.9123% | 2.929% | 0.7812% | 1.509% | 2.251% | 4.152%
| PDM3 |  | 7.002% | -3.902% | -9.313% | 6.002% | 6.127% | -7.484%
| Z-Score (Absolute) |  | 0.73 | 0.61 | 1.27 | 0.61 | 0.62 | 1.04
| Analytical results for Ni in OREAS 231b (Certified Value 73 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 1
| 1 | 1 | 74 | 72.6 | 62.33 | 70.75 | 77.4412 | 73.1 | 74.2 | 79 | 76.4 | 66 | 67.7 | 75.2 | 70.3 | 71 | 68.911 | 72.1 | 65 | 73.9471 | 72 | 72.2 | 73 | 69.97 | 71.4564 | 69 | 82.1364 | 73 | 69.2 | 79.1 | 68 | 74.9 |  | 1
| 1 | 2 | 72.9 | 73.9 | 58.68 | 71.03 | 77.3309 | 74.2 | 75 | 78 | 73.4 | 65 | 67.5 | 75.1 | 72.2 | 70.7 | 70.458 | 71.5 | 64 | 74.7782 | 73 | 75.4 | 71 | 69.27 | 67.8408 | 70 | 76.3698 | 72.4 | 68.9 | 77 | 68.9 | 73.6 |  | 28
| 1 | 3 | 73 | 73.5 | 59.73 | 70.93 | 78.0105 | 72.7 | 75.9 | 80 | 74.6 | 65 | 68 | 75.8 | 71.9 | 76.3 | 70.245 | 73.5 | 62 | 74.1499 | 73 | 72.1 | 71 | 71.04 | 67.5832 | 68 | 79.8 | 72.5 | 69.7 | 77 | 69.6 | 73.7 |  | 16
| 1 | 4 | 73.2 | 73.1 | 63.01 | 71.88 | 75.7581 | 71 | 77.4 | 78 | 72.6 | 65 | 68.5 | 73.3 | 75.3 | 75.3 | 74.244 | 72.5 | 62 | 74.6291 | 72 | 74.6 | 71 | 70 | 67.988 | 70 | 79.1298 | 73.1 | 68.2 | 79.5 | 66.5 | 73.4 |  | 72.8669
| 1 | 5 | 72.8 | 72.8 | 59.91 | 71.39 | 83.2154 | 72.5 | 72.6 | 80 | 73 | 64 | 69.1 | 73.2 | 74.2 | 75.4 | 69.587 | 72.6 | 64 | 76.045 | 72 | 70.9 | 72 | 69.12 | 66.5252 | 66 | 79.6495 | 73.8 | 69.5 | 78 | 67.7 | 73.9 |  | 107
| 1 | 6 | 72.1 | 72.7 | 63.57 | 71.69 | 75.2671 | 73.2 | 77 | 81 | 73.9 | 64 | 67.2 | 73.9 | 73.6 | 74.6 | 72.499 | 73.8 | 64 | 76.7785 | 74 | 71.1 | 71 | 70.09 | 67.0772 | 66 | 78.6288 | 73 | 69.5 | 75.6 | 69.2 | 75
| Mean |  | 73 | 73.1 | 61.205 | 71.2783 | 77.8372 | 72.7833 | 75.35 | 79.3333 | 73.9833 | 64.8333 | 68 | 74.4167 | 72.9167 | 73.8833 | 70.9907 | 72.6667 | 63.5 | 75.0546 | 72.6667 | 72.7167 | 71.5 | 69.915 | 68.0785 | 68.1667 | 79.2857 | 72.9667 | 69.1667 | 77.7 | 68.3167 | 74.0833
| Median |  | 72.95 | 72.95 | 61.12 | 71.21 | 77.3861 | 72.9 | 75.45 | 79.5 | 73.65 | 65 | 67.85 | 74.5 | 72.9 | 74.95 | 70.3515 | 72.55 | 64 | 74.7036 | 72.5 | 72.15 | 71 | 69.985 | 67.712 | 68.5 | 79.3897 | 73 | 69.35 | 77.5 | 68.45 | 73.8
| Std Dev. |  | 0.61644 | 0.5099 | 2.0172 | 0.44857 | 2.8389 | 1.0534 | 1.8019 | 1.2111 | 1.3747 | 0.75277 | 0.69857 | 1.0944 | 1.7994 | 2.4128 | 1.9997 | 0.85946 | 1.2247 | 1.1185 | 0.8165 | 1.8606 | 0.83666 | 0.68623 | 1.7399 | 1.8348 | 1.8699 | 0.50067 | 0.55015 | 1.4615 | 1.1444 | 0.69113
| Rel.Std.Dev. |  | 0.8444% | 0.6975% | 3.296% | 0.6293% | 3.647% | 1.447% | 2.391% | 1.527% | 1.858% | 1.161% | 1.027% | 1.471% | 2.468% | 3.266% | 2.817% | 1.183% | 1.929% | 1.49% | 1.124% | 2.559% | 1.17% | 0.9815% | 2.556% | 2.692% | 2.358% | 0.6862% | 0.7954% | 1.881% | 1.675% | 0.9329%
| PDM3 |  | 0.1827% | 0.3199% | -16% | -2.18% | 6.821% | -0.1146% | 3.408% | 8.874% | 1.532% | -11.02% | -6.679% | 2.127% | 0.06834% | 1.395% | -2.575% | -0.2748% | -12.85% | 3.002% | -0.2748% | -0.2061% | -1.876% | -4.051% | -6.571% | -6.45% | 8.809% | 0.137% | -5.078% | 6.633% | -6.245% | 1.669%
| Z-Score (Absolute) |  | 0.08 | 0.12 | 3.83 | 0.49 | 1.69 | 0.01 | 0.86 | 2.18 | 0.41 | 2.63 | 1.58 | 0.55 | 0.06 | 0.38 | 0.58 | 0.03 | 3.07 | 0.76 | 0.03 | 0.01 | 0.41 | 0.94 | 1.55 | 1.52 | 2.17 | 0.07 | 1.19 | 1.64 | 1.47 | 0.44
| Analytical results for P in OREAS 231b (Certified Value 0.043 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.0449 | 0.0429 | 0.036246 | 0.04404 | 0.04 | 0.0401 | 0.0444 | 0.0407 | 0.0444 | 0.05 | 0.044 | 0.044 | 0.041 | 0.04462 | 0.043 | 0.0405 | 0.045959 | 0.043 | 0.042 | 0.05 | 0.043 | 0.042 | 0.041 | 0.0447 | 0.0408 | 0.042 | [5 empty] | 1
| 1 | 2 | 0.0448 | 0.0448 | 0.031854 | 0.044229 | 0.04 | 0.0406 | 0.0434 | 0.0435 | 0.0442 | 0.047 | 0.043 | 0.044 | 0.041 | 0.046639 | 0.044 | 0.0389 | 0.046894 | 0.045 | 0.042 | 0.05 | 0.0445 | 0.042 | 0.041 | 0.0447 | 0.0423 | 0.042 | [5 empty] | #N/A
| 1 | 3 | 0.0446 | 0.0441 | 0.032198 | 0.044458 | 0.04 | 0.0406 | 0.0438 | 0.0428 | 0.0435 | 0.048 | 0.043 | 0.044 | 0.041 | 0.043971 | 0.044 | 0.0382 | 0.045068 | 0.043 | 0.041 | 0.05 | 0.0425 | 0.042 | 0.042 | 0.0444 | 0.0412 | 0.043 | [5 empty] | 16
| 1 | 4 | 0.0447 | 0.0421 | 0.03579 | 0.043678 | 0.04 | 0.0401 | 0.044 | 0.0414 | 0.043 | 0.046 | 0.044 | 0.044 | 0.041 | 0.047769 | 0.044 | 0.038 | 0.046595 | 0.045 | 0.042 | 0.04 | 0.0435 | 0.041 | 0.041 | 0.0444 | 0.0407 | 0.043 | [5 empty] | 0.043109
| 1 | 5 | 0.0446 | 0.045 | 0.031879 | 0.047223 | 0.04 | 0.041 | 0.0448 | 0.0403 | 0.0441 | 0.049 | 0.043 | 0.043 | 0.041 | 0.043114 | 0.043 | 0.0393 | 0.047443 | 0.042 | 0.041 | 0.04 | 0.0425 | 0.042 | 0.042 | 0.0448 | 0.0408 | 0.042 | [5 empty] | 108
| 1 | 6 | 0.0443 | 0.0424 | 0.03628 | 0.042679 | 0.04 | 0.0415 | 0.0446 | 0.0439 | 0.0431 | 0.047 | 0.044 | 0.044 | 0.041 | 0.046298 | 0.044 | 0.0398 | 0.047059 | 0.042 | 0.041 | 0.04 | 0.0435 | 0.042 | 0.041 | 0.0437 | 0.0408 | 0.044
| Mean |  | 0.04465 | 0.04355 | 0.034041 | 0.044384 | 0.04 | 0.04065 | 0.044167 | 0.0421 | 0.043717 | 0.047833 | 0.0435 | 0.043833 | 0.041 | 0.045402 | 0.043667 | 0.039117 | 0.046503 | 0.043333 | 0.0415 | 0.045 | 0.04325 | 0.041833 | 0.041333 | 0.04445 | 0.0411 | 0.042667
| Median |  | 0.04465 | 0.0435 | 0.033994 | 0.044135 | 0.04 | 0.0406 | 0.0442 | 0.0421 | 0.0438 | 0.0475 | 0.0435 | 0.044 | 0.041 | 0.045459 | 0.044 | 0.0391 | 0.046744 | 0.043 | 0.0415 | 0.045 | 0.04325 | 0.042 | 0.041 | 0.04455 | 0.0408 | 0.0425
| Std Dev. |  | 0.00020736 | 0.0012502 | 0.002271 | 0.001524 | 0 | 0.00053944 | 0.00052789 | 0.0015086 | 0.00059805 | 0.001472 | 0.00054772 | 0.00040825 | 0 | 0.0017794 | 0.0005164 | 0.00095376 | 0.00086128 | 0.0013663 | 0.00054772 | 0.0054772 | 0.00075829 | 0.00040825 | 0.0005164 | 0.00040373 | 0.00061319 | 0.0008165
| Rel.Std.Dev. |  | 0.4644% | 2.871% | 6.671% | 3.434% | 0% | 1.327% | 1.195% | 3.583% | 1.368% | 3.077% | 1.259% | 0.9314% | 0% | 3.919% | 1.183% | 2.438% | 1.852% | 3.153% | 1.32% | 12.17% | 1.753% | 0.9759% | 1.249% | 0.9083% | 1.492% | 1.914%
| PDM3 |  | 3.573% | 1.022% | -21.04% | 2.957% | -7.213% | -5.705% | 2.452% | -2.342% | 1.408% | 10.96% | 0.9059% | 1.679% | -4.893% | 5.317% | 1.292% | -9.262% | 7.872% | 0.5193% | -3.733% | 4.385% | 0.3259% | -2.96% | -4.12% | 3.11% | -4.661% | -1.027%
| Z-Score (Absolute) |  | 0.57 | 0.06 | 4.36 | 0.45 | 1.59 | 1.28 | 0.35 | 0.61 | 0.14 | 2.05 | 0.04 | 0.19 | 1.12 | 0.92 | 0.12 | 2 | 1.44 | 0.04 | 0.89 | 0.74 | 0.08 | 0.74 | 0.97 | 0.48 | 1.08 | 0.35
| Analytical results for Pb in OREAS 231b (Certified Value 13.3 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 2
| 1 | 1 | 13.9 | 13.2 | 15.73 | 12.84 | 13.5477 | 9.8 | 13.8 | 13.1 | 13.7 | 14 | 14.6 | 13.4 | 13.3 | 12.6 | 12.8 | 13 | 12.0664 | 13 | 12.22 | 14 | 15.54 | 14.5638 | 15 | 13.5766 | 13.1 | 18.1948 | 13 | 14.2 | 13.5 | 13.3 |  | 1
| 1 | 2 | 13.7 | 14 | 18.56 | 12.25 | 14.2476 | 8.6 | 14 | 12.8 | 13.8 | 13 | 14.4 | 13.4 | 13.2 | 12.6 | 12.5 | 10 | 10.2408 | 12.8 | 12.66 | 13.8 | 15.63 | 13.9035 | 15 | 13.2707 | 13 | 18.443 | 12.8 | 13.2 | 13.5 | 13.8 |  | 29
| 1 | 3 | 13.7 | 13.7 | 18.91 | 12.71 | 13.6279 | 8.5 | 13.6 | 12.9 | 14 | 13 | 14.3 | 13.3 | 13.4 | 13.6 | 13.2 | 10 | 10.0761 | 13.2 | 12.05 | 14.2 | 23.28 | 12.6201 | 15 | 13.6589 | 12.8 | 17.132 | 13.1 | 13.5 | 13.4 | 13.9 |  | 16
| 1 | 4 | 13.8 | 13.2 | 14.36 | 12.71 | 13.2976 | 8.6 | 13.5 | 12.7 | 13.7 | 13 | 14.4 | 13.3 | 13.4 | 13.1 | 12.7 | 13 | 10.6675 | 13.1 | 12.22 | 14.2 | 16.86 | 12.2202 | 15 | 13.4105 | 13 | 17.09 | 12.9 | 13.9 | 13 | 13.8 |  | 13.3381
| 1 | 5 | 13.7 | 13.5 | 13.9 | 12.08 | 15.2004 | 8 | 14.1 | 12.9 | 13.8 | 13 | 14.2 | 12.8 | 13.2 | 13.5 | 12.6 | 16 | 11.1179 | 13 | 12.12 | 13.4 | 16.88 | 12.8991 | 15 | 13.6326 | 13 | 16.306 | 13 | 13.6 | 13.1 | 13.6 |  | 109
| 1 | 6 | 13.9 | 13.4 | 12.72 | 12.48 | 13.5306 | 8.7 | 14.1 | 12.7 | 14 | 13 | 14.1 | 12.7 | 13.2 | 13.2 | 12.9 | 14 | 10.6518 | 13.4 | 11.79 | 13.5 | 22.89 | 14.3127 | 15 | 13.9311 | 13 | 23.068 | 13.1 | 13.9 | 13.2 | 14
| Mean |  | 13.7833 | 13.5 | 15.6967 | 12.5117 | 13.9086 | 8.7 | 13.85 | 12.85 | 13.8333 | 13.1667 | 14.3333 | 13.15 | 13.2833 | 13.1 | 12.7833 | 12.6667 | 10.8034 | 13.0833 | 12.1767 | 13.85 | 18.5133 | 13.4199 | 15 | 13.5801 | 12.9833 | 18.3723 | 12.9833 | 13.7167 | 13.2833 | 13.7333
| Median |  | 13.75 | 13.45 | 15.045 | 12.595 | 13.5878 | 8.6 | 13.9 | 12.85 | 13.8 | 13 | 14.35 | 13.3 | 13.25 | 13.15 | 12.75 | 13 | 10.6596 | 13.05 | 12.17 | 13.9 | 16.87 | 13.4013 | 15 | 13.6046 | 13 | 17.6634 | 13 | 13.75 | 13.3 | 13.8
| Std Dev. |  | 0.098319 | 0.30984 | 2.5456 | 0.29742 | 0.70836 | 0.5933 | 0.25884 | 0.15166 | 0.13663 | 0.40825 | 0.17512 | 0.31464 | 0.098319 | 0.42895 | 0.24833 | 2.3381 | 0.71846 | 0.20412 | 0.285 | 0.34496 | 3.5898 | 0.96847 | 0 | 0.22656 | 0.098319 | 2.4306 | 0.1169 | 0.35449 | 0.2137 | 0.25033
| Rel.Std.Dev. |  | 0.7133% | 2.295% | 16.22% | 2.377% | 5.093% | 6.819% | 1.869% | 1.18% | 0.9877% | 3.101% | 1.222% | 2.393% | 0.7402% | 3.274% | 1.943% | 18.46% | 6.65% | 1.56% | 2.341% | 2.491% | 19.39% | 7.217% | 0% | 1.668% | 0.7573% | 13.23% | 0.9004% | 2.584% | 1.609% | 1.823%
| PDM3 |  | 3.338% | 1.214% | 17.68% | -6.196% | 4.277% | -34.77% | 3.838% | -3.66% | 3.713% | -1.285% | 7.461% | -1.41% | -0.4107% | -1.785% | -4.159% | -5.034% | -19% | -1.91% | -8.708% | 3.838% | 38.8% | 0.6131% | 12.46% | 1.814% | -2.66% | 37.74% | -2.66% | 2.838% | -0.4107% | 2.963%
| Z-Score (Absolute) |  | 0.57 | 0.13 | 3.57 | 1.42 | 0.77 | 7.4 | 0.67 | 0.89 | 0.65 | NA | 1.43 | 0.42 | 0.21 | 0.5 | 1 | NA | 4.1 | 0.53 | 1.95 | 0.67 | 7.99 | 0 | NA | 0.25 | 0.68 | 7.76 | 0.68 | 0.47 | 0.21 | 0.49
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 231b (Indicative Value 16.4 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 1
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 1
| 1 | 1 | 15 | 22 | 12 | < 50 | 24.21 | 43 | [25 empty] | 1
| 1 | 2 | 13 | 18 | 12 | < 50 | 20.55 | 41 | [25 empty] | 1
| 1 | 3 | 15 | 13 | 12 | < 50 | < 20 | 38 | [25 empty] | 16
| 1 | 4 | 16 | 19 | 12 | < 50 | 28.81 | 27 | [25 empty] | 16.4333
| 1 | 5 | 16 | 20 | 12 | < 50 | < 20 | 38 | [25 empty] | 12
| 1 | 6 | 16 | 18 | 13 | < 50 | 26.83 | 50
| Mean |  | 15.1667 | 18.3333 | 12.1667 |  | 25.1 | 39.5
| Median |  | 15.5 | 18.5 | 12 |  | 25.52 | 39.5
| Std Dev. |  | 1.169 | 3.0111 | 0.40825 |  | 3.5708 | 7.5565
| Rel.Std.Dev. |  | 7.708% | 16.42% | 3.355% |  | 14.23% | 19.13%
| PDM3 |  | -7.708% | 11.56% | -25.96% |  | 52.74% | 140.4%
| Z-Score (Absolute) |  | 0.55 | 0.12 | 0.96 | 0.8 | 0.12 | 2.78
| Analytical results for Pr in OREAS 231b (Certified Value 1.07 ppm) | [32 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.115 | 1.022 | 1.0151 | 1.2 | 1.17 | 1.1 | [25 empty] | 1
| 1 | 2 | 1.139 | 0.999 | 0.99807 | 1.1 | 1.09 | 1.1 | [25 empty] | 30
| 1 | 3 | 1.148 | 0.949 | 1.026 | 1.2 | 1.12 | 1.1 | [25 empty] | 16
| 1 | 4 | 1.113 | 0.983 | 1.0126 | 1.2 | 1.12 | 1 | [25 empty] | 1.0724
| 1 | 5 | 1.125 | 1.005 | 1.0097 | 1.1 | 1.1 | 1 | [25 empty] | 110
| 1 | 6 | 1.125 | 0.926 | 0.99557 | 1.1 | 1.1 | 1
| Mean |  | 1.1275 | 0.98067 | 1.0095 | 1.15 | 1.1167 | 1.05
| Median |  | 1.125 | 0.991 | 1.0112 | 1.15 | 1.11 | 1.05
| Std Dev. |  | 0.013649 | 0.036423 | 0.011292 | 0.054772 | 0.028752 | 0.054772
| Rel.Std.Dev. |  | 1.211% | 3.714% | 1.119% | 4.763% | 2.575% | 5.216%
| PDM3 |  | 5.139% | -8.553% | -5.863% | 7.237% | 4.129% | -2.088%
| Z-Score (Absolute) |  | 0.54 | 1.25 | 0.9 | 0.81 | 0.41 | 0.41
| Analytical results for Pt in OREAS 231b (Indicative Value 8.77 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 2
| 1 | 1 | 8 | 11 | 6 | < 10 | 11.43 | < 10 | [25 empty] | 1
| 1 | 2 | 9 | 11 | 5 | < 10 | 9.56 | < 10 | [25 empty] | 1
| 1 | 3 | 8 | 7 | 6 | < 10 | 14.21 | < 10 | [25 empty] | 16
| 1 | 4 | 8 | 9 | 5 | < 10 | 9.23 | < 10 | [25 empty] | 8.7733
| 1 | 5 | 10 | 9 | 6 | < 10 | 11.64 | < 10 | [25 empty] | 13
| 1 | 6 | 8 | 12 | 7 | < 10 | 9.49 | < 10
| Mean |  | 8.5 | 9.8333 | 5.8333 |  | 10.9267
| Median |  | 8 | 10 | 6 |  | 10.495
| Std Dev. |  | 0.83666 | 1.8348 | 0.75277 |  | 1.9159
| Rel.Std.Dev. |  | 9.843% | 18.66% | 12.9% |  | 17.53%
| PDM3 |  | -3.116% | 12.08% | -33.51% |  | 24.54%
| Z-Score (Absolute) |  | 0.41 | 0.83 | 0.41 | 0.67 | 1.17 | 0.67
| Analytical results for Rb in OREAS 231b (Certified Value 3.99 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 2
| 1 | 1 | 3.9 | 3.92 | 3.7404 | 4.0636 | 3.2 | 4.68 | 4.95 | 4.42 | 4.3 | 3.7 | 3.9 | 3.4 | 3.7 | 3.8 | 4.3 | 4 | 4.1 | 4 | 4.1 | 4.4 | 3.8 | [10 empty] | 1
| 1 | 2 | 3.8 | 4.11 | 3.7758 | 4.0162 | 3.1 | 4.73 | 4.83 | 4.17 | 4.3 | 3.8 | 4 | 3.4 | 3.6 | 3.71 | 4.2 | 4 | 4.1 | 3.9 | 3.9 | 4.3 | 3.7 | [10 empty] | 31
| 1 | 3 | 3.8 | 4.07 | 3.7649 | 4.0068 | 3.2 | 4.75 | 4.97 | 4.22 | 4.3 | 3.7 | 4.1 | 3.7 | 3.9 | 3.92 | 4.3 | 3.5 | 4.1 | 3.9 | 4 | 4.3 | 3.7 | [10 empty] | 16
| 1 | 4 | 3.7 | 3.99 | 3.8442 | 4.0796 | 3.1 | 4.8 | 4.91 | 4.04 | 4.4 | 3.6 | 4.2 | 3.6 | 3.8 | 3.79 | 4.3 | 3.5 | 4.2 | 4 | 4.1 | 4.2 | 3.7 | [10 empty] | 3.9903
| 1 | 5 | 3.7 | 4.03 | 3.8269 | 4.3569 | 3.1 | 4.89 | 4.87 | 4.21 | 4.5 | 3.6 | 4.1 | 3.7 | 3.7 | 3.87 | 4.3 | 3.5 | 4.1 | 4 | 4.1 | 4.3 | 3.7 | [10 empty] | 111
| 1 | 6 | 3.5 | 3.98 | 3.7059 | 3.9654 | 3.1 | 4.82 | 5.34 | 4.25 | 4.3 | 3.5 | 4.1 | 3.6 | 3.8 | 3.75 | 4.2 | 3.5 | 4.1 | 3.7 | 3.9 | 4.3 | 3.8
| Mean |  | 3.7333 | 4.0167 | 3.7763 | 4.0814 | 3.1333 | 4.7783 | 4.9783 | 4.2183 | 4.35 | 3.65 | 4.0667 | 3.5667 | 3.75 | 3.8067 | 4.2667 | 3.6667 | 4.1167 | 3.9167 | 4.0167 | 4.3 | 3.7333
| Median |  | 3.75 | 4.01 | 3.7703 | 4.0399 | 3.1 | 4.775 | 4.93 | 4.215 | 4.3 | 3.65 | 4.1 | 3.6 | 3.75 | 3.795 | 4.3 | 3.5 | 4.1 | 3.95 | 4.05 | 4.3 | 3.7
| Std Dev. |  | 0.13663 | 0.06802 | 0.052029 | 0.14107 | 0.05164 | 0.07414 | 0.18444 | 0.12319 | 0.083666 | 0.10488 | 0.10328 | 0.13663 | 0.10488 | 0.077115 | 0.05164 | 0.2582 | 0.040825 | 0.1169 | 0.098319 | 0.063246 | 0.05164
| Rel.Std.Dev. |  | 3.66% | 1.693% | 1.378% | 3.456% | 1.648% | 1.552% | 3.705% | 2.92% | 1.923% | 2.873% | 2.54% | 3.831% | 2.797% | 2.026% | 1.21% | 7.042% | 0.9917% | 2.985% | 2.448% | 1.471% | 1.383%
| PDM3 |  | -6.439% | 0.6619% | -5.361% | 2.285% | -21.48% | 19.75% | 24.76% | 5.716% | 9.016% | -8.527% | 1.915% | -10.62% | -6.021% | -4.601% | 6.927% | -8.109% | 3.168% | -1.844% | 0.6619% | 7.763% | -6.439%
| Z-Score (Absolute) |  | 0.72 | 0 | 0.61 | 0.16 | 2.23 | 1.93 | 2.43 | 0.51 | 0.84 | 0.93 | 0.13 | 1.14 | 0.67 | 0.53 | 0.63 | 0.89 | 0.25 | 0.25 | 0 | 0.72 | 0.72
| Analytical results for Re in OREAS 231b (Certified Value 0.002 ppm) | [32 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.001 | 0.002 | 0.002 | 0.001 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.004 | < 0.002 | 0.002 | 0.005 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.002 | < 0.002 | 0.003 | 0.006 | [13 empty] | 32
| 1 | 3 | < 0.1 | 0.002 | 0.002 | < 0.05 | < 0.001 | < 0.005 | 0.001 | 0.002 | 0.003 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.004 | < 0.002 | 0.002 | 0.005 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.006 | < 0.002 | 0.002 | 0.006 | [13 empty] | 0.0019444
| 1 | 5 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.001 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.005 | < 0.002 | 0.003 | 0.005 | [13 empty] | 112
| 1 | 6 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.001 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.005 | < 0.002 | 0.003 | 0.005
| Mean |  |  | 0.002 | 0.002 |  | 0.002 |  | 0.0016667 | 0.002 | 0.002 | 0.0016667 |  |  |  | 0.0016667 | 0.0043333 |  | 0.0025 | 0.0053333
| Median |  |  | 0.002 | 0.002 |  | 0.002 |  | 0.002 | 0.002 | 0.002 | 0.0015 |  |  |  | 0.002 | 0.0045 |  | 0.0025 | 0.005
| Std Dev. |  |  | 0 | 0 |  | 0 |  | 0.0005164 | 0 | 0.00063246 | 0.0008165 |  |  |  | 0.0005164 | 0.0013663 |  | 0.00054772 | 0.0005164
| Rel.Std.Dev. |  |  | 0% | 0% |  | 0% |  | 30.98% | 0% | 31.62% | 48.99% |  |  |  | 30.98% | 31.53% |  | 21.91% | 9.682%
| PDM3 |  |  | 2.857% | 2.857% |  | 2.857% |  | -14.29% | 2.857% | 2.857% | -14.29% |  |  |  | -14.29% | 122.9% |  | 28.57% | 174.3%
| Z-Score (Absolute) |  | 55.2 | 0.29 | 0.29 | 26.3 | 0.58 | 0.29 | 0.67 | 0.29 | 0.29 | 0.67 | 2887.29 | 26.3 | 26.3 | 0.67 | 2.41 | 1.44 | 0.29 | 3.56
| Analytical results for S in OREAS 231b (Certified Value 0.262 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.27 | 0.26 | 0.37601 | 0.26409 | 0.29484 | 0.246 | 0.282 | 0.27 | 0.26 | 0.27 | 0.26 | 0.26 | 0.25 | 0.26 | 0.27 | 0.24 | 0.24972 | 0.25 | 0.24 | 0.27 | 0.27257 | 0.2705 | 0.24 | 0.27 | 0.28 | 0.26 | 0.27 |  |  |  |  | 1
| 1 | 2 | 0.28 | 0.26 | 0.40364 | 0.27667 | 0.28388 | 0.248 | 0.278 | 0.26 | 0.27 | 0.27 | 0.27 | 0.26 | 0.25 | 0.26 | 0.27 | 0.23 | 0.24612 | 0.27 | 0.24 | 0.26 | 0.25447 | 0.275 | 0.24 | 0.25 | 0.27 | 0.26 | 0.27 |  |  |  |  | 17
| 1 | 3 | 0.28 | 0.26 | 0.35816 | 0.26445 | 0.28553 | 0.245 | 0.281 | 0.26 | 0.24 | 0.26 | 0.27 | 0.26 | 0.26 | 0.26 | 0.27 | 0.22 | 0.25316 | 0.26 | 0.23 | 0.26 | 0.25067 | 0.2645 | 0.24 | 0.26 | 0.28 | 0.26 | 0.27 |  |  |  |  | 16
| 1 | 4 | 0.28 | 0.25 | 0.40379 | 0.26747 | 0.2868 | 0.248 | 0.277 | 0.26 | 0.24 | 0.25 | 0.27 | 0.26 | 0.26 | 0.26 | 0.27 | 0.22 | 0.24862 | 0.26 | 0.24 | 0.27 | 0.2532 | 0.273 | 0.24 | 0.26 | 0.27 | 0.25 | 0.27 |  |  |  |  | 0.2615
| 1 | 5 | 0.27 | 0.26 | 0.32607 | 0.26579 | 0.29933 | 0.244 | 0.28 | 0.26 | 0.25 | 0.25 | 0.27 | 0.26 | 0.26 | 0.26 | 0.27 | 0.23 | 0.25418 | 0.27 | 0.23 | 0.27 | 0.24977 | 0.2665 | 0.24 | 0.25 | 0.28 | 0.25 | 0.27 |  |  |  |  | 113
| 1 | 6 | 0.27 | 0.26 | 0.40885 | 0.26249 | 0.27296 | 0.249 | 0.282 | 0.26 | 0.26 | 0.25 | 0.27 | 0.26 | 0.25 | 0.26 | 0.27 | 0.24 | 0.26046 | 0.26 | 0.23 | 0.26 | 0.2544 | 0.266 | 0.24 | 0.27 | 0.28 | 0.25 | 0.28
| Mean |  | 0.275 | 0.25833 | 0.37942 | 0.26683 | 0.28722 | 0.24667 | 0.28 | 0.26167 | 0.25333 | 0.25833 | 0.26833 | 0.26 | 0.255 | 0.26 | 0.27 | 0.23 | 0.25204 | 0.26167 | 0.235 | 0.265 | 0.25585 | 0.26925 | 0.24 | 0.26 | 0.27667 | 0.255 | 0.27167
| Median |  | 0.275 | 0.26 | 0.38982 | 0.26512 | 0.28616 | 0.247 | 0.2805 | 0.26 | 0.255 | 0.255 | 0.27 | 0.26 | 0.255 | 0.26 | 0.27 | 0.23 | 0.25144 | 0.26 | 0.235 | 0.265 | 0.2538 | 0.2685 | 0.24 | 0.26 | 0.28 | 0.255 | 0.27
| Std Dev. |  | 0.0054772 | 0.0040825 | 0.032732 | 0.0051056 | 0.0091914 | 0.0019664 | 0.0020976 | 0.0040825 | 0.012111 | 0.0098319 | 0.0040825 | 0 | 0.0054772 | 0 | 0 | 0.0089443 | 0.005078 | 0.0075277 | 0.0054772 | 0.0054772 | 0.0084196 | 0.0042279 | 0 | 0.0089443 | 0.005164 | 0.0054772 | 0.0040825
| Rel.Std.Dev. |  | 1.992% | 1.58% | 8.627% | 1.913% | 3.2% | 0.7972% | 0.7491% | 1.56% | 4.781% | 3.806% | 1.521% | 0% | 2.148% | 0% | 0% | 3.889% | 2.015% | 2.877% | 2.331% | 2.067% | 3.291% | 1.57% | 0% | 3.44% | 1.866% | 2.148% | 1.503%
| PDM3 |  | 5.162% | -1.212% | 45.09% | 2.036% | 9.836% | -5.673% | 7.074% | 0.06307% | -3.124% | -1.212% | 2.612% | -0.5743% | -2.486% | -0.5743% | 3.25% | -12.05% | -3.617% | 0.06307% | -10.13% | 1.338% | -2.162% | 2.963% | -8.222% | -0.5743% | 5.799% | -2.486% | 3.887%
| Z-Score (Absolute) |  | 1.27 | 0.14 | 10.12 | 0.58 | 2.31 | 1.13 | 1.7 | 0.14 | 0.57 | 0.14 | 0.71 | 0 | 0.42 | 0 | 0.85 | 2.54 | 0.67 | 0.14 | 2.12 | 0.42 | 0.35 | 0.78 | 1.7 | 0 | 1.41 | 0.42 | 0.99
| Analytical results for Sb in OREAS 231b (Certified Value 0.5 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 0.65 | 0.58 | 4.98 | 0.58087 | 0.5 | 1.44 | 0.16 | 0.28 | < 5 | 0.7 | 0.49 | 0.53 | 0.51 | 61.698 | 0.49 | < 5 | < 5 | 0.48 | 0.42 | 0.38 | 7.35 | < 0.5 | 0.59521 | 0.53 | 0.19 | 0.6 | 0.36 | 0.38 | 0.49 |  |  | 1
| 1 | 2 | 0.62 | 0.55 | 3.89 | 0.54774 | 0.51 | 0.77 | 0.21 | 0.27 | < 5 | 0.72 | 0.47 | 0.5 | 0.5 | 61.342 | 0.47 | < 5 | < 5 | 0.46 | 0.4 | 0.4 | 6.52 | < 0.5 | 0.59144 | 0.53 | 0.169 | 0.48 | 0.33 | 0.37 | 0.46 |  |  | 33
| 1 | 3 | 0.63 | 0.57 | 4.51 | 0.54965 | 0.5 | 0.7 | 0.2 | 0.27 | < 5 | 0.71 | 0.49 | 0.52 | 0.59 | 59.242 | 0.48 | < 5 | < 5 | 0.51 | 0.37 | 0.41 | 8.64 | < 0.5 | 0.61878 | 0.53 | 0.151 | 0.47 | 0.32 | 0.39 | 0.48 |  |  | 16
| 1 | 4 | 0.66 | 0.53 | 4.56 | 0.56054 | 0.47 | 0.66 | 0.18 | 0.3 | < 5 | 0.67 | 0.5 | 0.52 | 0.54 | 66.365 | 0.49 | < 5 | < 5 | 0.48 | 0.41 | 0.41 | 5.42 | < 0.5 | 0.64234 | 0.56 | 0.143 | 0.43 | 0.34 | 0.38 | 0.56 |  |  | 0.50231
| 1 | 5 | 0.65 | 0.57 | 5.84 | 0.59564 | 0.48 | 0.63 | 0.23 | 0.27 | < 5 | 0.66 | 0.47 | 0.52 | 0.57 | 60.738 | 0.47 | < 5 | < 5 | 0.51 | 0.37 | 0.41 | 7.08 | < 0.5 | 0.64893 | 0.58 | 0.134 | 0.4 | 0.35 | 0.38 | 0.48 |  |  | 114
| 1 | 6 | 0.63 | 0.57 | 4.65 | 0.54409 | 0.46 | 0.64 | 0.2 | 0.27 | < 5 | 0.68 | 0.69 | 0.53 | 0.56 | 62.423 | 0.49 | < 5 | < 5 | 0.47 | 0.38 | 0.41 | 7.74 | < 0.5 | 0.61187 | 0.57 | 0.14 | 0.41 | 0.33 | 0.39 | 0.48
| Mean |  | 0.64 | 0.56167 | 4.7383 | 0.56309 | 0.48667 | 0.80667 | 0.19667 | 0.27667 |  | 0.69 | 0.51833 | 0.52 | 0.545 | 61.968 | 0.48167 |  |  | 0.485 | 0.39167 | 0.40333 | 7.125 |  | 0.61809 | 0.55 | 0.1545 | 0.465 | 0.33833 | 0.38167 | 0.49167
| Median |  | 0.64 | 0.57 | 4.605 | 0.5551 | 0.49 | 0.68 | 0.2 | 0.27 |  | 0.69 | 0.49 | 0.52 | 0.55 | 61.52 | 0.485 |  |  | 0.48 | 0.39 | 0.41 | 7.215 |  | 0.61533 | 0.545 | 0.147 | 0.45 | 0.335 | 0.38 | 0.48
| Std Dev. |  | 0.015492 | 0.018348 | 0.64552 | 0.020781 | 0.019664 | 0.31443 | 0.024221 | 0.012111 |  | 0.023664 | 0.084951 | 0.010954 | 0.035071 | 2.4062 | 0.0098319 |  |  | 0.020736 | 0.02137 | 0.012111 | 1.0954 |  | 0.023715 | 0.022804 | 0.021192 | 0.073417 | 0.01472 | 0.0075277 | 0.034881
| Rel.Std.Dev. |  | 2.421% | 3.267% | 13.62% | 3.691% | 4.041% | 38.98% | 12.32% | 4.377% |  | 3.43% | 16.39% | 2.107% | 6.435% | 3.883% | 2.041% |  |  | 4.276% | 5.456% | 3.003% | 15.37% |  | 3.837% | 4.146% | 13.72% | 15.79% | 4.351% | 1.972% | 7.094%
| PDM3 |  | 27.41% | 11.82% | 843.3% | 12.1% | -3.114% | 60.59% | -60.85% | -44.92% |  | 37.37% | 3.19% | 3.522% | 8.499% | 12237% | -4.11% |  |  | -3.446% | -22.03% | -19.7% | 1318% |  | 23.05% | 9.494% | -69.24% | -7.428% | -32.64% | -24.02% | -2.119%
| Z-Score (Absolute) |  | 0.63 | 0.22 | 22.16 | 0.23 | 0.18 | 1.51 | 1.7 | 1.28 | 10.4 | 0.89 | 0.01 | 0 | 0.13 | 322.87 | 0.2 | 10.4 | 10.4 | 0.18 | 0.67 | 0.61 | 34.71 | 1.42 | 0.52 | 0.16 | 1.92 | 0.29 | 0.95 | 0.73 | 0.15
| Analytical results for Sc in OREAS 231b (Certified Value 4.46 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 4.1 | 4.49 | 4.751 | 5.96 | 4.6858 | 5.9 | 4 | 6 | 5.6 | 5 | 4.1 | 4.6 | 4.2 | 4.2 | 4.3 | 4 | 5 | 3.5 | 4.5 | 3 | 4.4 | 1.5 | 4.5 | 4.5 | 3.3 | [6 empty] | 1
| 1 | 2 | 3.9 | 4.65 | 4.641 | 6.01 | 4.7306 | 6 | 4 | 6 | 5.1 | 5 | 4.1 | 4.8 | 4.4 | 4.1 | 4.2 | 4 | 5 | 3.7 | 4.6 | 4 | 4.5 | 1.3 | 4.7 | 4.4 | 3.5 | [6 empty] | 34
| 1 | 3 | 3.8 | 4.53 | 4.407 | 5.9 | 4.7609 | 6.3 | 4 | 6 | 5.1 | 5 | 4.2 | 4.7 | 4.3 | 4.5 | 4.5 | 4 | 5 | 3.5 | 4.5 | 3 | 4.4 | 1.2 | 4.4 | 4.5 | 3.3 | [6 empty] | 16
| 1 | 4 | 3.7 | 4.57 | 4.458 | 5.84 | 4.6932 | 5.6 | 4 | 6 | 5 | 5 | 4.2 | 4.5 | 4.2 | 4.4 | 4.2 | 3 | 5 | 3.4 | 4.5 | 3 | 4.4 | 1 | 4.4 | 4.3 | 3.1 | [6 empty] | 4.4618
| 1 | 5 | 3.6 | 4.64 | 4.809 | 5.72 | 5.0867 | 5.5 | 4 | 6 | 5.1 | 5 | 4.4 | 4.5 | 4.1 | 4.6 | 4.3 | 4 | 5 | 3.4 | 4.6 | 3 | 4.5 | 0.9 | 4.5 | 4.3 | 3.4 | [6 empty] | 115
| 1 | 6 | 3.6 | 4.53 | 4.958 | 5.98 | 4.5549 | 5.5 | 4 | 6 | 5.4 | 5 | 4.3 | 4.5 | 4.4 | 4.3 | 4.4 | 4 | 5 | 3.5 | 4.6 | 3 | 4.5 | 0.9 | 4.6 | 4.5 | 3.6
| Mean |  | 3.7833 | 4.5683 | 4.6707 | 5.9017 | 4.752 | 5.8 | 4 | 6 | 5.2167 | 5 | 4.2167 | 4.6 | 4.2667 | 4.35 | 4.3167 | 3.8333 | 5 | 3.5 | 4.55 | 3.1667 | 4.45 | 1.1333 | 4.5167 | 4.4167 | 3.3667
| Median |  | 3.75 | 4.55 | 4.696 | 5.93 | 4.7119 | 5.75 | 4 | 6 | 5.1 | 5 | 4.2 | 4.55 | 4.25 | 4.35 | 4.3 | 4 | 5 | 3.5 | 4.55 | 3 | 4.45 | 1.1 | 4.5 | 4.45 | 3.35
| Std Dev. |  | 0.19408 | 0.064627 | 0.21155 | 0.10778 | 0.17848 | 0.32249 | 0 | 0 | 0.23166 | 0 | 0.1169 | 0.12649 | 0.12111 | 0.18708 | 0.1169 | 0.40825 | 0 | 0.10954 | 0.054772 | 0.40825 | 0.054772 | 0.24221 | 0.1169 | 0.098319 | 0.17512
| Rel.Std.Dev. |  | 5.13% | 1.415% | 4.529% | 1.826% | 3.756% | 5.56% | 0% | 0% | 4.441% | 0% | 2.772% | 2.75% | 2.838% | 4.301% | 2.708% | 10.65% | 0% | 3.13% | 1.204% | 12.89% | 1.231% | 21.37% | 2.588% | 2.226% | 5.202%
| PDM3 |  | -15.21% | 2.387% | 4.68% | 32.27% | 6.504% | 29.99% | -10.35% | 34.47% | 16.92% | 12.06% | -5.495% | 3.096% | -4.374% | -2.507% | -3.254% | -14.09% | 12.06% | -21.56% | 1.976% | -29.03% | -0.2653% | -74.6% | 1.229% | -1.012% | -24.55%
| Z-Score (Absolute) |  | 2.04 | 0.36 | 0.67 | 4.44 | 0.92 | 4.13 | NA | NA | 2.34 | NA | 0.71 | 0.46 | 0.56 | 0.31 | 0.41 | NA | NA | 2.9 | 0.31 | NA | 0 | 10.14 | 0.2 | 0.1 | 3.31
Results from laboratories 1.08, 1.10, 1.12, 1.19, 1.22 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 231b (Indicative Value 0.52 ppm) | [32 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | 0.41 | < 1 | < 1.5 | 0.7 | < 1 | < 1 | < 1 | < 5 | 2.2 | 0.8 | 0.5 | 0.5 | 0.5 | < 5 | 6.476 | < 1 | 0.5 | 4.6 | < 1 | 0.30749 | 0.8 | 0.49 | 1.9 | < 1 | 0.4 | 0.4 |  |  |  |  | 1
| 1 | 2 | < 1 | 0.42 | < 1 | < 1.5 | 0.8 | 2 | < 1 | < 1 | < 5 | 2.2 | 0.5 | 0.5 | 0.4 | 0.6 | < 5 | < 5 | < 1 | 0.5 | 1.68 | < 1 | 0.46689 | 0.6 | 0.417 | 1.6 | < 1 | 0.4 | 0.4 |  |  |  |  | 4
| 1 | 3 | < 1 | 0.4 | < 1 | < 1.5 | 0.7 | 4 | < 1 | < 1 | < 5 | 1.9 | 0.5 | 0.5 | 0.9 | 0.5 | < 5 | < 5 | < 1 | 0.5 | 3.45 | < 1 | 0.40345 | 0.5 | 0.392 | 1.4 | < 1 | 0.4 | 0.4 |  |  |  |  | 16
| 1 | 4 | < 1 | 0.39 | < 1 | < 1.5 | 0.7 | 7 | < 1 | < 1 | < 5 | 2.9 | 0.8 | 0.5 | 0.6 | 0.8 | < 5 | < 5 | < 1 | 0.5 | 2.15 | < 1 | 0.36583 | 0.7 | 0.369 | 2.4 | < 1 | 0.4 | 0.4 |  |  |  |  | 0.52439
| 1 | 5 | < 1 | 0.38 | < 1 | < 1.5 | 0.7 | < 1 | < 1 | < 1 | < 5 | 2.2 | 0.6 | 0.5 | 0.8 | 0.8 | < 5 | < 5 | < 1 | 0.4 | 0.87 | < 1 | 0.38612 | 0.6 | 0.346 | 3.3 | < 1 | 0.4 | 0.4 |  |  |  |  | 14
| 1 | 6 | < 1 | 0.38 | < 1 | < 1.5 | 0.7 | < 1 | < 1 | < 1 | < 5 | 2.6 | 0.5 | 0.5 | 1.2 | 0.9 | < 5 | 6.4577 | < 1 | 0.4 | 0.92 | < 1 | 0.48968 | 0.8 | 1.584 | 2.5 | < 1 | 0.4 | 0.4
| Mean |  |  | 0.39667 |  |  | 0.71667 | 4.3333 |  |  |  | 2.3333 | 0.61667 | 0.5 | 0.73333 | 0.68333 |  | 6.4668 |  | 0.46667 | 2.2783 |  | 0.40324 | 0.66667 | 0.59967 | 2.1833 |  | 0.4 | 0.4
| Median |  |  | 0.395 |  |  | 0.7 | 4 |  |  |  | 2.2 | 0.55 | 0.5 | 0.7 | 0.7 |  | 6.4668 |  | 0.5 | 1.915 |  | 0.39478 | 0.65 | 0.4045 | 2.15 |  | 0.4 | 0.4
| Std Dev. |  |  | 0.01633 |  |  | 0.040825 | 2.5166 |  |  |  | 0.3559 | 0.1472 | 0 | 0.29439 | 0.17224 |  | 0.012966 |  | 0.05164 | 1.4809 |  | 0.066909 | 0.12111 | 0.48476 | 0.6969 |  | 6.0809e-17 | 6.0809e-17
| Rel.Std.Dev. |  |  | 4.117% |  |  | 5.696% | 58.08% |  |  |  | 15.25% | 23.87% | 0% | 40.14% | 25.21% |  | 0.2005% |  | 11.07% | 65% |  | 16.59% | 18.17% | 80.84% | 31.92% |  | 1.52e-14% | 1.52e-14%
| PDM3 |  |  | -24.36% |  |  | 36.67% | 726.4% |  |  |  | 345% | 17.6% | -4.652% | 39.84% | 30.31% |  | 1133% |  | -11.01% | 334.5% |  | -23.1% | 27.13% | 14.35% | 316.4% |  | -23.72% | -23.72%
| Z-Score (Absolute) |  | 0.51 | 1.03 | 0.51 | 0.76 | 0.59 | 9.21 | 0.51 | 0.51 | 9.63 | 8.79 | 0.09 | 0.51 | 0.68 | 0.42 | 9.63 | 16.34 | 0.51 | 0.67 | 8.51 | 0.51 | 1 | 0.34 | 0 | 8.03 | 0.51 | 1.01 | 1.01
| Analytical results for Si in OREAS 231b (Indicative Value 0.081 wt.%) | [32 empty] | Indicative
Si | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.25 | 1.27 | [29 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | [29 empty] | 3
|  |  | 0.5g | 25g | [29 empty] | 3
| 1 | 1 | 0.080646 | 0.085 | [29 empty] | 1
| 1 | 2 | 0.072854 | 0.071 | [29 empty] | 6
| 1 | 3 | 0.10836 | 0.079 | [29 empty] | 16
| 1 | 4 | 0.081276 | 0.075 | [29 empty] | 0.080817
| 1 | 5 | 0.077155 | 0.074 | [29 empty] | 12
| 1 | 6 | 0.081514 | 0.084
| Mean |  | 0.083634 | 0.078
| Median |  | 0.080961 | 0.077
| Std Dev. |  | 0.012557 | 0.0056569
| Rel.Std.Dev. |  | 15.01% | 7.252%
| PDM3 |  | 3.486% | -3.486%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sm in OREAS 231b (Certified Value 1.53 ppm) | [32 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.503 | 1.429 | 1.4611 | 1.5 | 1.73 | 1.6 | [25 empty] | 1
| 1 | 2 | 1.515 | 1.42 | 1.4477 | 1.5 | 1.67 | 1.7 | [25 empty] | 19
| 1 | 3 | 1.532 | 1.362 | 1.4175 | 1.9 | 1.68 | 1.5 | [25 empty] | 16
| 1 | 4 | 1.516 | 1.411 | 1.4245 | 1.5 | 1.68 | 1.6 | [25 empty] | 1.5312
| 1 | 5 | 1.516 | 1.473 | 1.3737 | 1.7 | 1.64 | 1.5 | [25 empty] | 116
| 1 | 6 | 1.505 | 1.34 | 1.3468 | 1.6 | 1.63 | 1.5
| Mean |  | 1.5145 | 1.4058 | 1.4119 | 1.6167 | 1.6717 | 1.5667
| Median |  | 1.5155 | 1.4155 | 1.421 | 1.55 | 1.675 | 1.55
| Std Dev. |  | 0.01033 | 0.048023 | 0.043794 | 0.16021 | 0.035449 | 0.08165
| Rel.Std.Dev. |  | 0.682% | 3.416% | 3.102% | 9.91% | 2.121% | 5.212%
| PDM3 |  | -1.091% | -8.188% | -7.793% | 5.582% | 9.174% | 2.316%
| Z-Score (Absolute) |  | 0.17 | 0.89 | 0.85 | 0.5 | 0.86 | 0.17
| Analytical results for Sn in OREAS 231b (Certified Value 0.52 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.6 | 0.53 | < 1 | 0.43748 | 0.42858 | 0.62 | 0.52 | 0.8 | 0.58 | < 10 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | < 10 | < 5 | 0.5 | 0.5 | 40.28 | < 0.5 | 0.5226 | 0.5 | 1.7 | 0.6 | 0.6 | 0.3 |  |  |  |  | 1
| 1 | 2 | 0.5 | 0.52 | < 1 | 0.46454 | 0.4327 | 0.61 | 0.5 | 0.8 | 0.55 | < 10 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | < 10 | < 5 | 0.6 | 0.5 | 39.22 | < 0.5 | 0.57834 | 0.5 | 1.9 | 0.5 | 0.6 | 0.3 |  |  |  |  | 20
| 1 | 3 | 0.5 | 0.5 | < 1 | 0.47061 | 0.3681 | 0.59 | 0.5 | 0.9 | 0.6 | < 10 | 0.5 | 0.5 | 0.6 | 0.5 | 0.5 | < 10 | < 5 | 0.5 | 0.5 | 43.81 | < 0.5 | 0.51678 | 0.5 | 1.8 | 0.5 | 0.6 | 0.2 |  |  |  |  | 16
| 1 | 4 | 0.5 | 0.54 | < 1 | 0.48607 | 0.39924 | 0.6 | 0.49 | 0.8 | 0.53 | < 10 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | < 10 | < 5 | 0.6 | 0.5 | 41.56 | < 0.5 | 0.54419 | 0.5 | 1.5 | 0.6 | 0.5 | 0.3 |  |  |  |  | 0.52088
| 1 | 5 | 1.3 | 0.52 | < 1 | 0.48334 | 0.39641 | 0.59 | 0.49 | 0.8 | 0.58 | < 10 | 0.5 | 0.5 | 0.5 | 0.6 | 0.5 | < 10 | < 5 | 0.6 | 0.5 | 41.81 | < 0.5 | 0.55797 | 0.5 | 1.3 | 0.6 | 0.6 | 0.3 |  |  |  |  | 117
| 1 | 6 | 0.5 | 0.52 | < 1 | 0.47713 | 0.42898 | 0.61 | 0.51 | 0.7 | 0.57 | < 10 | 0.4 | 0.5 | 0.6 | 0.5 | 0.5 | < 10 | < 5 | 0.6 | 0.5 | 41.98 | < 0.5 | 0.57244 | 0.5 | 1.2 | 0.5 | 0.6 | 0.3
| Mean |  | 0.65 | 0.52167 |  | 0.46986 | 0.409 | 0.60333 | 0.50167 | 0.8 | 0.56833 |  | 0.48333 | 0.5 | 0.53333 | 0.51667 | 0.5 |  |  | 0.56667 | 0.5 | 41.4433 |  | 0.54872 | 0.5 | 1.5667 | 0.55 | 0.58333 | 0.28333
| Median |  | 0.5 | 0.52 |  | 0.47387 | 0.41391 | 0.605 | 0.5 | 0.8 | 0.575 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 0.6 | 0.5 | 41.685 |  | 0.55108 | 0.5 | 1.6 | 0.55 | 0.6 | 0.3
| Std Dev. |  | 0.32094 | 0.013292 |  | 0.017746 | 0.02558 | 0.012111 | 0.01169 | 0.063246 | 0.024833 |  | 0.040825 | 0 | 0.05164 | 0.040825 | 0 |  |  | 0.05164 | 0 | 1.5706 |  | 0.025489 | 0 | 0.28048 | 0.054772 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 49.37% | 2.548% |  | 3.777% | 6.254% | 2.007% | 2.33% | 7.906% | 4.369% |  | 8.447% | 0% | 9.682% | 7.902% | 0% |  |  | 9.113% | 0% | 3.79% |  | 4.645% | 0% | 17.9% | 9.959% | 6.999% | 14.41%
| PDM3 |  | 24.79% | 0.1502% |  | -9.795% | -21.48% | 15.83% | -3.689% | 53.58% | 9.109% |  | -7.209% | -4.009% | 2.39% | -0.8097% | -4.009% |  |  | 8.789% | -4.009% | 7856% |  | 5.344% | -4.009% | 200.8% | 5.59% | 11.99% | -45.61%
| Z-Score (Absolute) |  | 1.57 | 0.16 | 0.45 | 0.86 | 1.68 | 0.94 | 0.43 | 3.6 | 0.47 | 60.24 | 0.67 | 0.45 | 0 | 0.22 | 0.45 | 60.24 | 26.52 | 0.45 | 0.45 | 551.72 | 3.82 | 0.21 | 0.45 | 13.94 | 0.22 | 0.67 | 3.37
| Analytical results for Sr in OREAS 231b (Certified Value 32 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 18.2 | 33.39 | 143.84 | 44.5012 | 33.7499 | 50.3 | 37 | 45 | 45.94 | 38 | 29 | 33.8 | 32.3 | 27.4 | 31.5 | 29 | 36.8062 | 32.8 | 25.6 | 29.5 | 23.6 | 33.2196 | 21 | 31.5 | 25.9 | 27.6 | 32.2 | 21.7 |  |  |  | 1
| 1 | 2 | 15.4 | 32.43 | 161.78 | 42.4326 | 34.4588 | 48.6 | 40 | 45.6 | 41.74 | 39 | 30 | 33.9 | 32.7 | 28 | 29.8 | 29 | 33.2307 | 33.2 | 26.4 | 30.4 | 23.57 | 30.8481 | 22 | 31.4 | 25.2 | 26.4 | 32.2 | 21.7 |  |  |  | 21
| 1 | 3 | 13.6 | 31.33 | 138.37 | 45.4327 | 34.329 | 50.7 | 40 | 46.4 | 41.89 | 38 | 31 | 33.7 | 33.2 | 30.4 | 33.2 | 28 | 37.2184 | 32.6 | 25.9 | 31 | 25.14 | 32.3175 | 21 | 31.3 | 24.6 | 26.8 | 32.5 | 20.8 |  |  |  | 16
| 1 | 4 | 12.5 | 31.13 | 158.09 | 44.6853 | 34.3681 | 49.1 | 39 | 46 | 38.71 | 37 | 31 | 32.1 | 34.2 | 29.5 | 31.2 | 28 | 35.1883 | 32.8 | 24.5 | 31.4 | 23.18 | 33.0243 | 22 | 32.1 | 23 | 27.5 | 30.8 | 20.1 |  |  |  | 32.046
| 1 | 5 | 13.9 | 31.62 | 126.73 | 43.8304 | 36.8887 | 49.4 | 42 | 43 | 41.47 | 38 | 32 | 32.5 | 34 | 30.4 | 31.1 | 30 | 39.2867 | 33 | 24.8 | 31.1 | 21.88 | 31.8711 | 20 | 31.6 | 22.4 | 28 | 31.3 | 21.8 |  |  |  | 118
| 1 | 6 | 14.2 | 31.75 | 161.39 | 44.1725 | 33.3016 | 49.4 | 39 | 44 | 44.47 | 39 | 32 | 32.7 | 33.3 | 30.4 | 31.3 | 29 | 40.782 | 33.2 | 25.4 | 31.2 | 22.27 | 30.6807 | 20 | 32.1 | 23.6 | 26.4 | 32.3 | 22.6
| Mean |  | 14.6333 | 31.9417 | 148.3667 | 44.1758 | 34.516 | 49.5833 | 39.5 | 45 | 42.37 | 38.1667 | 30.8333 | 33.1167 | 33.2833 | 29.35 | 31.35 | 28.8333 | 37.0854 | 32.9333 | 25.4333 | 30.7667 | 23.2733 | 31.9935 | 21 | 31.6667 | 24.1167 | 27.1167 | 31.8833 | 21.45
| Median |  | 14.05 | 31.685 | 150.965 | 44.3369 | 34.3485 | 49.4 | 39.5 | 45.3 | 41.815 | 38 | 31 | 33.2 | 33.25 | 29.95 | 31.25 | 29 | 37.0123 | 32.9 | 25.5 | 31.05 | 23.375 | 32.0943 | 21 | 31.55 | 24.1 | 27.15 | 32.2 | 21.7
| Std Dev. |  | 1.9826 | 0.83777 | 14.3711 | 1.0103 | 1.2451 | 0.77825 | 1.6432 | 1.29 | 2.5293 | 0.75277 | 1.169 | 0.77567 | 0.73052 | 1.3383 | 1.0895 | 0.75277 | 2.7226 | 0.24221 | 0.70048 | 0.70616 | 1.152 | 1.0697 | 0.89443 | 0.35024 | 1.3452 | 0.67651 | 0.67355 | 0.8735
| Rel.Std.Dev. |  | 13.55% | 2.623% | 9.686% | 2.287% | 3.607% | 1.57% | 4.16% | 2.867% | 5.97% | 1.972% | 3.791% | 2.342% | 2.195% | 4.56% | 3.475% | 2.611% | 7.342% | 0.7355% | 2.754% | 2.295% | 4.95% | 3.343% | 4.259% | 1.106% | 5.578% | 2.495% | 2.113% | 4.072%
| PDM3 |  | -54.34% | -0.3257% | 363% | 37.85% | 7.708% | 54.73% | 23.26% | 40.42% | 32.22% | 19.1% | -3.784% | 3.341% | 3.861% | -8.413% | -2.172% | -10.03% | 15.73% | 2.769% | -20.64% | -3.992% | -27.38% | -0.1638% | -34.47% | -1.184% | -24.74% | -15.38% | -0.5078% | -33.07%
| Z-Score (Absolute) |  | 2.34 | 0 | 15.75 | 1.66 | 0.35 | 2.39 | 1.03 | 1.77 | 1.41 | 0.85 | 0.15 | 0.16 | 0.19 | 0.35 | 0.08 | 0.42 | 0.7 | 0.14 | 0.88 | 0.16 | 1.17 | 0.01 | 1.48 | 0.03 | 1.05 | 0.65 | 0 | 1.42
| Analytical results for Ta in OREAS 231b (Certified Value < 0.01 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.24 | 0.22 | 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.32 | 0.12 | 0.01 | < 0.01 | [11 empty] | 22
| 1 | 3 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.05 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.46 | 0.14 | 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | 0.01193 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.19 | 0.13 | < 0.01 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.27 | 0.12 | < 0.01 | < 0.01 | [11 empty] | 119
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.2 | 0.18 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01193 |  | 0.04 |  |  | 5 |  |  | 0.01 | [5 empty] | 0.28 | 0.15167 | 0.01
| Median |  |  |  | 0.01193 |  | 0.04 |  |  | 5 |  |  | 0.01 | [5 empty] | 0.255 | 0.135 | 0.01
| Std Dev. | [5 empty] | 0.0063246 |  |  | 0 |  |  | 0 | [5 empty] | 0.1002 | 0.040208 | 0
| Rel.Std.Dev. | [5 empty] | 15.81% |  |  | 0% |  |  | 0% | [5 empty] | 35.79% | 26.51% | 0%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.61 | 0.4 | 1.21 | 0.4 | 0.67 | 268.78 | 1.75 | 0.67 | 0.4 | 0.67 | 0.67 | 0.4 | 12.54 | 0.67 | 14.16 | 7.24 | 0.54 | 0.67
| Analytical results for Tb in OREAS 231b (Certified Value 0.31 ppm) | [32 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 3
| 1 | 1 | 0.31 | 0.311 | 0.31 | 0.31143 | 0.4 | 0.31 | 0.25 | 0.33 | [23 empty] | 1
| 1 | 2 | 0.29 | 0.31 | 0.302 | 0.31396 | 0.4 | 0.3 | 0.25 | 0.31 | [23 empty] | 23
| 1 | 3 | 0.3 | 0.312 | 0.289 | 0.31247 | 0.4 | 0.31 | 0.3 | 0.31 | [23 empty] | 16
| 1 | 4 | 0.29 | 0.301 | 0.298 | 0.31699 | 0.4 | 0.31 | 0.25 | 0.32 | [23 empty] | 0.3066
| 1 | 5 | 0.29 | 0.307 | 0.317 | 0.31753 | 0.3 | 0.32 | 0.25 | 0.32 | [23 empty] | 120
| 1 | 6 | 0.28 | 0.302 | 0.287 | 0.29908 | 0.4 | 0.31 | 0.25 | 0.31
| Mean |  | 0.29333 | 0.30717 | 0.3005 | 0.31191 | 0.38333 | 0.31 | 0.25833 | 0.31667
| Median |  | 0.29 | 0.3085 | 0.3 | 0.31322 | 0.4 | 0.31 | 0.25 | 0.315
| Std Dev. |  | 0.010328 | 0.0047081 | 0.011709 | 0.0067347 | 0.040825 | 0.0063246 | 0.020412 | 0.008165
| Rel.Std.Dev. |  | 3.521% | 1.533% | 3.896% | 2.159% | 10.65% | 2.04% | 7.902% | 2.578%
| PDM3 |  | -4.326% | 0.1861% | -1.988% | 1.733% | 25.03% | 1.11% | -15.74% | 3.285%
| Z-Score (Absolute) |  | 1.4 | 0 | 0.67 | 0.48 | NA | 0.29 | 4.94 | 0.96
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 231b (Certified Value 0.06 ppm) | [32 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.05 | 0.06 | 0.10953 | 0.04 | 0.36 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.05 | 0.06 | 0.07 | 0.06 | < 5 | 0.05 | 0.05 | < 0.1 | 0.05605 | 0.06 | 0.283 | 0.13 | 0.07 | 0.06 | 0.08 | [7 empty] | 1
| 1 | 2 | < 0.05 | 0.06 | 0.11148 | 0.06 | 0.16 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.05 | 0.06 | 0.05 | 0.06 | < 5 | 0.06 | 0.06 | < 0.1 | 0.06314 | 0.05 | 0.206 | 0.11 | 0.06 | 0.05 | 0.06 | [7 empty] | 24
| 1 | 3 | 0.06 | 0.06 | 0.085323 | 0.06 | 0.12 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.05 | 0.07 | 0.06 | 0.06 | < 5 | 0.06 | 0.06 | < 0.1 | 0.12059 | 0.06 | 0.103 | 0.07 | 0.06 | 0.05 | 0.05 | [7 empty] | 16
| 1 | 4 | < 0.05 | 0.06 | 0.097327 | 0.09 | 0.09 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.05 | 0.06 | 0.07 | 0.07 | < 5 | 0.05 | 0.07 | < 0.1 | 0.04559 | 0.07 | 0.103 | 0.15 | 0.06 | 0.05 | 0.04 | [7 empty] | 0.059653
| 1 | 5 | < 0.05 | 0.06 | 0.09795 | 0.09 | 0.08 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.06 | 0.06 | 0.07 | 0.05 | < 5 | 0.05 | 0.06 | < 0.1 | 0.08802 | 0.06 | < 0.03 | < 0.01 | 0.05 | 0.06 | 0.07 | [7 empty] | 121
| 1 | 6 | 0.07 | 0.06 | 0.086166 | 0.05 | 0.08 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.04 | 0.07 | 0.06 | 0.06 | < 5 | 0.06 | 0.06 | < 0.1 | 0.10944 | 0.06 | 0.72 | 0.09 | 0.05 | 0.07 | 0.07
| Mean |  | 0.065 | 0.06 | 0.097962 | 0.065 | 0.14833 |  |  |  |  | 0.05 | 0.063333 | 0.063333 | 0.06 |  | 0.055 | 0.06 |  | 0.080472 | 0.06 | 0.283 | 0.11 | 0.058333 | 0.056667 | 0.061667
| Median |  | 0.065 | 0.06 | 0.097639 | 0.06 | 0.105 |  |  |  |  | 0.05 | 0.06 | 0.065 | 0.06 |  | 0.055 | 0.06 |  | 0.07558 | 0.06 | 0.206 | 0.11 | 0.06 | 0.055 | 0.065
| Std Dev. |  | 0.0070711 | 0 | 0.011098 | 0.020736 | 0.10815 |  |  |  |  | 0.0063246 | 0.005164 | 0.008165 | 0.0063246 |  | 0.0054772 | 0.0063246 |  | 0.030395 | 0.0063246 | 0.25578 | 0.031623 | 0.0075277 | 0.008165 | 0.01472
| Rel.Std.Dev. |  | 10.88% | 0% | 11.33% | 31.9% | 72.91% |  |  |  |  | 12.65% | 8.154% | 12.89% | 10.54% |  | 9.959% | 10.54% |  | 37.77% | 10.54% | 90.38% | 28.75% | 12.9% | 14.41% | 23.87%
| PDM3 |  | 8.964% | 0.5821% | 64.22% | 8.964% | 148.7% |  |  |  |  | -16.18% | 6.17% | 6.17% | 0.5821% |  | -7.8% | 0.5821% |  | 34.9% | 0.5821% | 374.4% | 84.4% | -2.212% | -5.006% | 3.376%
| Z-Score (Absolute) |  | 1.46 | 0 | 2.56 | 0.34 | 5.96 | 2.36 | 0.67 | 164.53 | 0.67 | 0.67 | 0.22 | 0.22 | 0 | 164.53 | 0.34 | 0 | 0.67 | 1.38 | 0 | 12.03 | 2.19 | 0.11 | 0.22 | 0.11
| Analytical results for Th in OREAS 231b (Certified Value 0.46 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.6 | 0.48 | 0.415 | 0.3556 | 0.50879 | 0.5 | 1.2 | 0.72 | 0.47 | 0.5 | 0.5 | 0.5 | 0.4 | < 100 | 0.4 | 0.45 | 0.4 | 0.5 | 0.4 | 0.4 | 2.1 | 0.5 | 0.5 | 0.4 | [7 empty] | 1
| 1 | 2 | 0.3 | 0.482 | 0.438 | 0.3581 | 0.49313 | 0.5 | 0.6 | 0.7 | 0.46 | 0.5 | 0.5 | 0.5 | 0.4 | < 100 | 0.4 | 0.43 | 0.4 | 0.5 | 0.5 | 0.4 | 3 | 0.4 | 0.5 | 0.5 | [7 empty] | 25
| 1 | 3 | 0.4 | 0.478 | 0.507 | 0.34693 | 0.49897 | 0.5 | 0.53 | 0.7 | 0.47 | 0.5 | 0.5 | 0.4 | 0.5 | < 100 | 0.5 | 0.45 | 0.4 | 0.5 | 0.4 | 0.4 | 2.4 | 0.5 | 0.5 | 0.5 | [7 empty] | 16
| 1 | 4 | 0.2 | 0.481 | 0.407 | 0.35389 | 0.48848 | 0.5 | 0.53 | 0.65 | 0.45 | 0.5 | 0.4 | 0.5 | 0.5 | < 100 | 0.4 | 0.44 | 0.4 | 0.5 | 0.4 | 0.5 | 2 | 0.5 | 0.5 | 0.4 | [7 empty] | 0.45898
| 1 | 5 | 0.1 | 0.486 | 0.42 | 0.36365 | 0.55537 | 0.5 | 0.54 | 0.68 | 0.48 | 0.5 | 0.4 | 0.4 | 0.5 | < 100 | 0.4 | 0.44 | 0.4 | 0.5 | 0.4 | 0.5 | 1.8 | 0.4 | 0.5 | 0.4 | [7 empty] | 122
| 1 | 6 | 0.2 | 0.47 | 0.397 | 0.34833 | 0.49185 | 0.5 | 0.52 | 0.75 | 0.47 | 0.5 | 0.4 | 0.5 | 0.5 | < 100 | 0.4 | 0.43 | 0.4 | 0.5 | 0.4 | 0.5 | 1.6 | 0.5 | 0.5 | 0.5
| Mean |  | 0.3 | 0.4795 | 0.43067 | 0.35441 | 0.5061 | 0.5 | 0.65333 | 0.7 | 0.46667 | 0.5 | 0.45 | 0.46667 | 0.46667 |  | 0.41667 | 0.44 | 0.4 | 0.5 | 0.41667 | 0.45 | 2.15 | 0.46667 | 0.5 | 0.45
| Median |  | 0.25 | 0.4805 | 0.4175 | 0.35474 | 0.49605 | 0.5 | 0.535 | 0.7 | 0.47 | 0.5 | 0.45 | 0.5 | 0.5 |  | 0.4 | 0.44 | 0.4 | 0.5 | 0.4 | 0.45 | 2.05 | 0.5 | 0.5 | 0.45
| Std Dev. |  | 0.17889 | 0.0053572 | 0.039833 | 0.0062224 | 0.025171 | 0 | 0.26934 | 0.034059 | 0.010328 | 0 | 0.054772 | 0.05164 | 0.05164 |  | 0.040825 | 0.0089443 | 6.0809e-17 | 0 | 0.040825 | 0.054772 | 0.49699 | 0.05164 | 0 | 0.054772
| Rel.Std.Dev. |  | 59.63% | 1.117% | 9.249% | 1.756% | 4.974% | 0% | 41.23% | 4.866% | 2.213% | 0% | 12.17% | 11.07% | 11.07% |  | 9.798% | 2.033% | 1.52e-14% | 0% | 9.798% | 12.17% | 23.12% | 11.07% | 0% | 12.17%
| PDM3 |  | -34.64% | 4.471% | -6.168% | -22.78% | 10.27% | 8.938% | 42.35% | 52.51% | 1.675% | 8.938% | -1.956% | 1.675% | 1.675% |  | -9.219% | -4.135% | -12.85% | 8.938% | -9.219% | -1.956% | 368.4% | 1.675% | 8.938% | -1.956%
| Z-Score (Absolute) |  | 3.37 | 0.26 | 0.73 | 2.27 | 0.8 | 0.67 | 3.78 | 4.72 | 0 | 0.67 | 0.34 | 0 | 0 | 1002.02 | 1.01 | 0.54 | 1.35 | 0.67 | 1.01 | 0.34 | 34.05 | 0 | 0.67 | 0.34
| Analytical results for Ti in OREAS 231b (Certified Value 0.329 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.38 | 0.3051 | 0.38909 | 0.47741 | 0.36494 | 0.39 | 0.4178 | 0.474 | 0.4278 | 0.34 | 0.27 | 0.342 | 0.363 | 0.326 | < 0.01 | 0.324 | 0.26 | 0.43574 | 0.4015 | 0.235 | 0.31 | 0.2481 | 0.34954 | 0.1215 | 0.308 | 0.302 | 0.32 | 0.301 | 0.222 |  |  | 1
| 1 | 2 | 0.32 | 0.3012 | 0.33426 | 0.47927 | 0.36855 | 0.39 | 0.443 | 0.466 | 0.4093 | 0.34 | 0.278 | 0.343 | 0.359 | 0.326 | < 0.01 | 0.321 | 0.27 | 0.37463 | 0.4194 | 0.256 | 0.32 | 0.25146 | 0.33188 | 0.132 | 0.312 | 0.293 | 0.31 | 0.301 | 0.223 |  |  | 26
| 1 | 3 | 0.32 | 0.3007 | 0.33615 | 0.47043 | 0.37133 | 0.39 | 0.4454 | 0.471 | 0.4142 | 0.34 | 0.285 | 0.339 | 0.361 | 0.332 | < 0.01 | 0.339 | 0.27 | 0.43894 | 0.4025 | 0.244 | 0.31 | 0.26703 | 0.34771 | 0.125 | 0.296 | 0.281 | 0.31 | 0.301 | 0.208 |  |  | 16
| 1 | 4 | 0.31 | 0.302 | 0.36853 | 0.4706 | 0.36625 | 0.39 | 0.4364 | 0.474 | 0.3873 | 0.35 | 0.294 | 0.344 | 0.355 | 0.319 | < 0.01 | 0.318 | 0.26 | 0.40343 | 0.4171 | 0.216 | 0.31 | 0.244 | 0.35046 | 0.1265 | 0.315 | 0.258 | 0.32 | 0.29 | 0.191 |  |  | 0.32883
| 1 | 5 | 0.3 | 0.3021 | 0.3439 | 0.47348 | 0.39544 | 0.38 | 0.4568 | 0.475 | 0.3997 | 0.34 | 0.3 | 0.335 | 0.359 | 0.334 | < 0.01 | 0.327 | 0.28 | 0.44996 | 0.4173 | 0.225 | 0.33 | 0.2277 | 0.33395 | 0.118 | 0.314 | 0.241 | 0.32 | 0.294 | 0.23 |  |  | 123
| 1 | 6 | 0.29 | 0.3004 | 0.3989 | 0.47024 | 0.35929 | 0.39 | 0.4352 | 0.469 | 0.4207 | 0.34 | 0.304 | 0.334 | 0.358 | 0.337 | < 0.01 | 0.326 | 0.26 | 0.46022 | 0.4015 | 0.237 | 0.32 | 0.23084 | 0.33446 | 0.1145 | 0.312 | 0.259 | 0.32 | 0.303 | 0.238
| Mean |  | 0.32 | 0.30192 | 0.3618 | 0.47357 | 0.37097 | 0.38833 | 0.4391 | 0.4715 | 0.40983 | 0.34167 | 0.2885 | 0.3395 | 0.35917 | 0.329 |  | 0.32583 | 0.26667 | 0.42715 | 0.40988 | 0.2355 | 0.31667 | 0.24486 | 0.34133 | 0.12292 | 0.3095 | 0.27233 | 0.31667 | 0.29833 | 0.21867
| Median |  | 0.315 | 0.3016 | 0.35621 | 0.47204 | 0.3674 | 0.39 | 0.4397 | 0.4725 | 0.41175 | 0.34 | 0.2895 | 0.3405 | 0.359 | 0.329 |  | 0.325 | 0.265 | 0.43734 | 0.4098 | 0.236 | 0.315 | 0.24605 | 0.34108 | 0.12325 | 0.312 | 0.27 | 0.32 | 0.301 | 0.2225
| Std Dev. |  | 0.031623 | 0.0017011 | 0.027939 | 0.0039246 | 0.012646 | 0.0040825 | 0.012995 | 0.0035071 | 0.014633 | 0.0040825 | 0.013172 | 0.0042308 | 0.0027142 | 0.0065727 |  | 0.0072503 | 0.008165 | 0.032084 | 0.0088626 | 0.014039 | 0.008165 | 0.014406 | 0.0087437 | 0.0062723 | 0.0070356 | 0.023441 | 0.005164 | 0.0051251 | 0.016801
| Rel.Std.Dev. |  | 9.882% | 0.5634% | 7.722% | 0.8287% | 3.409% | 1.051% | 2.96% | 0.7438% | 3.571% | 1.195% | 4.566% | 1.246% | 0.7557% | 1.998% |  | 2.225% | 3.062% | 7.511% | 2.162% | 5.961% | 2.578% | 5.884% | 2.562% | 5.103% | 2.273% | 8.607% | 1.631% | 1.718% | 7.683%
| PDM3 |  | -2.684% | -8.184% | 10.03% | 44.02% | 12.81% | 18.1% | 33.54% | 43.39% | 24.63% | 3.905% | -12.26% | 3.246% | 9.227% | 0.05265% |  | -0.9104% | -18.9% | 29.9% | 24.65% | -28.38% | -3.698% | -25.54% | 3.804% | -62.62% | -5.878% | -17.18% | -3.698% | -9.273% | -33.5%
| Z-Score (Absolute) |  | 0.09 | 0.36 | 0.54 | 2.21 | 0.67 | 0.93 | 1.69 | 2.18 | 1.26 | 0.24 | 0.56 | 0.2 | 0.5 | 0.05 | 4.79 | 0 | 0.88 | 1.51 | 1.26 | 1.35 | 0.14 | 1.21 | 0.23 | 3.03 | 0.24 | 0.8 | 0.14 | 0.41 | 1.6
| Analytical results for Tl in OREAS 231b (Certified Value 0.064 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.06 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.07 | < 10 | 0.08 | 0.07 | 0.07 | 0.07 | 0.06 | < 5 | 0.06 | 0.06 | 0.07 | 5.57 | 0.05 | 0.06326 | 0.06 | 0.083 | 0.05 | 0.07 | 0.07 | 0.06 | [6 empty] | 1
| 1 | 2 | 0.06 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.06 | < 10 | 0.07 | 0.06 | 0.07 | 0.06 | 0.06 | < 5 | 0.06 | 0.06 | 0.07 | 2.52 | 0.05 | 0.0583 | 0.07 | 0.069 | 0.02 | 0.06 | 0.06 | 0.06 | [6 empty] | 27
| 1 | 3 | 0.06 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.06 | < 10 | 0.07 | 0.06 | 0.07 | 0.07 | 0.06 | < 5 | 0.07 | 0.06 | 0.07 | 1.67 | < 0.05 | 0.06674 | 0.07 | 0.072 | 0.02 | 0.06 | 0.07 | 0.06 | [6 empty] | 16
| 1 | 4 | 0.06 | 0.06 | < 0.1 | 0.07 | 0.07 | 0.06 | < 10 | 0.08 | 0.06 | 0.06 | 0.07 | 0.06 | < 5 | 0.06 | 0.06 | 0.07 | 5.09 | < 0.05 | 0.06827 | 0.06 | 0.065 | < 0.01 | 0.06 | 0.06 | 0.06 | [6 empty] | 0.064382
| 1 | 5 | 0.06 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.06 | < 10 | 0.07 | 0.05 | 0.06 | 0.08 | 0.06 | < 5 | 0.07 | 0.06 | 0.07 | 3.97 | < 0.05 | 0.06271 | 0.06 | 0.069 | < 0.01 | 0.06 | 0.07 | 0.06 | [6 empty] | 124
| 1 | 6 | 0.06 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.06 | < 10 | 0.08 | 0.06 | 0.07 | 0.06 | 0.06 | < 5 | 0.07 | 0.06 | 0.07 | 2.08 | < 0.05 | 0.06696 | 0.06 | 0.087 | < 0.01 | 0.06 | 0.06 | 0.06
| Mean |  | 0.06 | 0.068333 |  | 0.066667 | 0.07 | 0.061667 |  | 0.075 | 0.06 | 0.066667 | 0.068333 | 0.06 |  | 0.065 | 0.06 | 0.07 | 3.4833 | 0.05 | 0.064373 | 0.063333 | 0.074167 | 0.03 | 0.061667 | 0.065 | 0.06
| Median |  | 0.06 | 0.07 |  | 0.07 | 0.07 | 0.06 |  | 0.075 | 0.06 | 0.07 | 0.07 | 0.06 |  | 0.065 | 0.06 | 0.07 | 3.245 | 0.05 | 0.065 | 0.06 | 0.0705 | 0.02 | 0.06 | 0.065 | 0.06
| Std Dev. |  | 0 | 0.0040825 |  | 0.005164 | 0 | 0.0040825 |  | 0.0054772 | 0.0063246 | 0.005164 | 0.0075277 | 0 |  | 0.0054772 | 0 | 0 | 1.6345 | 0 | 0.0036978 | 0.005164 | 0.0087731 | 0.017321 | 0.0040825 | 0.0054772 | 0
| Rel.Std.Dev. |  | 0% | 5.974% |  | 7.746% | 0% | 6.62% |  | 7.303% | 10.54% | 7.746% | 11.02% | 0% |  | 8.427% | 0% | 0% | 46.92% | 0% | 5.744% | 8.154% | 11.83% | 57.74% | 6.62% | 8.427% | 0%
| PDM3 |  | -6.806% | 6.137% |  | 3.549% | 8.726% | -4.218% |  | 16.49% | -6.806% | 3.549% | 6.137% | -6.806% |  | 0.9599% | -6.806% | 8.726% | 5310% | -22.34% | -0.01346% | -1.629% | 15.2% | -53.4% | -4.218% | 0.9599% | -6.806%
| Z-Score (Absolute) |  | 0.67 | 0.45 | 2.02 | 0.22 | 0.67 | 0.45 | 665.54 | 1.35 | 0.67 | 0.22 | 0.45 | 0.67 | 328.39 | 0 | 0.67 | 0.67 | 461 | 4.27 | 0.08 | 0.22 | 1.24 | 6.41 | 0.45 | 0 | 0.67
| Analytical results for Tm in OREAS 231b (Certified Value 0.17 ppm) | [32 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | [26 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [26 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [26 empty] | 2
| 1 | 1 | 0.169 | 0.164 | 0.15027 | 0.2 | 0.15 | [26 empty] | 1
| 1 | 2 | 0.168 | 0.163 | 0.1519 | 0.2 | 0.15 | [26 empty] | 28
| 1 | 3 | 0.176 | 0.152 | 0.14707 | 0.2 | 0.15 | [26 empty] | 16
| 1 | 4 | 0.168 | 0.16 | 0.14703 | 0.2 | 0.15 | [26 empty] | 0.16578
| 1 | 5 | 0.175 | 0.163 | 0.15317 | 0.2 | 0.15 | [26 empty] | 125
| 1 | 6 | 0.168 | 0.152 | 0.14587 | 0.2 | 0.15
| Mean |  | 0.17067 | 0.159 | 0.14922 | 0.2 | 0.15
| Median |  | 0.1685 | 0.1615 | 0.14867 | 0.2 | 0.15
| Std Dev. |  | 0.0037771 | 0.0055857 | 0.0029857 | 3.0405e-17 | 0
| Rel.Std.Dev. |  | 2.213% | 3.513% | 2.001% | 1.52e-14% | 0%
| PDM3 |  | 2.949% | -4.088% | -9.988% | 20.64% | -9.517%
| Z-Score (Absolute) |  | 0.8 | 0 | 0.67 | 2.83 | 0.62
| Analytical results for U in OREAS 231b (Certified Value 0.13 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.14 | 0.135 | 0.154 | 0.10936 | 0.18195 | 0.1 | 0.26 | 0.17 | 0.14 | 0.1 | 0.13 | 0.14 | 0.12 | 1469.811 | 0.12 | < 10 | 0.13 | 0.1 | 0.13 | 0.1 | 0.12 | 0.18 | 0.14 | 0.13 | 0.13 | 0.11 | [5 empty] | 1
| 1 | 2 | 0.12 | 0.133 | 0.153 | 0.10594 | 0.16591 | 0.1 | 0.27 | 0.17 | 0.14 | 0.1 | 0.13 | 0.13 | 0.12 | 1542.453 | 0.12 | < 10 | 0.12 | 0.1 | 0.13 | 0.1 | 0.13 | 0.135 | 0.13 | 0.12 | 0.13 | 0.11 | [5 empty] | 29
| 1 | 3 | 0.12 | 0.136 | 0.152 | 0.10297 | 0.14874 | 0.2 | 0.26 | 0.18 | 0.15 | 0.1 | 0.13 | 0.14 | 0.14 | 1556.744 | 0.14 | < 10 | 0.13 | 0.1 | 0.14 | 0.1 | 0.13 | 0.138 | 0.13 | 0.13 | 0.13 | 0.11 | [5 empty] | 16
| 1 | 4 | 0.12 | 0.133 | 0.146 | 0.10639 | 0.15365 | 0.1 | 0.27 | 0.17 | 0.14 | 0.1 | 0.15 | 0.14 | 0.13 | 1694.164 | 0.13 | < 10 | 0.14 | 0.1 | 0.13 | 0.1 | 0.17 | 0.124 | 0.11 | 0.13 | 0.13 | 0.11 | [5 empty] | 0.13068
| 1 | 5 | 0.12 | 0.134 | 0.152 | 0.10644 | 0.16947 | 0.2 | 0.27 | 0.17 | 0.14 | 0.1 | 0.11 | 0.13 | 0.13 | 1622.764 | 0.12 | < 10 | 0.13 | 0.1 | 0.13 | < 0.1 | 0.12 | 0.134 | 0.12 | 0.13 | 0.13 | 0.11 | [5 empty] | 126
| 1 | 6 | 0.12 | 0.133 | 0.155 | 0.11285 | 0.152 | 0.2 | 0.28 | 0.17 | 0.14 | 0.1 | 0.13 | 0.13 | 0.13 | 1681.553 | 0.13 | < 10 | 0.14 | 0.1 | 0.13 | 0.1 | 0.13 | 0.157 | 0.11 | 0.13 | 0.13 | 0.11
| Mean |  | 0.12333 | 0.134 | 0.152 | 0.10732 | 0.16195 | 0.15 | 0.26833 | 0.17167 | 0.14167 | 0.1 | 0.13 | 0.135 | 0.12833 | 1594.5815 | 0.12667 |  | 0.13167 | 0.1 | 0.13167 | 0.1 | 0.13333 | 0.14467 | 0.12333 | 0.12833 | 0.13 | 0.11
| Median |  | 0.12 | 0.1335 | 0.1525 | 0.10641 | 0.15978 | 0.15 | 0.27 | 0.17 | 0.14 | 0.1 | 0.13 | 0.135 | 0.13 | 1589.754 | 0.125 |  | 0.13 | 0.1 | 0.13 | 0.1 | 0.13 | 0.1365 | 0.125 | 0.13 | 0.13 | 0.11
| Std Dev. |  | 0.008165 | 0.0012649 | 0.0031623 | 0.0033842 | 0.012766 | 0.054772 | 0.0075277 | 0.0040825 | 0.0040825 | 1.5202e-17 | 0.012649 | 0.0054772 | 0.0075277 | 87.1685 | 0.008165 |  | 0.0075277 | 1.5202e-17 | 0.0040825 | 0 | 0.018619 | 0.020393 | 0.012111 | 0.0040825 | 0 | 0
| Rel.Std.Dev. |  | 6.62% | 0.944% | 2.08% | 3.153% | 7.882% | 36.51% | 2.805% | 2.378% | 2.882% | 1.52e-14% | 9.73% | 4.057% | 5.866% | 5.467% | 6.446% |  | 5.717% | 1.52e-14% | 3.101% | 0% | 13.96% | 14.1% | 9.819% | 3.181% | 0% | 0%
| PDM3 |  | -5.618% | 2.544% | 16.32% | -17.87% | 23.94% | 14.79% | 105.3% | 31.37% | 8.411% | -23.47% | -0.5165% | 3.31% | -1.792% | 1220165% | -3.067% |  | 0.7589% | -23.47% | 0.7589% | -23.47% | 2.034% | 10.71% | -5.618% | -1.792% | -0.5165% | -15.82%
| Z-Score (Absolute) |  | 0.67 | 0.11 | 1.43 | 1.85 | 2.17 | NA | 9.99 | 2.88 | 0.67 | NA | 0.18 | 0.18 | 0.31 | 117289.2 | 0.43 | 358.06 | 0.06 | NA | 0.06 | NA | 0.06 | 0.89 | 0.67 | 0.31 | 0.18 | 1.66
| Analytical results for V in OREAS 231b (Certified Value 141 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 157 | 144.2 | 136.84 | 160.578 | 156.3221 | 159 | 176 | 191 | 180 | 145 | 143 | 142 | 145 | 133 | < 10 | 138 | 123 | 174.614 | 151 | 123 | 141 | 117.45 | 153.4376 | 100 | 135 | 126 | 131 | 140 | 106 |  |  | 1
| 1 | 2 | 142 | 148.6 | 120.1 | 160.6927 | 158.5641 | 156 | 178 | 187 | 175 | 145 | 143 | 142 | 145 | 134 | < 10 | 138 | 124 | 160.5245 | 157 | 128 | 145 | 117.27 | 143.5292 | 100 | 136 | 124 | 131 | 142 | 105 |  |  | 30
| 1 | 3 | 140 | 147.8 | 122.22 | 160.2613 | 160.1714 | 160 | 177 | 189 | 178 | 146 | 146 | 139 | 144 | 135 | < 10 | 143 | 122 | 176.1655 | 153 | 119 | 144 | 123.28 | 149.0952 | 95 | 132 | 122 | 131 | 141 | 99.9 |  |  | 16
| 1 | 4 | 137 | 144.2 | 131.52 | 162.031 | 156.7908 | 155 | 183 | 190 | 167 | 149 | 149 | 141 | 145 | 132 | < 10 | 137 | 120 | 168.6866 | 155 | 116 | 142 | 114.95 | 150.7328 | 100 | 136 | 114 | 130 | 137 | 95.9 |  |  | 140.6777
| 1 | 5 | 134 | 147.5 | 123.91 | 161.6124 | 170.4396 | 150 | 181 | 191 | 172 | 146 | 152 | 139 | 145 | 137 | < 10 | 140 | 128 | 182.9793 | 157 | 116 | 142 | 109.57 | 146.5284 | 95 | 137 | 110 | 132 | 138 | 107 |  |  | 127
| 1 | 6 | 132 | 145.3 | 139.93 | 154.684 | 154.8169 | 155 | 185 | 188 | 180 | 146 | 149 | 140 | 144 | 137 | < 10 | 140 | 122 | 184.9422 | 153 | 117 | 147 | 111.61 | 139.2972 | 95 | 136 | 114 | 131 | 141 | 109
| Mean |  | 140.3333 | 146.2667 | 129.0867 | 159.9766 | 159.5175 | 155.8333 | 180 | 189.3333 | 175.3333 | 146.1667 | 147 | 140.5 | 144.6667 | 134.6667 |  | 139.3333 | 123.1667 | 174.652 | 154.3333 | 119.8333 | 143.5 | 115.6883 | 147.1034 | 97.5 | 135.3333 | 118.3333 | 131 | 139.8333 | 103.8
| Median |  | 138.5 | 146.4 | 127.715 | 160.6354 | 157.6775 | 155.5 | 179.5 | 189.5 | 176.5 | 146 | 147.5 | 140.5 | 145 | 134.5 |  | 139 | 122.5 | 175.3898 | 154 | 118 | 143 | 116.11 | 147.8118 | 97.5 | 136 | 118 | 131 | 140.5 | 105.5
| Std Dev. |  | 8.9592 | 1.9387 | 8.2258 | 2.6782 | 5.6634 | 3.5449 | 3.5777 | 1.633 | 5.1251 | 1.472 | 3.6332 | 1.3784 | 0.5164 | 2.0656 |  | 2.1602 | 2.7142 | 9.0851 | 2.4221 | 4.7924 | 2.2583 | 4.8549 | 5.1245 | 2.7386 | 1.7512 | 6.5013 | 0.63246 | 1.9408 | 4.923
| Rel.Std.Dev. |  | 6.384% | 1.325% | 6.372% | 1.674% | 3.55% | 2.275% | 1.988% | 0.8625% | 2.923% | 1.007% | 2.472% | 0.9811% | 0.357% | 1.534% |  | 1.55% | 2.204% | 5.202% | 1.569% | 3.999% | 1.574% | 4.197% | 3.484% | 2.809% | 1.294% | 5.494% | 0.4828% | 1.388% | 4.743%
| PDM3 |  | -0.2448% | 3.973% | -8.239% | 13.72% | 13.39% | 10.77% | 27.95% | 34.59% | 24.63% | 3.902% | 4.494% | -0.1263% | 2.836% | -4.273% |  | -0.9557% | -12.45% | 24.15% | 9.707% | -14.82% | 2.006% | -17.76% | 4.568% | -30.69% | -3.799% | -15.88% | -6.879% | -0.6002% | -26.21%
| Z-Score (Absolute) |  | 0.01 | 0.28 | 0.56 | 0.95 | 0.93 | 0.75 | 1.93 | 2.38 | 1.7 | 0.28 | 0.32 | 0 | 0.2 | 0.28 | 6.6 | 0.06 | 0.84 | 1.66 | 0.67 | 1.01 | 0.15 | 1.21 | 0.32 | 2.1 | 0.25 | 1.08 | 0.46 | 0.03 | 1.79
| Analytical results for W in OREAS 231b (Certified Value 11.9 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 12.8 | 10.39 | 16.48 | 14.4334 | 10.1 | 9.3 | 9 | 9.12 | 10 | 12.9 | 12.3 | 12.15 | 11.15 | < 100 | 12.3 | < 10 | 11.6 | 8.5 | 13.5 | 9 | 12.35 | 13.8 | 13.5 | 10.67 | 14.2 | [6 empty] | 1
| 1 | 2 | 13 | 10.68 | 15.72 | 14.1074 | 10.4 | 9.5 | 8.5 | 8.97 | 10 | 12.9 | 12.5 | 11.5 | 11.6 | < 100 | 11.6 | < 10 | 11.8 | 8.8 | 13.1 | 10 | 11.9 | 13.9 | 12.9 | 10.4 | 14.6 | [6 empty] | 31
| 1 | 3 | 13.1 | 10.75 | 15.04 | 14.0917 | 10 | 9.1 | 8.8 | 9 | 10 | 12.5 | 12.35 | 12.4 | 12.7 | < 100 | 12.3 | < 10 | 11.1 | 8.3 | 13.4 | 10 | 11.8 | 14.2 | 13 | 10.7 | 14.7 | [6 empty] | 16
| 1 | 4 | 13.4 | 10.31 | 15.61 | 13.7731 | 10.3 | 9.2 | 8.2 | 9.13 | 10 | 12.5 | 12.45 | 11.9 | 12.1 | < 100 | 11.95 | < 10 | 11.2 | 9 | 13.7 | 10 | 12.65 | 13.8 | 13.5 | 10.27 | 14.9 | [6 empty] | 11.8727
| 1 | 5 | 13.4 | 10.41 | 14.16 | 15.6303 | 10 | 9.4 | 9 | 9 | 10 | 12.6 | 12.1 | 11.85 | 13.4 | < 100 | 11.55 | < 10 | 11.9 | 8.7 | 12.5 | 10 | 12.25 | 13.8 | 13.2 | 10.59 | 14.4 | [6 empty] | 128
| 1 | 6 | 13.1 | 10.62 | 15.62 | 14.2996 | 9.7 | 9.3 | 8.6 | 8.63 | 10 | 12.2 | 12.3 | 12.25 | 12.45 | < 100 | 12.35 | < 10 | 11.2 | 8.5 | 13.2 | 10 | 12.05 | 14.4 | 12.7 | 10.73 | 14.9
| Mean |  | 13.1333 | 10.5267 | 15.4383 | 14.3892 | 10.0833 | 9.3 | 8.6833 | 8.975 | 10 | 12.6 | 12.3333 | 12.0083 | 12.2333 |  | 12.0083 |  | 11.4667 | 8.6333 | 13.2333 | 9.8333 | 12.1667 | 13.9833 | 13.1333 | 10.56 | 14.6167
| Median |  | 13.1 | 10.515 | 15.615 | 14.2035 | 10.05 | 9.3 | 8.7 | 9 | 10 | 12.55 | 12.325 | 12.025 | 12.275 |  | 12.125 |  | 11.4 | 8.6 | 13.3 | 10 | 12.15 | 13.85 | 13.1 | 10.63 | 14.65
| Std Dev. |  | 0.23381 | 0.17963 | 0.77713 | 0.64772 | 0.24833 | 0.14142 | 0.31252 | 0.18185 | 0 | 0.26833 | 0.14024 | 0.32468 | 0.80229 |  | 0.36526 |  | 0.34448 | 0.25033 | 0.41793 | 0.40825 | 0.31411 | 0.25626 | 0.3266 | 0.18504 | 0.27869
| Rel.Std.Dev. |  | 1.78% | 1.706% | 5.034% | 4.501% | 2.463% | 1.521% | 3.599% | 2.026% | 0% | 2.13% | 1.137% | 2.704% | 6.558% |  | 3.042% |  | 3.004% | 2.9% | 3.158% | 4.152% | 2.582% | 1.833% | 2.487% | 1.752% | 1.907%
| PDM3 |  | 10.62% | -11.34% | 30.03% | 21.2% | -15.07% | -21.67% | -26.86% | -24.41% | -15.77% | 6.126% | 3.88% | 1.142% | 3.037% |  | 1.142% |  | -3.42% | -27.28% | 11.46% | -17.18% | 2.476% | 17.78% | 10.62% | -11.06% | 23.11%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 1.35 | 0.91 | 0.86 | 1.18 | 1.43 | 1.31 | NA | 0.18 | 0.07 | 0.07 | 0.03 | 15.56 | 0.07 | 2.95 | 0.29 | 1.45 | 0.44 | NA | 0 | 0.75 | 0.4 | 0.66 | 1.01
Results from laboratories 1.12 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 231b (Certified Value 10.8 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 11.05 | 11.31 | 10.75 | 10.6635 | 11.2546 | 11.8 | 11 | 14.6 | 11.09 | 10 | 10.8 | 10.7 | 10.65 | 9.9 | 10.85 | 10 | 11.28 | 10.7 | 9.1 | 11.45 | 9.78 | 10.3 | 11.33 | 9.45 | [7 empty] | 1
| 1 | 2 | 10.74 | 11.65 | 9.87 | 10.7664 | 11.3034 | 11.5 | 11 | 14.4 | 10.64 | 10 | 10.9 | 10.75 | 10.9 | 9.75 | 10.5 | 9 | 11.91 | 10.5 | 9.5 | 11.6 | 9.76 | 9.9 | 11.41 | 9.41 | [7 empty] | 32
| 1 | 3 | 10.48 | 11.43 | 10.05 | 10.7552 | 11.2914 | 11.7 | 11.3 | 14.2 | 10.93 | 10 | 10.9 | 10.65 | 10.65 | 10.7 | 11.2 | 9 | 11.38 | 10.6 | 8.9 | 11.55 | 9.75 | 10 | 11.54 | 9.31 | [7 empty] | 16
| 1 | 4 | 10.47 | 11.23 | 10.64 | 10.7739 | 11.1075 | 11.3 | 11.3 | 14.3 | 10.35 | 10 | 11.1 | 10.25 | 10.7 | 10.95 | 10.95 | 9 | 11.82 | 10.7 | 9.2 | 11.65 | 9.46 | 10.4 | 11.04 | 9.08 | [7 empty] | 10.777
| 1 | 5 | 10.32 | 11.5 | 9.98 | 10.7565 | 12.1676 | 11.3 | 11.3 | 14.4 | 10.61 | 9 | 11.3 | 10.45 | 10.7 | 10.75 | 10.85 | 10 | 11.76 | 10.8 | 8.6 | 11.5 | 9.37 | 10.2 | 11.17 | 9.51 | [7 empty] | 129
| 1 | 6 | 10.13 | 11.22 | 10.95 | 10.6507 | 11.0697 | 11.3 | 11.2 | 14.6 | 11.06 | 9 | 11 | 10.45 | 11.1 | 10.5 | 11 | 9 | 11.55 | 10.7 | 8.8 | 11.65 | 9.36 | 9.8 | 11.61 | 9.68
| Mean |  | 10.5317 | 11.39 | 10.3733 | 10.7277 | 11.3657 | 11.4833 | 11.1833 | 14.4167 | 10.78 | 9.6667 | 11 | 10.5417 | 10.7833 | 10.425 | 10.8917 | 9.3333 | 11.6167 | 10.6667 | 9.0167 | 11.5667 | 9.58 | 10.1 | 11.35 | 9.4067
| Median |  | 10.475 | 11.37 | 10.345 | 10.7559 | 11.273 | 11.4 | 11.25 | 14.4 | 10.785 | 10 | 10.95 | 10.55 | 10.7 | 10.6 | 10.9 | 9 | 11.655 | 10.7 | 9 | 11.575 | 9.605 | 10.1 | 11.37 | 9.43
| Std Dev. |  | 0.32394 | 0.16864 | 0.46003 | 0.055273 | 0.40475 | 0.22286 | 0.1472 | 0.16021 | 0.29339 | 0.5164 | 0.17889 | 0.19083 | 0.18074 | 0.48862 | 0.23112 | 0.5164 | 0.25367 | 0.10328 | 0.31885 | 0.08165 | 0.20406 | 0.23664 | 0.21735 | 0.20166
| Rel.Std.Dev. |  | 3.076% | 1.481% | 4.435% | 0.5152% | 3.561% | 1.941% | 1.316% | 1.111% | 2.722% | 5.342% | 1.626% | 1.81% | 1.676% | 4.687% | 2.122% | 5.533% | 2.184% | 0.9682% | 3.536% | 0.7059% | 2.13% | 2.343% | 1.915% | 2.144%
| PDM3 |  | -2.276% | 5.688% | -3.745% | -0.4573% | 5.463% | 6.554% | 3.771% | 33.77% | 0.02799% | -10.3% | 2.069% | -2.184% | 0.05892% | -3.266% | 1.064% | -13.4% | 7.791% | -1.024% | -16.33% | 7.327% | -11.11% | -6.282% | 5.317% | -12.72%
| Z-Score (Absolute) |  | 0.35 | 0.84 | 0.56 | 0.07 | 0.81 | 0.97 | 0.55 | 5.02 | 0 | NA | 0.3 | 0.33 | 0 | 0.49 | 0.15 | NA | 1.15 | 0.16 | 2.44 | 1.08 | 1.66 | 0.94 | 0.78 | 1.9
Results from laboratories 1.12 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 231b (Certified Value 0.98 ppm) | [32 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 2
| 1 | 1 | 1 | 1.069 | 0.88 | 0.9708 | 1.1 | 1.23 | 0.8 | 0.9 | [23 empty] | 1
| 1 | 2 | 0.9 | 1.06 | 0.865 | 0.95943 | 1.1 | 1.17 | 0.85 | 1 | [23 empty] | 33
| 1 | 3 | 0.9 | 1.064 | 0.83 | 0.96806 | 1.1 | 1.19 | 0.8 | 0.9 | [23 empty] | 16
| 1 | 4 | 1 | 1.081 | 0.851 | 0.95878 | 1.1 | 1.23 | 0.8 | 0.8 | [23 empty] | 0.98121
| 1 | 5 | 0.9 | 1.097 | 0.897 | 0.98394 | 1.1 | 1.26 | 0.75 | 0.9 | [23 empty] | 130
| 1 | 6 | 0.9 | 1.09 | 0.84 | 0.95324 | 1.1 | 1.25 | 0.75 | 0.9
| Mean |  | 0.93333 | 1.0768 | 0.8605 | 0.96571 | 1.1 | 1.2217 | 0.79167 | 0.9
| Median |  | 0.9 | 1.075 | 0.858 | 0.96375 | 1.1 | 1.23 | 0.8 | 0.9
| Std Dev. |  | 0.05164 | 0.014878 | 0.025193 | 0.011014 | 2.4324e-16 | 0.034881 | 0.037639 | 0.063246
| Rel.Std.Dev. |  | 5.533% | 1.382% | 2.928% | 1.14% | 2.211e-14% | 2.855% | 4.754% | 7.027%
| PDM3 |  | -4.88% | 9.745% | -12.3% | -1.58% | 12.11% | 24.51% | -19.32% | -8.277%
| Z-Score (Absolute) |  | 0.1 | 0.79 | 0.55 | 0.1 | 0.94 | 1.7 | 0.98 | 0.31
| Analytical results for Zn in OREAS 231b (Certified Value 102 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 105 | 106.2 | 95.14 | 102.16 | 107.5623 | 103 | 104 | 107 | 105 | 99 | 112 | 107 | 99 | 105 | < 100 | 112 | 98 | 97.5641 | 104 | 105.3 | 101 | 95.61 | 100.08 | 95 | 108.6236 | 106 | 93.625 | 97.4 | 110 | 96 | 105 | 1
| 1 | 2 | 105 | 104.3 | 91.32 | 103.2 | 107.3478 | 102 | 101 | 105 | 102 | 99 | 112 | 105 | 101 | 104 | < 100 | 113 | 94 | 98.5462 | 105 | 101.8 | 99 | 96.57 | 98.08 | 100 | 111.6143 | 106 | 90.552 | 97.3 | 106 | 98 | 103 | 34
| 1 | 3 | 104 | 103.3 | 87.57 | 99.6 | 108.4406 | 106 | 102 | 107 | 103 | 98 | 112 | 105 | 101 | 104 | < 100 | 114 | 92 | 96.1503 | 102 | 101.4 | 99 | 95.89 | 97.75 | 95 | 112.2249 | 106 | 90.262 | 99 | 107 | 97 | 103 | 16
| 1 | 4 | 104 | 101.9 | 93.35 | 100.18 | 105.7859 | 100 | 102 | 107 | 100 | 97 | 115 | 106 | 100 | 104 | < 100 | 112 | 92 | 98.0939 | 105 | 100.7 | 99 | 96.14 | 98.17 | 100 | 111.6262 | 106 | 91.357 | 96.3 | 112 | 96 | 103 | 102.0616
| 1 | 5 | 104 | 102.7 | 85.76 | 101.8 | 105.8953 | 99.6 | 107 | 107 | 102 | 98 | 115 | 104 | 102 | 105 | < 100 | 113 | 95 | 99.9215 | 105 | 98.9 | 100 | 96.29 | 97.43 | 95 | 114.9506 | 106 | 91.482 | 98.3 | 109 | 95 | 104 | 131
| 1 | 6 | 103 | 102.4 | 93.42 | 101 | 104.2789 | 103 | 107 | 108 | 104 | 99 | 113 | 106 | 99 | 105 | < 100 | 114 | 96 | 101.5199 | 102 | 102.2 | 98 | 98.25 | 97.78 | 95 | 110.3433 | 106 | 92.788 | 99 | 105 | 96 | 105
| Mean |  | 104.1667 | 103.4667 | 91.0933 | 101.3233 | 106.5518 | 102.2667 | 103.8333 | 106.8333 | 102.6667 | 98.3333 | 113.1667 | 105.5 | 100.3333 | 104.5 |  | 113 | 94.5 | 98.6326 | 103.8333 | 101.7167 | 99.3333 | 96.4583 | 98.215 | 96.6667 | 111.5638 | 106 | 91.6777 | 97.8833 | 108.1667 | 96.3333 | 103.8333
| Median |  | 104 | 103 | 92.335 | 101.4 | 106.6215 | 102.5 | 103 | 107 | 102.5 | 98.5 | 112.5 | 105.5 | 100.5 | 104.5 |  | 113 | 94.5 | 98.32 | 104.5 | 101.6 | 99 | 96.215 | 97.93 | 95 | 111.6202 | 106 | 91.4195 | 97.85 | 108 | 96 | 103.5
| Std Dev. |  | 0.75277 | 1.5731 | 3.6821 | 1.329 | 1.5104 | 2.3381 | 2.6394 | 0.98319 | 1.7512 | 0.8165 | 1.472 | 1.0488 | 1.2111 | 0.54772 |  | 0.89443 | 2.3452 | 1.8761 | 1.472 | 2.1047 | 1.0328 | 0.93754 | 0.9508 | 2.582 | 2.0999 | 0 | 1.299 | 1.0722 | 2.6394 | 1.0328 | 0.98319
| Rel.Std.Dev. |  | 0.7227% | 1.52% | 4.042% | 1.312% | 1.417% | 2.286% | 2.542% | 0.9203% | 1.706% | 0.8303% | 1.301% | 0.9941% | 1.207% | 0.5241% |  | 0.7915% | 2.482% | 1.902% | 1.418% | 2.069% | 1.04% | 0.972% | 0.9681% | 2.671% | 1.882% | 0% | 1.417% | 1.095% | 2.44% | 1.072% | 0.9469%
| PDM3 |  | 2.063% | 1.377% | -10.75% | -0.7234% | 4.399% | 0.2009% | 1.736% | 4.675% | 0.5928% | -3.653% | 10.88% | 3.369% | -1.693% | 2.389% |  | 10.72% | -7.409% | -3.36% | 1.736% | -0.338% | -2.673% | -5.49% | -3.769% | -5.286% | 9.31% | 3.859% | -10.17% | -4.094% | 5.982% | -5.613% | 1.736%
| Z-Score (Absolute) |  | 0.33 | 0.21 | 1.92 | 0.16 | 0.73 | 0 | 0.27 | 0.78 | 0.07 | 0.67 | 1.87 | 0.55 | 0.33 | 0.38 | 8.96 | 1.84 | 1.33 | 0.62 | 0.27 | 0.09 | 0.5 | 1 | 0.69 | 0.96 | 1.59 | 0.64 | 1.82 | 0.75 | 1.01 | 1.02 | 0.27
| Analytical results for Zr in OREAS 231b (Certified Value 12.1 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 12.3 | 12.63 | 18.22 | 13.5664 | 10.8 | 11.9 | 14 | 16.5 | 11 | 10.1 | 12.2 | 13.3 | 11 | 12.9 | 11 | 11.8 | 10.8 | 16.002 | 7 | 13.2 | 9.3 | 11.8 | 13 | 9.3 | [7 empty] | 1
| 1 | 2 | 9.8 | 12.85 | 16.28 | 13.6955 | 11.9 | 10.8 | 14.7 | 15.8 | 12 | 10.1 | 12.2 | 13 | 11.4 | 12.1 | 11 | 10.7 | 10.7 | 14.679 | 7 | 13.5 | 10.1 | 11.2 | 12.5 | 9.3 | [7 empty] | 35
| 1 | 3 | 10.3 | 12.5 | 16.46 | 13.8475 | 12.2 | 11.3 | 14.8 | 16.1 | 12 | 10.2 | 11.5 | 13.7 | 12.2 | 13.2 | 11 | 11.7 | 10.7 | 15.453 | 7 | 12.7 | 10.1 | 11.7 | 13.1 | 8.8 | [7 empty] | 16
| 1 | 4 | 9.9 | 12.13 | 17.64 | 13.4272 | 11.2 | 10.8 | 14.1 | 15 | 11 | 9.9 | 11.5 | 13.4 | 11.8 | 12.8 | 11 | 11.9 | 11 | 15.489 | 7 | 13.7 | 9.7 | 10.6 | 12.4 | 8.4 | [7 empty] | 12.1224
| 1 | 5 | 9.6 | 12.86 | 16.74 | 14.7398 | 10.9 | 10.6 | 14.7 | 15.8 | 12 | 10.5 | 11.4 | 13.2 | 12.7 | 12.6 | 11 | 10.4 | 11.2 | 14.85 | 6 | 13.4 | 9.5 | 11.7 | 12.8 | 9.8 | [7 empty] | 132
| 1 | 6 | 9.6 | 12.63 | 18.82 | 13.5059 | 10.7 | 11.2 | 14.3 | 16.5 | 11 | 10 | 11.3 | 14.1 | 12 | 12.9 | 11 | 11.1 | 11 | 14.13 | 6 | 13.2 | 9.5 | 11.3 | 13.4 | 10
| Mean |  | 10.25 | 12.6 | 17.36 | 13.797 | 11.2833 | 11.1 | 14.4333 | 15.95 | 11.5 | 10.1333 | 11.6833 | 13.45 | 11.85 | 12.75 | 11 | 11.2667 | 10.9 | 15.1005 | 6.6667 | 13.2833 | 9.7 | 11.3833 | 12.8667 | 9.2667
| Median |  | 9.85 | 12.63 | 17.19 | 13.6309 | 11.05 | 11 | 14.5 | 15.95 | 11.5 | 10.1 | 11.5 | 13.35 | 11.9 | 12.85 | 11 | 11.4 | 10.9 | 15.1515 | 7 | 13.3 | 9.6 | 11.5 | 12.9 | 9.3
| Std Dev. |  | 1.0368 | 0.26937 | 1.0306 | 0.48508 | 0.62423 | 0.47329 | 0.34448 | 0.56125 | 0.54772 | 0.20656 | 0.40702 | 0.3937 | 0.59917 | 0.37283 | 0 | 0.62823 | 0.2 | 0.67382 | 0.5164 | 0.34303 | 0.33466 | 0.4535 | 0.37771 | 0.59889
| Rel.Std.Dev. |  | 10.12% | 2.138% | 5.936% | 3.516% | 5.532% | 4.264% | 2.387% | 3.519% | 4.763% | 2.038% | 3.484% | 2.927% | 5.056% | 2.924% | 0% | 5.576% | 1.835% | 4.462% | 7.746% | 2.582% | 3.45% | 3.984% | 2.936% | 6.463%
| PDM3 |  | -15.45% | 3.939% | 43.21% | 13.81% | -6.922% | -8.434% | 19.06% | 31.57% | -5.135% | -16.41% | -3.622% | 10.95% | -2.247% | 5.177% | -9.259% | -7.059% | -10.08% | 24.57% | -45.01% | 9.576% | -19.98% | -6.097% | 6.139% | -23.56%
| Z-Score (Absolute) |  | 0.75 | 0.35 | 2.59 | 0.92 | 0.27 | 0.35 | 1.22 | 1.93 | NA | 0.81 | 0.08 | 0.75 | 0 | 0.42 | NA | 0.27 | 0.45 | 1.53 | NA | 0.67 | 1.01 | 0.22 | 0.48 | 1.22
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C568: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1149: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1204: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 231b (Indicative Value 13.15 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.15 | 1
| 1 | 2 | 13.14 | 10
| Mean |  | 13.145 | 16
| Median |  | 13.145 | 13.145
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.05379%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 231b (Indicative Value 9.38 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 9.38 | 1
| 1 | 2 | 9.37 | 11
| Mean |  | 9.375 | 16
| Median |  | 9.375 | 9.375
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.07542%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 231b (Indicative Value 12.31 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.3 | 1
| 1 | 2 | 12.31 | 12
| Mean |  | 12.305 | 16
| Median |  | 12.305 | 12.305
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.05746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 231b (Indicative Value 0.55 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.551 | 1
| 1 | 2 | 0.549 | 13
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0.0014142 | 19
| Rel.Std.Dev. |  | 0.2571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 231b (Indicative Value 7.05 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 7.04 | 1
| 1 | 2 | 7.05 | 10
| Mean |  | 7.045 | 16
| Median |  | 7.045 | 7.045
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.1004%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 231b (Indicative Value 0.192 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.192 | 1
| 1 | 2 | 0.192 | 11
| Mean |  | 0.192 | 16
| Median |  | 0.192 | 0.192
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 231b (Indicative Value 2.74 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.73 | 1
| 1 | 2 | 2.74 | 12
| Mean |  | 2.735 | 16
| Median |  | 2.735 | 2.735
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.2585%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 231b (Indicative Value 0.097 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.097 | 1
| 1 | 2 | 0.096 | 13
| Mean |  | 0.0965 | 16
| Median |  | 0.0965 | 0.0965
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 231b (Indicative Value 0.219 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.219 | 1
| 1 | 2 | 0.219 | 17
| Mean |  | 0.219 | 16
| Median |  | 0.219 | 0.21905
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.02107%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 231b (Indicative Value 50.17 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 50.16 | 1
| 1 | 2 | 50.18 | 11
| Mean |  | 50.17 | 16
| Median |  | 50.17 | 50.17
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.02819%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 231b (Indicative Value 1.11 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.113 | 1
| 1 | 2 | 1.112 | 12
| Mean |  | 1.1125 | 16
| Median |  | 1.1125 | 1.1125
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.06356%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 231b (Indicative Value 2.98 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.98 | 1
| 1 | 2 | 2.97 | 15
| Mean |  | 2.975 | 16
| Median |  | 2.975 | 2.975
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.2377%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 231b (Indicative Value 0.1 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 17
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.1
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 231b (Indicative Value 0.245 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.24 | 17
| Mean |  | 0.245 | 16
| Median |  | 0.245 | 0.245
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 231b (Indicative Value 0.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 19
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 231b (Indicative Value 33.2 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.2 | 1
| 1 | 2 | 33.2 | 20
| Mean |  | 33.2 | 16
| Median |  | 33.2 | 33.2
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 231b (Indicative Value 295 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 294 | 1
| 1 | 2 | 295 | 21
| Mean |  | 294.5 | 16
| Median |  | 294.5 | 294.5
| Std Dev. |  | 0.70711 | 27
| Rel.Std.Dev. |  | 0.2401%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 231b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 22
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 231b (Indicative Value 0.04 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.04 | 23
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 231b (Indicative Value 0.4 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 24
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 231b (Indicative Value 11 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.1 | 1
| 1 | 2 | 10.9 | 25
| Mean |  | 11 | 16
| Median |  | 11 | 11
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 1.286%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 231b (Indicative Value 47.8 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47.6 | 1
| 1 | 2 | 48 | 26
| Mean |  | 47.8 | 16
| Median |  | 47.8 | 47.8
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 0.5917%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 231b (Indicative Value 190 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 193 | 1
| 1 | 2 | 186 | 27
| Mean |  | 189.5 | 16
| Median |  | 189.5 | 189.5
| Std Dev. |  | 4.9497 | 33
| Rel.Std.Dev. |  | 2.612%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 231b (Indicative Value 0.55 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.53 | 1
| 1 | 2 | 0.56 | 28
| Mean |  | 0.545 | 16
| Median |  | 0.545 | 0.545
| Std Dev. |  | 0.021213 | 34
| Rel.Std.Dev. |  | 3.892%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 231b (Indicative Value 160 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 156 | 29
| Mean |  | 160 | 16
| Median |  | 160 | 160
| Std Dev. |  | 5.6569 | 35
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 231b (Indicative Value 3.86 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.9 | 1
| 1 | 2 | 3.82 | 30
| Mean |  | 3.86 | 16
| Median |  | 3.86 | 3.86
| Std Dev. |  | 0.056569 | 36
| Rel.Std.Dev. |  | 1.466%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 231b (Indicative Value 2.46 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.37 | 1
| 1 | 2 | 2.54 | 31
| Mean |  | 2.455 | 16
| Median |  | 2.455 | 2.455
| Std Dev. |  | 0.12021 | 37
| Rel.Std.Dev. |  | 4.896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 231b (Indicative Value 0.89 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.87 | 32
| Mean |  | 0.885 | 16
| Median |  | 0.885 | 0.885
| Std Dev. |  | 0.021213 | 38
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 231b (Indicative Value 15.1 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.1 | 1
| 1 | 2 | 15 | 33
| Mean |  | 15.05 | 16
| Median |  | 15.05 | 15.05
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.4698%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 231b (Indicative Value 3.21 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 3.22 | 34
| Mean |  | 3.21 | 16
| Median |  | 3.21 | 3.21
| Std Dev. |  | 0.014142 | 40
| Rel.Std.Dev. |  | 0.4406%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 231b (Indicative Value 1.53 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.45 | 1
| 1 | 2 | 1.6 | 4
| Mean |  | 1.525 | 16
| Median |  | 1.525 | 1.525
| Std Dev. |  | 0.10607 | 41
| Rel.Std.Dev. |  | 6.955%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 231b (Indicative Value 1.88 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.84 | 1
| 1 | 2 | 1.92 | 19
| Mean |  | 1.88 | 16
| Median |  | 1.88 | 1.88
| Std Dev. |  | 0.056569 | 25
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 231b (Indicative Value 0.84 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.83 | 20
| Mean |  | 0.835 | 16
| Median |  | 0.835 | 0.835
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 231b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 21
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 231b (Indicative Value 4.87 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.09 | 1
| 1 | 2 | 4.64 | 22
| Mean |  | 4.865 | 16
| Median |  | 4.865 | 4.865
| Std Dev. |  | 0.3182 | 28
| Rel.Std.Dev. |  | 6.541%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 231b (Indicative Value 0.35 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.35 | 23
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 231b (Indicative Value 0.168 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.168 | 1
| 1 | 2 | 0.168 | 24
| Mean |  | 0.168 | 16
| Median |  | 0.168 | 0.168
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 231b (Indicative Value 1.3 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.2 | 25
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 231b (Indicative Value 3.59 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.66 | 1
| 1 | 2 | 3.52 | 8
| Mean |  | 3.59 | 16
| Median |  | 3.59 | 3.59
| Std Dev. |  | 0.098995 | 32
| Rel.Std.Dev. |  | 2.758%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 231b (Indicative Value 8.24 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.27 | 1
| 1 | 2 | 8.21 | 27
| Mean |  | 8.24 | 16
| Median |  | 8.24 | 8.24
| Std Dev. |  | 0.042426 | 33
| Rel.Std.Dev. |  | 0.5149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 231b (Indicative Value 115 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 120 | 1
| 1 | 2 | 110 | 28
| Mean |  | 115 | 16
| Median |  | 115 | 115
| Std Dev. |  | 7.0711 | 34
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 231b (Indicative Value 13.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13 | 1
| 1 | 2 | 14 | 29
| Mean |  | 13.5 | 16
| Median |  | 13.5 | 13.5
| Std Dev. |  | 0.70711 | 35
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 231b (Indicative Value 1.74 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.73 | 1
| 1 | 2 | 1.74 | 30
| Mean |  | 1.735 | 16
| Median |  | 1.735 | 1.735
| Std Dev. |  | 0.0070711 | 36
| Rel.Std.Dev. |  | 0.4076%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 231b (Indicative Value 9.9 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 10 | 1
| 1 | 2 | 9.8 | 31
| Mean |  | 9.9 | 16
| Median |  | 9.9 | 9.9
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 231b (Indicative Value 0.018 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.03 | 1
| 1 | 2 | < 0.01 | 32
| Mean |  | 0.03 | 16
| Median |  | 0.03 | 0.0175
| Std Dev. |  |  | 38
Rel.Std.Dev.
| PDM3 |  | 71.43%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 231b (Indicative Value 0.8 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.8 | 33
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 231b (Indicative Value 41.3 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.9 | 1
| 1 | 2 | 41.6 | 34
| Mean |  | 41.25 | 16
| Median |  | 41.25 | 41.25
| Std Dev. |  | 0.49497 | 40
| Rel.Std.Dev. |  | 1.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 231b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 4
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 231b (Indicative Value 2.61 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.59 | 1
| 1 | 2 | 2.62 | 19
| Mean |  | 2.605 | 16
| Median |  | 2.605 | 2.605
| Std Dev. |  | 0.021213 | 25
| Rel.Std.Dev. |  | 0.8143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 231b (Indicative Value 0.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.8 | 20
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 231b (Indicative Value 197 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 197 | 1
| 1 | 2 | 196 | 21
| Mean |  | 196.5 | 16
| Median |  | 196.5 | 196.5
| Std Dev. |  | 0.70711 | 27
| Rel.Std.Dev. |  | 0.3599%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 231b (Indicative Value 0.21 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.19 | 1
| 1 | 2 | 0.22 | 22
| Mean |  | 0.205 | 16
| Median |  | 0.205 | 0.205
| Std Dev. |  | 0.021213 | 28
| Rel.Std.Dev. |  | 10.35%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 231b (Indicative Value 0.58 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.56 | 23
| Mean |  | 0.575 | 16
| Median |  | 0.575 | 0.575
| Std Dev. |  | 0.021213 | 29
| Rel.Std.Dev. |  | 3.689%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 231b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 231b (Indicative Value 0.65 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.68 | 1
| 1 | 2 | 0.61 | 25
| Mean |  | 0.645 | 16
| Median |  | 0.645 | 0.645
| Std Dev. |  | 0.049497 | 31
| Rel.Std.Dev. |  | 7.674%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 231b (Indicative Value 0.667 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.67 | 1
| 1 | 2 | 0.663 | 26
| Mean |  | 0.6665 | 16
| Median |  | 0.6665 | 0.6665
| Std Dev. |  | 0.0049497 | 32
| Rel.Std.Dev. |  | 0.7426%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 231b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 231b (Indicative Value 0.35 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.34 | 28
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 231b (Indicative Value 0.17 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.19 | 29
| Mean |  | 0.17 | 16
| Median |  | 0.17 | 0.17
| Std Dev. |  | 0.028284 | 35
| Rel.Std.Dev. |  | 16.64%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 231b (Indicative Value 298 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 298 | 1
| 1 | 2 | 297 | 30
| Mean |  | 297.5 | 16
| Median |  | 297.5 | 297.5
| Std Dev. |  | 0.70711 | 36
| Rel.Std.Dev. |  | 0.2377%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 231b (Indicative Value 16.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.5 | 1
| 1 | 2 | 16.5 | 31
| Mean |  | 16.5 | 16
| Median |  | 16.5 | 16.5
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 231b (Indicative Value 21.5 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.5 | 1
| 1 | 2 | 21.5 | 32
| Mean |  | 21.5 | 16
| Median |  | 21.5 | 21.5
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 231b (Indicative Value 2.44 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.44 | 1
| 1 | 2 | 2.44 | 33
| Mean |  | 2.44 | 16
| Median |  | 2.44 | 2.44
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 231b (Indicative Value 120 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 120 | 1
| 1 | 2 | 120 | 34
| Mean |  | 120 | 16
| Median |  | 120 | 120
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 231b (Indicative Value 65 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 65 | 1
| 1 | 2 | 64.5 | 35
| Mean |  | 64.75 | 16
| Median |  | 64.75 | 64.75
| Std Dev. |  | 0.35355 | 41
| Rel.Std.Dev. |  | 0.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.