## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 232b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.94567 | 0.037434 | 0.8708 | 1.0205 | 0.83337 | 1.058 | 3.958% | 7.917% | 11.88% | 0.89839 | 0.99296
Aqua Regia Digestion (sample mass 10-50g)
Au, ppm | 0.89819 | 0.052013 | 0.79417 | 1.0022 | 0.74215 | 1.0542 | 5.791% | 11.58% | 17.37% | 0.85328 | 0.9431
Cyanide Leach
Au, ppm | 0.81009 | 0.045023 | 0.72004 | 0.90014 | 0.67502 | 0.94516 | 5.558% | 11.12% | 16.67% | 0.76958 | 0.85059
PhotonAssay (recommended gross mass* 510±15 g)
Au, ppm | 0.95867 | 0.054591 | 0.84949 | 1.0679 | 0.7949 | 1.1224 | 5.694% | 11.39% | 17.08% | 0.91074 | 1.0066
4-Acid Digestion
Ag, ppm | 0.1126 | 0.028597 | 0.055406 | 0.16979 | 0.02681 | 0.19839 | 25.4% | 50.79% | 76.19% | 0.10697 | 0.11823
Al, wt.% | 7.0119 | 0.1787 | 6.6545 | 7.3693 | 6.4758 | 7.548 | 2.549% | 5.097% | 7.646% | 6.6613 | 7.3625
As, ppm | 441.1417 | 22.4949 | 396.1518 | 486.1315 | 373.6569 | 508.6265 | 5.099% | 10.2% | 15.3% | 419.0846 | 463.1987
Ba, ppm | 693.8677 | 17.3103 | 659.2471 | 728.4883 | 641.9368 | 745.7986 | 2.495% | 4.99% | 7.484% | 659.1743 | 728.5611
Be, ppm | 2.3482 | 0.18991 | 1.9684 | 2.7281 | 1.7785 | 2.918 | 8.087% | 16.17% | 24.26% | 2.2308 | 2.4657
Bi, ppm | 0.30852 | 0.025161 | 0.2582 | 0.35884 | 0.23304 | 0.384 | 8.155% | 16.31% | 24.47% | 0.2931 | 0.32395
Ca, wt.% | 0.9952 | 0.023431 | 0.94833 | 1.0421 | 0.9249 | 1.0655 | 2.354% | 4.709% | 7.063% | 0.94544 | 1.045
Ce, ppm | 79.2791 | 3.5267 | 72.2256 | 86.3325 | 68.6989 | 89.8592 | 4.448% | 8.897% | 13.35% | 75.3151 | 83.243
Co, ppm | 16.6474 | 0.96457 | 14.7183 | 18.5765 | 13.7537 | 19.5411 | 5.794% | 11.59% | 17.38% | 15.815 | 17.4798
Cr, ppm | 121.7424 | 11.726 | 98.2904 | 145.1944 | 86.5644 | 156.9204 | 9.632% | 19.26% | 28.9% | 115.6553 | 127.8295
Cs, ppm | 8.6274 | 0.34895 | 7.9295 | 9.3253 | 7.5805 | 9.6742 | 4.045% | 8.089% | 12.13% | 8.196 | 9.0588
Cu, ppm | 26.7063 | 1.6808 | 23.3448 | 30.0679 | 21.664 | 31.7487 | 6.294% | 12.59% | 18.88% | 25.371 | 28.0417
Dy, ppm | 3.3802 | 0.4955 | 2.3892 | 4.3712 | 1.8937 | 4.8667 | 14.66% | 29.32% | 43.98% | 3.2112 | 3.5492
Er, ppm | 1.7773 | 0.2719 | 1.2335 | 2.3211 | 0.9616 | 2.593 | 15.3% | 30.6% | 45.9% | 1.6884 | 1.8662
Eu, ppm | 1.2394 | 0.076688 | 1.086 | 1.3927 | 1.0093 | 1.4694 | 6.188% | 12.38% | 18.56% | 1.1774 | 1.3013
Fe, wt.% | 3.8519 | 0.12571 | 3.6005 | 4.1034 | 3.4748 | 4.2291 | 3.264% | 6.527% | 9.791% | 3.6593 | 4.0445
Ga, ppm | 18.4039 | 0.95074 | 16.5024 | 20.3054 | 15.5517 | 21.2561 | 5.166% | 10.33% | 15.5% | 17.4837 | 19.3241
Gd, ppm | 4.9634 | 0.26401 | 4.4354 | 5.4914 | 4.1713 | 5.7554 | 5.319% | 10.64% | 15.96% | 4.7152 | 5.2115
Hf, ppm | 4.064 | 0.30247 | 3.459 | 4.6689 | 3.1566 | 4.9714 | 7.443% | 14.89% | 22.33% | 3.8608 | 4.2672
Ho, ppm | 0.61915 | 0.097277 | 0.42459 | 0.8137 | 0.32732 | 0.91098 | 15.71% | 31.42% | 47.13% | 0.58819 | 0.6501
In, ppm | 0.063353 | 0.0066906 | 0.049972 | 0.076735 | 0.043282 | 0.083425 | 10.56% | 21.12% | 31.68% | 0.060186 | 0.066521
K, wt.% | 2.5063 | 0.068723 | 2.3688 | 2.6437 | 2.3001 | 2.7124 | 2.742% | 5.484% | 8.226% | 2.381 | 2.6316
La, ppm | 37.5965 | 2.6323 | 32.3319 | 42.8611 | 29.6996 | 45.4934 | 7.001% | 14% | 21% | 35.7167 | 39.4763
Li, ppm | 50.4211 | 3.1799 | 44.0613 | 56.7809 | 40.8814 | 59.9607 | 6.307% | 12.61% | 18.92% | 47.9 | 52.9421
Lu, ppm | 0.27659 | 0.030056 | 0.21647 | 0.3367 | 0.18642 | 0.36675 | 10.87% | 21.73% | 32.6% | 0.26276 | 0.29041
Mg, wt.% | 1.6239 | 0.056424 | 1.511 | 1.7367 | 1.4546 | 1.7931 | 3.475% | 6.949% | 10.42% | 1.5427 | 1.705
Mn, wt.% | 0.041416 | 0.0015891 | 0.038238 | 0.044594 | 0.036649 | 0.046183 | 3.837% | 7.674% | 11.51% | 0.039345 | 0.043487
Mo, ppm | 1.0517 | 0.087886 | 0.87593 | 1.2275 | 0.78805 | 1.3154 | 8.356% | 16.71% | 25.07% | 0.99912 | 1.1043
Na, wt.% | 0.81004 | 0.024227 | 0.76159 | 0.8585 | 0.73736 | 0.88272 | 2.991% | 5.982% | 8.973% | 0.76954 | 0.85054
Nb, ppm | 13.4954 | 1.4456 | 10.6041 | 16.3867 | 9.1585 | 17.8323 | 10.71% | 21.42% | 32.14% | 12.8206 | 14.1702
Nd, ppm | 33.4711 | 1.106 | 31.2592 | 35.683 | 30.1532 | 36.789 | 3.304% | 6.608% | 9.913% | 31.7976 | 35.1447
Ni, ppm | 63.2 | 2.8081 | 57.5839 | 68.8161 | 54.7758 | 71.6242 | 4.443% | 8.886% | 13.33% | 60.04 | 66.36
P, wt.% | 0.066666 | 0.0025222 | 0.061621 | 0.07171 | 0.059099 | 0.074232 | 3.783% | 7.567% | 11.35% | 0.063332 | 0.069999
Pb, ppm | 19.2316 | 1.1949 | 16.8418 | 21.6213 | 15.647 | 22.8162 | 6.213% | 12.43% | 18.64% | 18.27 | 20.1931
Pr, ppm | 8.9629 | 0.31587 | 8.3311 | 9.5946 | 8.0152 | 9.9105 | 3.524% | 7.048% | 10.57% | 8.5147 | 9.411
Rb, ppm | 144.3516 | 8.3949 | 127.5618 | 161.1414 | 119.1669 | 169.5363 | 5.816% | 11.63% | 17.45% | 137.134 | 151.5692
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.17235 | 0.01062 | 0.15111 | 0.19359 | 0.14049 | 0.20421 | 6.162% | 12.32% | 18.49% | 0.16373 | 0.18097
Sb, ppm | 186.6604 | 9.5384 | 167.5835 | 205.7372 | 158.045 | 215.2757 | 5.11% | 10.22% | 15.33% | 177.3273 | 195.9934
Sc, ppm | 13.624 | 0.62122 | 12.3816 | 14.8665 | 11.7604 | 15.4877 | 4.56% | 9.12% | 13.68% | 12.9428 | 14.3052
Sm, ppm | 6.438 | 0.45198 | 5.534 | 7.3419 | 5.082 | 7.7939 | 7.021% | 14.04% | 21.06% | 6.1161 | 6.7599
Sn, ppm | 3.436 | 0.22301 | 2.99 | 3.882 | 2.767 | 4.105 | 6.49% | 12.98% | 19.47% | 3.2642 | 3.6078
Sr, ppm | 133.574 | 4.2019 | 125.1703 | 141.9778 | 120.9684 | 146.1796 | 3.146% | 6.291% | 9.437% | 126.8953 | 140.2527
Ta, ppm | 1.0078 | 0.062834 | 0.88213 | 1.1335 | 0.8193 | 1.1963 | 6.235% | 12.47% | 18.7% | 0.95741 | 1.0582
Tb, ppm | 0.67386 | 0.027312 | 0.61924 | 0.72848 | 0.59193 | 0.7558 | 4.053% | 8.106% | 12.16% | 0.64017 | 0.70755
Th, ppm | 14.2812 | 0.76328 | 12.7546 | 15.8077 | 11.9913 | 16.571 | 5.345% | 10.69% | 16.03% | 13.5671 | 14.9952
Ti, wt.% | 0.44255 | 0.010835 | 0.42088 | 0.46422 | 0.41004 | 0.47505 | 2.448% | 4.897% | 7.345% | 0.42042 | 0.46467
Tl, ppm | 0.76917 | 0.040771 | 0.68763 | 0.85071 | 0.64685 | 0.89148 | 5.301% | 10.6% | 15.9% | 0.73071 | 0.80763
Tm, ppm | 0.24925 | 0.041541 | 0.16617 | 0.33234 | 0.12463 | 0.37388 | 16.67% | 33.33% | 50% | 0.23679 | 0.26172
U, ppm | 2.7067 | 0.14928 | 2.4081 | 3.0053 | 2.2589 | 3.1545 | 5.515% | 11.03% | 16.55% | 2.5714 | 2.842
V, ppm | 98.0314 | 5.0969 | 87.8375 | 108.2253 | 82.7405 | 113.3222 | 5.199% | 10.4% | 15.6% | 93.1298 | 102.933
W, ppm | 1.8998 | 0.31326 | 1.2733 | 2.5263 | 0.96 | 2.8395 | 16.49% | 32.98% | 49.47% | 1.8048 | 1.9948
Y, ppm | 15.8747 | 2.0315 | 11.8117 | 19.9376 | 9.7802 | 21.9691 | 12.8% | 25.59% | 38.39% | 15.0809 | 16.6684
Yb, ppm | 1.7021 | 0.26203 | 1.1781 | 2.2262 | 0.91603 | 2.4882 | 15.39% | 30.79% | 46.18% | 1.617 | 1.7872
Zn, ppm | 90.5943 | 2.8137 | 84.967 | 96.2217 | 82.1534 | 99.0353 | 3.106% | 6.212% | 9.317% | 86.0646 | 95.1241
Zr, ppm | 139.9245 | 7.1123 | 125.6999 | 154.1491 | 118.5876 | 161.2614 | 5.083% | 10.17% | 15.25% | 132.9283 | 146.9207
Aqua Regia Digestion
Ag, ppm | 0.10232 | 0.01101 | 0.080295 | 0.12434 | 0.069285 | 0.13535 | 10.76% | 21.52% | 32.28% | 0.0972 | 0.10743
Al, wt.% | 2.7163 | 0.20571 | 2.3049 | 3.1277 | 2.0992 | 3.3334 | 7.573% | 15.15% | 22.72% | 2.5805 | 2.8521
As, ppm | 455.0471 | 24.2926 | 406.4619 | 503.6324 | 382.1692 | 527.925 | 5.338% | 10.68% | 16.02% | 432.2948 | 477.7995
Ba, ppm | 117.8455 | 8.4139 | 101.0178 | 134.6732 | 92.6039 | 143.0871 | 7.14% | 14.28% | 21.42% | 111.9532 | 123.7378
Be, ppm | 1.2557 | 0.11909 | 1.0175 | 1.4939 | 0.89844 | 1.613 | 9.484% | 18.97% | 28.45% | 1.1929 | 1.3185
Bi, ppm | 0.29887 | 0.021676 | 0.25552 | 0.34222 | 0.23384 | 0.3639 | 7.253% | 14.51% | 21.76% | 0.28393 | 0.31381
Ca, wt.% | 0.33697 | 0.017852 | 0.30126 | 0.37267 | 0.28341 | 0.39052 | 5.298% | 10.6% | 15.89% | 0.32012 | 0.35381
Cd, ppm | 0.044859 | 0.016526 | 0.011807 | 0.077911 | 0 | 0.094437 | 36.84% | 73.68% | 110.5% | 0.042616 | 0.047102
Ce, ppm | 52.2172 | 2.4281 | 47.361 | 57.0735 | 44.9328 | 59.5016 | 4.65% | 9.3% | 13.95% | 49.6064 | 54.8281
Co, ppm | 15.3523 | 0.97336 | 13.4056 | 17.2991 | 12.4323 | 18.2724 | 6.34% | 12.68% | 19.02% | 14.5847 | 16.12
Cr, ppm | 108.6305 | 4.8953 | 98.8398 | 118.4212 | 93.9445 | 123.3165 | 4.506% | 9.013% | 13.52% | 103.199 | 114.062
Cs, ppm | 6.7465 | 0.3794 | 5.9877 | 7.5053 | 5.6084 | 7.8847 | 5.624% | 11.25% | 16.87% | 6.4092 | 7.0839
Cu, ppm | 25.4034 | 1.4144 | 22.5746 | 28.2321 | 21.1603 | 29.6464 | 5.568% | 11.14% | 16.7% | 24.1332 | 26.6735
Fe, wt.% | 3.3021 | 0.11428 | 3.0735 | 3.5306 | 2.9592 | 3.6449 | 3.461% | 6.922% | 10.38% | 3.137 | 3.4672
Ga, ppm | 8.5777 | 0.39311 | 7.7915 | 9.3639 | 7.3983 | 9.757 | 4.583% | 9.166% | 13.75% | 8.1488 | 9.0065
Ge, ppm | 0.123 | 0.011188 | 0.10062 | 0.14538 | 0.089436 | 0.15656 | 9.096% | 18.19% | 27.29% | 0.11685 | 0.12915
Hf, ppm | 0.4998 | 0.086164 | 0.32748 | 0.67213 | 0.24131 | 0.7583 | 17.24% | 34.48% | 51.72% | 0.47481 | 0.52479
In, ppm | 0.033608 | 0.004124 | 0.02536 | 0.041856 | 0.021236 | 0.04598 | 12.27% | 24.54% | 36.81% | 0.031927 | 0.035288
K, wt.% | 0.86889 | 0.042657 | 0.78358 | 0.95421 | 0.74092 | 0.99687 | 4.909% | 9.819% | 14.73% | 0.82545 | 0.91234
La, ppm | 25.5034 | 1.3342 | 22.8349 | 28.1718 | 21.5006 | 29.5061 | 5.232% | 10.46% | 15.69% | 24.2282 | 26.7785
Li, ppm | 41.4797 | 3.3556 | 34.7685 | 48.1909 | 31.4129 | 51.5465 | 8.09% | 16.18% | 24.27% | 39.4057 | 43.5537
Lu, ppm | 0.11561 | 0.025059 | 0.065497 | 0.16573 | 0.040438 | 0.19079 | 21.67% | 43.35% | 65.02% | 0.10983 | 0.1214
Mg, wt.% | 1.3332 | 0.08092 | 1.1713 | 1.495 | 1.0904 | 1.5759 | 6.07% | 12.14% | 18.21% | 1.2665 | 1.3998
Mn, wt.% | 0.029529 | 0.0016605 | 0.026208 | 0.03285 | 0.024548 | 0.034511 | 5.623% | 11.25% | 16.87% | 0.028053 | 0.031006
Mo, ppm | 0.93558 | 0.065834 | 0.80391 | 1.0672 | 0.73808 | 1.1331 | 7.037% | 14.07% | 21.11% | 0.8888 | 0.98236
Na, wt.% | 0.10035 | 0.011994 | 0.07636 | 0.12434 | 0.064366 | 0.13633 | 11.95% | 23.91% | 35.86% | 0.095331 | 0.10537
Nb, ppm | 0.3622 | 0.075396 | 0.21141 | 0.51299 | 0.13601 | 0.58839 | 20.82% | 41.63% | 62.45% | 0.34409 | 0.38031
Ni, ppm | 58.8159 | 3.2308 | 52.3544 | 65.2775 | 49.1236 | 68.5083 | 5.493% | 10.99% | 16.48% | 55.8751 | 61.7567
P, wt.% | 0.058827 | 0.0019225 | 0.054982 | 0.062672 | 0.053059 | 0.064594 | 3.268% | 6.536% | 9.804% | 0.055885 | 0.061768
Pb, ppm | 8.6164 | 0.65009 | 7.3162 | 9.9165 | 6.6661 | 10.5666 | 7.545% | 15.09% | 22.63% | 8.1855 | 9.0472
Rb, ppm | 83.6152 | 4.3199 | 74.9755 | 92.2549 | 70.6557 | 96.5748 | 5.166% | 10.33% | 15.5% | 79.4345 | 87.796
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.18064 | 0.016914 | 0.14681 | 0.21446 | 0.12989 | 0.23138 | 9.364% | 18.73% | 28.09% | 0.1716 | 0.18967
Sb, ppm | 140.5588 | 23.7514 | 93.0561 | 188.0615 | 69.3047 | 211.8129 | 16.9% | 33.8% | 50.69% | 133.5308 | 147.5867
Sc, ppm | 6.7317 | 0.46265 | 5.8064 | 7.657 | 5.3437 | 8.1196 | 6.873% | 13.75% | 20.62% | 6.3951 | 7.0683
Sn, ppm | 1.6152 | 0.095108 | 1.425 | 1.8054 | 1.3299 | 1.9005 | 5.888% | 11.78% | 17.66% | 1.5344 | 1.696
Sr, ppm | 28.9839 | 1.5232 | 25.9375 | 32.0303 | 24.4143 | 33.5535 | 5.255% | 10.51% | 15.77% | 27.5347 | 30.4331
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 11.6413 | 0.55986 | 10.5216 | 12.761 | 9.9617 | 13.3209 | 4.809% | 9.618% | 14.43% | 11.0592 | 12.2234
Ti, wt.% | 0.16552 | 0.017052 | 0.13141 | 0.19962 | 0.11436 | 0.21667 | 10.3% | 20.6% | 30.91% | 0.15724 | 0.17379
Tl, ppm | 0.49302 | 0.032938 | 0.42715 | 0.5589 | 0.39421 | 0.59184 | 6.681% | 13.36% | 20.04% | 0.46837 | 0.51768
U, ppm | 1.3951 | 0.1163 | 1.1625 | 1.6277 | 1.0462 | 1.744 | 8.336% | 16.67% | 25.01% | 1.3254 | 1.4649
V, ppm | 66.6288 | 2.8792 | 60.8704 | 72.3871 | 57.9913 | 75.2662 | 4.321% | 8.642% | 12.96% | 63.2973 | 69.9602
W, ppm | 0.33663 | 0.038153 | 0.26033 | 0.41294 | 0.22218 | 0.45109 | 11.33% | 22.67% | 34% | 0.3198 | 0.35347
Y, ppm | 8.8357 | 0.55112 | 7.7335 | 9.938 | 7.1823 | 10.4891 | 6.237% | 12.47% | 18.71% | 8.3939 | 9.2775
Yb, ppm | 0.82623 | 0.082697 | 0.66084 | 0.99163 | 0.57814 | 1.0743 | 10.01% | 20.02% | 30.03% | 0.78492 | 0.86754
Zn, ppm | 81.1161 | 3.2268 | 74.6624 | 87.5697 | 71.4356 | 90.7966 | 3.978% | 7.956% | 11.93% | 77.0603 | 85.1719
Zr, ppm | 19.9978 | 2.8622 | 14.2733 | 25.7223 | 11.4111 | 28.5845 | 14.31% | 28.63% | 42.94% | 18.9979 | 20.9977
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 232b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.94567 | 0.93265 | 0.95869 | 0.93713 | 0.95422
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.89819 | 0.87449 | 0.92189 | 0.8893 | 0.90708
Cyanide Leach
Au, Gold (ppm) | 0.81009 | 0.7876 | 0.83258 | 0.80277 | 0.81741
PhotonAssay
Au, Gold (ppm) | 0.95867 | 0.91835 | 0.99899 | 0.95614 | 0.96121
4-Acid Digestion
Ag, Silver (ppm) | 0.1126 | 0.085493 | 0.13971 | 0.10186 | 0.12334
Al, Aluminium (wt.%) | 7.0119 | 6.8175 | 7.2063 | 6.8958 | 7.128
As, Arsenic (ppm) | 441.1417 | 419.5473 | 462.7361 | 426.9056 | 455.3777
Ba, Barium (ppm) | 693.8677 | 673.9155 | 713.8199 | 681.8299 | 705.9056
Be, Beryllium (ppm) | 2.3482 | 2.2373 | 2.4592 | 2.2407 | 2.4558
Bi, Bismuth (ppm) | 0.30852 | 0.27244 | 0.34461 | 0.27051 | 0.34653
Ca, Calcium (wt.%) | 0.9952 | 0.96848 | 1.0219 | 0.97481 | 1.0156
Ce, Cerium (ppm) | 79.2791 | 75.1889 | 83.3692 | 77.2978 | 81.2603
Co, Cobalt (ppm) | 16.6474 | 15.6698 | 17.625 | 16.1921 | 17.1027
Cr, Chromium (ppm) | 121.7424 | 112.9323 | 130.5524 | 117.5907 | 125.8941
Cs, Caesium (ppm) | 8.6274 | 8.1764 | 9.0784 | 8.3969 | 8.8578
Cu, Copper (ppm) | 26.7063 | 25.3447 | 28.068 | 25.6512 | 27.7615
Dy, Dysprosium (ppm) | 3.3802 | 2.9993 | 3.7612 | 3.217 | 3.5434
Er, Erbium (ppm) | 1.7773 | 1.5502 | 2.0044 | 1.6646 | 1.89
Eu, Europium (ppm) | 1.2394 | 1.1279 | 1.3508 | 1.1859 | 1.2928
Fe, Iron (wt.%) | 3.8519 | 3.7589 | 3.945 | 3.7861 | 3.9178
Ga, Gallium (ppm) | 18.4039 | 17.3931 | 19.4147 | 17.8114 | 18.9964
Gd, Gadolinium (ppm) | 4.9634 | 4.6924 | 5.2344 | 4.7569 | 5.1698
Hf, Hafnium (ppm) | 4.064 | 3.8139 | 4.314 | 3.8745 | 4.2534
Ho, Holmium (ppm) | 0.61915 | 0.5274 | 0.71089 | 0.57575 | 0.66255
In, Indium (ppm) | 0.063353 | 0.052811 | 0.073896 | 0.060403 | 0.066304
K, Potassium (wt.%) | 2.5063 | 2.4311 | 2.5814 | 2.4554 | 2.5571
La, Lanthanum (ppm) | 37.5965 | 35.9604 | 39.2326 | 36.3262 | 38.8668
Li, Lithium (ppm) | 50.4211 | 48.2742 | 52.5679 | 49.0615 | 51.7806
Lu, Lutetium (ppm) | 0.27659 | 0.2374 | 0.31577 | 0.26408 | 0.2891
Mg, Magnesium (wt.%) | 1.6239 | 1.5761 | 1.6716 | 1.5946 | 1.6531
Mn, Manganese (wt.%) | 0.041416 | 0.040117 | 0.042714 | 0.040712 | 0.04212
Mo, Molybdenum (ppm) | 1.0517 | 0.95273 | 1.1507 | 0.97521 | 1.1282
Na, Sodium (wt.%) | 0.81004 | 0.78408 | 0.83601 | 0.79118 | 0.82891
Nb, Niobium (ppm) | 13.4954 | 12.493 | 14.4977 | 13.0171 | 13.9737
Nd, Neodymium (ppm) | 33.4711 | 32.019 | 34.9232 | 32.6116 | 34.3306
Ni, Nickel (ppm) | 63.2 | 61.1116 | 65.2883 | 61.5885 | 64.8115
P, Phosphorus (wt.%) | 0.066666 | 0.064538 | 0.068793 | 0.065634 | 0.067698
Pb, Lead (ppm) | 19.2316 | 18.034 | 20.4292 | 18.6477 | 19.8155
Pr, Praseodymium (ppm) | 8.9629 | 8.5143 | 9.4114 | 8.6102 | 9.3155
Rb, Rubidium (ppm) | 144.3516 | 137.5055 | 151.1978 | 139.547 | 149.1562
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.17235 | 0.16245 | 0.18225 | 0.165 | 0.1797
Sb, Antimony (ppm) | 186.6604 | 178.6193 | 194.7015 | 180.7319 | 192.5888
Sc, Scandium (ppm) | 13.624 | 12.8875 | 14.3605 | 13.2893 | 13.9588
Sm, Samarium (ppm) | 6.438 | 6.0402 | 6.8357 | 6.1781 | 6.6978
Sn, Tin (ppm) | 3.436 | 3.1989 | 3.673 | 3.2476 | 3.6244
Sr, Strontium (ppm) | 133.574 | 129.7472 | 137.4009 | 130.7547 | 136.3934
Ta, Tantalum (ppm) | 1.0078 | 0.93944 | 1.0762 | 0.95743 | 1.0582
Tb, Terbium (ppm) | 0.67386 | 0.62685 | 0.72087 | 0.6299 | 0.71782
Th, Thorium (ppm) | 14.2812 | 13.7369 | 14.8255 | 13.9275 | 14.6349
Ti, Titanium (wt.%) | 0.44255 | 0.43269 | 0.4524 | 0.42897 | 0.45612
Tl, Thallium (ppm) | 0.76917 | 0.72672 | 0.81161 | 0.74498 | 0.79335
Tm, Thulium (ppm) | 0.24925 | 0.21532 | 0.28319 | 0.23432 | 0.26419
U, Uranium (ppm) | 2.7067 | 2.5596 | 2.8538 | 2.5853 | 2.8281
V, Vanadium (ppm) | 98.0314 | 95.1769 | 100.8859 | 95.7075 | 100.3552
W, Tungsten (ppm) | 1.8998 | 1.6882 | 2.1113 | 1.7634 | 2.0361
Y, Yttrium (ppm) | 15.8747 | 14.7358 | 17.0136 | 15.1514 | 16.598
Yb, Ytterbium (ppm) | 1.7021 | 1.5141 | 1.8901 | 1.589 | 1.8152
Zn, Zinc (ppm) | 90.5943 | 88.4042 | 92.7845 | 88.4787 | 92.7099
Zr, Zirconium (ppm) | 139.9245 | 133.6096 | 146.2394 | 135.4279 | 144.4211
Aqua Regia Digestion
Ag, Silver (ppm) | 0.10232 | 0.084627 | 0.12 | IND | IND
Al, Aluminium (wt.%) | 2.7163 | 2.6027 | 2.8299 | 2.6693 | 2.7633
As, Arsenic (ppm) | 455.0471 | 438.8067 | 471.2876 | 444.3263 | 465.7679
Ba, Barium (ppm) | 117.8455 | 112.2561 | 123.4349 | 114.462 | 121.229
Be, Beryllium (ppm) | 1.2557 | 1.169 | 1.3424 | 1.2021 | 1.3093
Bi, Bismuth (ppm) | 0.29887 | 0.26579 | 0.33195 | 0.26233 | 0.33541
Ca, Calcium (wt.%) | 0.33697 | 0.32503 | 0.3489 | 0.33086 | 0.34308
Cd, Cadmium (ppm) | 0.044859 | 0.031187 | 0.058531 | 0.042209 | 0.047509
Ce, Cerium (ppm) | 52.2172 | 49.5669 | 54.8676 | 50.3921 | 54.0424
Co, Cobalt (ppm) | 15.3523 | 14.6862 | 16.0184 | 14.857 | 15.8477
Cr, Chromium (ppm) | 108.6305 | 105.5046 | 111.7564 | 106.3818 | 110.8792
Cs, Caesium (ppm) | 6.7465 | 6.4517 | 7.0414 | 6.4994 | 6.9937
Cu, Copper (ppm) | 25.4034 | 24.4716 | 26.3351 | 24.0526 | 26.7541
Fe, Iron (wt.%) | 3.3021 | 3.2272 | 3.377 | 3.2571 | 3.3471
Ga, Gallium (ppm) | 8.5777 | 8.1379 | 9.0174 | 8.3066 | 8.8487
Ge, Germanium (ppm) | 0.123 | 0.10254 | 0.14346 | IND | IND
Hf, Hafnium (ppm) | 0.4998 | 0.43147 | 0.56814 | 0.47852 | 0.52109
In, Indium (ppm) | 0.033608 | 0.028957 | 0.038258 | 0.030183 | 0.037033
K, Potassium (wt.%) | 0.86889 | 0.84265 | 0.89514 | 0.85751 | 0.88028
La, Lanthanum (ppm) | 25.5034 | 24.1972 | 26.8096 | 24.4751 | 26.5316
Li, Lithium (ppm) | 41.4797 | 38.9377 | 44.0217 | 40.3972 | 42.5622
Lu, Lutetium (ppm) | 0.11561 | 0.084873 | 0.14636 | IND | IND
Mg, Magnesium (wt.%) | 1.3332 | 1.2881 | 1.3783 | 1.3102 | 1.3561
Mn, Manganese (wt.%) | 0.029529 | 0.028514 | 0.030544 | 0.028993 | 0.030065
Mo, Molybdenum (ppm) | 0.93558 | 0.88047 | 0.99069 | 0.87068 | 1.0005
Na, Sodium (wt.%) | 0.10035 | 0.093142 | 0.10756 | 0.09708 | 0.10362
Nb, Niobium (ppm) | 0.3622 | 0.27905 | 0.44535 | 0.32546 | 0.39894
Ni, Nickel (ppm) | 58.8159 | 56.3257 | 61.3061 | 57.6903 | 59.9416
P, Phosphorus (wt.%) | 0.058827 | 0.056935 | 0.060718 | 0.057703 | 0.059951
Pb, Lead (ppm) | 8.6164 | 8.0046 | 9.2282 | 8.1101 | 9.1227
Rb, Rubidium (ppm) | 83.6152 | 80.2074 | 87.0231 | 81.43 | 85.8004
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.18064 | 0.16952 | 0.19175 | 0.17555 | 0.18572
Sb, Antimony (ppm) | 140.5588 | 128.3563 | 152.7613 | 137.172 | 143.9456
Sc, Scandium (ppm) | 6.7317 | 6.3842 | 7.0792 | 6.5899 | 6.8735
Sn, Tin (ppm) | 1.6152 | 1.5029 | 1.7275 | 1.5652 | 1.6652
Sr, Strontium (ppm) | 28.9839 | 27.7559 | 30.2119 | 28.3914 | 29.5763
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Th, Thorium (ppm) | 11.6413 | 10.9817 | 12.301 | 11.1336 | 12.149
Ti, Titanium (wt.%) | 0.16552 | 0.15613 | 0.17491 | 0.16167 | 0.16936
Tl, Thallium (ppm) | 0.49302 | 0.46228 | 0.52377 | 0.47427 | 0.51178
U, Uranium (ppm) | 1.3951 | 1.3114 | 1.4788 | 1.3366 | 1.4536
V, Vanadium (ppm) | 66.6288 | 64.774 | 68.4835 | 65.3761 | 67.8814
W, Tungsten (ppm) | 0.33663 | 0.29897 | 0.3743 | 0.29395 | 0.37932
Y, Yttrium (ppm) | 8.8357 | 8.4158 | 9.2556 | 8.6032 | 9.0682
Yb, Ytterbium (ppm) | 0.82623 | 0.70412 | 0.94834 | 0.78925 | 0.86322
Zn, Zinc (ppm) | 81.1161 | 78.4859 | 83.7463 | 79.0577 | 83.1744
Zr, Zirconium (ppm) | 19.9978 | 18.2455 | 21.7501 | 19.3942 | 20.6014
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 232b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 0.60747 | Pt | ppb | 0.45174
4-Acid Digestion
Au | ppm | 0.91773 | Hg | ppm | 0.02865 | Rh | ppm | 0.0013
B | ppm | 26.5533 | Ir | ppm | 0.0025167 | Ru | ppm | 0.0028333
Cd | ppm | 0.07153 | Pd | ppb | 26.6667 | Se | ppm | 0.66521
Ge | ppm | 0.19286 | Pt | ppb | 2.8667 | Te | ppm | 0.060429
Aqua Regia Digestion
B | ppm | 14.3392 | Ir | ppm | 0.00088333 | Se | ppm | 0.23362
Dy | ppm | 2.0888 | Nd | ppm | 22.942 | Si | wt.% | 0.40472
Er | ppm | 0.90798 | Pd | ppb | 15.9167 | Sm | ppm | 4.3109
Eu | ppm | 0.57985 | Pr | ppm | 6.2389 | Tb | ppm | 0.45098
Gd | ppm | 4.0497 | Pt | ppb | 2.2167 | Te | ppm | 0.021167
Hg | ppm | 0.0095611 | Rh | ppm | 0.00076667 | Tm | ppm | 0.11471
Ho | ppm | 0.34103 | Ru | ppm | < 0.0028
Borate Fusion XRF
Al2O3 | wt.% | 13.845 | MgO | wt.% | 2.745 | SiO2 | wt.% | 69.185
CaO | wt.% | 1.38 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.4505
Fe2O3 | wt.% | 5.575 | Na2O | wt.% | 1.1 | TiO2 | wt.% | 0.81
K2O | wt.% | 3.095 | P2O5 | wt.% | 0.1515
Thermogravimetry
LOI1000 | wt.% | 1.795
Infrared Combustion
C | wt.% | 0.06 | S | wt.% | 0.17917
Laser Ablation ICP-MS
Ag | ppm | 0.125 | Hf | ppm | 7.085 | Sm | ppm | 7.05
As | ppm | 446.5 | Ho | ppm | 1.155 | Sn | ppm | 3.5
Ba | ppm | 700 | In | ppm | < 0.05 | Sr | ppm | 132.5
Be | ppm | 2.9 | La | ppm | 39.05 | Ta | ppm | 1.23
Bi | ppm | 0.34 | Lu | ppm | 0.465 | Tb | ppm | 0.94
Cd | ppm | 0.125 | Mn | wt.% | 0.04415 | Te | ppm | < 0.2
Ce | ppm | 77.35 | Mo | ppm | 1.2 | Th | ppm | 14.55
Co | ppm | 17.45 | Nb | ppm | 15.9 | Ti | wt.% | 0.489
Cr | ppm | 150 | Nd | ppm | 35.3 | Tl | ppm | 0.4
Cs | ppm | 8.47 | Ni | ppm | 67 | Tm | ppm | 0.515
Cu | ppm | 29 | Pb | ppm | 20 | U | ppm | 3.14
Dy | ppm | 5.615 | Pr | ppm | 9.5 | V | ppm | 103
Er | ppm | 3.325 | Rb | ppm | 146 | W | ppm | 2.75
Eu | ppm | 1.335 | Re | ppm | < 0.01 | Y | ppm | 30.6
Ga | ppm | 18.15 | Sb | ppm | 190 | Yb | ppm | 3.32
Gd | ppm | 6.08 | Sc | ppm | 13.75 | Zn | ppm | 87.5
Ge | ppm | 1.45 | Se | ppm | < 5 | Zr | ppm | 252
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 232b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 232b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Carsurin, Cikarang, West Java, Indonesia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Kalgoorlie, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 232b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 4 | 4 | 219474 | 0.085737 | 0.97448 | 0.99629
2 | 1 | ANSLu | 1 | 6 | 6 | 219475 | 0.08836 | 0.9713 | 0.99612
2 | 1 | ANSLu | 1 | 3 | 3 | 219476 | 0.086093 | 1.0128 | 0.99835
2 | 1 | ANSLu | 1 | 18 | 18 | 219477 | 0.084026 | 0.95613 | 0.99531
2 | 1 | ANSLu | 1 | 19 | 19 | 219478 | 0.085922 | 0.98894 | 0.99707
2 | 1 | ANSLu | 1 | 16 | 16 | 219479 | 0.0872 | 0.94684 | 0.99481
2 | 1 | ANSLu | 1 | 15 | 15 | 219480 | 0.082405 | 0.97661 | 0.99641
2 | 1 | ANSLu | 1 | 11 | 11 | 219481 | 0.086196 | 0.97386 | 0.99626
2 | 1 | ANSLu | 1 | 12 | 12 | 219482 | 0.083335 | 0.98099 | 0.99664
2 | 1 | ANSLu | 1 | 10 | 10 | 219483 | 0.084344 | 0.98608 | 0.99691
2 | 1 | ANSLu | 1 | 1 | 1 | 219484 | 0.086788 | 0.97716 | 0.99644
2 | 1 | ANSLu | 1 | 8 | 8 | 219485 | 0.086529 | 0.99124 | 0.99719
2 | 1 | ANSLu | 1 | 20 | 20 | 219486 | 0.087924 | 0.99804 | 0.99755
2 | 1 | ANSLu | 1 | 2 | 2 | 219487 | 0.085305 | 1.0392 | 0.99975
2 | 1 | ANSLu | 1 | 5 | 5 | 219488 | 0.085109 | 0.99525 | 0.99741
2 | 1 | ANSLu | 1 | 9 | 9 | 219489 | 0.086936 | 0.93375 | 0.99411
2 | 1 | ANSLu | 1 | 17 | 17 | 219490 | 0.087242 | 0.98564 | 0.99689
2 | 1 | ANSLu | 1 | 13 | 13 | 219491 | 0.083033 | 1.0141 | 0.99841
2 | 1 | ANSLu | 1 | 14 | 14 | 219492 | 0.087126 | 1.062 | 1.001
2 | 1 | ANSLu | 1 | 7 | 7 | 219493 | 0.08815 | 1.1862 | 1.0076
Mean | [6 empty] | 0.085888 | 0.99753 | 0.99753
Median | [7 empty] | 0.98586 | 0.9969
SD | [7 empty] | 0.053129 | 0.0028427
RSD | [7 empty] | 5.326% | 0.285%
RSD Upscale Factor | [7 empty] | 0.053506
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 232b (Certified Value 0.946 ppm) | [32 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*MS | 3
|  |  | 0.085g | 25g | 50g | 50g | 40g | 30g | 40g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 25g | 25g | 3
| 1 | 1 | 0.97716 | 0.943 | 1.02 | 0.911 | 0.94 | 1.09 | 0.94 | 0.91 | 0.92 | 0.95 | 0.964 | 0.966 | 0.888 | 0.955 | 0.96 | 0.961 | 0.90013 | 0.954 | 0.901 | 0.92 | 0.95 | 0.94 | 0.959 | 0.93 | 0.911 | 0.96 | 0.99 | 1.003 | 0.91816 | 0.95 | 0.92755 | 1
| 1 | 2 | 1.0392 | 0.95 | 1.01 | 0.912 | 0.93 | 0.923 | 0.93 | 0.9 | 0.97 | 0.95 | 0.921 | 0.948 | 0.901 | 0.937 | 0.948 | 0.969 | 0.91664 | 0.946 | 1.02 | 0.92 | 0.97 | 0.94 | 0.956 | 0.91 | 0.89 | 0.96 | 1.02 | 1 | 0.92832 | 0.92 | 0.91496 | 3
| 1 | 3 | 1.0128 | 0.947 | 1.01 | 0.924 | 0.95 | 1.06 | 0.95 | 0.93 | 0.97 | 0.95 | 0.974 | 0.97 | 0.873 | 0.945 | 0.965 | 0.98 | 0.89411 | 0.958 | 0.945 | 0.94 | 0.97 | 0.93 | 0.95 | 0.93 | 0.892 | 0.94 | 0.98 | 1.013 | 0.92928 | 0.92 | 0.89021 | 16
| 1 | 4 | 0.97448 | 0.973 | 0.99 | 0.889 | 0.95 | 0.906 | 0.94 | 0.91 | 0.94 | 0.96 | 0.963 | 0.96 | 0.896 | 0.95 | 0.949 | 0.979 | 0.89932 | 0.982 | 1.01 | 0.99 | 0.95 | 0.95 | 0.95 | 0.83 | 0.916 | 0.94 | 1.02 | 0.998 | 0.92832 | 0.92 | 0.90404 | 0.94567
| 1 | 5 | 0.99525 | 0.961 | 1.02 | 0.868 | 0.93 | 0.941 | 0.95 | 0.92 | 0.94 | 0.96 | 0.943 | 0.959 | 0.919 | 0.963 | 0.959 | 0.974 | 0.89689 | 0.988 | 0.99 | 1.04 | 0.97 | 0.94 | 0.925 | 0.92 | 0.907 | 0.93 | 0.93 | 1.006 | 0.93008 | 0.9 | 0.91255 | 7
| 1 | 6 | 0.9713 | 0.994 | 1.03 | 0.876 | 0.94 | 0.926 | 0.96 | 0.92 | 0.98 | 0.97 | 0.976 | 0.96 | 0.908 | 0.946 | 0.936 | 0.956 | 0.90717 | 0.984 | 1.04 | 1.016 | 0.97 | 0.93 | 0.915 | 0.87 | 0.895 | 0.95 | 0.95 | 1 | 0.9216 | 0.92 | 0.88944
|  | 7 | 1.1862
|  | 8 | 0.99124
|  | 9 | 0.93375
|  | 10 | 0.98608
|  | 11 | 0.97386
|  | 12 | 0.98099
|  | 13 | 1.0141
|  | 14 | 1.062
|  | 15 | 0.97661
|  | 16 | 0.94684
|  | 17 | 0.98564
|  | 18 | 0.95613
|  | 19 | 0.98894
|  | 20 | 0.99804
| Mean |  | 0.99753 | 0.96133 | 1.0133 | 0.89667 | 0.94 | 0.97433 | 0.945 | 0.915 | 0.95333 | 0.95667 | 0.95683 | 0.9605 | 0.8975 | 0.94933 | 0.95283 | 0.96983 | 0.90238 | 0.96867 | 0.98433 | 0.971 | 0.96333 | 0.93833 | 0.9425 | 0.89833 | 0.90183 | 0.94667 | 0.98167 | 1.0033 | 0.92596 | 0.92167 | 0.90646
| Median |  | 0.98586 | 0.9555 | 1.015 | 0.9 | 0.94 | 0.9335 | 0.945 | 0.915 | 0.955 | 0.955 | 0.9635 | 0.96 | 0.8985 | 0.948 | 0.954 | 0.9715 | 0.89972 | 0.97 | 1 | 0.965 | 0.97 | 0.94 | 0.95 | 0.915 | 0.901 | 0.945 | 0.985 | 1.0015 | 0.92832 | 0.92 | 0.90829
| Std Dev. |  | 0.053129 | 0.019356 | 0.013663 | 0.02234 | 0.0089443 | 0.079334 | 0.010488 | 0.010488 | 0.023381 | 0.008165 | 0.021104 | 0.0074766 | 0.015984 | 0.0089592 | 0.010572 | 0.0097451 | 0.0082358 | 0.018052 | 0.052064 | 0.051595 | 0.010328 | 0.0075277 | 0.018053 | 0.040208 | 0.010907 | 0.012111 | 0.03656 | 0.0055015 | 0.0048782 | 0.016021 | 0.014925
| Rel.Std.Dev. |  | 5.326% | 2.013% | 1.348% | 2.491% | 0.9515% | 8.142% | 1.11% | 1.146% | 2.453% | 0.8535% | 2.206% | 0.7784% | 1.781% | 0.9437% | 1.11% | 1.005% | 0.9127% | 1.864% | 5.289% | 5.314% | 1.072% | 0.8022% | 1.915% | 4.476% | 1.209% | 1.279% | 3.724% | 0.5483% | 0.5268% | 1.738% | 1.646%
| PDM3 |  | 5.483% | 1.656% | 7.155% | -5.182% | -0.5998% | 3.031% | -0.07106% | -3.243% | 0.8101% | 1.163% | 1.18% | 1.568% | -5.094% | 0.3872% | 0.7573% | 2.555% | -4.578% | 2.432% | 4.088% | 2.678% | 1.868% | -0.776% | -0.3354% | -5.006% | -4.636% | 0.1052% | 3.806% | 6.097% | -2.084% | -2.538% | -4.147%
| Z-Score (Absolute) |  | NA | 0.36 | 2.17 | 1.9 | 0.39 | 0.81 | 0.21 | 1.26 | 0.08 | 0.19 | 0.2 | 0.33 | 1.87 | 0.06 | 0.06 | 0.65 | 1.7 | 0.61 | 1.16 | 0.69 | 0.43 | 0.44 | 0.3 | 1.84 | 1.72 | 0.15 | 1.07 | 1.82 | 0.88 | 1.03 | 1.56
| Analytical results for Pd in OREAS 232b (Indicative Value 0.607 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.63504 | [30 empty] | 1
| 1 | 2 | 0.61866 | [30 empty] | 1
| 1 | 3 | 0.58693 | [30 empty] | 16
| 1 | 4 | 0.57399 | [30 empty] | 0.60747
| 1 | 5 | 0.61024 | [30 empty] | 7
| 1 | 6 | 0.61998
| Mean |  | 0.60747
| Median |  | 0.61445
| Std Dev. |  | 0.02277
| Rel.Std.Dev. |  | 3.748%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 232b (Indicative Value 0.452 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.48919 | [30 empty] | 1
| 1 | 2 | 0.42095 | [30 empty] | 1
| 1 | 3 | 0.43875 | [30 empty] | 16
| 1 | 4 | 0.45644 | [30 empty] | 0.45174
| 1 | 5 | 0.43357 | [30 empty] | 7
| 1 | 6 | 0.47155
| Mean |  | 0.45174
| Median |  | 0.44759
| Std Dev. |  | 0.02555
| Rel.Std.Dev. |  | 5.656%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 232b (Certified Value 0.898 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 0.085g | 25g | 50g | 30g | 40g | 30g | 30g | 25g | 30g | 25g | 25g | 25g | 25g | 15g | 50g | 25g | 50g | 25g | 10g | 25g | 20g | 15g | 15g | 50g | 3
| 1 | 1 | 0.97716 | 0.95 | 0.94 | 0.9221 | 0.85 | 0.91 | 0.93 | > 0.5 | 0.85 | 0.832 | 0.841 | 0.865 | 0.93 | 0.927 | 0.88 | 1.05 | 0.82 | 0.5 | 0.87 | 0.925 | 0.9661 | 0.8636 | 0.8035 | 0.94 | 1
| 1 | 2 | 1.0392 | 0.94 | 0.9 | 0.9215 | 0.842 | 0.957 | 0.95 | > 0.5 | 0.85 | 0.863 | 0.851 | 0.841 | 0.94 | 0.9173 | 0.94 | 1.03 | 0.84 | 0.51 | 0.88 | 0.918 | 0.9593 | 0.8742 | 0.8069 | 0.9242 | 3
| 1 | 3 | 1.0128 | 0.94 | 0.96 | 0.9318 | 0.834 | 0.888 | 0.96 | > 0.5 | 0.85 | 0.848 | 0.849 | 0.904 | 0.94 | 0.9341 | 0.92 | 0.974 | 0.79 | 0.53 | 0.86 | 0.905 | 0.9529 | 0.8857 | 0.8237 | 0.9094 | 16
| 1 | 4 | 0.97448 | 0.94 | 0.9 | 0.9115 | 0.865 | 0.931 | 0.95 | > 0.5 | 0.85 | 0.848 | 0.879 | 0.833 | 0.93 | 0.9534 | 0.9 | 1 | 0.86 | 0.5 | 0.89 | 0.915 | 0.9492 | 0.8577 | 0.8262 | 0.9176 | 0.89819
| 1 | 5 | 0.99525 | 0.95 | 0.94 | 0.923 | 0.855 | 0.94 | 0.93 | > 0.5 | 0.86 | 0.855 | 0.85 | 0.811 | 0.93 | 0.8831 | 0.9 | 0.986 | 0.81 | 0.53 | 0.87 | 0.925 | 0.9448 | 0.8694 | 0.8137 | 0.9242 | 9
| 1 | 6 | 0.9713 | 0.94 | 0.91 | 0.9327 | 0.834 | 0.944 | 0.93 | > 0.5 | 0.84 | 0.855 | 0.877 | 0.803 | 0.92 | 0.9855 | 0.9 | 0.994 | 0.81 | 0.53 | 0.9 | 0.904 | 0.9405 | 0.8409 | 0.8249 | 0.9397
|  | 7 | 1.1862
|  | 8 | 0.99124
|  | 9 | 0.93375
|  | 10 | 0.98608
|  | 11 | 0.97386
|  | 12 | 0.98099
|  | 13 | 1.0141
|  | 14 | 1.062
|  | 15 | 0.97661
|  | 16 | 0.94684
|  | 17 | 0.98564
|  | 18 | 0.95613
|  | 19 | 0.98894
|  | 20 | 0.99804
| Mean |  | 0.99753 | 0.94333 | 0.925 | 0.92377 | 0.84667 | 0.92833 | 0.94167 |  | 0.85 | 0.85017 | 0.85783 | 0.84283 | 0.93167 | 0.9334 | 0.90667 | 1.0057 | 0.82167 | 0.51667 | 0.87833 | 0.91533 | 0.95213 | 0.86525 | 0.81648 | 0.92585
| Median |  | 0.98586 | 0.94 | 0.925 | 0.92255 | 0.846 | 0.9355 | 0.94 |  | 0.85 | 0.8515 | 0.8505 | 0.837 | 0.93 | 0.93055 | 0.9 | 0.997 | 0.815 | 0.52 | 0.875 | 0.9165 | 0.95105 | 0.8665 | 0.8187 | 0.9242
| Std Dev. |  | 0.053129 | 0.005164 | 0.0251 | 0.0077886 | 0.012323 | 0.025193 | 0.013292 |  | 0.0063246 | 0.010496 | 0.01603 | 0.037237 | 0.0075277 | 0.034482 | 0.020656 | 0.028689 | 0.024833 | 0.015055 | 0.01472 | 0.0092664 | 0.0094312 | 0.015285 | 0.0098538 | 0.012134
| Rel.Std.Dev. |  | 5.326% | 0.5474% | 2.713% | 0.8431% | 1.456% | 2.714% | 1.411% |  | 0.7441% | 1.235% | 1.869% | 4.418% | 0.808% | 3.694% | 2.278% | 2.853% | 3.022% | 2.914% | 1.676% | 1.012% | 0.9905% | 1.767% | 1.207% | 1.311%
| PDM3 |  | 11.06% | 5.026% | 2.985% | 2.847% | -5.737% | 3.356% | 4.84% |  | -5.366% | -5.347% | -4.493% | -6.163% | 3.727% | 3.92% | 0.9434% | 11.97% | -8.52% | -42.48% | -2.211% | 1.908% | 6.005% | -3.668% | -9.097% | 3.079%
| Z-Score (Absolute) |  | NA | 0.59 | 0.26 | 0.23 | 1.18 | 0.32 | 0.56 | NA | 1.11 | 1.11 | 0.97 | 1.25 | 0.38 | 0.41 | 0.08 | 1.73 | 1.63 | 7.21 | 0.6 | 0.08 | 0.75 | 0.84 | 1.73 | 0.27
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 232b (Certified Value 0.81 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*OES | 3
|  |  | 0.085g | 200g | 200g | 200g | 50g | 50g | 20g | 200g | 05g | 30g | 30g | 30g | 30g | 30g | 200g | 50g | 30g | 15g | 3
| 1 | 1 | 0.97716 | 0.78 | 0.85 | 0.88 | 0.775 | 0.84 | 0.81 | 0.81 | 0.9 | 0.88 | 0.79 | 0.85 | 0.77 | 0.79 | 0.72353 | 0.89 | 0.74 | 0.7695 | 1
| 1 | 2 | 1.0392 | 0.81 | 0.81 | 0.92 | 0.724 | 0.84 | 0.79 | 0.82 | 0.9 | 0.83 | 0.81 | 0.82 | 0.76 | 0.76 | 0.72683 | 0.87 | 0.77 | 0.78375 | 3
| 1 | 3 | 1.0128 | 0.79 | 0.8 | 0.9 | 0.764 | 0.87 | 0.84 | 0.81 | 0.9 | 0.85 | 0.81 | 0.85 | 0.76 | 0.77 | 0.70863 | 0.862 | 0.88 | 0.7809 | 16
| 1 | 4 | 0.97448 | 0.82 | 0.89 | 0.88 | 0.762 | 0.85 | 0.87 | 0.8 | 0.9 | 0.84 | 0.81 | 0.82 | 0.78 | 0.77 | 0.71791 | 0.9 | 0.73 | 0.79705 | 0.81009
| 1 | 5 | 0.99525 | 0.81 | 0.82 | 0.84 | 0.745 | 0.84 | 0.85 | 0.8 | 0.9 | 0.84 | 0.84 | 0.82 | 0.78 | 0.79 | 0.73213 | 0.86 | 0.8 | 0.798 | 11
| 1 | 6 | 0.9713 | 0.81 | 0.81 | 0.84 | 0.76 | 0.85 | 0.88 | 0.81 | 0.9 | 0.83 | 0.85 | 0.81 | 0.8 | 0.76 | 0.7573 | 0.848 | 0.77 | 0.81985
|  | 7 | 1.1862
|  | 8 | 0.99124
|  | 9 | 0.93375
|  | 10 | 0.98608
|  | 11 | 0.97386
|  | 12 | 0.98099
|  | 13 | 1.0141
|  | 14 | 1.062
|  | 15 | 0.97661
|  | 16 | 0.94684
|  | 17 | 0.98564
|  | 18 | 0.95613
|  | 19 | 0.98894
|  | 20 | 0.99804
| Mean |  | 0.99753 | 0.80333 | 0.83 | 0.87667 | 0.755 | 0.84833 | 0.84 | 0.80833 | 0.9 | 0.845 | 0.81833 | 0.82833 | 0.775 | 0.77333 | 0.72772 | 0.87167 | 0.78167 | 0.79151
| Median |  | 0.98586 | 0.81 | 0.815 | 0.88 | 0.761 | 0.845 | 0.845 | 0.81 | 0.9 | 0.84 | 0.81 | 0.82 | 0.775 | 0.77 | 0.72518 | 0.866 | 0.77 | 0.7904
| Std Dev. |  | 0.053129 | 0.015055 | 0.034059 | 0.032042 | 0.017978 | 0.01169 | 0.034641 | 0.0075277 | 0 | 0.018708 | 0.022286 | 0.017224 | 0.015166 | 0.013663 | 0.016575 | 0.019654 | 0.054191 | 0.017511
| Rel.Std.Dev. |  | 5.326% | 1.874% | 4.103% | 3.655% | 2.381% | 1.378% | 4.124% | 0.9313% | 0% | 2.214% | 2.723% | 2.079% | 1.957% | 1.767% | 2.278% | 2.255% | 6.933% | 2.212%
| PDM3 |  | 23.14% | -0.834% | 2.458% | 8.219% | -6.8% | 4.721% | 3.692% | -0.2168% | 11.1% | 4.309% | 1.018% | 2.252% | -4.332% | -4.537% | -10.17% | 7.601% | -3.509% | -2.294%
| Z-Score (Absolute) |  | NA | 0.21 | 0.35 | 1.35 | 1.24 | 0.75 | 0.57 | 0.11 | NA | 0.67 | 0.11 | 0.32 | 0.82 | 0.85 | 1.82 | 1.24 | 0.67 | 0.46
Results from laboratory 1.10 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 232b (Certified Value 0.959 ppm) | [6 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | 510±15 g | 510±15 g | 510±15 g | 510±15 g | 3
| 1 | 1 | 0.97716 | 0.95 | 0.95 | 0.86 | 1.1 | 1
| 1 | 2 | 1.0392 | 0.91 | 1 | 0.86 | 0.95 | 3
| 1 | 3 | 1.0128 | 0.84 | 1.03 | 0.86 | 0.92 | 16
| 1 | 4 | 0.97448 | 0.94 | 0.94 | 0.86 | 0.97 | 0.95867
| 1 | 5 | 0.99525 | 0.87 | 1.01 | 0.86 | 0.94 | 13
| 1 | 6 | 0.9713 | 0.93 | 1.01 | 0.86 | 1.02
| 2 | 7 | 1.1862 | 0.981 | 1 | 0.93 | 0.93
| 2 | 8 | 0.99124 | 1.043 | 0.99 | 0.93 | 0.98
| 2 | 9 | 0.93375 | 0.936 | 1.04 | 0.93 | 1.02
| 2 | 10 | 0.98608 | 0.884 | 0.98 | 0.93 | 1.04
| 2 | 11 | 0.97386 | 1.036 | 0.96 | 0.93 | 1.02
| 2 | 12 | 0.98099 | 1.018 | 0.98 | 0.93 | 0.92
| 3 | 13 | 1.0141 |  | 0.92 | 0.92
| 3 | 14 | 1.062 |  | 0.92 | 0.92
| 3 | 15 | 0.97661 |  | 0.89 | 0.92
| 3 | 16 | 0.94684 |  | 1.05 | 0.92
| 3 | 17 | 0.98564 |  | 0.9 | 0.92
| 3 | 18 | 0.95613 |  | 1.04 | 0.92
| 4 | 19 | 0.98894 |  | 0.924 | 0.934
| 4 | 20 | 0.99804 |  | 0.967 | 1.011
| 4 | 21 |  |  | 1.015 | 0.878
| 4 | 22 |  |  | 0.93 | 0.97
| 4 | 23 |  |  | 1.04 | 0.989
| 4 | 24 |  |  | 0.968 | 0.942
| 5 | 25 |  |  | 0.912 | 0.959
| 5 | 26 |  |  | 0.944 | 0.98
| 5 | 27 |  |  | 0.994 | 1.013
| 5 | 28 |  |  | 0.991 | 0.975
| 5 | 29 |  |  | 0.982 | 0.89
| 5 | 30 |  |  | 0.897 | 0.913
| 6 | 31 |  |  | 1.0414 | 0.921
| 6 | 32 |  |  | 0.95946 | 0.96
| 6 | 33 |  |  | 0.95432 | 1.001
| 6 | 34 |  |  | 1.0338 | 0.977
| 6 | 35 |  |  | 0.98926 | 1.037
| 6 | 36 |  |  | 0.95863 | 0.884
| Mean |  | 0.99753 | 0.94483 | 0.9753 | 0.93039 | 0.98417
| Median |  | 0.98586 | 0.938 | 0.98 | 0.93 | 0.975
| Std Dev. |  | 0.053129 | 0.064991 | 0.04576 | 0.047941 | 0.05632
| Rel.Std.Dev. |  | 5.326% | 6.879% | 4.692% | 5.153% | 5.723%
| PDM3 |  | 4.053% | -1.444% | 1.734% | -2.95% | 2.659%
| Z-Score (Absolute) |  | NA | 0.52 | 0.52 | 1.02 | 0.83
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 232b (Certified Value 0.113 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.13 | < 0.2 | 0.13 | 0.7804 | 0.1 | 0.21 | 0.2 | 0.3 | 0.1 | 0.11 | 0.1 | 0.12 | 0.08 | < 1 | 0.28 | 0.19 | < 0.5 | 0.06 | 0.12 | 0.232 | 0.31306 | 1.4661 |  |  |  |  | 1
| 1 | 2 | 0.13 | < 0.2 | 0.13 | 0.77 | 0.1 | 0.22 | 0.2 | 0.2 | 0.09 | 0.08 | 0.1 | 0.12 | 0.08 | < 1 | 0.28 | 0.18 | < 0.5 | 0.08 | 0.11 | 0.12 | 0.28738 | 1.3409 |  |  |  |  | 25
| 1 | 3 | 0.12 | < 0.2 | 0.11 | 0.77577 | 0.1 | 0.2 | 0.2 | 0.2 | 0.09 | 0.09 | 0.12 | 0.12 | 0.07 | < 1 | 0.26 | 0.17 | < 0.5 | 0.09 | 0.11 | 0.119 | 0.27511 | 1.2984 |  |  |  |  | 16
| 1 | 4 | 0.11 | < 0.2 | 0.13 | 0.8008 | 0.1 | 0.23 | 0.2 | 0.2 | 0.08 | 0.09 | 0.11 | 0.11 | 0.06 | < 1 | 0.26 | 0.21 | < 0.5 | 0.11 | 0.12 | 0.122 | 0.24253 | 1.2379 |  |  |  |  | 0.1126
| 1 | 5 | 0.12 | < 0.2 | 0.15 | 0.79313 | 0.1 | 0.26 | 0.2 | 0.2 | 0.1 | 0.08 | 0.12 | 0.12 | 0.07 | < 1 | 0.28 | 0.17 | < 0.5 | 0.12 | 0.11 | 0.119 | 0.28711 | 1.2379 |  |  |  |  | 15
| 1 | 6 | 0.11 | < 0.2 | 0.12 | 0.78271 | 0.1 | 0.24 | 0.2 | 0.2 | 0.09 | 0.09 | 0.13 | 0.11 | 0.06 | < 1 | 0.25 | 0.2 | < 0.5 | 0.11 | 0.12 | 0.103 | 0.2551 | 1.2157
| Mean |  | 0.12 |  | 0.12833 | 0.7838 | 0.1 | 0.22667 | 0.2 | 0.21667 | 0.091667 | 0.09 | 0.11333 | 0.11667 | 0.07 |  | 0.26833 | 0.18667 |  | 0.095 | 0.115 | 0.13583 | 0.27671 | 1.2995
| Median |  | 0.12 |  | 0.13 | 0.78155 | 0.1 | 0.225 | 0.2 | 0.2 | 0.09 | 0.09 | 0.115 | 0.12 | 0.07 |  | 0.27 | 0.185 |  | 0.1 | 0.115 | 0.1195 | 0.28111 | 1.2681
| Std Dev. |  | 0.0089443 |  | 0.013292 | 0.011348 | 1.5202e-17 | 0.021602 | 3.0405e-17 | 0.040825 | 0.0075277 | 0.010954 | 0.012111 | 0.005164 | 0.0089443 |  | 0.013292 | 0.01633 |  | 0.022583 | 0.0054772 | 0.047613 | 0.025224 | 0.093896
| Rel.Std.Dev. |  | 7.454% |  | 10.36% | 1.448% | 1.52e-14% | 9.531% | 1.52e-14% | 18.84% | 8.212% | 12.17% | 10.69% | 4.426% | 12.78% |  | 4.953% | 8.748% |  | 23.77% | 4.763% | 35.05% | 9.115% | 7.226%
| PDM3 |  | 6.572% |  | 13.97% | 596.1% | -11.19% | 101.3% | 77.62% | 92.42% | -18.59% | -20.07% | 0.6513% | 3.612% | -37.83% |  | 138.3% | 65.78% |  | -15.63% | 2.131% | 20.63% | 145.8% | 1054%
| Z-Score (Absolute) |  | 0.15 | 0.5 | 0 | 11.53 | NA | 1.73 | NA | NA | 0.64 | 0.67 | 0.26 | 0.21 | 1.03 | 6.54 | 2.46 | 1.03 | 2.14 | 0.59 | 0.23 | 0.13 | 2.61 | 20.6
Results from laboratories 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 232b (Certified Value 7.01 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 7.0125 | 7.0228 | 6.76 | 6.69 | 6.987 | 7.121 | 4.71 | 7.07 | 6.81 | 7.07 | 6.83 | 7.11 | 7.22 | 6.567 | 6.8899 | 7.092 | 7.14 | 7.07 | 6.76 | 7.14 | 7.33 | 7.0475 | 7.0234 | 7.5185 |  |  | 1
| 1 | 2 | 7.0417 | 7.0425 | 6.84 | 6.68 | 7.02 | 7.156 | 4.32 | 7.2 | 6.79 | 6.88 | 6.95 | 7.5 | 7.07 | 6.4377 | 6.8049 | 6.933 | 7.12 | 7.08 | 6.82 | 7.14 | 7.24 | 6.9258 | 7.1759 | 7.4961 |  |  | #N/A
| 1 | 3 | 6.9405 | 7.0446 | 6.82 | 6.46 | 6.98 | 7.059 | 4.95 | 7.18 | 6.63 | 6.78 | 6.82 | 7.21 | 7.09 | 6.6408 | 6.8732 | 7.039 | 6.97 | 7 | 6.65 | 7.12 | 7.23 | 7.0181 | 7.0999 | 7.5197 |  |  | 16
| 1 | 4 | 6.9623 | 6.9957 | 6.79 | 6.4 | 6.9892 | 7.04 | 4.9 | 7.09 | 6.61 | 6.93 | 7.39 | 7.01 | 7.05 | 6.8296 | 6.9619 | 7.092 | 7.29 | 7.07 | 7 | 7.16 | 7.33 | 7.1717 | 7.3302 | 7.4444 |  |  | 7.0119
| 1 | 5 | 7.1475 | 7.0407 | 6.89 | 6.15 | 7.02 | 7.191 | 5.1 | 7.27 | 6.59 | 6.95 | 6.64 | 7.28 | 7.15 | 6.8609 | 6.8458 | 6.88 | 7.24 | 7.01 | 6.99 | 7.14 | 7.11 | 6.9698 | 6.9402 | 7.517 |  |  | 16
| 1 | 6 | 6.9103 | 7.0143 | 6.73 | 6.71 | 7.0176 | 7.067 | 4.36 | 7.3 | 6.79 | 6.95 | 7.21 | 6.95 | 7.2 | 6.6367 | 6.937 | 7.198 | 7.17 | 7.1 | 6.96 | 7.11 | 6.63 | 6.9871 | 7.0584 | 7.5219
| Mean |  | 7.0025 | 7.0268 | 6.805 | 6.515 | 7.0023 | 7.1057 | 4.7233 | 7.185 | 6.7033 | 6.9267 | 6.9733 | 7.1767 | 7.13 | 6.6621 | 6.8855 | 7.039 | 7.155 | 7.055 | 6.8633 | 7.135 | 7.145 | 7.02 | 7.1047 | 7.5029
| Median |  | 6.9874 | 7.0317 | 6.805 | 6.57 | 7.0034 | 7.094 | 4.805 | 7.19 | 6.71 | 6.94 | 6.89 | 7.16 | 7.12 | 6.6387 | 6.8815 | 7.0655 | 7.155 | 7.07 | 6.89 | 7.14 | 7.235 | 7.0026 | 7.0792 | 7.5178
| Std Dev. |  | 0.085583 | 0.019476 | 0.057533 | 0.22152 | 0.018781 | 0.060059 | 0.32228 | 0.092682 | 0.10328 | 0.095219 | 0.27761 | 0.19997 | 0.070711 | 0.16005 | 0.057838 | 0.11612 | 0.11077 | 0.040373 | 0.14292 | 0.017607 | 0.26501 | 0.085146 | 0.13547 | 0.030179
| Rel.Std.Dev. |  | 1.222% | 0.2772% | 0.8454% | 3.4% | 0.2682% | 0.8452% | 6.823% | 1.29% | 1.541% | 1.375% | 3.981% | 2.786% | 0.9917% | 2.402% | 0.84% | 1.65% | 1.548% | 0.5723% | 2.082% | 0.2468% | 3.709% | 1.213% | 1.907% | 0.4022%
| PDM3 |  | -0.1343% | 0.2121% | -2.95% | -7.086% | -0.1367% | 1.337% | -32.64% | 2.469% | -4.4% | -1.215% | -0.5498% | 2.35% | 1.685% | -4.988% | -1.803% | 0.3867% | 2.041% | 0.6149% | -2.119% | 1.756% | 1.898% | 0.1158% | 1.323% | 7.003%
| Z-Score (Absolute) |  | 0.12 | 0.02 | 1.26 | 2.94 | 0.12 | 0.48 | 13.3 | 0.93 | 1.85 | 0.56 | 0.29 | 0.89 | 0.62 | 2.09 | 0.8 | 0.09 | 0.76 | 0.18 | 0.93 | 0.65 | 0.7 | 0.02 | 0.47 | 2.77
| Analytical results for As in OREAS 232b (Certified Value 441 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 0
0
| 1 | 1 | 456.8 | 414 | 430 | 449.0255 | 437 | 435 | 441 | 45.1 | 458 | 444 | 424 | 481 | 453.6 | 417.2426 | 400 | 452 | 449.2 | 473.4 | 181.8 | 444.9 | 729.4345 | [5 empty] | 1
| 1 | 2 | 459.5 | 412 | 403 | 445.107 | 424 | 455 | 423 | 45.1 | 445 | 437 | 448 | 471 | 452.9 | 406.3555 | 431 | 450.3 | 432.4 | 463.4 | 163.9 | 431.7 | 674.093 | [5 empty] | 26
| 1 | 3 | 464.3 | 413 | 377 | 443.89 | 445 | 460 | 439 | 50.2 | 429 | 424 | 436 | 467 | 451.3 | 415.0918 | 411 | 456.6 | 415.4 | 459.8 | 190.4 | 434.8 | 662.5026 | [5 empty] | 16
| 1 | 4 | 443.6 | 422 | 396 | 447.3 | 430 | 458 | 449 | 52.8 | 429 | 467 | 415 | 482 | 459.4 | 389.6652 | 438 | 459.9 | 419.3 | 494.9 | 175.9 | 442 | 664.5799 | [5 empty] | 441.1417
| 1 | 5 | 455.5 | 404 | 348 | 449.7414 | 443 | 460 | 427 | 44.2 | 426 | 409 | 451 | 463 | 463.6 | 409.5823 | 427 | 462.6 | 430.7 | 484.9 | 189.2 | 436.4 | 687.2985 | [5 empty] | 17
| 1 | 6 | 452.4 | 417 | 464 | 450.32 | 454 | 454 | 405 | 50.4 | 442 | 467 | 442 | 482 | 464.9 | 404.8786 | 450 | 450.5 | 459.9 | 469.1 | 211.8 | 439.8 | 648.1288
| Mean |  | 455.35 | 413.6667 | 403 | 447.564 | 438.8333 | 453.6667 | 430.6667 | 47.9667 | 438.1667 | 441.3333 | 436 | 474.3333 | 457.6167 | 407.136 | 426.1667 | 455.3167 | 434.4833 | 474.25 | 185.5 | 438.2667 | 677.6729
| Median |  | 456.15 | 413.5 | 399.5 | 448.1628 | 440 | 456.5 | 433 | 47.65 | 435.5 | 440.5 | 439 | 476 | 456.5 | 407.9689 | 429 | 454.3 | 431.55 | 471.25 | 185.5 | 438.1 | 669.3365
| Std Dev. |  | 7.0173 | 5.9554 | 40.4969 | 2.6107 | 10.8336 | 9.4798 | 15.7692 | 3.6026 | 12.4164 | 23.192 | 14.0712 | 8.4301 | 5.8362 | 9.822 | 18.1485 | 5.1975 | 17.1811 | 13.3687 | 16.1357 | 4.8767 | 28.4926
| Rel.Std.Dev. |  | 1.541% | 1.44% | 10.05% | 0.5833% | 2.469% | 2.09% | 3.662% | 7.511% | 2.834% | 5.255% | 3.227% | 1.777% | 1.275% | 2.412% | 4.259% | 1.142% | 3.954% | 2.819% | 8.698% | 1.113% | 4.204%
| PDM3 |  | 3.221% | -6.228% | -8.646% | 1.456% | -0.5233% | 2.839% | -2.375% | -89.13% | -0.6744% | 0.04345% | -1.166% | 7.524% | 3.735% | -7.709% | -3.395% | 3.213% | -1.509% | 7.505% | -57.95% | -0.6517% | 53.62%
| Z-Score (Absolute) |  | 0.68 | 0.97 | 1.39 | 0.37 | 0.02 | 0.61 | 0.3 | 15.44 | 0 | 0.12 | 0.09 | 1.43 | 0.77 | 1.23 | 0.48 | 0.67 | 0.15 | 1.42 | 10 | 0 | 9.47
| Analytical results for Au in OREAS 232b (Indicative Value 0.918 ppm) | [27 empty] | Indicative
Au | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.8944 | [25 empty] | 1
| 1 | 2 | 0.88488 | [25 empty] | 3
| 1 | 3 | 0.9472 | [25 empty] | 16
| 1 | 4 | 0.95888 | [25 empty] | 0.91773
| 1 | 5 | 0.89784 | [25 empty] | 9
| 1 | 6 | 0.9232
| Mean |  | 0.91773
| Median |  | 0.91052
| Std Dev. |  | 0.030367
| Rel.Std.Dev. |  | 3.309%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for B in OREAS 232b (Indicative Value 26.6 ppm) | [27 empty] | Indicative
B | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.07 | [25 empty] | ppm
|  |  | 4A*OES | [25 empty] | 1
1
| 1 | 1 | 27.7 | [25 empty] | 1
| 1 | 2 | 28.27 | [25 empty] | 4
| 1 | 3 | 25.81 | [25 empty] | 16
| 1 | 4 | 25.28 | [25 empty] | 26.5533
| 1 | 5 | 24.84 | [25 empty] | 10
| 1 | 6 | 27.42
| Mean |  | 26.5533
| Median |  | 26.615
| Std Dev. |  | 1.4228
| Rel.Std.Dev. |  | 5.358%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 232b (Certified Value 694 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 695.9 | 690.9981 | 686 | 751 | 662.82 | 720 | 690 | 701 | 678 | 710 | 680 | 700 | 710 | 665 | 665.5505 | 691 | 595.7 | 710 | 695.4 | 682 | 690 | 691 | 680.7 | 692 | 717.3981 | 684.9141 | 1
| 1 | 2 | 695.8 | 692.6628 | 690 | 767 | 675.14 | 729 | 700 | 693 | 677 | 690 | 710 | 740 | 700 | 672 | 682.3221 | 696 | 605.6 | 696 | 704.2 | 666 | 689 | 690 | 661 | 697 | 704.7561 | 665.1235 | 27
| 1 | 3 | 701.6 | 705.0766 | 718 | 733 | 671.104 | 716 | 683 | 727 | 678 | 690 | 680 | 710 | 700 | 671 | 688.8523 | 701 | 622.7 | 697 | 701.7 | 635 | 692 | 699 | 688.2 | 690 | 701.6754 | 685.4696 | 16
| 1 | 4 | 694.3 | 697.6597 | 693 | 754 | 664.28 | 724 | 695 | 729 | 680 | 700 | 730 | 690 | 700 | 683 | 678.548 | 699 | 620.8 | 723 | 698.2 | 673 | 693 | 672 | 695.7 | 708 | 704.8175 | 695.1024 | 693.8677
| 1 | 5 | 700.3 | 698.2845 | 720 | 730 | 671.84 | 730 | 708 | 713 | 677 | 700 | 660 | 720 | 710 | 676 | 658.5802 | 682 | 622.9 | 716 | 691.9 | 680 | 691 | 686 | 674.1 | 679 | 715.6798 | 689.5119 | 18
| 1 | 6 | 696.4 | 696.9835 | 707 | 740 | 666.519 | 722 | 706 | 772 | 672 | 700 | 720 | 680 | 710 | 660 | 666.8122 | 716 | 621.4 | 690 | 699.9 | 702 | 688 | 672 | 682 | 688 | 726.2843 | 685.6852
| Mean |  | 697.3833 | 696.9442 | 702.3333 | 745.8333 | 668.6172 | 723.5 | 697 | 722.5 | 677 | 698.3333 | 696.6667 | 706.6667 | 705 | 671.1667 | 673.4442 | 697.5 | 614.85 | 705.3333 | 698.55 | 673 | 690.5 | 685 | 680.2833 | 692.3333 | 711.7685 | 684.3011
| Median |  | 696.15 | 697.3216 | 700 | 745.5 | 668.8115 | 723 | 697.5 | 720 | 677.5 | 700 | 695 | 705 | 705 | 671.5 | 672.6801 | 697.5 | 621.1 | 703.5 | 699.05 | 676.5 | 690.5 | 688 | 681.35 | 691 | 710.2486 | 685.5774
| Std Dev. |  | 2.8799 | 4.9423 | 14.7332 | 14.0772 | 4.8154 | 5.3572 | 9.5917 | 28.0553 | 2.6833 | 7.5277 | 27.3252 | 21.6025 | 5.4772 | 8.085 | 11.5716 | 11.327 | 11.4631 | 12.9563 | 4.4257 | 22.1991 | 1.8708 | 10.9179 | 11.9461 | 9.6885 | 9.561 | 10.1514
| Rel.Std.Dev. |  | 0.413% | 0.7091% | 2.098% | 1.887% | 0.7202% | 0.7405% | 1.376% | 3.883% | 0.3963% | 1.078% | 3.922% | 3.057% | 0.7769% | 1.205% | 1.718% | 1.624% | 1.864% | 1.837% | 0.6336% | 3.299% | 0.2709% | 1.594% | 1.756% | 1.399% | 1.343% | 1.483%
| PDM3 |  | 0.5067% | 0.4434% | 1.22% | 7.489% | -3.639% | 4.271% | 0.4514% | 4.126% | -2.431% | 0.6436% | 0.4034% | 1.845% | 1.604% | -3.272% | -2.943% | 0.5235% | -11.39% | 1.652% | 0.6748% | -3.007% | -0.4854% | -1.278% | -1.958% | -0.2211% | 2.58% | -1.379%
| Z-Score (Absolute) |  | 0.03 | 0 | 0.33 | 3.04 | 1.76 | 1.65 | 0 | 1.59 | 1.24 | 0.08 | 0.02 | 0.6 | 0.5 | 1.61 | 1.46 | 0.03 | 5.11 | 0.52 | 0.1 | 1.49 | 0.4 | 0.75 | 1.04 | 0.29 | 0.92 | 0.79
| Analytical results for Be in OREAS 232b (Certified Value 2.35 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.41 | 2.5 | 2.4 | < 5 | 2.2 | 2.3 | 1.8 | 2 | 2.3 | 2.53 | 2.35 | 2.28 | 2 | 2.3565 | 2.9 | 2.047 | 2.5 | 3.05 | 2.53 | 2.33 | 2 | 2.3353 | 2.7755 | 2.3507 |  |  | 1
| 1 | 2 | 2.32 | 2.5 | 2.3 | < 5 | 2.1 | 2.29 | 3 | 2 | 2.22 | 2.48 | 2.38 | 2.26 | 3 | 2.3279 | 2.8 | 2.07 | 2.7 | 3.15 | 2.31 | 2.38 | 2 | 2.2638 | 2.8311 | 2.3445 |  |  | 28
| 1 | 3 | 2.37 | 2.5 | 2.2 | < 5 | 2.2 | 2.25 | 2.6 | 2 | 2.25 | 2.46 | 2.32 | 2.18 | 2 | 2.3361 | 3 | 2.121 | 2.6 | 3.16 | 2.47 | 2.28 | 3 | 2.3584 | 2.7755 | 2.3318 |  |  | 16
| 1 | 4 | 2.31 | 2.5 | 2.2 | < 5 | 2.1 | 2.38 | 1.9 | 2 | 2.25 | 2.69 | 2.46 | 2.14 | 3 | 2.3595 | 2.9 | 2.04 | 2.7 | 3.1 | 2.22 | 2.28 | 2 | 2.3859 | 2.4417 | 2.3443 |  |  | 2.3482
| 1 | 5 | 2.36 | 2.5 | 2.1 | < 5 | 2.1 | 2.33 | 3 | 2 | 2.16 | 2.47 | 2.54 | 2.21 | 3 | 2.3592 | 3 | 2.076 | 2.7 | 3.13 | 2.68 | 2.3 | 2 | 2.3111 | 2.7755 | 2.2971 |  |  | 19
| 1 | 6 | 2.33 | 2.5 | 2.4 | < 5 | 2.1 | 2.43 | 2.5 | 2 | 2.24 | 2.55 | 2.46 | 2.11 | 3 | 2.331 | 3 | 2.038 | 2.8 | 3.15 | 2.21 | 2.34 | 2 | 2.3287 | 2.3861 | 2.4034
| Mean |  | 2.35 | 2.5 | 2.2667 |  | 2.1333 | 2.33 | 2.4667 | 2 | 2.2367 | 2.53 | 2.4183 | 2.1967 | 2.6667 | 2.345 | 2.9333 | 2.0653 | 2.6667 | 3.1233 | 2.4033 | 2.3183 | 2.1667 | 2.3305 | 2.6642 | 2.3453
| Median |  | 2.345 | 2.5 | 2.25 |  | 2.1 | 2.315 | 2.55 | 2 | 2.245 | 2.505 | 2.42 | 2.195 | 3 | 2.3463 | 2.95 | 2.0585 | 2.7 | 3.14 | 2.39 | 2.315 | 2 | 2.332 | 2.7755 | 2.3444
| Std Dev. |  | 0.037417 | 0 | 0.12111 |  | 0.05164 | 0.065422 | 0.52026 | 0 | 0.045898 | 0.086023 | 0.082563 | 0.066533 | 0.5164 | 0.014898 | 0.08165 | 0.031494 | 0.10328 | 0.041793 | 0.18801 | 0.0392 | 0.40825 | 0.041711 | 0.19589 | 0.034411
| Rel.Std.Dev. |  | 1.592% | 0% | 5.343% |  | 2.421% | 2.808% | 21.09% | 0% | 2.052% | 3.4% | 3.414% | 3.029% | 19.36% | 0.6353% | 2.784% | 1.525% | 3.873% | 1.338% | 7.823% | 1.691% | 18.84% | 1.79% | 7.353% | 1.467%
| PDM3 |  | 0.0747% | 6.462% | -3.474% |  | -9.152% | -0.777% | 5.043% | -14.83% | -4.752% | 7.74% | 2.985% | -6.455% | 13.56% | -0.1366% | 24.92% | -12.05% | 13.56% | 33.01% | 2.346% | -1.274% | -7.733% | -0.7543% | 13.46% | -0.126%
| Z-Score (Absolute) |  | 0 | 0.87 | 0.48 | 0.87 | 1.25 | 0.12 | 0.67 | NA | 0.66 | 1.04 | 0.39 | 0.89 | NA | 0.03 | 3.37 | 1.65 | 1.83 | 4.47 | 0.31 | 0.18 | NA | 0.11 | 1.82 | 0.03
Results from laboratories 1.11, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 232b (Certified Value 0.31 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 0.31 | 0.29755 | 0.4 | 0.31 | < 5 | 0.29 | 0.27 | 0.33 | 0.5 | 0.31 | 0.3 | 0.34 | 0.33 | 0.29 | 0.31284 | 0.5 | 0.296 | 0.25 | < 5 | 0.34 | 0.34 | 0.31 | 0.6575 | < 5 | 1.3862 | 0.34049 | 1
| 1 | 2 | 0.35 | 0.3184 | 0.3 | 0.32 | < 5 | 0.3 | 0.26 | 0.34 | 0.4 | 0.3 | 0.28 | 0.32 | 0.32 | 0.29 | 0.30664 | 0.4 | 0.283 | 0.26 | < 5 | 0.31 | 0.34 | 0.33 | 0.6881 | < 5 | 1.1005 | 0.29257 | 29
| 1 | 3 | 0.32 | 0.3228 | 0.3 | 0.38 | < 5 | 0.31 | 0.25 | 0.36 | 0.4 | 0.28 | 0.26 | 0.31 | 0.28 | 0.28 | 0.30507 | 0.4 | 0.266 | 0.26 | < 5 | 0.29 | 0.3 | 0.37 | 0.7051 | < 5 | 1.8471 | 0.32242 | 16
| 1 | 4 | 0.32 | 0.29696 | 0.4 | 0.29 | < 5 | 0.31 | 0.27 | 0.3 | 0.3 | 0.29 | 0.34 | 0.35 | 0.28 | 0.29 | 0.31421 | 0.4 | 0.271 | 0.27 | < 5 | 0.3 | 0.31 | 0.3 | 0.6098 | < 5 | 1.9102 | 0.32365 | 0.30852
| 1 | 5 | 0.36 | 0.33119 | 0.3 | 0.33 | < 5 | 0.29 | 0.33 | 0.31 | 0.4 | 0.28 | 0.28 | 0.3 | 0.36 | 0.28 | 0.31522 | 0.4 | 0.289 | 0.26 | < 5 | 0.31 | 0.3 | 0.29 | 0.5755 | < 5 | 1.6403 | 0.31295 | 20
| 1 | 6 | 0.32 | 0.31726 | 0.3 | 0.29 | < 5 | 0.32 | 0.32 | 0.55 | 0.4 | 0.29 | 0.37 | 0.36 | 0.3 | 0.3 | 0.3055 | 0.4 | 0.282 | 0.26 | < 5 | 0.31 | 0.3 | 0.31 | 0.633 | < 5 | 1.9043 | 0.30045
| Mean |  | 0.33 | 0.31403 | 0.33333 | 0.32 |  | 0.30333 | 0.28333 | 0.365 | 0.4 | 0.29167 | 0.305 | 0.33 | 0.31167 | 0.28833 | 0.30991 | 0.41667 | 0.28117 | 0.26 |  | 0.31 | 0.315 | 0.31833 | 0.64483 |  | 1.6314 | 0.31542
| Median |  | 0.32 | 0.31783 | 0.3 | 0.315 |  | 0.305 | 0.27 | 0.335 | 0.4 | 0.29 | 0.29 | 0.33 | 0.31 | 0.29 | 0.30974 | 0.4 | 0.2825 | 0.26 |  | 0.31 | 0.305 | 0.31 | 0.64525 |  | 1.7437 | 0.31768
| Std Dev. |  | 0.02 | 0.013885 | 0.05164 | 0.033466 |  | 0.012111 | 0.033267 | 0.093113 | 0.063246 | 0.01169 | 0.041833 | 0.023664 | 0.031252 | 0.0075277 | 0.0046653 | 0.040825 | 0.011125 | 0.0063246 |  | 0.016733 | 0.019748 | 0.028577 | 0.04866 |  | 0.32887 | 0.017306
| Rel.Std.Dev. |  | 6.061% | 4.421% | 15.49% | 10.46% |  | 3.993% | 11.74% | 25.51% | 15.81% | 4.008% | 13.72% | 7.171% | 10.03% | 2.611% | 1.505% | 9.798% | 3.957% | 2.433% |  | 5.398% | 6.269% | 8.977% | 7.546% |  | 20.16% | 5.487%
| PDM3 |  | 6.962% | 1.785% | 8.042% | 3.721% |  | -1.682% | -8.164% | 18.31% | 29.65% | -5.463% | -1.141% | 6.962% | 1.019% | -6.543% | 0.4512% | 35.05% | -8.866% | -15.73% |  | 0.4793% | 2.1% | 3.18% | 109% |  | 428.8% | 2.237%
| Z-Score (Absolute) |  | 0.67 | 0.04 | NA | 0.22 | 98.22 | 0.52 | 1.42 | 2.25 | NA | 1.05 | 0.45 | 0.67 | 0.15 | 1.2 | 0.23 | NA | 1.52 | 2.47 | 98.22 | 0.22 | 0 | 0.15 | 14.83 | 98.22 | 59.18 | 0.02
Results from laboratories 1.05, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 232b (Certified Value 0.995 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.9813 | 0.98422 | 0.98 | 0.9 | 0.99 | 1.036 | 0.83 | 0.962 | 1.03 | 0.99 | 0.98 | 1.03 | 1.05 | 0.918 | 0.99216 | 0.908 | 0.99 | 1.04 | 0.97 | 0.99 | 1 | 0.9823 | 1.0117 |  |  |  | 1
| 1 | 2 | 0.9948 | 0.9899 | 0.96 | 0.9 | 1.0133 | 1.025 | 0.78 | 0.983 | 1.03 | 0.97 | 0.99 | 1.08 | 1.02 | 0.9149 | 0.99373 | 0.886 | 0.98 | 1.05 | 0.98 | 0.98 | 1 | 0.94959 | 1.0292 |  |  |  | #N/A
| 1 | 3 | 0.984 | 0.9871 | 0.97 | 0.93 | 1 | 1.02 | 0.78 | 0.985 | 0.99 | 0.97 | 0.97 | 1.04 | 1.03 | 0.9353 | 0.96892 | 0.908 | 0.96 | 1.05 | 0.96 | 0.98 | 0.99 | 0.99349 | 1.0181 |  |  |  | 16
| 1 | 4 | 0.9941 | 0.986 | 0.96 | 0.88 | 1 | 1.018 | 0.81 | 0.967 | 0.97 | 0.99 | 1.04 | 1.01 | 1.02 | 0.9537 | 0.98588 | 0.915 | 1.02 | 1.04 | 1 | 0.99 | 1.01 | 1.0075 | 1.0389 |  |  |  | 0.9952
| 1 | 5 | 0.9881 | 0.9882 | 0.96 | 0.85 | 1.03 | 1.03 | 0.91 | 0.99 | 0.98 | 0.99 | 0.95 | 1.04 | 1.04 | 0.9639 | 0.97791 | 0.886 | 1 | 1.03 | 1 | 0.98 | 0.96 | 0.97939 | 0.9925 |  |  |  | 21
| 1 | 6 | 0.9792 | 0.9836 | 0.96 | 0.99 | 0.98667 | 1.037 | 0.84 | 0.999 | 1 | 0.99 | 1.02 | 0.99 | 1.04 | 0.9272 | 0.97685 | 0.922 | 1 | 1.04 | 1 | 0.98 | 0.92 | 0.98333 | 1.0033
| Mean |  | 0.98692 | 0.9865 | 0.965 | 0.90833 | 1.0033 | 1.0277 | 0.825 | 0.981 | 1 | 0.98333 | 0.99167 | 1.0317 | 1.0333 | 0.9355 | 0.98257 | 0.90417 | 0.99167 | 1.0417 | 0.985 | 0.98333 | 0.98 | 0.9826 | 1.0156
| Median |  | 0.98605 | 0.98655 | 0.96 | 0.9 | 1 | 1.0275 | 0.82 | 0.984 | 0.995 | 0.99 | 0.985 | 1.035 | 1.035 | 0.93125 | 0.9819 | 0.908 | 0.995 | 1.04 | 0.99 | 0.98 | 0.995 | 0.98282 | 1.0149
| Std Dev. |  | 0.0065536 | 0.0023945 | 0.0083666 | 0.047924 | 0.016055 | 0.0080166 | 0.048477 | 0.014014 | 0.025298 | 0.010328 | 0.033116 | 0.030605 | 0.012111 | 0.019684 | 0.0096744 | 0.014999 | 0.020412 | 0.0075277 | 0.017607 | 0.005164 | 0.034059 | 0.019177 | 0.016936
| Rel.Std.Dev. |  | 0.664% | 0.2427% | 0.867% | 5.276% | 1.6% | 0.7801% | 5.876% | 1.429% | 2.53% | 1.05% | 3.339% | 2.967% | 1.172% | 2.104% | 0.9846% | 1.659% | 2.058% | 0.7227% | 1.787% | 0.5252% | 3.475% | 1.952% | 1.668%
| PDM3 |  | -0.832% | -0.8735% | -3.034% | -8.728% | 0.8176% | 3.263% | -17.1% | -1.427% | 0.4826% | -1.192% | -0.3547% | 3.665% | 3.832% | -5.998% | -1.268% | -9.147% | -0.3547% | 4.669% | -1.025% | -1.192% | -1.527% | -1.266% | 2.052%
| Z-Score (Absolute) |  | 0.09 | 0.07 | 0.9 | 3.45 | 0.82 | 1.92 | 7.19 | 0.18 | 0.67 | 0.07 | 0.3 | 2.1 | 2.17 | 2.23 | 0.11 | 3.64 | 0.3 | 2.55 | 0 | 0.07 | 0.22 | 0.11 | 1.38
| Analytical results for Cd in OREAS 232b (Indicative Value 0.072 ppm) | [27 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.06 | 1.25 | < 0.5 | 0.4 | 0.79 | 0.05 | 0.11 | 0.34 | 0.2 | 0.04 | 0.03 | 0.04 | 0.04 | 0.11 | 0.11709 | < 0.1 | 0.12 | 0.06 | 4.8 | 0.05 | 0.05 | 0.07 | 0.1205 | < 5 | 0.2385 | 0.06242 | 1
| 1 | 2 | 0.04 | 1.29 | < 0.5 | < 0.3 | 0.824 | 0.06 | 0.11 | 0.29 | 0.2 | 0.05 | 0.03 | 0.04 | 0.04 | 0.11 | 0.10883 | 0.1 | 0.105 | 0.06 | 4.38 | 0.09 | 0.05 | 0.04 | 0.1392 | < 5 | 0.242 | 0.04956 | 5
| 1 | 3 | 0.06 | 1.28 | < 0.5 | 0.4 | 0.83 | 0.07 | 0.1 | 0.31 | 0.3 | 0.05 | 0.05 | 0.03 | 0.05 | 0.09 | 0.12133 | < 0.1 | 0.101 | 0.05 | 4.47 | 0.06 | 0.05 | 0.05 | 0.1459 | < 5 | 0.2385 | 0.05638 | 16
| 1 | 4 | 0.07 | 1.22 | < 0.5 | < 0.3 | 0.82 | 0.06 | 0.11 | 0.24 | 0.3 | 0.05 | 0.05 | 0.04 | 0.04 | 0.11 | 0.12719 | < 0.1 | < 0.1 | 0.06 | 4.79 | 0.06 | 0.05 | 0.05 | 0.1433 | < 5 | 0.2759 | 0.06386 | 0.07153
| 1 | 5 | 0.06 | 1.21 | < 0.5 | 0.3 | 0.79167 | 0.05 | 0.11 | 0.28 | 0.2 | 0.04 | 0.05 | 0.03 | 0.04 | 0.11 | 0.11031 | 0.1 | < 0.1 | 0.06 | 5.03 | 0.07 | 0.05 | 0.03 | 0.1486 | < 5 | 0.2314 | 0.05847 | 11
| 1 | 6 | 0.06 | 1.32 | < 0.5 | 0.6 | 0.78 | 0.07 | 0.11 | 0.22 | 0.3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.11 | 0.12093 | < 0.1 | 0.101 | 0.06 | 4.73 | 0.04 | 0.04 | 0.07 | 0.1378 | < 5 | 0.2438 | 0.06228
| Mean |  | 0.058333 | 1.2617 |  | 0.425 | 0.80594 | 0.06 | 0.10833 | 0.28 | 0.25 | 0.045 | 0.041667 | 0.036667 | 0.041667 | 0.10667 | 0.11761 | 0.1 | 0.10675 | 0.058333 | 4.7 | 0.061667 | 0.048333 | 0.051667 | 0.13922 |  | 0.24502 | 0.058828
| Median |  | 0.06 | 1.265 |  | 0.4 | 0.80583 | 0.06 | 0.11 | 0.285 | 0.25 | 0.045 | 0.045 | 0.04 | 0.04 | 0.11 | 0.11901 | 0.1 | 0.103 | 0.06 | 4.76 | 0.06 | 0.05 | 0.05 | 0.14125 |  | 0.24025 | 0.060375
| Std Dev. |  | 0.0098319 | 0.042622 |  | 0.12583 | 0.021135 | 0.0089443 | 0.0040825 | 0.044272 | 0.054772 | 0.0054772 | 0.0098319 | 0.005164 | 0.0040825 | 0.008165 | 0.0070317 | 0 | 0.0090323 | 0.0040825 | 0.23799 | 0.017224 | 0.0040825 | 0.016021 | 0.010017 |  | 0.015714 | 0.0053321
| Rel.Std.Dev. |  | 16.85% | 3.378% |  | 29.61% | 2.622% | 14.91% | 3.768% | 15.81% | 21.91% | 12.17% | 23.6% | 14.08% | 9.798% | 7.655% | 5.979% | 0% | 8.461% | 6.999% | 5.064% | 27.93% | 8.447% | 31.01% | 7.195% |  | 6.414% | 9.064%
| PDM3 |  | -18.45% | 1664% |  | 494.2% | 1027% | -16.12% | 51.45% | 291.4% | 249.5% | -37.09% | -41.75% | -48.74% | -41.75% | 49.12% | 64.42% | 39.8% | 49.24% | -18.45% | 6471% | -13.79% | -32.43% | -27.77% | 94.63% |  | 242.5% | -17.76%
| Z-Score (Absolute) |  | 0.46 | 18.48 | 2.55 | NA | 11.3 | 0.44 | 0.32 | 3.03 | NA | 0.67 | 0.73 | 0.81 | 0.73 | 0.3 | 0.47 | NA | 0 | 0.46 | 72.61 | 0.41 | 0.62 | 0.57 | 0.81 | 37.97 | 2.47 | 0.46
| Analytical results for Ce in OREAS 232b (Certified Value 79 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 81.41 | 88.2708 | 74.5 | 84.3 | 75.45 | 79 | 81.7 | 81.1 | 76.4 | 80.7 | 78.25 | 76.6479 | 63 | 67.98 | 84.19 | 77.95 | 88.67 | 76.03 | 66.2 | 73.3752 | 80.9213 | [5 empty] | 1
| 1 | 2 | 80.4 | 88.7464 | 75.7 | 84.3 | 74.45 | 79 | 79.2 | 79.9 | 85 | 76.3 | 77.59 | 77.9286 | 65 | 73.8 | 90.69 | 77.05 | 80.14 | 77.4 | 66.3 | 77.1027 | 77.7559 | [5 empty] | 31
| 1 | 3 | 80.85 | 88.6341 | 78.7 | 80 | 74.19 | 79 | 78.6 | 84.4 | 78.1 | 77.5 | 77.54 | 77.6128 | 64 | 73.52 | 87.04 | 73.01 | 81.83 | 77.3 | 68.3 | 76.1147 | 80.1211 | [5 empty] | 16
| 1 | 4 | 80.26 | 88.8596 | 74.5 | 82.7 | 79.33 | 79 | 83.7 | 90 | 77.3 | 76.3 | 81.01 | 76.8679 | 65 | 75.35 | 90.01 | 77.95 | 82.83 | 76.31 | 69.8 | 78.6745 | 81.4915 | [5 empty] | 79.2791
| 1 | 5 | 82.36 | 88.2315 | 74.5 | 80.8 | 83.25 | 79 | 79.6 | 80.2 | 82.7 | 77.9 | 78.35 | 78.375 | 65 | 73.12 | 86.52 | 77.39 | 82.07 | 75.75 | 67.8 | 75.3961 | 81.258 | [5 empty] | 22
| 1 | 6 | 81.56 | 88.1352 | 78.5 | 83.2 | 78.09 | 77 | 82.2 | 85.3 | 79.8 | 73.6 | 74.6 | 78.8206 | 65 | 75.26 | 87.22 | 80.02 | 78.98 | 77.25 | 67.3 | 77.5069 | 80.2483
| Mean |  | 81.14 | 88.4796 | 76.0667 | 82.55 | 77.46 | 78.6667 | 80.8333 | 83.4833 | 79.8833 | 77.05 | 77.89 | 77.7088 | 64.5 | 73.1717 | 87.6117 | 77.2283 | 82.42 | 76.6733 | 67.6167 | 76.3617 | 80.2993
| Median |  | 81.13 | 88.4525 | 75.1 | 82.95 | 76.77 | 79 | 80.65 | 82.75 | 78.95 | 76.9 | 77.92 | 77.7707 | 65 | 73.66 | 87.13 | 77.67 | 81.95 | 76.78 | 67.55 | 76.6087 | 80.5848
| Std Dev. |  | 0.79294 | 0.30437 | 2.0176 | 1.7964 | 3.4989 | 0.8165 | 2.0007 | 3.8964 | 3.3499 | 2.3356 | 2.0549 | 0.84498 | 0.83666 | 2.7048 | 2.392 | 2.3103 | 3.3685 | 0.72825 | 1.3497 | 1.8505 | 1.3587
| Rel.Std.Dev. |  | 0.9773% | 0.344% | 2.652% | 2.176% | 4.517% | 1.038% | 2.475% | 4.667% | 4.193% | 3.031% | 2.638% | 1.087% | 1.297% | 3.696% | 2.73% | 2.992% | 4.087% | 0.9498% | 1.996% | 2.423% | 1.692%
| PDM3 |  | 2.347% | 11.61% | -4.052% | 4.126% | -2.295% | -0.7725% | 1.961% | 5.303% | 0.7622% | -2.812% | -1.752% | -1.981% | -18.64% | -7.704% | 10.51% | -2.587% | 3.962% | -3.287% | -14.71% | -3.68% | 1.287%
| Z-Score (Absolute) |  | 0.91 | 2.96 | 0.51 | 1.3 | 0.12 | 0.22 | 0.82 | 1.57 | 0.56 | 0.24 | 0 | 0.05 | 3.75 | 1.32 | 2.72 | 0.19 | 1.27 | 0.34 | 2.88 | 0.43 | 0.67
| Analytical results for Co in OREAS 232b (Certified Value 16.6 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 16.7 | 14.4373 | 16 | 17 | 20.48 | 16.7 | 15 | 16 | 15.9 | 18 | 15.4 | 16 | 17.7 | 17.5 | 17.2014 | 17.4 | 15.67 | 16.2 | 23.06 | 16.5 | 19.3 | 17.3 | 19.475 | 19 | 20.0781 | 16.9147 | 1
| 1 | 2 | 16.8 | 14.7537 | 16 | 17 | 21.92 | 16.7 | 17 | 16 | 16.4 | 16.7 | 16.5 | 17.6 | 17.4 | 18.1 | 17.1613 | 17.5 | 14.74 | 16.6 | 23.27 | 16.3 | 17.2 | 17.6 | 18.525 | 19 | 21.1471 | 16.8088 | 32
| 1 | 3 | 16.6 | 14.7475 | 16 | 17 | 21.7267 | 16.5 | 16 | 17 | 16.3 | 17 | 16.7 | 16 | 16.6 | 16.9 | 16.7718 | 18 | 14.92 | 16.2 | 22.64 | 15.3 | 17.8 | 17.7 | 19.76 | 19 | 19.8108 | 16.8627 | 16
| 1 | 4 | 16.4 | 14.5534 | 16 | 17 | 22.0767 | 16.8 | 15 | 17 | 16.2 | 17.3 | 18 | 16.3 | 16.8 | 18.2 | 16.6652 | 18.5 | 14.52 | 16.7 | 23.17 | 16.1 | 17.8 | 17.4 | 19.475 | 19 | 19.1427 | 16.8614 | 16.6474
| 1 | 5 | 16.9 | 14.7955 | 16 | 17 | 20.8533 | 16.8 | 16 | 16 | 16.2 | 16.9 | 15.3 | 17 | 17.2 | 17.8 | 16.3437 | 17.7 | 15.03 | 16.3 | 23.06 | 16.1 | 17.7 | 17.1 | 19.285 | 18 | 20.479 | 16.7506 | 23
| 1 | 6 | 16.4 | 14.1344 | 16 | 17 | 20.9533 | 16.6 | 15 | 18 | 16 | 17.4 | 17.3 | 16.1 | 16.3 | 17.3 | 16.6956 | 17.1 | 14.59 | 16.7 | 22.92 | 16.8 | 16.9 | 17.4 | 18.715 | 19 | 21.1471 | 16.7052
| Mean |  | 16.6333 | 14.5703 | 16 | 17 | 21.335 | 16.6833 | 15.6667 | 16.6667 | 16.1667 | 17.2167 | 16.5333 | 16.5 | 17 | 17.6333 | 16.8065 | 17.7 | 14.9117 | 16.45 | 23.02 | 16.1833 | 17.7833 | 17.4167 | 19.2058 | 18.8333 | 20.3008 | 16.8172
| Median |  | 16.65 | 14.6505 | 16 | 17 | 21.34 | 16.7 | 15.5 | 16.5 | 16.2 | 17.15 | 16.6 | 16.2 | 17 | 17.65 | 16.7337 | 17.6 | 14.83 | 16.45 | 23.06 | 16.2 | 17.75 | 17.4 | 19.38 | 19 | 20.2785 | 16.8351
| Std Dev. |  | 0.20656 | 0.25447 | 0 | 0 | 0.65641 | 0.1169 | 0.8165 | 0.8165 | 0.18619 | 0.46224 | 1.0558 | 0.65727 | 0.52536 | 0.49666 | 0.32542 | 0.49396 | 0.41845 | 0.2429 | 0.22027 | 0.50761 | 0.82805 | 0.2137 | 0.48223 | 0.40825 | 0.78678 | 0.078187
| Rel.Std.Dev. |  | 1.242% | 1.746% | 0% | 0% | 3.077% | 0.7007% | 5.212% | 4.899% | 1.152% | 2.685% | 6.386% | 3.983% | 3.09% | 2.817% | 1.936% | 2.791% | 2.806% | 1.477% | 0.9569% | 3.137% | 4.656% | 1.227% | 2.511% | 2.168% | 3.876% | 0.4649%
| PDM3 |  | -0.08446% | -12.48% | -3.889% | 2.118% | 28.16% | 0.2159% | -5.891% | 0.1158% | -2.888% | 3.42% | -0.6852% | -0.8854% | 2.118% | 5.922% | 0.9558% | 6.323% | -10.43% | -1.186% | 38.28% | -2.788% | 6.824% | 4.621% | 15.37% | 13.13% | 21.95% | 1.02%
| Z-Score (Absolute) |  | 0.2 | 2.39 | NA | NA | 4.81 | 0.14 | NA | NA | 0.69 | 0.42 | 0.3 | 0.34 | 0.19 | 0.87 | 0.01 | 0.94 | 2.03 | 0.39 | 6.6 | 0.67 | 1.03 | 0.64 | 2.54 | NA | 3.71 | 0
Results from laboratories 1.05, 1.06, 1.09, 1.10 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 232b (Certified Value 122 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 153 | 110 | 122 | 114.705 | 120 | 118 | 85 | 115 | 129 | 125 | 127 | 126 | 99 | 130.1988 | 131 | 118.9 | 144.3 | 132 | 122 | 118 | 133.1 | 118 | 87.0333 | 149.0834 |  |  | 1
| 1 | 2 | 146 | 110 | 96 | 114.07 | 114 | 126 | 108 | 111 | 126 | 132 | 132 | 122 | 97 | 130.3529 | 132 | 120.9 | 145.7 | 129 | 121 | 121 | 129.1 | 116 | 92.2226 | 148.1123 |  |  | 33
| 1 | 3 | 137 | 110 | 107 | 113.9867 | 113 | 130 | 108 | 109 | 127 | 125 | 127 | 121 | 105 | 129.0698 | 135 | 119 | 143.3 | 122 | 120 | 128 | 134.8 | 116 | 93.6358 | 148.951 |  |  | 16
| 1 | 4 | 131 | 110 | 114 | 116.7467 | 119 | 129 | 120 | 106 | 127 | 135 | 125 | 122 | 97 | 129.3801 | 135 | 117 | 141.4 | 128 | 119 | 122 | 136.3 | 117 | 93.0658 | 152.2169 |  |  | 121.7424
| 1 | 5 | 135 | 100 | 100 | 118.4935 | 120 | 119 | 120 | 101 | 126 | 124 | 128 | 125 | 96 | 127.6125 | 129 | 119.1 | 145.2 | 121 | 122 | 122 | 132 | 119 | 92.4245 | 149.6981 |  |  | 24
| 1 | 6 | 147 | 100 | 105 | 118.31 | 117 | 118 | 113 | 104 | 126 | 129 | 122 | 121 | 94 | 130.567 | 138 | 118.3 | 144.2 | 119 | 123 | 111 | 133.5 | 121 | 95.417 | 150.1432
| Mean |  | 141.5 | 106.6667 | 107.3333 | 116.052 | 117.1667 | 123.3333 | 109 | 107.6667 | 126.8333 | 128.3333 | 126.8333 | 122.8333 | 98 | 129.5302 | 133.3333 | 118.8667 | 144.0167 | 125.1667 | 121.1667 | 120.3333 | 133.1333 | 117.8333 | 92.2998 | 149.7008
| Median |  | 141.5 | 110 | 106 | 115.7258 | 118 | 122.5 | 110.5 | 107.5 | 126.5 | 127 | 127 | 122 | 97 | 129.7895 | 133.5 | 118.95 | 144.25 | 125 | 121.5 | 121.5 | 133.3 | 117.5 | 92.7451 | 149.3908
| Std Dev. |  | 8.4321 | 5.164 | 9.4587 | 2.0761 | 3.0605 | 5.6451 | 12.9306 | 5.0465 | 1.169 | 4.4572 | 3.3116 | 2.137 | 3.7947 | 1.1055 | 3.266 | 1.266 | 1.5303 | 5.1929 | 1.472 | 5.6095 | 2.4695 | 1.9408 | 2.823 | 1.4137
| Rel.Std.Dev. |  | 5.959% | 4.841% | 8.812% | 1.789% | 2.612% | 4.577% | 11.86% | 4.687% | 0.9217% | 3.473% | 2.611% | 1.74% | 3.872% | 0.8535% | 2.449% | 1.065% | 1.063% | 4.149% | 1.215% | 4.662% | 1.855% | 1.647% | 3.058% | 0.9444%
| PDM3 |  | 16.23% | -12.38% | -11.84% | -4.674% | -3.759% | 1.307% | -10.47% | -11.56% | 4.182% | 5.414% | 4.182% | 0.8961% | -19.5% | 6.397% | 9.521% | -2.362% | 18.3% | 2.813% | -0.4729% | -1.157% | 9.357% | -3.211% | -24.18% | 22.97%
| Z-Score (Absolute) |  | 1.9 | 1.49 | 1.43 | 0.58 | 0.47 | 0.13 | 1.26 | 1.39 | 0.47 | 0.62 | 0.47 | 0.08 | 2.33 | 0.73 | 1.1 | 0.3 | 2.14 | 0.31 | 0.08 | 0.16 | 1.08 | 0.41 | 2.89 | 2.69
| Analytical results for Cs in OREAS 232b (Certified Value 8.63 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 8.72 | 7.5368 | 8.1 | 9.17 | 8.8 | 7.22 | 8.6 | 9.2 | 8.65 | 8.22 | 8.06 | 9.04 | 8.4 | 9.8 | 6.859 | 10 | 8.87 | 9.1 | 8.6 | 9.2813 | 8.5174 | [5 empty] | 1
| 1 | 2 | 8.6 | 7.5764 | 8.3 | 9.04 | 8.9 | 7.52 | 8.4 | 9 | 8.38 | 8.32 | 8.67 | 8.94 | 9.1 | 9.8 | 7.363 | 11 | 8.28 | 8.28 | 8.8 | 8.7563 | 8.5092 | [5 empty] | 34
| 1 | 3 | 8.79 | 7.6636 | 8.6 | 8.93 | 8.7 | 8.07 | 8.9 | 8.7 | 8.26 | 8.38 | 8.15 | 8.76 | 9 | 10.2 | 7.541 | 10 | 8.31 | 8.45 | 8.8 | 9.1851 | 8.4793 | [5 empty] | 16
| 1 | 4 | 8.79 | 7.614 | 8.2 | 8.95 | 8.8 | 7.73 | 9 | 9.1 | 8.68 | 9.02 | 8.06 | 8.71 | 9 | 9.9 | 7.597 | 11 | 8.72 | 8.49 | 8.6 | 9.1048 | 8.5496 | [5 empty] | 8.6274
| 1 | 5 | 8.71 | 7.4911 | 8.6 | 8.92 | 8.9 | 7.77 | 9 | 8.9 | 8.44 | 8.18 | 8.47 | 9.02 | 8.4 | 10 | 7.44 | 11 | 8.64 | 8.48 | 8.8 | 9.0047 | 8.5409 | [5 empty] | 25
| 1 | 6 | 8.58 | 7.526 | 8.4 | 8.86 | 8.8 | 8.47 | 10.1 | 8.7 | 8.52 | 8.69 | 8.2 | 8.77 | 9.3 | 9.7 | 7.561 | 11 | 9.01 | 8.08 | 8.3 | 8.9062 | 8.5913
| Mean |  | 8.6983 | 7.568 | 8.3667 | 8.9783 | 8.8167 | 7.7967 | 9 | 8.9333 | 8.4883 | 8.4683 | 8.2683 | 8.8733 | 8.8667 | 9.9 | 7.3935 | 10.6667 | 8.6383 | 8.48 | 8.65 | 9.0397 | 8.5313
| Median |  | 8.715 | 7.5566 | 8.35 | 8.94 | 8.8 | 7.75 | 8.95 | 8.95 | 8.48 | 8.35 | 8.175 | 8.855 | 9 | 9.85 | 7.4905 | 11 | 8.68 | 8.465 | 8.7 | 9.0547 | 8.5292
| Std Dev. |  | 0.090646 | 0.063204 | 0.20656 | 0.11053 | 0.075277 | 0.43413 | 0.58992 | 0.20656 | 0.1613 | 0.32499 | 0.24799 | 0.14418 | 0.37771 | 0.17889 | 0.27568 | 0.5164 | 0.29485 | 0.34217 | 0.19748 | 0.19137 | 0.038502
| Rel.Std.Dev. |  | 1.042% | 0.8351% | 2.469% | 1.231% | 0.8538% | 5.568% | 6.555% | 2.312% | 1.9% | 3.838% | 2.999% | 1.625% | 4.26% | 1.807% | 3.729% | 4.841% | 3.413% | 4.035% | 2.283% | 2.117% | 0.4513%
| PDM3 |  | 0.8223% | -12.28% | -3.022% | 4.068% | 2.194% | -9.629% | 4.319% | 3.546% | -1.612% | -1.844% | -4.162% | 2.851% | 2.773% | 14.75% | -14.3% | 23.64% | 0.1268% | -1.708% | 0.262% | 4.779% | -1.114%
| Z-Score (Absolute) |  | 0.14 | 2.86 | 0.74 | 0.89 | 0.46 | 2.26 | 0.95 | 0.77 | 0.41 | 0.47 | 1 | 0.61 | 0.59 | 3.34 | 3.33 | NA | 0.02 | 0.44 | 0.02 | 1.05 | 0.3
Results from laboratory 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 232b (Certified Value 26.7 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 27.6 | 23.9516 | 26 | 25 | 33.62 | 30 | 24 | 22 | 29.9 | 26.1 | 24.2 | 26.2 | 27.2 | 25.7 | 28.2323 | 26.2 | 28.09 | 28.2 | 29.29 | 27.8 | 26.7 | 26.4 | 28.74 | 27 | 4 | 26.1525 | 1
| 1 | 2 | 28 | 24.0988 | 28 | 25 | 30.61 | 29 | 24.9 | 23 | 29.5 | 26.1 | 24.7 | 28.1 | 25.9 | 25.1 | 28.3371 | 26.7 | 28.01 | 28.1 | 30.17 | 26 | 26.7 | 26.9 | 27.53 | 27 | 3 | 26.5756 | 35
| 1 | 3 | 28.5 | 25.8192 | 28 | 25 | 31.455 | 29 | 23.7 | 23 | 30.2 | 25.2 | 26.2 | 26.7 | 25.9 | 24.3 | 28.8458 | 26.1 | 26.26 | 28.1 | 30.96 | 25.8 | 25.4 | 27.4 | 28.99 | 27 | 4 | 26.2016 | 16
| 1 | 4 | 27.5 | 26.27 | 26 | 25 | 32.045 | 28 | 25.4 | 20 | 26.2 | 26.3 | 27.3 | 26.7 | 25.5 | 25.2 | 28.2108 | 24.9 | 27.26 | 28.4 | 28.54 | 26.3 | 25.6 | 26.7 | 29.04 | 26 | 5 | 25.4731 | 26.7063
| 1 | 5 | 28.9 | 24.6324 | 28 | 25 | 31.38 | 28 | 31.4 | 22 | 27.6 | 25.4 | 25.7 | 28 | 26.1 | 25.4 | 29.0436 | 26.1 | 26.37 | 29.6 | 28.44 | 26.8 | 25.4 | 26.6 | 28.42 | 26 | 3 | 26.2207 | 26
| 1 | 6 | 27.4 | 24.1908 | 26 | 25 | 33.14 | 29 | 27.3 | 27 | 27.3 | 26.3 | 26.9 | 26.4 | 25.2 | 24.9 | 28.4853 | 26 | 26.53 | 27.3 | 30.45 | 27.4 | 26.5 | 28.1 | 28.71 | 26 | 4 | 25.7227
| Mean |  | 27.9833 | 24.8271 | 27 | 25 | 32.0417 | 28.8333 | 26.1167 | 22.8333 | 28.45 | 25.9 | 25.8333 | 27.0167 | 25.9667 | 25.1 | 28.5258 | 26 | 27.0867 | 28.2833 | 29.6417 | 26.6833 | 26.05 | 27.0167 | 28.5717 | 26.5 | 3.8333 | 26.0577
| Median |  | 27.8 | 24.4116 | 27 | 25 | 31.75 | 29 | 25.15 | 22.5 | 28.55 | 26.1 | 25.95 | 26.7 | 25.9 | 25.15 | 28.4112 | 26.1 | 26.895 | 28.15 | 29.73 | 26.55 | 26.05 | 26.8 | 28.725 | 26.5 | 4 | 26.1771
| Std Dev. |  | 0.6047 | 0.9804 | 1.0954 | 0 | 1.1427 | 0.75277 | 2.8854 | 2.3166 | 1.6355 | 0.47749 | 1.216 | 0.82321 | 0.68605 | 0.47749 | 0.34443 | 0.5933 | 0.82415 | 0.74677 | 1.0442 | 0.79603 | 0.6473 | 0.63061 | 0.55676 | 0.54772 | 0.75277 | 0.39459
| Rel.Std.Dev. |  | 2.161% | 3.949% | 4.057% | 0% | 3.566% | 2.611% | 11.05% | 10.15% | 5.749% | 1.844% | 4.707% | 3.047% | 2.642% | 1.902% | 1.207% | 2.282% | 3.043% | 2.64% | 3.523% | 2.983% | 2.485% | 2.334% | 1.949% | 2.067% | 19.64% | 1.514%
| PDM3 |  | 4.782% | -7.037% | 1.1% | -6.389% | 19.98% | 7.964% | -2.208% | -14.5% | 6.529% | -3.019% | -3.269% | 1.162% | -2.77% | -6.015% | 6.813% | -2.645% | 1.424% | 5.905% | 10.99% | -0.08617% | -2.458% | 1.162% | 6.985% | -0.7726% | -85.65% | -2.429%
| Z-Score (Absolute) |  | 0.87 | 1.11 | 0.26 | 1 | 3.42 | 1.41 | 0.3 | 2.36 | 1.17 | 0.43 | 0.48 | 0.27 | 0.39 | 0.94 | 1.21 | 0.37 | 0.31 | 1.06 | 1.91 | 0.06 | 0.34 | 0.27 | 1.24 | 0.06 | 14.28 | 0.33
| Analytical results for Dy in OREAS 232b (Certified Value 3.38 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 3.1031 | 3.7 | 3.8 | 3.2 | 3.88 | 2.914 | 3.61 | 3.1 | 2.62 | 5.4681 | [16 empty] | 1
| 1 | 2 | 3.1254 | 3.6 | 3.8 | 3.5 | 4.02 | 2.957 | 3.56 | 3.2 | 2.64 | 5.1875 | [16 empty] | 9
| 1 | 3 | 3.0807 | 3.65 | 3.7 | 3 | 4.06 | 2.942 | 4.48 | 3.2 | 2.75 | 5.2871 | [16 empty] | 16
| 1 | 4 | 2.9828 | 3.7 | 3.8 | 4.1 | 3.94 | 3.267 | 3.6 | 3.1 | 2.65 | 5.4341 | [16 empty] | 3.3802
| 1 | 5 | 3.1274 | 3.8 | 3.7 | 3.3 | 3.97 | 2.867 | 4.81 | 3.1 | 2.71 | 5.5482 | [16 empty] | 27
| 1 | 6 | 3.1719 | 3.65 | 3.8 | 3.1 | 3.89 | 2.879 | 3.9 | 3 | 2.66 | 5.5311
| Mean |  | 3.0986 | 3.6833 | 3.7667 | 3.3667 | 3.96 | 2.971 | 3.9933 | 3.1167 | 2.6717 | 5.4094
| Median |  | 3.1143 | 3.675 | 3.8 | 3.25 | 3.955 | 2.928 | 3.755 | 3.1 | 2.655 | 5.4511
| Std Dev. |  | 0.06431 | 0.068313 | 0.05164 | 0.39833 | 0.071274 | 0.14912 | 0.52952 | 0.075277 | 0.048751 | 0.14308
| Rel.Std.Dev. |  | 2.075% | 1.855% | 1.371% | 11.83% | 1.8% | 5.019% | 13.26% | 2.415% | 1.825% | 2.645%
| PDM3 |  | -8.333% | 8.967% | 11.43% | -0.4011% | 17.15% | -12.11% | 18.14% | -7.797% | -20.96% | 60.03%
| Z-Score (Absolute) |  | 0.67 | 0.25 | 0.38 | 0.25 | 0.68 | 0.87 | 0.73 | 0.64 | 1.34 | 2.95
| Analytical results for Er in OREAS 232b (Certified Value 1.78 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 1.5894 | 1.95 | 2.2 | 1.6 | 2.07 | 1.481 | 1.95 | 1.5 | 1.4225 | 3.1644 | [16 empty] | 1
| 1 | 2 | 1.6057 | 2 | 2.2 | 1.9 | 2.14 | 1.631 | 1.96 | 1.5 | 1.366 | 3.0571 | [16 empty] | 10
| 1 | 3 | 1.5568 | 1.9 | 2 | 1.6 | 2.16 | 1.516 | 1.79 | 1.6 | 1.4443 | 3.0814 | [16 empty] | 16
| 1 | 4 | 1.4748 | 1.95 | 1.8 | 2.1 | 2.07 | 1.553 | 1.98 | 1.5 | 1.4534 | 3.1977 | [16 empty] | 1.7773
| 1 | 5 | 1.644 | 2.05 | 2 | 1.7 | 2.13 | 1.47 | 2 | 1.5 | 1.4314 | 3.2784 | [16 empty] | 28
| 1 | 6 | 1.5416 | 2 | 2.2 | 2 | 2.07 | 1.546 | 2.06 | 1.5 | 1.4168 | 3.2542
| Mean |  | 1.5687 | 1.975 | 2.0667 | 1.8167 | 2.1067 | 1.5328 | 1.9567 | 1.5167 | 1.4224 | 3.1722
| Median |  | 1.5731 | 1.975 | 2.1 | 1.8 | 2.1 | 1.531 | 1.97 | 1.5 | 1.427 | 3.1811
| Std Dev. |  | 0.058569 | 0.05244 | 0.1633 | 0.2137 | 0.041312 | 0.05854 | 0.09048 | 0.040825 | 0.03077 | 0.089692
| Rel.Std.Dev. |  | 3.734% | 2.655% | 7.902% | 11.76% | 1.961% | 3.819% | 4.624% | 2.692% | 2.163% | 2.827%
| PDM3 |  | -11.73% | 11.12% | 16.28% | 2.215% | 18.53% | -13.75% | 10.09% | -14.66% | -19.97% | 78.49%
| Z-Score (Absolute) |  | 0.8 | 0.22 | 0.45 | 0.18 | 0.55 | 0.89 | 0.18 | 0.93 | 1.16 | 3.22
| Analytical results for Eu in OREAS 232b (Certified Value 1.24 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 1.5609 | 1.3 | 1.22 | 1.2 | 1.32 | 1.184 | 1.2 | 1.2 | 1.1065 | 1.3177 | [16 empty] | 1
| 1 | 2 | 1.6208 | 1.25 | 1.22 | 1.2 | 1.34 | 1.257 | 1.23 | 1.2 | 1.0905 | 1.2426 | [16 empty] | 11
| 1 | 3 | 1.6554 | 1.25 | 1.14 | 1.1 | 1.36 | 1.291 | 1.12 | 1.2 | 1.1309 | 1.2677 | [16 empty] | 16
| 1 | 4 | 1.5703 | 1.35 | 1.22 | 1.1 | 1.34 | 1.27 | 1.18 | 1.2 | 1.1299 | 1.3046 | [16 empty] | 1.2394
| 1 | 5 | 1.5756 | 1.25 | 1.17 | 1.2 | 1.34 | 1.258 | 1.23 | 1.2 | 1.1196 | 1.3173 | [16 empty] | 29
| 1 | 6 | 1.6177 | 1.2 | 1.19 | 1.2 | 1.33 | 1.332 | 1.24 | 1.2 | 1.1375 | 1.3061
| Mean |  | 1.6001 | 1.2667 | 1.1933 | 1.1667 | 1.3383 | 1.2653 | 1.2 | 1.2 | 1.1192 | 1.2927
| Median |  | 1.5966 | 1.25 | 1.205 | 1.2 | 1.34 | 1.264 | 1.215 | 1.2 | 1.1247 | 1.3054
| Std Dev. |  | 0.03692 | 0.05164 | 0.033267 | 0.05164 | 0.013292 | 0.04871 | 0.045166 | 0 | 0.017728 | 0.030604
| Rel.Std.Dev. |  | 2.307% | 4.077% | 2.788% | 4.426% | 0.9931% | 3.85% | 3.764% | 0% | 1.584% | 2.368%
| PDM3 |  | 29.11% | 2.204% | -3.713% | -5.865% | 7.986% | 2.096% | -3.175% | -3.175% | -9.699% | 4.302%
| Z-Score (Absolute) |  | 3.26 | 0.01 | 0.71 | NA | 0.71 | 0.01 | 0.64 | NA | 1.43 | 0.26
Results from laboratories 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 232b (Certified Value 3.85 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 3.96 | 3.7879 | 3.77 | 3.66 | 3.93 | 4.15 | 3.67 | 3.81 | 3.82 | 3.84 | 3.77 | 3.84 | 3.78 | 3.6904 | 3.9106 | 3.784 | 3.86 | 3.98 | 3.92 | 3.95 | 3.89 | 3.9556 | 3.7204 | 3.9168 |  |  | 1
| 1 | 2 | 3.91 | 3.8384 | 3.79 | 3.63 | 3.935 | 4.17 | 3.63 | 3.86 | 3.77 | 3.75 | 3.91 | 4.04 | 3.7 | 3.7301 | 3.926 | 3.693 | 3.88 | 3.98 | 3.88 | 3.97 | 3.91 | 3.815 | 3.7592 | 3.8532 |  |  | #N/A
| 1 | 3 | 3.95 | 3.8118 | 3.84 | 3.56 | 3.98 | 4.1 | 3.51 | 3.87 | 3.78 | 3.76 | 3.76 | 3.9 | 3.72 | 3.7046 | 3.9026 | 3.784 | 3.76 | 3.97 | 3.94 | 3.95 | 3.96 | 4.0007 | 3.7364 | 3.9887 |  |  | 16
| 1 | 4 | 3.94 | 3.7645 | 3.82 | 3.49 | 3.98 | 4.1 | 3.74 | 3.83 | 3.7 | 3.78 | 4.07 | 3.81 | 3.7 | 3.7924 | 3.8995 | 3.798 | 3.93 | 3.99 | 3.98 | 3.97 | 3.95 | 4.0349 | 3.8315 | 3.8773 |  |  | 3.8519
| 1 | 5 | 3.88 | 3.8351 | 3.83 | 3.43 | 3.935 | 4.17 | 3.7 | 3.92 | 3.7 | 3.77 | 3.69 | 3.91 | 3.75 | 3.772 | 3.904 | 3.7 | 3.93 | 3.95 | 3.99 | 3.96 | 3.82 | 3.9045 | 3.6716 | 3.9446 |  |  | 30
| 1 | 6 | 3.86 | 3.8061 | 3.73 | 3.79 | 3.9333 | 4.17 | 3.76 | 3.96 | 3.72 | 3.78 | 4 | 3.73 | 3.76 | 3.7312 | 3.9123 | 3.861 | 3.87 | 4.04 | 3.93 | 3.95 | 3.83 | 3.9233 | 3.7259 | 3.929
| Mean |  | 3.9167 | 3.8073 | 3.7967 | 3.5933 | 3.9489 | 4.1433 | 3.6683 | 3.875 | 3.7483 | 3.78 | 3.8667 | 3.8717 | 3.735 | 3.7368 | 3.9092 | 3.77 | 3.8717 | 3.985 | 3.94 | 3.9583 | 3.8933 | 3.939 | 3.7408 | 3.9183
| Median |  | 3.925 | 3.8089 | 3.805 | 3.595 | 3.935 | 4.16 | 3.685 | 3.865 | 3.745 | 3.775 | 3.84 | 3.87 | 3.735 | 3.7306 | 3.9073 | 3.784 | 3.875 | 3.98 | 3.935 | 3.955 | 3.9 | 3.9394 | 3.7312 | 3.9229
| Std Dev. |  | 0.040332 | 0.028179 | 0.041793 | 0.12879 | 0.024168 | 0.034448 | 0.090646 | 0.056125 | 0.04916 | 0.031623 | 0.15029 | 0.10534 | 0.033317 | 0.038978 | 0.0095961 | 0.063696 | 0.062423 | 0.030166 | 0.040497 | 0.0098319 | 0.058878 | 0.077578 | 0.052936 | 0.04834
| Rel.Std.Dev. |  | 1.03% | 0.7401% | 1.101% | 3.584% | 0.612% | 0.8314% | 2.471% | 1.448% | 1.312% | 0.8366% | 3.887% | 2.721% | 0.892% | 1.043% | 0.2455% | 1.69% | 1.612% | 0.757% | 1.028% | 0.2484% | 1.512% | 1.969% | 1.415% | 1.234%
| PDM3 |  | 1.681% | -1.159% | -1.435% | -6.713% | 2.517% | 7.565% | -4.766% | 0.599% | -2.689% | -1.867% | 0.3827% | 0.5125% | -3.036% | -2.989% | 1.486% | -2.127% | 0.5125% | 3.455% | 2.286% | 2.762% | 1.075% | 2.26% | -2.884% | 1.722%
| Z-Score (Absolute) |  | 0.4 | 0.57 | 0.66 | 2.47 | 0.68 | 2.41 | 1.8 | 0.03 | 1.09 | 0.81 | 0.04 | 0 | 1.21 | 1.2 | 0.33 | 0.9 | 0 | 1 | 0.61 | 0.77 | 0.19 | 0.6 | 1.16 | 0.41
| Analytical results for Ga in OREAS 232b (Certified Value 18.4 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 18.38 | 14.2966 | 17.4 | 19.5 | 20.4067 | 17.9 | 17.97 | 21.2 | 19 | 18.2 | 18.65 | 18.15 | 18.6 | 17.95 | 17.7359 | 17 | 18.3 | 17.52 | 19.88 | 17.92 | 19.8505 | 23 | 18.0686 |  |  |  | 1
| 1 | 2 | 18.49 | 14.4361 | 17.8 | 19.7 | 20.735 | 18.3 | 18.07 | 19.4 | 19 | 17.6 | 19.3 | 19.55 | 18.45 | 18.55 | 17.6398 | 16.7 | 18.3 | 16.83 | 18.72 | 17.97 | 19.2947 | 22 | 18.0228 |  |  |  | 39
| 1 | 3 | 18.44 | 14.5851 | 18.2 | 21 | 21.29 | 18.2 | 17.55 | 21.4 | 19 | 17.45 | 19.75 | 17.95 | 18.15 | 16.78 | 17.6586 | 17.4 | 17.6 | 16.06 | 18.65 | 18.42 | 20.0911 | 24 | 18.241 |  |  |  | 16
| 1 | 4 | 18.35 | 14.0922 | 18.2 | 20.3 | 22.02 | 18 | 18.43 | 20.6 | 19 | 18.15 | 20.9 | 18.15 | 17.95 | 17.82 | 17.558 | 17 | 19.4 | 17.07 | 19.68 | 18.07 | 20.0993 | 22 | 18.0522 |  |  |  | 18.4039
| 1 | 5 | 18.43 | 13.8575 | 17.4 | 20.9 | 20.98 | 18.3 | 18.96 | 19 | 19 | 17.45 | 18.9 | 19.4 | 17.6 | 17.91 | 17.6242 | 17.2 | 18.6 | 17.07 | 19.29 | 17.74 | 20.0022 | 21 | 17.9381 |  |  |  | 31
| 1 | 6 | 18.2 | 14.2757 | 18 | 22.7 | 20.4567 | 18.5 | 19.32 | 20.3 | 19 | 17.75 | 20.4 | 18.3 | 18.05 | 17.91 | 17.8166 | 16.9 | 19.4 | 18.4 | 18.13 | 17.4 | 20.838 | 22 | 18.5331
| Mean |  | 18.3817 | 14.2572 | 17.8333 | 20.6833 | 20.9814 | 18.2 | 18.3833 | 20.3167 | 19 | 17.7667 | 19.65 | 18.5833 | 18.1333 | 17.82 | 17.6722 | 17.0333 | 18.6 | 17.1583 | 19.0583 | 17.92 | 20.0293 | 22.3333 | 18.1426
| Median |  | 18.405 | 14.2861 | 17.9 | 20.6 | 20.8575 | 18.25 | 18.25 | 20.45 | 19 | 17.675 | 19.525 | 18.225 | 18.1 | 17.91 | 17.6492 | 17 | 18.45 | 17.07 | 19.005 | 17.945 | 20.0466 | 22 | 18.0604
| Std Dev. |  | 0.10147 | 0.25629 | 0.36697 | 1.16 | 0.60668 | 0.21909 | 0.65963 | 0.96003 | 0 | 0.33566 | 0.87407 | 0.70119 | 0.3587 | 0.57404 | 0.091082 | 0.24221 | 0.70143 | 0.77489 | 0.67205 | 0.33994 | 0.497 | 1.0328 | 0.21539
| Rel.Std.Dev. |  | 0.552% | 1.798% | 2.058% | 5.609% | 2.892% | 1.204% | 3.588% | 4.725% | 0% | 1.889% | 4.448% | 3.773% | 1.978% | 3.221% | 0.5154% | 1.422% | 3.771% | 4.516% | 3.526% | 1.897% | 2.481% | 4.624% | 1.187%
| PDM3 |  | -0.1207% | -22.53% | -3.1% | 12.39% | 14.01% | -1.108% | -0.1117% | 10.39% | 3.239% | -3.462% | 6.771% | 0.9751% | -1.47% | -3.173% | -3.976% | -7.447% | 1.066% | -6.768% | 3.556% | -2.629% | 8.832% | 21.35% | -1.42%
| Z-Score (Absolute) |  | 0 | 4.5 | 0.6 | 2.51 | 2.84 | 0.2 | 0 | 2.11 | 0.67 | 0.67 | 1.38 | 0.22 | 0.27 | 0.61 | 0.77 | 1.47 | 0.24 | 1.33 | 0.74 | 0.5 | 1.8 | 4.31 | 0.26
| Analytical results for Gd in OREAS 232b (Certified Value 4.96 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.6914 | 5.2 | 4.9 | 4.4 | 5 | 4.745 | 4.89 | 4.6 | 4.9323 | 0.90431 | [16 empty] | 1
| 1 | 2 | 6.6654 | 5.2 | 5.2 | 5.3 | 5 | 5.075 | 4.96 | 4.7 | 4.7697 | 0.86496 | [16 empty] | 12
| 1 | 3 | 6.6763 | 5.4 | 5.2 | 4.5 | 5.1 | 5.308 | 4.62 | 4.7 | 4.9458 | 0.87637 | [16 empty] | 16
| 1 | 4 | 6.7121 | 5.2 | 5.2 | 5 | 5 | 5.068 | 4.86 | 4.5 | 5.088 | 0.90707 | [16 empty] | 4.9634
| 1 | 5 | 6.7555 | 5.4 | 4.9 | 4.7 | 5.1 | 5 | 4.87 | 4.6 | 4.87 | 0.91217 | [16 empty] | 32
| 1 | 6 | 6.737 | 5.2 | 5 | 5.5 | 5 | 5.077 | 5.29 | 4.4 | 4.8627 | 0.91567
| Mean |  | 6.7063 | 5.2667 | 5.0667 | 4.9 | 5.0333 | 5.0455 | 4.915 | 4.5833 | 4.9114 | 0.89676
| Median |  | 6.7017 | 5.2 | 5.1 | 4.85 | 5 | 5.0715 | 4.88 | 4.6 | 4.9011 | 0.90569
| Std Dev. |  | 0.035213 | 0.10328 | 0.15055 | 0.44272 | 0.05164 | 0.18091 | 0.21696 | 0.1169 | 0.10674 | 0.020906
| Rel.Std.Dev. |  | 0.5251% | 1.961% | 2.971% | 9.035% | 1.026% | 3.586% | 4.414% | 2.551% | 2.173% | 2.331%
| PDM3 |  | 35.12% | 6.111% | 2.081% | -1.277% | 1.409% | 1.655% | -0.9747% | -7.657% | -1.047% | -81.93%
| Z-Score (Absolute) |  | 14.02 | 2.37 | 0.75 | 0.6 | 0.48 | 0.58 | 0.48 | 3.16 | 0.51 | 32.99
| Analytical results for Ge in OREAS 232b (Indicative Value 0.19 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 1.7 | 0.1 | 1.2 | 0.15 | 1.6 | 3.3 | 0.18 | 0.2 | 0.09 | 0.13 | 0.31 | < 0.05 | 0.13 | 0.2903 | [12 empty] | 1
| 1 | 2 | 1.6 | 0.1 | 1.1 | 0.16 | 1.9 | 3.7 | 0.17 | 0.23 | 0.07 | 0.13 | 0.37 | < 0.05 | 0.12 | 0.3231 | [12 empty] | 6
| 1 | 3 | 1.5 | 0.3 | 1.2 | 0.17 | 1.9 | 3.5 | 0.17 | 0.24 | 0.09 | 0.15 | 0.4 | < 0.05 | 0.12 | 0.3397 | [12 empty] | 16
| 1 | 4 | 1.5 | 0.3 | 1.1 | 0.17 | 1.2 | 3.6 | 0.19 | 0.25 | 0.08 | 0.16 | 0.36 | < 0.05 | 0.13 | 0.2361 | [12 empty] | 0.19286
| 1 | 5 | 1.5 | 0.2 | 1.2 | 0.16 | 1.8 | 3.7 | 0.17 | 0.22 | 0.08 | 0.16 | 0.31 | < 0.05 | 0.13 | 0.3375 | [12 empty] | 12
| 1 | 6 | 1.7 | 0.1 | 1.1 | 0.17 | 1.6 | 3.8 | 0.16 | 0.23 | 0.08 | 0.16 | 0.27 | < 0.05 | 0.12 | 0.2475
| Mean |  | 1.5833 | 0.18333 | 1.15 | 0.16333 | 1.6667 | 3.6 | 0.17333 | 0.22833 | 0.081667 | 0.14833 | 0.33667 |  | 0.125 | 0.2957
| Median |  | 1.55 | 0.15 | 1.15 | 0.165 | 1.7 | 3.65 | 0.17 | 0.23 | 0.08 | 0.155 | 0.335 |  | 0.125 | 0.3067
| Std Dev. |  | 0.098319 | 0.098319 | 0.054772 | 0.008165 | 0.26583 | 0.17889 | 0.010328 | 0.017224 | 0.0075277 | 0.01472 | 0.048028 |  | 0.0054772 | 0.045472
| Rel.Std.Dev. |  | 6.21% | 53.63% | 4.763% | 4.999% | 15.95% | 4.969% | 5.958% | 7.543% | 9.218% | 9.923% | 14.27% |  | 4.382% | 15.38%
| PDM3 |  | 721% | -4.937% | 496.3% | -15.31% | 764.2% | 1767% | -10.12% | 18.4% | -57.65% | -23.09% | 74.57% |  | -35.18% | 53.33%
| Z-Score (Absolute) |  | 8.68 | 0.14 | 5.95 | 0.27 | 9.2 | 21.39 | 0.2 | 0.14 | 0.78 | 0.36 | 0.82 | 1.14 | 0.51 | 0.57
| Analytical results for Hf in OREAS 232b (Certified Value 4.06 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4.4 | 4.4115 | 7.6 | 3.6 | 4.2 | 3.6 | 4.3 | 4.8 | 4.2 | 4 | 3.9 | 4 | 4.12 | 4.7 | 3.79 | 4.33 | 4.5 | 3.56 | 3.7721 | 2.5966 | 6.6986 | [5 empty] | 1
| 1 | 2 | 4.22 | 4.2373 | 8.4 | 3.1 | 4.1 | 3.6 | 4 | 4.4 | 4.1 | 4.1 | 4.1 | 3.9 | 4.05 | 4.6 | 4 | 4.15 | 4.1 | 3.56 | 3.4169 | 2.7796 | 6.4413 | [5 empty] | 25
| 1 | 3 | 4.22 | 4.4289 | 8 | 3.1 | 4.3 | 3.5 | 4.3 | 4.3 | 4.1 | 4 | 4.1 | 3.9 | 3.9 | 4.8 | 3.91 | 3.96 | 4.2 | 3.66 | 3.4773 | 2.5634 | 6.4975 | [5 empty] | 16
| 1 | 4 | 4.17 | 4.2367 | 8 | 3.3 | 4.1 | 3.6 | 4 | 4.5 | 4.2 | 4.3 | 4 | 3.8 | 4.2 | 4.6 | 3.93 | 4.2 | 4.1 | 3.56 | 3.639 | 2.9209 | 6.5567 | [5 empty] | 4.064
| 1 | 5 | 4.32 | 4.2251 | 8.2 | 3 | 4.1 | 3.8 | 4.1 | 4.5 | 4 | 3.8 | 4.2 | 3.9 | 4.04 | 4.6 | 4.06 | 4.11 | 4.1 | 3.6 | 3.5277 | 2.6964 | 6.6532 | [5 empty] | 33
| 1 | 6 | 4.18 | 4.1494 | 8 | 2.3 | 4 | 3.7 | 4.4 | 4.3 | 4.1 | 4.1 | 4 | 3.9 | 3.99 | 4.6 | 3.86 | 4.2 | 4 | 3.52 | 3.5323 | 2.58 | 6.6356
| Mean |  | 4.2517 | 4.2815 | 8.0333 | 3.0667 | 4.1333 | 3.6333 | 4.1833 | 4.4667 | 4.1167 | 4.05 | 4.05 | 3.9 | 4.05 | 4.65 | 3.925 | 4.1583 | 4.1667 | 3.5767 | 3.5609 | 2.6895 | 6.5805
| Median |  | 4.22 | 4.237 | 8 | 3.1 | 4.1 | 3.6 | 4.2 | 4.45 | 4.1 | 4.05 | 4.05 | 3.9 | 4.045 | 4.6 | 3.92 | 4.175 | 4.1 | 3.56 | 3.53 | 2.6465 | 6.5961
| Std Dev. |  | 0.089981 | 0.11245 | 0.26583 | 0.43205 | 0.10328 | 0.10328 | 0.17224 | 0.18619 | 0.075277 | 0.16432 | 0.10488 | 0.063246 | 0.10354 | 0.083666 | 0.096488 | 0.12222 | 0.17512 | 0.048028 | 0.12677 | 0.14017 | 0.099092
| Rel.Std.Dev. |  | 2.116% | 2.626% | 3.309% | 14.09% | 2.499% | 2.843% | 4.117% | 4.168% | 1.829% | 4.057% | 2.59% | 1.622% | 2.556% | 1.799% | 2.458% | 2.939% | 4.203% | 1.343% | 3.56% | 5.212% | 1.506%
| PDM3 |  | 4.619% | 5.353% | 97.67% | -24.54% | 1.707% | -10.6% | 2.937% | 9.909% | 1.297% | -0.3436% | -0.3436% | -4.035% | -0.3436% | 14.42% | -3.419% | 2.322% | 2.527% | -11.99% | -12.38% | -33.82% | 61.92%
| Z-Score (Absolute) |  | 0.47 | 0.58 | 13.78 | 3.69 | 0.06 | 1.7 | 0.23 | 1.23 | 0 | 0.23 | 0.23 | 0.76 | 0.23 | 1.88 | 0.67 | 0.15 | 0.18 | 1.9 | 1.96 | 5.02 | 8.67
| Analytical results for Hg in OREAS 232b (Indicative Value 0.029 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 1.09 | 1.30 | [24 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [24 empty] | 3
3
| 1 | 1 | < 2 | 0.0329 | [24 empty] | 1
| 1 | 2 | < 2 | 0.0323 | [24 empty] | 13
| 1 | 3 | < 2 | 0.0309 | [24 empty] | 16
| 1 | 4 | < 2 | 0.0269 | [24 empty] | 0.02865
| 1 | 5 | < 2 | 0.024 | [24 empty] | 9
| 1 | 6 | < 2 | 0.0249
| Mean |  |  | 0.02865
| Median |  |  | 0.0289
| Std Dev. |  |  | 0.003878
| Rel.Std.Dev. |  |  | 13.54%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 232b (Certified Value 0.62 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.56178 | 0.66 | 0.8 | 0.7 | 0.68 | 0.529 | 0.71 | 0.5 | 0.4913 | 1.0878 | [16 empty] | 1
| 1 | 2 | 0.52763 | 0.68 | 0.7 | 0.6 | 0.7 | 0.554 | 0.68 | 0.5 | 0.4819 | 1.0496 | [16 empty] | 14
| 1 | 3 | 0.57057 | 0.66 | 0.7 | 0.6 | 0.7 | 0.529 | 0.79 | 0.5 | 0.5065 | 1.066 | [16 empty] | 16
| 1 | 4 | 0.5407 | 0.66 | 0.7 | 0.7 | 0.67 | 0.604 | 0.68 | 0.5 | 0.5218 | 1.1024 | [16 empty] | 0.61915
| 1 | 5 | 0.55798 | 0.74 | 0.7 | 0.6 | 0.69 | 0.52 | 0.7 | 0.5 | 0.501 | 1.1231 | [16 empty] | 34
| 1 | 6 | 0.5605 | 0.7 | 0.8 | 0.8 | 0.67 | 0.54 | 0.73 | 0.5 | 0.5059 | 1.12
| Mean |  | 0.55319 | 0.68333 | 0.73333 | 0.66667 | 0.685 | 0.546 | 0.715 | 0.5 | 0.5014 | 1.0915
| Median |  | 0.55924 | 0.67 | 0.7 | 0.65 | 0.685 | 0.5345 | 0.705 | 0.5 | 0.50345 | 1.0951
| Std Dev. |  | 0.015883 | 0.032042 | 0.05164 | 0.08165 | 0.013784 | 0.030718 | 0.041352 | 0 | 0.013747 | 0.029497
| Rel.Std.Dev. |  | 2.871% | 4.689% | 7.042% | 12.25% | 2.012% | 5.626% | 5.784% | 0% | 2.742% | 2.702%
| PDM3 |  | -10.65% | 10.37% | 18.44% | 7.675% | 10.64% | -11.81% | 15.48% | -19.24% | -19.02% | 76.29%
| Z-Score (Absolute) |  | 0.91 | 0.06 | 0.44 | 0.06 | 0.07 | 0.97 | 0.3 | 1.31 | 1.3 | 3.12
| Analytical results for In in OREAS 232b (Certified Value 0.063 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 0.06 | 0.1 | < 0.1 | < 2 | 0.06 | 0.041 | 0.07 | 0.06 | 0.062 | 0.064 | 0.062 | 0.062 | 0.07 | 0.07 | 0.077 | 0.061 | 0.06 | 0.0795 | 0.06323 | [7 empty] | 1
| 1 | 2 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.044 | < 0.05 | 0.06 | 0.063 | 0.066 | 0.065 | 0.064 | 0.07 | 0.07 | 0.069 | 0.058 | 0.06 | 0.0769 | 0.06303 | [7 empty] | 27
| 1 | 3 | 0.05 | 0.05 | < 0.1 | < 2 | 0.07 | 0.044 | 0.09 | 0.05 | 0.067 | 0.064 | 0.062 | 0.061 | 0.06 | 0.07 | 0.08 | 0.059 | 0.06 | 0.0873 | 0.06312 | [7 empty] | 16
| 1 | 4 | 0.07 | 0.05 | < 0.1 | < 2 | 0.06 | 0.043 | 0.09 | 0.07 | 0.063 | 0.065 | 0.071 | 0.057 | 0.07 | 0.08 | 0.068 | 0.063 | 0.06 | 0.0849 | 0.0631 | [7 empty] | 0.063353
| 1 | 5 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.045 | 0.08 | 0.05 | 0.072 | 0.063 | 0.059 | 0.068 | 0.07 | 0.07 | 0.047 | 0.06 | 0.06 | 0.0792 | 0.06276 | [7 empty] | 35
| 1 | 6 | 0.06 | < 0.05 | < 0.1 | < 2 | 0.07 | 0.041 | 0.09 | 0.05 | 0.064 | 0.065 | 0.064 | 0.059 | 0.07 | 0.08 | 0.074 | 0.057 | 0.07 | 0.0876 | 0.06577
| Mean |  | 0.063333 | 0.066667 |  |  | 0.063333 | 0.043 | 0.084 | 0.056667 | 0.065167 | 0.0645 | 0.063833 | 0.061833 | 0.068333 | 0.073333 | 0.069167 | 0.059667 | 0.061667 | 0.082567 | 0.063502
| Median |  | 0.065 | 0.05 |  |  | 0.06 | 0.0435 | 0.09 | 0.055 | 0.0635 | 0.0645 | 0.063 | 0.0615 | 0.07 | 0.07 | 0.0715 | 0.0595 | 0.06 | 0.0822 | 0.06311
| Std Dev. |  | 0.008165 | 0.028868 |  |  | 0.005164 | 0.0016733 | 0.0089443 | 0.008165 | 0.0037639 | 0.0010488 | 0.0040702 | 0.0038687 | 0.0040825 | 0.005164 | 0.011788 | 0.0021602 | 0.0040825 | 0.0046051 | 0.0011224
| Rel.Std.Dev. |  | 12.89% | 43.3% |  |  | 8.154% | 3.891% | 10.65% | 14.41% | 5.776% | 1.626% | 6.376% | 6.257% | 5.974% | 7.042% | 17.04% | 3.621% | 6.62% | 5.577% | 1.767%
| PDM3 |  | -0.03172% | 5.23% |  |  | -0.03172% | -32.13% | 32.59% | -10.55% | 2.862% | 1.81% | 0.7575% | -2.399% | 7.861% | 15.75% | 9.176% | -5.819% | -2.662% | 30.33% | 0.234%
| Z-Score (Absolute) |  | 0.02 | 2.47 | 1.88 | 130.7 | 0.02 | 2.86 | 1.49 | 0.95 | 0.23 | 0.14 | 0.05 | 0.23 | 0.67 | 1.37 | 0.79 | 0.54 | 0.26 | 2.66 | 0
| Analytical results for Ir in OREAS 232b (Indicative Value 0.003 ppm) | [27 empty] | Indicative
Ir | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | < 0.001 | [25 empty] | 1
| 1 | 2 | 0.0006 | [25 empty] | 4
| 1 | 3 | 0.0022 | [25 empty] | 16
| 1 | 4 | 0.0041 | [25 empty] | 0.0025167
| 1 | 5 | 0.0045 | [25 empty] | 10
| 1 | 6 | 0.0032
| Mean |  | 0.00292
| Median |  | 0.0032
| Std Dev. |  | 0.0015707
| Rel.Std.Dev. |  | 53.79%
| PDM3 |  | 16.03%
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 232b (Certified Value 2.51 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 2.5766 | 2.5064 | 2.39 | 1.78 | 2.4823 | 2.528 | 2.38 | 2.5 | 2.59 | 2.52 | 2.47 | 2.45 | 2.46 | 2.3888 | 2.476 | 2.5154 | 2.48 | 2.46 | 2.53 | 2.54 | 2.62 | 2.5497 | 2.6392 | 3.1581 |  |  | 1
| 1 | 2 | 2.5941 | 2.5567 | 2.42 | 1.8 | 2.505 | 2.565 | 2.45 | 2.54 | 2.56 | 2.46 | 2.55 | 2.61 | 2.45 | 2.4205 | 2.4782 | 2.4406 | 2.48 | 2.47 | 2.48 | 2.54 | 2.6 | 2.4558 | 2.6311 | 3.1544 |  |  | #N/A
| 1 | 3 | 2.5786 | 2.5431 | 2.42 | 2.42 | 2.5 | 2.567 | 2.4 | 2.54 | 2.52 | 2.43 | 2.47 | 2.48 | 2.42 | 2.4215 | 2.4537 | 2.4822 | 2.44 | 2.45 | 2.38 | 2.54 | 2.64 | 2.5766 | 2.6487 | 3.1659 |  |  | 16
| 1 | 4 | 2.5786 | 2.519 | 2.44 | 2.34 | 2.47 | 2.526 | 2.43 | 2.57 | 2.52 | 2.47 | 2.61 | 2.48 | 2.47 | 2.497 | 2.4793 | 2.5154 | 2.57 | 2.46 | 2.54 | 2.56 | 2.64 | 2.5955 | 2.6977 | 3.1429 |  |  | 2.5063
| 1 | 5 | 2.6068 | 2.5706 | 2.45 | 2.33 | 2.46 | 2.559 | 2.42 | 2.56 | 2.52 | 2.48 | 2.39 | 2.57 | 2.44 | 2.3338 | 2.4507 | 2.4323 | 2.56 | 2.43 | 2.52 | 2.55 | 2.56 | 2.5214 | 2.5778 | 3.1599 |  |  | 36
| 1 | 6 | 2.5607 | 2.5487 | 2.41 | 2.12 | 2.51 | 2.55 | 2.42 | 2.59 | 2.52 | 2.46 | 2.59 | 2.43 | 2.47 | 2.0349 | 2.4622 | 2.5403 | 2.49 | 2.46 | 2.4 | 2.55 | 2.52 | 2.5075 | 2.6138 | 3.1581
| Mean |  | 2.5826 | 2.5407 | 2.4217 | 2.1317 | 2.4879 | 2.5492 | 2.4167 | 2.55 | 2.5383 | 2.47 | 2.5133 | 2.5033 | 2.4517 | 2.3494 | 2.4667 | 2.4877 | 2.5033 | 2.455 | 2.475 | 2.5467 | 2.5967 | 2.5344 | 2.6347 | 3.1565
| Median |  | 2.5786 | 2.5459 | 2.42 | 2.225 | 2.4912 | 2.5545 | 2.42 | 2.55 | 2.52 | 2.465 | 2.51 | 2.48 | 2.455 | 2.4047 | 2.4691 | 2.4988 | 2.485 | 2.46 | 2.5 | 2.545 | 2.61 | 2.5355 | 2.6351 | 3.1581
| Std Dev. |  | 0.015909 | 0.023953 | 0.02137 | 0.28273 | 0.020285 | 0.01817 | 0.024221 | 0.030984 | 0.029944 | 0.029665 | 0.084301 | 0.070899 | 0.019408 | 0.16294 | 0.012834 | 0.043862 | 0.050859 | 0.013784 | 0.06921 | 0.008165 | 0.048028 | 0.050633 | 0.03968 | 0.0076733
| Rel.Std.Dev. |  | 0.616% | 0.9427% | 0.8824% | 13.26% | 0.8154% | 0.7128% | 1.002% | 1.215% | 1.18% | 1.201% | 3.354% | 2.832% | 0.7916% | 6.935% | 0.5203% | 1.763% | 2.032% | 0.5615% | 2.796% | 0.3206% | 1.85% | 1.998% | 1.506% | 0.2431%
| PDM3 |  | 3.044% | 1.376% | -3.375% | -14.95% | -0.7335% | 1.712% | -3.575% | 1.745% | 1.28% | -1.447% | 0.282% | -0.117% | -2.178% | -6.258% | -1.58% | -0.7408% | -0.117% | -2.045% | -1.247% | 1.612% | 3.607% | 1.123% | 5.125% | 25.95%
| Z-Score (Absolute) |  | 1.2 | 0.57 | 1.24 | 5.62 | 0.23 | 0.69 | 1.31 | 0.71 | 0.53 | 0.5 | 0.15 | 0 | 0.78 | 2.33 | 0.55 | 0.24 | 0 | 0.73 | 0.43 | 0.66 | 1.41 | 0.47 | 1.99 | 9.88
| Analytical results for La in OREAS 232b (Certified Value 37.6 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 39.64 | 43.4146 | 35.9 | 42.1 | 33.07 | 32.9 | 39 | 40.1 | 38.9 | 36.9 | 40.1 | 40 | 36.2821 | 38 | 33.42 | 36.4 | 34.78 | 36.9 | 41.2 | 37.9 | 32.6 | 38.2047 |  |  |  |  | 1
| 1 | 2 | 38.84 | 42.8722 | 36.5 | 42.8 | 34.28 | 33.1 | 38 | 38.4 | 37.8 | 40.6 | 39.3 | 37.6 | 36.5449 | 38 | 35.84 | 36 | 36.46 | 35.9 | 37.3 | 38.6 | 31.7 | 36.8278 |  |  |  |  | 28
| 1 | 3 | 39.54 | 43.0304 | 37.9 | 40.7 | 33.57 | 33.2 | 39 | 38.7 | 40.2 | 37.4 | 38.3 | 39.8 | 36.4992 | 38 | 36.43 | 34.8 | 34.97 | 33.9 | 38 | 38.9 | 32.85 | 37.6221 |  |  |  |  | 16
| 1 | 4 | 38.97 | 42.9965 | 36.3 | 41.3 | 33.54 | 35 | 37 | 40.3 | 42.8 | 37.8 | 38.4 | 40.7 | 36.1155 | 38 | 36.2 | 34.7 | 36.33 | 36 | 38.5 | 37.9 | 33.52 | 38.6632 |  |  |  |  | 37.5965
| 1 | 5 | 39.82 | 43.6406 | 36.5 | 40.8 | 33.09 | 36.8 | 37 | 39.6 | 38.6 | 39.6 | 39.3 | 39 | 36.2957 | 37 | 35.77 | 35.6 | 35.06 | 36.9 | 38 | 38.9 | 32.52 | 38.6994 |  |  |  |  | 37
| 1 | 6 | 39.47 | 43.0982 | 37.5 | 42.4 | 32.07 | 39 | 38 | 40 | 40 | 37.8 | 37.9 | 38.4 | 35.5686 | 39 | 35.11 | 37.6 | 34.57 | 37.6 | 36.8 | 37.4 | 32.67 | 37.9935
| Mean |  | 39.38 | 43.1754 | 36.7667 | 41.6833 | 33.27 | 35 | 38 | 39.5167 | 39.7167 | 38.35 | 38.8833 | 39.25 | 36.2177 | 38 | 35.4617 | 35.85 | 35.3617 | 36.2 | 38.3 | 38.2667 | 32.6433 | 38.0018
| Median |  | 39.505 | 43.0643 | 36.5 | 41.7 | 33.315 | 34.1 | 38 | 39.8 | 39.45 | 37.8 | 38.85 | 39.4 | 36.2889 | 38 | 35.805 | 35.8 | 35.015 | 36.45 | 38 | 38.25 | 32.635 | 38.0991
| Std Dev. |  | 0.38848 | 0.29139 | 0.76594 | 0.87502 | 0.73422 | 2.4698 | 0.89443 | 0.78846 | 1.7555 | 1.4307 | 0.82077 | 1.138 | 0.35448 | 0.63246 | 1.0966 | 1.084 | 0.81896 | 1.2931 | 1.5414 | 0.62183 | 0.58633 | 0.70539
| Rel.Std.Dev. |  | 0.9865% | 0.6749% | 2.083% | 2.099% | 2.207% | 7.057% | 2.354% | 1.995% | 4.42% | 3.731% | 2.111% | 2.899% | 0.9788% | 1.664% | 3.092% | 3.024% | 2.316% | 3.572% | 4.025% | 1.625% | 1.796% | 1.856%
| PDM3 |  | 4.744% | 14.84% | -2.207% | 10.87% | -11.51% | -6.906% | 1.073% | 5.107% | 5.639% | 2.004% | 3.423% | 4.398% | -3.667% | 1.073% | -5.678% | -4.645% | -5.944% | -3.714% | 1.871% | 1.782% | -13.17% | 1.078%
| Z-Score (Absolute) |  | 0.58 | 2.16 | 0.52 | 1.54 | 1.97 | 1.25 | 0 | 0.63 | 0.72 | 0.15 | 0.37 | 0.52 | 0.74 | 0 | 1.06 | 0.9 | 1.1 | 0.75 | 0.12 | 0.11 | 2.24 | 0
| Analytical results for Li in OREAS 232b (Certified Value 50 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 49.7 | 54.08 | 50 | 55 | 52.75 | 48 | 50.4 | 48.4 | 71.8 | 52.7 | 47.4 | 49.5 | 45.9 | 50.6 | 52.9568 | 52 | 43.6 | 45.8 | 52.91 | 48.8 | 52.3 | 51.5 | 55.9 | 47.1671 |  |  | 1
| 1 | 2 | 50.5 | 55.02 | 50 | 55 | 55.035 | 48.7 | 51 | 48.8 | 68.5 | 51.4 | 47 | 50.9 | 44.6 | 51.1 | 53.3197 | 50 | 43.92 | 47.7 | 53.41 | 47.7 | 52 | 51.7 | 54 | 48.4543 |  |  | #N/A
| 1 | 3 | 49.3 | 55.44 | 50 | 53 | 52.4133 | 48.4 | 50.7 | 48.4 | 70.1 | 52.2 | 47.4 | 49.4 | 45.1 | 50.9 | 51.4457 | 51 | 44.38 | 45.1 | 52.61 | 45.2 | 49.8 | 52.7 | 56.8 | 47.6371 |  |  | 16
| 1 | 4 | 49.8 | 54.46 | 50 | 55 | 53.24 | 47.5 | 50.8 | 48.6 | 78.5 | 53.5 | 49.5 | 51.1 | 44.5 | 51.5 | 51.3289 | 52 | 43.5 | 47.7 | 53.76 | 47.4 | 49.8 | 52.4 | 57.5 | 47.3563 |  |  | 50.4211
| 1 | 5 | 50.4 | 54.38 | 50 | 55 | 54.5933 | 47.9 | 51.6 | 50.1 | 72.6 | 51.3 | 46.3 | 53.9 | 45.3 | 51.4 | 51.1222 | 50 | 43.77 | 49.6 | 51.34 | 49 | 49.9 | 51.2 | 55.5 | 46.8919 |  |  | 38
| 1 | 6 | 50.2 | 54.45 | 50 | 55 | 51.76 | 47.3 | 51.5 | 51.4 | 73 | 52.5 | 49.8 | 53 | 45.1 | 51 | 52.1274 | 52 | 43.24 | 47.8 | 52.32 | 51 | 51.7 | 50.4 | 55.7 | 48.1342
| Mean |  | 49.9833 | 54.6383 | 50 | 54.6667 | 53.2986 | 47.9667 | 51 | 49.2833 | 72.4167 | 52.2667 | 47.9 | 51.3 | 45.0833 | 51.0833 | 52.0501 | 51.1667 | 43.735 | 47.2833 | 52.725 | 48.1833 | 50.9167 | 51.65 | 55.9 | 47.6068
| Median |  | 50 | 54.455 | 50 | 55 | 52.995 | 47.95 | 50.9 | 48.7 | 72.2 | 52.35 | 47.4 | 51 | 45.1 | 51.05 | 51.7865 | 51.5 | 43.685 | 47.7 | 52.76 | 48.25 | 50.8 | 51.6 | 55.8 | 47.4967
| Std Dev. |  | 0.46224 | 0.49681 | 0 | 0.8165 | 1.2765 | 0.52789 | 0.46904 | 1.2172 | 3.4196 | 0.83106 | 1.417 | 1.8276 | 0.50761 | 0.33116 | 0.91503 | 0.98319 | 0.39242 | 1.6092 | 0.85676 | 1.9354 | 1.2024 | 0.83126 | 1.1983 | 0.59406
| Rel.Std.Dev. |  | 0.9248% | 0.9093% | 0% | 1.494% | 2.395% | 1.101% | 0.9197% | 2.47% | 4.722% | 1.59% | 2.958% | 3.563% | 1.126% | 0.6483% | 1.758% | 1.922% | 0.8973% | 3.403% | 1.625% | 4.017% | 2.361% | 1.609% | 2.144% | 1.248%
| PDM3 |  | -0.8682% | 8.364% | -0.8351% | 8.42% | 5.707% | -4.868% | 1.148% | -2.257% | 43.62% | 3.66% | -5% | 1.743% | -10.59% | 1.313% | 3.231% | 1.479% | -13.26% | -6.223% | 4.569% | -4.438% | 0.9829% | 2.437% | 10.87% | -5.582%
| Z-Score (Absolute) |  | 0.33 | 1.08 | NA | 1.09 | 0.67 | 0.95 | 0.03 | 0.55 | 6.49 | 0.36 | 0.97 | 0.07 | 1.83 | 0 | 0.29 | 0.03 | 2.24 | 1.16 | 0.5 | 0.88 | 0.05 | 0.17 | 1.47 | 1.06
Results from laboratory 1.05 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 232b (Certified Value 0.28 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.26083 | 0.3 | 0.3 | 0.3 | 0.225 | 0.28 | 0.31 | 0.2 | 0.1573 | 0.45054 | [16 empty] | 1
| 1 | 2 | 0.26309 | 0.32 | 0.3 | 0.3 | 0.232 | 0.28 | 0.32 | 0.2 | 0.1336 | 0.42509 | [16 empty] | 29
| 1 | 3 | 0.26502 | 0.32 | 0.3 | 0.3 | 0.238 | 0.27 | 0.28 | 0.2 | 0.1313 | 0.43366 | [16 empty] | 16
| 1 | 4 | 0.26593 | 0.28 | 0.3 | 0.3 | 0.236 | 0.28 | 0.31 | 0.2 | 0.1441 | 0.44776 | [16 empty] | 0.27659
| 1 | 5 | 0.26036 | 0.32 | 0.3 | 0.3 | 0.228 | 0.28 | 0.28 | 0.2 | 0.1632 | 0.46215 | [16 empty] | 39
| 1 | 6 | 0.26135 | 0.3 | 0.3 | 0.3 | 0.232 | 0.28 | 0.32 | 0.2 | 0.1508 | 0.45743
| Mean |  | 0.26276 | 0.30667 | 0.3 | 0.3 | 0.23183 | 0.27833 | 0.30333 | 0.2 | 0.14672 | 0.4461
| Median |  | 0.26222 | 0.31 | 0.3 | 0.3 | 0.232 | 0.28 | 0.31 | 0.2 | 0.14745 | 0.44915
| Std Dev. |  | 0.0023112 | 0.01633 | 0 | 0 | 0.0048339 | 0.0040825 | 0.018619 | 3.0405e-17 | 0.012782 | 0.014175
| Rel.Std.Dev. |  | 0.8796% | 5.325% | 0% | 0% | 2.085% | 1.467% | 6.138% | 1.52e-14% | 8.712% | 3.177%
| PDM3 |  | -4.998% | 10.88% | 8.465% | 8.465% | -16.18% | 0.6319% | 9.671% | -27.69% | -46.95% | 61.29%
| Z-Score (Absolute) |  | 0.37 | 0.67 | NA | NA | 1.11 | 0 | 0.59 | NA | 3.13 | 3.99
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 232b (Certified Value 1.62 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 1.5803 | 1.6804 | 1.52 | 1.53 | 1.59 | 1.672 | 1.58 | 1.64 | 1.65 | 1.61 | 1.54 | 1.59 | 1.59 | 1.567 | 1.6164 | 1.7006 | 1.61 | 1.7 | 1.7 | 1.62 | 1.68 | 1.6853 | 1.8464 | 1.5406 |  |  | 1
| 1 | 2 | 1.5587 | 1.6972 | 1.6 | 1.52 | 1.62 | 1.686 | 1.61 | 1.66 | 1.63 | 1.57 | 1.6 | 1.68 | 1.55 | 1.5866 | 1.6528 | 1.6645 | 1.63 | 1.71 | 1.69 | 1.62 | 1.67 | 1.6322 | 1.8668 | 1.6027 |  |  | #N/A
| 1 | 3 | 1.5539 | 1.6765 | 1.59 | 1.47 | 1.62 | 1.665 | 1.57 | 1.67 | 1.6 | 1.57 | 1.54 | 1.61 | 1.56 | 1.5749 | 1.6489 | 1.6886 | 1.56 | 1.69 | 1.71 | 1.62 | 1.69 | 1.7007 | 1.8583 | 1.5683 |  |  | 16
| 1 | 4 | 1.5765 | 1.654 | 1.62 | 1.48 | 1.59 | 1.659 | 1.62 | 1.64 | 1.59 | 1.57 | 1.66 | 1.57 | 1.56 | 1.5994 | 1.6482 | 1.7127 | 1.65 | 1.71 | 1.72 | 1.62 | 1.69 | 1.7186 | 1.8927 | 1.7352 |  |  | 1.6239
| 1 | 5 | 1.6305 | 1.6996 | 1.64 | 1.44 | 1.61 | 1.682 | 1.65 | 1.68 | 1.59 | 1.57 | 1.51 | 1.62 | 1.6 | 1.5876 | 1.6618 | 1.6584 | 1.62 | 1.69 | 1.73 | 1.62 | 1.64 | 1.6661 | 1.7998 | 1.585 |  |  | 40
| 1 | 6 | 1.5571 | 1.6777 | 1.54 | 1.57 | 1.59 | 1.677 | 1.67 | 1.7 | 1.61 | 1.58 | 1.62 | 1.55 | 1.58 | 1.5778 | 1.6513 | 1.7308 | 1.63 | 1.7 | 1.71 | 1.61 | 1.63 | 1.6834 | 1.8317 | 1.5758
| Mean |  | 1.5762 | 1.6809 | 1.585 | 1.5017 | 1.6033 | 1.6735 | 1.6167 | 1.665 | 1.6117 | 1.5783 | 1.5783 | 1.6033 | 1.5733 | 1.5822 | 1.6466 | 1.6926 | 1.6167 | 1.7 | 1.71 | 1.6183 | 1.6667 | 1.6811 | 1.8493 | 1.6013
| Median |  | 1.5676 | 1.6791 | 1.595 | 1.5 | 1.6 | 1.6745 | 1.615 | 1.665 | 1.605 | 1.57 | 1.57 | 1.6 | 1.57 | 1.5822 | 1.6501 | 1.6946 | 1.625 | 1.7 | 1.71 | 1.62 | 1.675 | 1.6843 | 1.8523 | 1.5804
| Std Dev. |  | 0.028753 | 0.016547 | 0.046368 | 0.047081 | 0.015055 | 0.010252 | 0.038816 | 0.023452 | 0.024014 | 0.016021 | 0.057417 | 0.045461 | 0.019664 | 0.011385 | 0.015562 | 0.027937 | 0.030768 | 0.0089443 | 0.014142 | 0.0040825 | 0.02582 | 0.029764 | 0.031747 | 0.068732
| Rel.Std.Dev. |  | 1.824% | 0.9844% | 2.925% | 3.135% | 0.939% | 0.6126% | 2.401% | 1.409% | 1.49% | 1.015% | 3.638% | 2.835% | 1.25% | 0.7195% | 0.9451% | 1.651% | 1.903% | 0.5261% | 0.827% | 0.2523% | 1.549% | 1.771% | 1.717% | 4.292%
| PDM3 |  | -2.937% | 3.513% | -2.393% | -7.524% | -1.264% | 3.057% | -0.4425% | 2.534% | -0.7504% | -2.803% | -2.803% | -1.264% | -3.111% | -2.564% | 1.398% | 4.234% | -0.4425% | 4.689% | 5.305% | -0.3399% | 2.637% | 3.522% | 13.88% | -1.391%
| Z-Score (Absolute) |  | 0.69 | 1.1 | 0.54 | 1.97 | 0.23 | 0.97 | 0 | 0.83 | 0.09 | 0.66 | 0.66 | 0.23 | 0.74 | 0.59 | 0.51 | 1.3 | 0 | 1.43 | 1.6 | 0.03 | 0.86 | 1.1 | 3.98 | 0.26
| Analytical results for Mn in OREAS 232b (Certified Value 0.041 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0439 | 0.041702 | 0.042 | 0.0392 | 0.043948 | 0.0406 | 0.0401 | 0.0386 | 0.0442 | 0.0416 | 0.04 | 0.0415 | 0.0419 | 0.0403 | 0.04188 | 0.0395 | 0.0398 | 0.04001 | 0.0423 | 0.0407 | 0.0436 | 0.044528 | 0.0419 | 0.0417 |  |  | 1
| 1 | 2 | 0.0436 | 0.042404 | 0.0412 | 0.0385 | 0.043657 | 0.0417 | 0.0391 | 0.0394 | 0.042 | 0.0402 | 0.0415 | 0.0432 | 0.0412 | 0.0412 | 0.041388 | 0.0387 | 0.0392 | 0.04002 | 0.042 | 0.0406 | 0.0431 | 0.042955 | 0.0431 | 0.0412 |  |  | 30
| 1 | 3 | 0.0444 | 0.042454 | 0.0418 | 0.0386 | 0.043686 | 0.0408 | 0.0387 | 0.0393 | 0.042 | 0.0405 | 0.0404 | 0.0417 | 0.0414 | 0.0405 | 0.041317 | 0.0395 | 0.0388 | 0.0398 | 0.0422 | 0.0407 | 0.0437 | 0.045155 | 0.0426 | 0.0422 |  |  | 16
| 1 | 4 | 0.0435 | 0.042212 | 0.0426 | 0.0371 | 0.043561 | 0.041 | 0.0409 | 0.0387 | 0.0412 | 0.0408 | 0.0428 | 0.0407 | 0.0412 | 0.0412 | 0.041807 | 0.041 | 0.0407 | 0.04006 | 0.0427 | 0.0407 | 0.0441 | 0.045078 | 0.0433 | 0.0417 |  |  | 0.041416
| 1 | 5 | 0.0448 | 0.042508 | 0.0422 | 0.0364 | 0.043833 | 0.0416 | 0.0409 | 0.0394 | 0.0401 | 0.0405 | 0.0392 | 0.0421 | 0.0408 | 0.0414 | 0.041909 | 0.0387 | 0.0399 | 0.03967 | 0.0428 | 0.0407 | 0.0424 | 0.044011 | 0.0413 | 0.0417 |  |  | 41
| 1 | 6 | 0.0439 | 0.042339 | 0.0432 | 0.0405 | 0.043237 | 0.0408 | 0.0402 | 0.0398 | 0.0415 | 0.0406 | 0.0423 | 0.04 | 0.0411 | 0.041 | 0.041804 | 0.0403 | 0.0397 | 0.04024 | 0.0423 | 0.0405 | 0.0424 | 0.044253 | 0.0415 | 0.0417
| Mean |  | 0.044017 | 0.04227 | 0.042167 | 0.038383 | 0.043653 | 0.041083 | 0.039983 | 0.0392 | 0.041833 | 0.0407 | 0.041033 | 0.041533 | 0.041267 | 0.040933 | 0.041684 | 0.039617 | 0.039683 | 0.039967 | 0.042383 | 0.04065 | 0.043217 | 0.04433 | 0.042283 | 0.0417
| Median |  | 0.0439 | 0.042372 | 0.0421 | 0.03855 | 0.043672 | 0.0409 | 0.04015 | 0.03935 | 0.04175 | 0.04055 | 0.04095 | 0.0416 | 0.0412 | 0.0411 | 0.041805 | 0.0395 | 0.03975 | 0.040015 | 0.0423 | 0.0407 | 0.04335 | 0.04439 | 0.04225 | 0.0417
| Std Dev. |  | 0.00049565 | 0.00029638 | 0.00068605 | 0.0014689 | 0.00024574 | 0.00045789 | 0.00091305 | 0.00046043 | 0.0013545 | 0.00048166 | 0.0013981 | 0.0011112 | 0.00036697 | 0.00043665 | 0.00026098 | 0.00090425 | 0.00064936 | 0.00020196 | 0.00030605 | 8.3666e-05 | 0.00070828 | 0.00080965 | 0.00084004 | 0.00031623
| Rel.Std.Dev. |  | 1.126% | 0.7012% | 1.627% | 3.827% | 0.5629% | 1.115% | 2.284% | 1.175% | 3.238% | 1.183% | 3.407% | 2.675% | 0.8893% | 1.067% | 0.6261% | 2.282% | 1.636% | 0.5053% | 0.7221% | 0.2058% | 1.639% | 1.826% | 1.987% | 0.7583%
| PDM3 |  | 6.28% | 2.062% | 1.813% | -7.322% | 5.403% | -0.8026% | -3.459% | -5.35% | 1.008% | -1.728% | -0.9233% | 0.284% | -0.3599% | -1.165% | 0.6481% | -4.344% | -4.183% | -3.499% | 2.336% | -1.849% | 4.348% | 7.037% | 2.095% | 0.6864%
| Z-Score (Absolute) |  | 2.01 | 0.67 | 0.59 | 2.32 | 1.73 | 0.24 | 1.09 | 1.69 | 0.33 | 0.54 | 0.28 | 0.1 | 0.1 | 0.36 | 0.22 | 1.36 | 1.32 | 1.1 | 0.76 | 0.58 | 1.4 | 2.25 | 0.68 | 0.23
| Analytical results for Mo in OREAS 232b (Certified Value 1.05 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 1.1 | 1 | < 1 | < 2 | 1 | 1.59 | 0.9 | 2.8 | 1.06 | 0.97 | 1.03 | 1.05 | 1.04 | 1.1263 | 1.19 | 1.267 | 1.24 | < 1 | 1.04 | 1.29 | 1.01 | 0.8922 | < 5 | 2.5133 | 1.0684 |  | 1
| 1 | 2 | 1.2 | 1 | < 1 | < 2 | 1 | 1.66 | 0.9 | 2.5 | 0.99 | 0.96 | 1.1 | 1.03 | 1.23 | 1.1234 | 1.03 | 1.032 | 1.02 | < 1 | 1.01 | 1.18 | 1.05 | 0.9246 | < 5 | 2.9583 | 1.0414 |  | 31
| 1 | 3 | 1.2 | 1 | < 1 | < 2 | 1 | 1.52 | 1.2 | 2.7 | 1.04 | 1 | 0.98 | 0.96 | 1.02 | 1.0788 | 1.19 | 1.376 | 0.95 | < 1 | 1.08 | 1.17 | 1.06 | 0.8695 | < 5 | 2.8471 | 1.0623 |  | 16
| 1 | 4 | 1.1 | 1 | < 1 | < 2 | 1 | 1.6 | 1.3 | 2.9 | 1.05 | 1.01 | 1 | 0.99 | 1.05 | 1.0759 | 1.08 | 1.323 | 1.07 | < 1 | 1.09 | 1.18 | 1.04 | 0.8922 | < 5 | 2.2352 | 1.0722 |  | 1.0517
| 1 | 5 | 1.1 | 1 | < 1 | < 2 | 1 | 1.47 | 0.8 | 2.9 | 1.01 | 0.99 | 1.16 | 1.06 | 1.04 | 1.1273 | 1.08 | 1.219 | 1.03 | < 1 | 0.96 | 1.17 | 0.98 | 0.8414 | < 5 | 2.3465 | 1.0161 |  | 42
| 1 | 6 | 1.1 | 1 | < 1 | < 2 | 1 | 1.5 | 1.1 | 3 | 1.02 | 1.01 | 1.14 | 0.94 | 1.03 | 1.1135 | 1.16 | 1.09 | 1.06 | < 1 | 1.09 | 1.09 | 1 | 0.8382 | < 5 | 2.3465 | 1.1956
| Mean |  | 1.1333 | 1 |  |  | 1 | 1.5567 | 1.0333 | 2.8 | 1.0283 | 0.99 | 1.0683 | 1.005 | 1.0683 | 1.1076 | 1.1217 | 1.2178 | 1.0617 |  | 1.045 | 1.18 | 1.0233 | 0.87635 |  | 2.5411 | 1.076
| Median |  | 1.1 | 1 |  |  | 1 | 1.555 | 1 | 2.85 | 1.03 | 0.995 | 1.065 | 1.01 | 1.04 | 1.1185 | 1.12 | 1.243 | 1.045 |  | 1.06 | 1.175 | 1.025 | 0.88085 |  | 2.4299 | 1.0653
| Std Dev. |  | 0.05164 | 0 |  |  | 0 | 0.07174 | 0.19664 | 0.17889 | 0.026394 | 0.020976 | 0.075476 | 0.049295 | 0.079854 | 0.023886 | 0.067355 | 0.13369 | 0.09704 |  | 0.05244 | 0.063875 | 0.031411 | 0.03333 |  | 0.2959 | 0.062209
| Rel.Std.Dev. |  | 4.556% | 0% |  |  | 0% | 4.609% | 19.03% | 6.389% | 2.567% | 2.119% | 7.065% | 4.905% | 7.475% | 2.157% | 6.005% | 10.98% | 9.14% |  | 5.018% | 5.413% | 3.07% | 3.803% |  | 11.64% | 5.781%
| PDM3 |  | 7.762% | -4.916% |  |  | -4.916% | 48.01% | -1.747% | 166.2% | -2.222% | -5.867% | 1.581% | -4.441% | 1.581% | 5.311% | 6.652% | 15.8% | 0.9472% |  | -0.6375% | 12.2% | -2.698% | -16.67% |  | 141.6% | 2.312%
| Z-Score (Absolute) |  | NA | NA | 5.61 | 0.67 | NA | 4.82 | NA | 17.09 | 0.39 | 0.77 | 0 | 0.62 | 0 | 0.39 | 0.53 | 1.48 | 0.07 | 5.61 | 0.23 | 1.1 | 0.44 | 1.89 | 14.13 | 14.53 | 0.08
Results from laboratories 1.01, 1.05, 1.08 and 1.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 232b (Certified Value 0.81 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.805 | 0.84396 | 0.79 | 0.79 | 0.88 | 0.796 | 0.78 | 0.83 | 0.82 | 0.81 | 0.79 | 0.8 | 0.83 | 0.772 | 0.79428 | 0.801 | 0.8 | 0.82 | 0.8 | 0.81 | 0.829 | 0.82928 | 0.8571 |  |  |  | 1
| 1 | 2 | 0.8091 | 0.84726 | 0.77 | 0.78 | 0.906 | 0.805 | 0.8 | 0.84 | 0.83 | 0.83 | 0.83 | 0.85 | 0.8 | 0.792 | 0.78015 | 0.786 | 0.8 | 0.82 | 0.81 | 0.81 | 0.836 | 0.80252 | 0.7907 |  |  |  | #N/A
| 1 | 3 | 0.8075 | 0.84464 | 0.78 | 0.77 | 0.88833 | 0.81 | 0.78 | 0.844 | 0.82 | 0.78 | 0.79 | 0.81 | 0.8 | 0.78 | 0.79331 | 0.801 | 0.79 | 0.82 | 0.8 | 0.81 | 0.85 | 0.83868 | 0.8853 |  |  |  | 16
| 1 | 4 | 0.8153 | 0.84065 | 0.78 | 0.78 | 0.89633 | 0.81 | 0.8 | 0.83 | 0.82 | 0.78 | 0.85 | 0.8 | 0.79 | 0.793 | 0.78252 | 0.809 | 0.84 | 0.82 | 0.82 | 0.81 | 0.849 | 0.84701 | 0.8086 |  |  |  | 0.81004
| 1 | 5 | 0.8087 | 0.85167 | 0.79 | 0.74 | 0.9125 | 0.79 | 0.81 | 0.847 | 0.82 | 0.78 | 0.77 | 0.83 | 0.83 | 0.799 | 0.80032 | 0.786 | 0.84 | 0.82 | 0.82 | 0.8 | 0.812 | 0.82183 | 0.9395 |  |  |  | 43
| 1 | 6 | 0.803 | 0.84718 | 0.77 | 0.83 | 0.887 | 0.79 | 0.8 | 0.854 | 0.81 | 0.79 | 0.83 | 0.78 | 0.8 | 0.786 | 0.7821 | 0.816 | 0.81 | 0.82 | 0.82 | 0.8 | 0.79 | 0.83006 | 0.8631
| Mean |  | 0.8081 | 0.84589 | 0.78 | 0.78167 | 0.89503 | 0.80017 | 0.795 | 0.84083 | 0.82 | 0.795 | 0.81 | 0.81167 | 0.80833 | 0.787 | 0.78878 | 0.79983 | 0.81333 | 0.82 | 0.81167 | 0.80667 | 0.82767 | 0.82823 | 0.85738
| Median |  | 0.8081 | 0.84591 | 0.78 | 0.78 | 0.89233 | 0.8005 | 0.8 | 0.842 | 0.82 | 0.785 | 0.81 | 0.805 | 0.8 | 0.789 | 0.78792 | 0.801 | 0.805 | 0.82 | 0.815 | 0.81 | 0.8325 | 0.82967 | 0.8601
| Std Dev. |  | 0.0042233 | 0.0037335 | 0.0089443 | 0.029269 | 0.012349 | 0.0093897 | 0.012247 | 0.0095586 | 0.0063246 | 0.020736 | 0.030984 | 0.024833 | 0.017224 | 0.009798 | 0.0082717 | 0.01209 | 0.021602 | 1.2162e-16 | 0.0098319 | 0.005164 | 0.023192 | 0.015287 | 0.053632
| Rel.Std.Dev. |  | 0.5226% | 0.4414% | 1.147% | 3.744% | 1.38% | 1.173% | 1.541% | 1.137% | 0.7713% | 2.608% | 3.825% | 3.059% | 2.131% | 1.245% | 1.049% | 1.512% | 2.656% | 1.483e-14% | 1.211% | 0.6402% | 2.802% | 1.846% | 6.255%
| PDM3 |  | -0.2397% | 4.426% | -3.709% | -3.503% | 10.49% | -1.219% | -1.857% | 3.801% | 1.229% | -1.857% | -0.005137% | 0.2006% | -0.2109% | -2.844% | -2.625% | -1.26% | 0.4064% | 1.229% | 0.2006% | -0.4166% | 2.176% | 2.245% | 5.844%
| Z-Score (Absolute) |  | 0.09 | 1.61 | 1.35 | 1.27 | 3.82 | 0.44 | 0.67 | 1.39 | 0.45 | 0.67 | 0 | 0.07 | 0.07 | 1.03 | 0.95 | 0.45 | 0.15 | 0.45 | 0.07 | 0.15 | 0.79 | 0.82 | 2.13
| Analytical results for Nb in OREAS 232b (Certified Value 13.5 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
1
| 1 | 1 | 14.78 | 11.9791 | 0.7 | 16.42 | 13.6 | 10 | 13.3 | 13.1 | 14.4 | 12.8 | 13 | 13.4 | 13.69 | 14.0752 | 15.4 | 10.6 | 14.93 | 15 | 11.42 | 11.8568 | 15.257 | [5 empty] | 1
| 1 | 2 | 14.42 | 12.058 | 0.5 | 15.36 | 13.4 | 10 | 12.3 | 13.9 | 13.7 | 13.5 | 13.4 | 13.1 | 14.5 | 14.2967 | 15.6 | 11.9 | 14.37 | 14 | 11.18 | 10.8682 | 15.558 | [5 empty] | 32
| 1 | 3 | 14.14 | 12.0725 | 0.2 | 16.28 | 12.8 | 10 | 12.5 | 13.5 | 13.6 | 13.4 | 13.2 | 12.8 | 14.72 | 13.8683 | 16 | 11 | 13.85 | 14.2 | 11.64 | 11.0708 | 15.225 | [5 empty] | 16
| 1 | 4 | 13.95 | 11.9436 | 0.3 | 16.11 | 12.4 | 10 | 12.3 | 12.6 | 14.3 | 14.8 | 13.2 | 12.9 | 14.61 | 14.4243 | 15.7 | 11.8 | 14.58 | 14.2 | 11.01 | 11.0998 | 15.1816 | [5 empty] | 13.4954
| 1 | 5 | 14.31 | 11.9104 | 0.4 | 15.37 | 12.5 | 9 | 11.8 | 13.2 | 13.6 | 12.7 | 13.3 | 13.1 | 14.93 | 14.104 | 15.5 | 11.6 | 14.55 | 14.1 | 11.11 | 10.8086 | 15.0491 | [5 empty] | 44
| 1 | 6 | 14.24 | 12.0228 | 1.1 | 16.45 | 12.1 | 9 | 13.1 | 12.6 | 14 | 14.2 | 13 | 12.6 | 13.95 | 13.8639 | 15.2 | 11.7 | 15.25 | 13.6 | 11.69 | 10.5743 | 15.3958
| Mean |  | 14.3067 | 11.9977 | 0.53333 | 15.9983 | 12.8 | 9.6667 | 12.55 | 13.15 | 13.9333 | 13.5667 | 13.1833 | 12.9833 | 14.4 | 14.1054 | 15.5667 | 11.4333 | 14.5883 | 14.1833 | 11.3417 | 11.0464 | 15.2777
| Median |  | 14.275 | 12.0009 | 0.45 | 16.195 | 12.65 | 10 | 12.4 | 13.15 | 13.85 | 13.45 | 13.2 | 13 | 14.555 | 14.0896 | 15.55 | 11.65 | 14.565 | 14.15 | 11.3 | 10.9695 | 15.241
| Std Dev. |  | 0.28154 | 0.06442 | 0.3266 | 0.50515 | 0.58992 | 0.5164 | 0.55767 | 0.50892 | 0.3559 | 0.81158 | 0.16021 | 0.27869 | 0.47833 | 0.22527 | 0.27325 | 0.5164 | 0.47935 | 0.45789 | 0.28506 | 0.44076 | 0.17731
| Rel.Std.Dev. |  | 1.968% | 0.5369% | 61.24% | 3.158% | 4.609% | 5.342% | 4.444% | 3.87% | 2.554% | 5.982% | 1.215% | 2.147% | 3.322% | 1.597% | 1.755% | 4.517% | 3.286% | 3.228% | 2.513% | 3.99% | 1.161%
| PDM3 |  | 6.012% | -11.1% | -96.05% | 18.55% | -5.153% | -28.37% | -7.005% | -2.559% | 3.245% | 0.5282% | -2.312% | -3.794% | 6.703% | 4.52% | 15.35% | -15.28% | 8.099% | 5.098% | -15.96% | -18.15% | 13.21%
| Z-Score (Absolute) |  | 0.64 | 0.95 | 8.85 | 1.81 | 0.4 | NA | 0.57 | 0.16 | 0.38 | 0.13 | 0.13 | 0.27 | 0.71 | 0.5 | 1.51 | 1.34 | 0.84 | 0.56 | 1.4 | 1.6 | 1.31
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 232b (Certified Value 33.5 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 1
2
| 1 | 1 | 33.0726 | 32.3 | 34.7 | 32.2 | 34.4 | 29.94 | 32.78 | 33.1 | 28.8 | 34.1882 | [16 empty] | 1
| 1 | 2 | 33.0026 | 33 | 35 | 33.8 | 35.2 | 31.76 | 32.15 | 33.7 | 28.27 | 33.366 | [16 empty] | 9
| 1 | 3 | 33.3138 | 34.3 | 32.6 | 29.8 | 36.4 | 31.93 | 30.42 | 34.2 | 29.2 | 34.5399 | [16 empty] | 16
| 1 | 4 | 32.8561 | 32 | 33.8 | 34.1 | 35.7 | 32.3 | 32.54 | 32.9 | 29.79 | 33.8353 | [16 empty] | 33.4711
| 1 | 5 | 32.7758 | 33.5 | 34.4 | 33.4 | 35.5 | 31.67 | 32.34 | 33.4 | 28.72 | 34.5546 | [16 empty] | 45
| 1 | 6 | 33.0407 | 34.1 | 35 | 33.2 | 34.9 | 31.98 | 33.27 | 33.2 | 28.88 | 34.2904
| Mean |  | 33.0103 | 33.2 | 34.25 | 32.75 | 35.35 | 31.5967 | 32.25 | 33.4167 | 28.9433 | 34.129
| Median |  | 33.0217 | 33.25 | 34.55 | 33.3 | 35.35 | 31.845 | 32.44 | 33.3 | 28.84 | 34.2393
| Std Dev. |  | 0.1875 | 0.93808 | 0.92466 | 1.5846 | 0.6892 | 0.84016 | 0.97697 | 0.47081 | 0.51196 | 0.4578
| Rel.Std.Dev. |  | 0.568% | 2.826% | 2.7% | 4.839% | 1.95% | 2.659% | 3.029% | 1.409% | 1.769% | 1.341%
| PDM3 |  | -1.377% | -0.81% | 2.327% | -2.154% | 5.613% | -5.6% | -3.648% | -0.1627% | -13.53% | 1.966%
| Z-Score (Absolute) |  | 0.07 | 0.07 | 0.82 | 0.25 | 1.61 | 1.08 | 0.61 | 0.22 | 2.99 | 0.73
| Analytical results for Ni in OREAS 232b (Certified Value 63 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 62.9 | 58.8 | 65 | 65 | 58.9392 | 63 | 64.9 | 60 | 65.4 | 64 | 60.2 | 62.3 | 62.8 | 58.8 | 66.0567 | 67 | 69.65 | 72 | 58.71 | 61.8 | 73.5 | 65.3 | 62.81 | 62 | 59.124 | 62.4069 | 1
| 1 | 2 | 63.3 | 60.7 | 65 | 65 | 59.4267 | 63 | 65.1 | 63 | 65.3 | 61.7 | 64.2 | 68.1 | 61.2 | 58.6 | 65.9706 | 68.5 | 64.71 | 67 | 60.13 | 59.8 | 66.9 | 66.3 | 60.94 | 65 | 61.4159 | 62.1476 | 34
| 1 | 3 | 62.9 | 58.37 | 65 | 64 | 58.14 | 62 | 63.9 | 64 | 65.3 | 62 | 62.6 | 62.8 | 60.6 | 58.1 | 66.5831 | 69.3 | 60.57 | 62 | 59.45 | 58 | 67.9 | 67.6 | 64.68 | 62 | 62.2727 | 63.3129 | 16
| 1 | 4 | 62.1 | 62.72 | 65 | 66 | 59.6733 | 63 | 69.5 | 61 | 65.6 | 62.8 | 66.5 | 62.7 | 60.7 | 59.2 | 66.2997 | 67.3 | 63.98 | 64 | 60.23 | 59.5 | 67.8 | 66.1 | 64.35 | 64 | 64.5004 | 64.1584 | 63.2
| 1 | 5 | 63.3 | 59.43 | 65 | 64 | 60.25 | 63 | 68.7 | 62 | 65.9 | 61.5 | 60.5 | 66.2 | 62.4 | 58.6 | 66.6186 | 69.1 | 61.26 | 65 | 59.24 | 59.9 | 67.8 | 64.7 | 62.37 | 62 | 61.7158 | 63.846 | 46
| 1 | 6 | 61.7 | 59.08 | 65 | 65 | 58.725 | 62 | 67.7 | 63 | 66.3 | 62.2 | 66.5 | 62.7 | 60.4 | 58.3 | 65.5974 | 65.7 | 61.66 | 66 | 59.75 | 63.2 | 64.9 | 65 | 63.03 | 62 | 62.4869 | 64.4555
| Mean |  | 62.7 | 59.85 | 65 | 64.8333 | 59.1924 | 62.6667 | 66.6333 | 62.1667 | 65.6333 | 62.3667 | 63.4167 | 64.1333 | 61.35 | 58.6 | 66.1877 | 67.8167 | 63.6383 | 66 | 59.585 | 60.3667 | 68.1333 | 65.8333 | 63.03 | 62.8333 | 61.9193 | 63.3879
| Median |  | 62.9 | 59.255 | 65 | 65 | 59.1829 | 63 | 66.4 | 62.5 | 65.5 | 62.1 | 63.4 | 62.75 | 60.95 | 58.6 | 66.1782 | 67.9 | 62.82 | 65.5 | 59.6 | 59.85 | 67.8 | 65.7 | 62.92 | 62 | 61.9943 | 63.5795
| Std Dev. |  | 0.65727 | 1.6138 | 0 | 0.75277 | 0.74764 | 0.5164 | 2.3001 | 1.472 | 0.39833 | 0.91797 | 2.7982 | 2.4188 | 1.0114 | 0.38471 | 0.39175 | 1.3949 | 3.3583 | 3.4059 | 0.57326 | 1.8425 | 2.8654 | 1.0652 | 1.3651 | 1.3292 | 1.7446 | 0.94336
| Rel.Std.Dev. |  | 1.048% | 2.696% | 0% | 1.161% | 1.263% | 0.824% | 3.452% | 2.368% | 0.6069% | 1.472% | 4.412% | 3.772% | 1.649% | 0.6565% | 0.5919% | 2.057% | 5.277% | 5.16% | 0.9621% | 3.052% | 4.206% | 1.618% | 2.166% | 2.115% | 2.817% | 1.488%
| PDM3 |  | -0.7911% | -5.301% | 2.848% | 2.584% | -6.341% | -0.8439% | 5.432% | -1.635% | 3.85% | -1.319% | 0.3428% | 1.477% | -2.927% | -7.278% | 4.727% | 7.305% | 0.6936% | 4.43% | -5.72% | -4.483% | 7.806% | 4.167% | -0.269% | -0.5802% | -2.026% | 0.2973%
| Z-Score (Absolute) |  | 0.19 | 1.24 | 0.66 | 0.6 | 1.48 | 0.2 | 1.27 | 0.39 | 0.9 | 0.31 | 0.08 | 0.34 | 0.69 | 1.7 | 1.1 | 1.7 | 0.16 | 1.03 | 1.34 | 1.05 | 1.82 | 0.97 | 0.07 | 0.14 | 0.48 | 0.07
| Analytical results for P in OREAS 232b (Certified Value 0.067 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0635 | 0.063763 | 0.065 | 0.062 | 0.065202 | 0.069 | 0.066 | 0.065 | 0.069 | 0.068 | 0.067 | 0.068 | 0.069 | 0.0646 | 0.067121 | 0.0711 | 0.07 | 0.07 | 0.0645 | 0.069 | 0.067 | 0.0672 | 0.0639 | 0.0728 |  |  | 1
| 1 | 2 | 0.0637 | 0.063403 | 0.06 | 0.059 | 0.06554 | 0.07 | 0.0668 | 0.066 | 0.068 | 0.066 | 0.069 | 0.071 | 0.067 | 0.0653 | 0.06929 | 0.0681 | 0.07 | 0.07 | 0.065 | 0.0691 | 0.067 | 0.064 | 0.0651 | 0.0731 |  |  | #N/A
| 1 | 3 | 0.0658 | 0.063776 | 0.06 | 0.058 | 0.06551 | 0.068 | 0.0643 | 0.066 | 0.068 | 0.066 | 0.066 | 0.07 | 0.068 | 0.065 | 0.068101 | 0.0689 | 0.06 | 0.07 | 0.0648 | 0.0689 | 0.069 | 0.0682 | 0.0642 | 0.074 |  |  | 16
| 1 | 4 | 0.0652 | 0.063313 | 0.06 | 0.058 | 0.065675 | 0.069 | 0.0671 | 0.065 | 0.067 | 0.067 | 0.072 | 0.067 | 0.067 | 0.0659 | 0.066978 | 0.0694 | 0.07 | 0.07 | 0.0671 | 0.0688 | 0.067 | 0.0686 | 0.0654 | 0.0751 |  |  | 0.066666
| 1 | 5 | 0.0666 | 0.063876 | 0.06 | 0.058 | 0.065635 | 0.07 | 0.0674 | 0.066 | 0.067 | 0.066 | 0.064 | 0.07 | 0.07 | 0.065 | 0.066946 | 0.0689 | 0.07 | 0.07 | 0.0656 | 0.0689 | 0.066 | 0.0653 | 0.0624 | 0.0732 |  |  | 47
| 1 | 6 | 0.0618 | 0.063861 | 0.05 | 0.06 | 0.065584 | 0.07 | 0.0657 | 0.067 | 0.068 | 0.067 | 0.07 | 0.066 | 0.068 | 0.0653 | 0.067402 | 0.0689 | 0.07 | 0.07 | 0.0647 | 0.069 | 0.065 | 0.0678 | 0.0634 | 0.0732
| Mean |  | 0.064433 | 0.063665 | 0.059167 | 0.059167 | 0.065524 | 0.069333 | 0.066217 | 0.065833 | 0.067833 | 0.066667 | 0.068 | 0.068667 | 0.068167 | 0.065183 | 0.067639 | 0.069217 | 0.068333 | 0.07 | 0.065283 | 0.06895 | 0.066833 | 0.06685 | 0.064067 | 0.073567
| Median |  | 0.06445 | 0.06377 | 0.06 | 0.0585 | 0.065562 | 0.0695 | 0.0664 | 0.066 | 0.068 | 0.0665 | 0.068 | 0.069 | 0.068 | 0.06515 | 0.067261 | 0.0689 | 0.07 | 0.07 | 0.0649 | 0.06895 | 0.067 | 0.0675 | 0.06405 | 0.0732
| Std Dev. |  | 0.0017603 | 0.00024417 | 0.004916 | 0.0016021 | 0.00016902 | 0.0008165 | 0.0011409 | 0.00075277 | 0.00075277 | 0.0008165 | 0.0028983 | 0.0019664 | 0.001169 | 0.00043551 | 0.00091438 | 0.0010128 | 0.0040825 | 0 | 0.00096626 | 0.00010488 | 0.0013292 | 0.001813 | 0.0011057 | 0.0008501
| Rel.Std.Dev. |  | 2.732% | 0.3835% | 8.309% | 2.708% | 0.2579% | 1.178% | 1.723% | 1.143% | 1.11% | 1.225% | 4.262% | 2.864% | 1.715% | 0.6681% | 1.352% | 1.463% | 5.974% | 0% | 1.48% | 0.1521% | 1.989% | 2.712% | 1.726% | 1.156%
| PDM3 |  | -3.349% | -4.501% | -11.25% | -11.25% | -1.712% | 4.001% | -0.6736% | -1.249% | 1.751% | 0.001416% | 2.001% | 3.001% | 2.251% | -2.224% | 1.461% | 3.826% | 2.501% | 5.001% | -2.074% | 3.426% | 0.2514% | 0.2764% | -3.899% | 10.35%
| Z-Score (Absolute) |  | 1.01 | 1.33 | 3.22 | 3.22 | 0.55 | 1.04 | 0.26 | 0.42 | 0.42 | 0.07 | 0.49 | 0.77 | 0.56 | 0.7 | 0.33 | 1 | 0.63 | 1.32 | 0.65 | 0.88 | 0 | 0 | 1.16 | 2.82
| Analytical results for Pb in OREAS 232b (Certified Value 19.2 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 20 | 20 | 16 | 18.2433 | 19 | 16 | 20 | 16.8 | 19.1 | 17.8 | 19 | 19.4 | 19.77 | 21.0865 | 20.1 | 20.49 | 17.1 | 20.11 | 20.1 | 19.3 | 20.66 | 18.8 | 14 | 19.2973 | 19.0468 |  | 1
| 1 | 2 | 20.4 | 19 | 19 | 18.5687 | 19 | 16.5 | 19 | 15.6 | 18.7 | 18.2 | 19.8 | 19.2 | 19.27 | 20.828 | 20.1 | 19.26 | 17.4 | 20.69 | 19.8 | 18.3 | 20.15 | 17.6 | 15 | 21.2373 | 18.8529 |  | 35
| 1 | 3 | 20.1 | 20 | 18 | 18.6233 | 19 | 14.5 | 20 | 16.1 | 18.7 | 18.2 | 18.8 | 18.8 | 19.05 | 21.1003 | 19 | 18.77 | 17.3 | 20.95 | 18.7 | 18.8 | 20.12 | 20.3 | 14 | 21.0135 | 19.4316 |  | 16
| 1 | 4 | 20.6 | 21 | 18 | 18.6947 | 19 | 15.8 | 20 | 18.1 | 19 | 19.6 | 18.9 | 19.8 | 19.28 | 21.0315 | 18.9 | 20.14 | 17.5 | 20.47 | 19.4 | 19.8 | 19.9 | 18.8 | 14 | 22.7048 | 20.0599 |  | 19.2316
| 1 | 5 | 20.4 | 19 | 18 | 16.5953 | 20 | 16.1 | 20 | 16.5 | 18.9 | 17.6 | 19.2 | 19.4 | 18.63 | 21.0163 | 19.6 | 18.31 | 17.5 | 19.58 | 19.8 | 18.7 | 20.16 | 17.6 | 13 | 29.495 | 19.3752 |  | 48
| 1 | 6 | 20.2 | 20 | 17 | 17.4547 | 19 | 13.8 | 21 | 15.4 | 19.2 | 19.2 | 19.1 | 19.4 | 20.05 | 20.7649 | 19.7 | 19.54 | 17.2 | 20.93 | 19.9 | 17.8 | 20.21 | 18.1 | 13 | 20.1927 | 19.9164
| Mean |  | 20.2833 | 19.8333 | 17.6667 | 18.03 | 19.1667 | 15.45 | 20 | 16.4167 | 18.9333 | 18.4333 | 19.1333 | 19.3333 | 19.3417 | 20.9712 | 19.5667 | 19.4183 | 17.3333 | 20.455 | 19.6167 | 18.7833 | 20.2 | 18.5333 | 13.8333 | 22.3234 | 19.4471
| Median |  | 20.3 | 20 | 18 | 18.406 | 19 | 15.9 | 20 | 16.3 | 18.95 | 18.2 | 19.05 | 19.4 | 19.275 | 21.0239 | 19.65 | 19.4 | 17.35 | 20.58 | 19.8 | 18.75 | 20.155 | 18.45 | 14 | 21.1254 | 19.4034
| Std Dev. |  | 0.22286 | 0.75277 | 1.0328 | 0.8389 | 0.40825 | 1.0559 | 0.63246 | 0.9786 | 0.20656 | 0.79415 | 0.3559 | 0.3266 | 0.50709 | 0.1405 | 0.52026 | 0.8195 | 0.1633 | 0.53118 | 0.50365 | 0.70828 | 0.24988 | 1.0191 | 0.75277 | 3.6924 | 0.4721
| Rel.Std.Dev. |  | 1.099% | 3.795% | 5.846% | 4.653% | 2.13% | 6.835% | 3.162% | 5.961% | 1.091% | 4.308% | 1.86% | 1.689% | 2.622% | 0.67% | 2.659% | 4.22% | 0.9421% | 2.597% | 2.567% | 3.771% | 1.237% | 5.499% | 5.442% | 16.54% | 2.428%
| PDM3 |  | 5.469% | 3.129% | -8.137% | -6.248% | -0.3375% | -19.66% | 3.996% | -14.64% | -1.551% | -4.151% | -0.5108% | 0.5291% | 0.5725% | 9.046% | 1.742% | 0.9711% | -9.87% | 6.362% | 2.002% | -2.331% | 5.036% | -3.631% | -28.07% | 16.08% | 1.121%
| Z-Score (Absolute) |  | 0.8 | 0.42 | 1.4 | 1.1 | 0.14 | 3.27 | 0.56 | 2.46 | 0.34 | 0.76 | 0.17 | 0 | 0.01 | 1.38 | 0.2 | 0.07 | 1.69 | 0.95 | 0.24 | 0.46 | 0.73 | 0.67 | 4.64 | 2.52 | 0.1
| Analytical results for Pd in OREAS 232b (Indicative Value 26.7 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 1
1
| 1 | 1 | 34.3 | [25 empty] | 1
| 1 | 2 | 36.5 | [25 empty] | 1
| 1 | 3 | 33.3 | [25 empty] | 16
| 1 | 4 | 23.6 | [25 empty] | 26.6667
| 1 | 5 | 25.8 | [25 empty] | 11
| 1 | 6 | 6.5
| Mean |  | 26.6667
| Median |  | 29.55
| Std Dev. |  | 11.0979
| Rel.Std.Dev. |  | 41.62%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 232b (Certified Value 8.96 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 8.6346 | 8.85 | 9.3 | 9.3 | 9.7 | 8.005 | 9.05 | 8.8 | 7.28 | 9.1811 | [16 empty] | 1
| 1 | 2 | 8.5721 | 9.35 | 9.2 | 9.9 | 9.8 | 8.549 | 8.92 | 9 | 7.22 | 8.7962 | [16 empty] | 11
| 1 | 3 | 8.6221 | 9.05 | 8.6 | 8.5 | 10.3 | 8.633 | 8.51 | 9.2 | 7.41 | 9.01 | [16 empty] | 16
| 1 | 4 | 8.7207 | 8.85 | 8.9 | 9.5 | 10 | 8.689 | 8.96 | 9.1 | 7.49 | 9.2298 | [16 empty] | 8.9629
| 1 | 5 | 8.6838 | 9.3 | 8.8 | 8.9 | 10.1 | 8.397 | 9.05 | 9 | 7.3 | 9.1402 | [16 empty] | 49
| 1 | 6 | 8.4629 | 8.95 | 9.1 | 9.5 | 9.8 | 8.659 | 9.26 | 9.3 | 7.14 | 9.1131
| Mean |  | 8.616 | 9.0583 | 8.9833 | 9.2667 | 9.95 | 8.4887 | 8.9583 | 9.0667 | 7.3067 | 9.0784
| Median |  | 8.6283 | 9 | 9 | 9.4 | 9.9 | 8.591 | 9.005 | 9.05 | 7.29 | 9.1267
| Std Dev. |  | 0.090901 | 0.22004 | 0.26394 | 0.49666 | 0.22583 | 0.25921 | 0.24911 | 0.17512 | 0.12675 | 0.1567
| Rel.Std.Dev. |  | 1.055% | 2.429% | 2.938% | 5.36% | 2.27% | 3.054% | 2.781% | 1.931% | 1.735% | 1.726%
| PDM3 |  | -3.87% | 1.065% | 0.2285% | 3.39% | 11.01% | -5.291% | -0.05041% | 1.158% | -18.48% | 1.289%
| Z-Score (Absolute) |  | 1.77 | 0.16 | 0.16 | 1.08 | 4.06 | 2.33 | 0.27 | 0.2 | 7.5 | 0.25
| Analytical results for Pt in OREAS 232b (Indicative Value 2.87 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 2
2
| 1 | 1 | 3.3 | [25 empty] | 1
| 1 | 2 | 3.1 | [25 empty] | 1
| 1 | 3 | 2.8 | [25 empty] | 16
| 1 | 4 | 2.9 | [25 empty] | 2.8667
| 1 | 5 | 1.9 | [25 empty] | 12
| 1 | 6 | 3.2
| Mean |  | 2.8667
| Median |  | 3
| Std Dev. |  | 0.50859
| Rel.Std.Dev. |  | 17.74%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 232b (Certified Value 144 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 153.23 | 129.0898 | 149 | 123 | 148.7 | 143.3 | 138 | 146.4 | 152 | 142.5 | 136 | 157.5 | 175.9 | 142.6042 | 147 | 126.7 | 136 | 143.72 | 158.1 | 145.7 | 164.309 | 138.3821 |  |  |  |  | 1
| 1 | 2 | 154.74 | 128.6373 | 142 | 125 | 151.6 | 144.2 | 137 | 149.2 | 148.5 | 138 | 145 | 151 | 200.6 | 142.9211 | 150 | 128.9 | 142 | 136.82 | 143 | 148.7 | 158.1554 | 143.6369 |  |  |  |  | 37
| 1 | 3 | 152.46 | 132.4732 | 146 | 141 | 150.6 | 146.8 | 144 | 141.4 | 145 | 147 | 140 | 148 | 172.8 | 138.8797 | 152 | 131.4 | 136 | 130.62 | 145.8 | 151.2 | 162.5846 | 138.1713 |  |  |  |  | 16
| 1 | 4 | 152.17 | 129.9146 | 152 | 142 | 149.8 | 148 | 138 | 142.6 | 154 | 160.5 | 134.5 | 149.5 | 190 | 139.3709 | 152 | 129.6 | 152 | 143.64 | 145.7 | 148.3 | 161.9399 | 140.6165 |  |  |  |  | 144.3516
| 1 | 5 | 155.36 | 129.6583 | 146 | 139 | 151.5 | 151.8 | 143 | 140.8 | 148 | 143 | 144.5 | 153.5 | 191.6 | 140.7465 | 150 | 131 | 137 | 144.29 | 145.3 | 144.6 | 161.2373 | 139.1158 |  |  |  |  | 50
| 1 | 6 | 149.35 | 127.1127 | 144 | 134 | 149.3 | 147.1 | 154 | 146.4 | 149.5 | 154.5 | 140 | 147 | 181.4 | 140.8795 | 149 | 130.1 | 148 | 142.83 | 139.5 | 146 | 160.044 | 139.5598
| Mean |  | 152.885 | 129.481 | 146.5 | 134 | 150.25 | 146.8667 | 142.3333 | 144.4667 | 149.5 | 147.5833 | 140 | 151.0833 | 185.3833 | 140.9003 | 150 | 129.6167 | 141.8333 | 140.32 | 146.2333 | 147.4167 | 161.3784 | 139.9137
| Median |  | 152.845 | 129.374 | 146 | 136.5 | 150.2 | 146.95 | 140.5 | 144.5 | 149 | 145 | 140 | 150.25 | 185.7 | 140.813 | 150 | 129.85 | 139.5 | 143.235 | 145.5 | 147.15 | 161.5886 | 139.3378
| Std Dev. |  | 2.1389 | 1.769 | 3.5637 | 8.2462 | 1.1845 | 3.0171 | 6.4083 | 3.353 | 3.1623 | 8.4108 | 4.2778 | 3.8912 | 10.5418 | 1.6391 | 1.8974 | 1.694 | 6.824 | 5.4952 | 6.2921 | 2.4343 | 2.1235 | 2.0248
| Rel.Std.Dev. |  | 1.399% | 1.366% | 2.433% | 6.154% | 0.7883% | 2.054% | 4.502% | 2.321% | 2.115% | 5.699% | 3.056% | 2.576% | 5.686% | 1.163% | 1.265% | 1.307% | 4.811% | 3.916% | 4.303% | 1.651% | 1.316% | 1.447%
| PDM3 |  | 5.912% | -10.3% | 1.488% | -7.171% | 4.086% | 1.742% | -1.398% | 0.07971% | 3.567% | 2.239% | -3.015% | 4.663% | 28.42% | -2.391% | 3.913% | -10.21% | -1.745% | -2.793% | 1.304% | 2.123% | 11.8% | -3.074%
| Z-Score (Absolute) |  | 0.95 | 2.46 | 0.02 | 1.8 | 0.57 | 0.07 | 0.59 | 0.28 | 0.46 | 0.18 | 0.93 | 0.69 | 5.69 | 0.8 | 0.53 | 2.44 | 0.66 | 0.88 | 0.02 | 0.15 | 2.19 | 0.94
| Analytical results for Re in OREAS 232b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.006 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.008 | < 0.002 | 0.005 | 0.004 | < 0.002 | < 0.002 | 0.0009 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.006 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.007 | < 0.002 | 0.008 | 0.002 | < 0.002 | < 0.002 | 0.0007 | [10 empty] | 38
| 1 | 3 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.005 | < 0.002 | 0.003 | 0.004 | < 0.002 | < 0.002 | [11 empty] | 16
| 1 | 4 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.006 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | 0.0007 | [10 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.005 | < 0.002 | 0.003 | 0.001 | < 0.002 | < 0.002 | < 0.0002 | [10 empty] | 51
| 1 | 6 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.005 | < 0.002 | 0.006 | 0.004 | < 0.002 | < 0.002 | 0.0016
| Mean | [5 empty] | 0.0051667 |  |  |  |  | 0.006 |  | 0.005 | 0.0028333 |  |  | 0.000975
| Median | [5 empty] | 0.005 |  |  |  |  | 0.0055 |  | 0.005 | 0.003 |  |  | 0.0008
| Std Dev. | [5 empty] | 0.00075277 |  |  |  |  | 0.0012649 |  | 0.0021213 | 0.0013292 |  |  | 0.0004272
| Rel.Std.Dev. | [5 empty] | 14.57% |  |  |  |  | 21.08% |  | 42.43% | 46.91% |  |  | 43.82%
PDM3
| Z-Score (Absolute) |  | 0.5 | 32 | 0.83 | 15.42 | 2.27 | 15.42 | 0.5 | 0.5 | 0.5 | 2.82 | 0.5 | 1.71 | 0.72 | 0.5 | 0.5 | 0.63
| Analytical results for Rh in OREAS 232b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Rh | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0017 | [25 empty] | 1
| 1 | 2 | 0.0014 | [25 empty] | 13
| 1 | 3 | 0.0006 | [25 empty] | 16
| 1 | 4 | 0.0015 | [25 empty] | 0.0013
| 1 | 5 | 0.0007 | [25 empty] | 9
| 1 | 6 | 0.0019
| Mean |  | 0.0013
| Median |  | 0.00145
| Std Dev. |  | 0.00053292
| Rel.Std.Dev. |  | 40.99%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 232b (Indicative Value 0.003 ppm) | [27 empty] | Indicative
Ru | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0037 | [25 empty] | 1
| 1 | 2 | < 0.0004 | [25 empty] | 14
| 1 | 3 | 0.0069 | [25 empty] | 16
| 1 | 4 | 0.0045 | [25 empty] | 0.0028333
| 1 | 5 | 0.0013 | [25 empty] | 10
| 1 | 6 | 0.0004
| Mean |  | 0.00336
| Median |  | 0.0037
| Std Dev. |  | 0.0025958
| Rel.Std.Dev. |  | 77.25%
| PDM3 |  | 18.59%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 232b (Certified Value 0.172 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.16 | 0.17 | 0.18 | 0.16 | 0.184 | 0.14 | 0.173 | 0.16 | 0.17 | 0.18 | 0.17 | 0.18 | 0.1741 | 0.176 | 0.19 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16644 | 0.1904 | [5 empty] | 1
| 1 | 2 | 0.17 | 0.18 | 0.17 | 0.175 | 0.186 | 0.14 | 0.177 | 0.15 | 0.17 | 0.18 | 0.18 | 0.17 | 0.1741 | 0.18 | 0.19 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16093 | 0.7008 | [5 empty] | 23
| 1 | 3 | 0.17 | 0.175 | 0.17 | 0.17 | 0.183 | 0.14 | 0.18 | 0.15 | 0.17 | 0.18 | 0.17 | 0.17 | 0.1761 | 0.177 | 0.18 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16606 | 0.3445 | [5 empty] | 16
| 1 | 4 | 0.16 | 0.175 | 0.17 | 0.175 | 0.192 | 0.14 | 0.173 | 0.16 | 0.17 | 0.19 | 0.17 | 0.17 | 0.1751 | 0.176 | 0.18 | 0.16 | 0.18 | 0.18 | 0.15 | 0.17072 | 2.1715 | [5 empty] | 0.17235
| 1 | 5 | 0.16 | 0.18 | 0.17 | 0.18 | 0.192 | 0.14 | 0.182 | 0.15 | 0.17 | 0.17 | 0.18 | 0.17 | 0.171 | 0.179 | 0.19 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16188 | 0.2807 | [5 empty] | 52
| 1 | 6 | 0.16 | 0.17 | 0.18 | 0.165 | 0.185 | 0.15 | 0.189 | 0.16 | 0.17 | 0.19 | 0.17 | 0.17 | 0.1813 | 0.182 | 0.19 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16615 | 0.2547
| Mean |  | 0.16333 | 0.175 | 0.17333 | 0.17083 | 0.187 | 0.14167 | 0.179 | 0.155 | 0.17 | 0.18167 | 0.17333 | 0.17167 | 0.17528 | 0.17833 | 0.18667 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16536 | 0.6571
| Median |  | 0.16 | 0.175 | 0.17 | 0.1725 | 0.1855 | 0.14 | 0.1785 | 0.155 | 0.17 | 0.18 | 0.17 | 0.17 | 0.1746 | 0.178 | 0.19 | 0.16 | 0.18 | 0.18 | 0.15 | 0.16611 | 0.3126
| Std Dev. |  | 0.005164 | 0.0044721 | 0.005164 | 0.0073598 | 0.004 | 0.0040825 | 0.0060992 | 0.0054772 | 0 | 0.0075277 | 0.005164 | 0.0040825 | 0.0034073 | 0.0024221 | 0.005164 | 0 | 3.0405e-17 | 3.0405e-17 | 0 | 0.0035412 | 0.76347
| Rel.Std.Dev. |  | 3.162% | 2.556% | 2.979% | 4.308% | 2.139% | 2.882% | 3.407% | 3.534% | 0% | 4.144% | 2.979% | 2.378% | 1.944% | 1.358% | 2.766% | 0% | 1.689e-14% | 1.689e-14% | 0% | 2.141% | 116.2%
| PDM3 |  | -5.23% | 1.539% | 0.5718% | -0.8788% | 8.501% | -17.8% | 3.86% | -10.07% | -1.362% | 5.407% | 0.5718% | -0.3953% | 1.703% | 3.473% | 8.308% | -7.165% | 4.44% | 4.44% | -12.97% | -4.053% | 281.3%
| Z-Score (Absolute) |  | 1.01 | 0.17 | 0 | 0.25 | 1.38 | 3.2 | 0.57 | 1.85 | 0.34 | 0.84 | 0 | 0.17 | 0.2 | 0.51 | 1.35 | 1.35 | 0.67 | 0.67 | 2.36 | 0.81 | 48.93
| Analytical results for Sb in OREAS 232b (Certified Value 187 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 201.35 | 7.3 | 182.43 | 191 | 185.23 | 168 | 188.9 | 196 | 207 | 182 | 185.5 | 181.35 | 192.6268 | 155.4 | 196 | 177.9 | 167.66 | 180.5 | 181.81 | 200.09 | 202 | 329.297 | 188.5349 |  |  |  | 1
| 1 | 2 | 200.94 | 12.1 | 186.9333 | 196.8 | 192.6 | 167 | 183.4 | 186.5 | 213 | 199 | 185 | 185.55 | 186.3589 | 155.1 | 206 | 178.1 | 164.09 | 180.3 | 186.37 | 193.242 | 197 | 357.9202 | 190.0737 |  |  |  | 39
| 1 | 3 | 201.38 | 31.3 | 184.935 | 197.5 | 188.62 | 174 | 180.9 | 185 | 205 | 186.5 | 184 | 179.5 | 180.5112 | 164.8 | 189 | 177.6 | 163.93 | 185.3 | 188.26 | 199.876 | 199 | 370.943 | 184.8359 |  |  |  | 16
| 1 | 4 | 203.05 | 31.6 | 188.3967 | 192.6 | 185.47 | 168 | 183.7 | 184 | 216 | 186.5 | 182 | 181.19 | 185.9907 | 167.1 | 219 | 182.3 | 170.26 | 184.8 | 184.66 | 206.403 | 201 | 369.0689 | 190.404 |  |  |  | 186.6604
| 1 | 5 | 201.06 | 22.6 | 187.72 | 196.6 | 172.19 | 175 | 179.6 | 188.5 | 200 | 195 | 187 | 184.73 | 184.2384 | 167.9 | 200 | 179.8 | 171.27 | 183.5 | 179.78 | 195.275 | 193 | 370.9878 | 188.7761 |  |  |  | 53
| 1 | 6 | 199.24 | 9.7 | 182.72 | 191.7 | 192.42 | 167 | 177.3 | 185 | 224 | 178 | 188 | 176.85 | 176.5136 | 166 | 205 | 182.1 | 178.28 | 186.1 | 187.31 | 199.341 | 198 | 373.8277 | 190.8545
| Mean |  | 201.17 | 19.1 | 185.5225 | 194.3667 | 186.0883 | 169.8333 | 182.3 | 187.5 | 210.8333 | 187.8333 | 185.25 | 181.5283 | 184.3733 | 162.7167 | 202.5 | 179.6333 | 169.2483 | 183.4167 | 184.6983 | 199.0378 | 198.3333 | 362.0074 | 188.9132
| Median |  | 201.205 | 17.35 | 185.9342 | 194.6 | 187.045 | 168 | 182.15 | 185.75 | 210 | 186.5 | 185.25 | 181.27 | 185.1146 | 165.4 | 202.5 | 178.95 | 168.96 | 184.15 | 185.515 | 199.6085 | 198.5 | 370.0059 | 189.4249
| Std Dev. |  | 1.2166 | 10.899 | 2.5629 | 2.9084 | 7.5257 | 3.656 | 4.0244 | 4.4497 | 8.612 | 7.8782 | 2.1389 | 3.2426 | 5.5025 | 5.8779 | 10.2127 | 2.1314 | 5.3684 | 2.4855 | 3.3116 | 4.5585 | 3.2042 | 16.9514 | 2.1966
| Rel.Std.Dev. |  | 0.6048% | 57.06% | 1.381% | 1.496% | 4.044% | 2.153% | 2.208% | 2.373% | 4.085% | 4.194% | 1.155% | 1.786% | 2.984% | 3.612% | 5.043% | 1.187% | 3.172% | 1.355% | 1.793% | 2.29% | 1.616% | 4.683% | 1.163%
| PDM3 |  | 7.773% | -89.77% | -0.6096% | 4.129% | -0.3065% | -9.015% | -2.336% | 0.4498% | 12.95% | 0.6284% | -0.7556% | -2.749% | -1.225% | -12.83% | 8.486% | -3.765% | -9.328% | -1.738% | -1.051% | 6.631% | 6.254% | 93.94% | 1.207%
| Z-Score (Absolute) |  | 1.79 | 19.06 | 0 | 1.01 | 0.06 | 1.8 | 0.37 | 0.23 | 2.9 | 0.26 | 0.03 | 0.46 | 0.13 | 2.61 | 1.94 | 0.67 | 1.86 | 0.24 | 0.09 | 1.55 | 1.47 | 20.21 | 0.39
| Analytical results for Sc in OREAS 232b (Certified Value 13.6 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 13.6 | 12.673 | 14 | 14 | 10.758 | 13 | 9 | 13 | 14 | 14.8 | 14.2 | 13.4 | 14.2 | 11.8 | 13.5597 | 13.4 | 11.3 | 12.4 | 13.4 | 14.1 | 13.3 | 13.8915 | 15.2559 | 13.202 |  |  | 1
| 1 | 2 | 13.7 | 12.6988 | 14 | 14 | 10.306 | 13 | 9 | 13 | 13.6 | 13.9 | 14.4 | 14.4 | 13.6 | 11.7 | 13.3749 | 13.6 | 11.79 | 13.1 | 13 | 13.2 | 13.4 | 13.4715 | 14.8559 | 13.0126 |  |  | 40
| 1 | 3 | 13.6 | 12.6164 | 15 | 13 | 10.712 | 13 | 10 | 13 | 13.4 | 13.9 | 15 | 13.6 | 13.9 | 11.9 | 13.5617 | 14.5 | 11.83 | 11.6 | 12.1 | 13.2 | 13.8 | 14.0385 | 14.1416 | 13.0061 |  |  | 16
| 1 | 4 | 13.6 | 12.6647 | 14 | 14 | 10.1213 | 14 | 9 | 13 | 13.1 | 14.7 | 15.2 | 13.7 | 13.8 | 12.2 | 13.301 | 14.4 | 13.6 | 12.7 | 13.1 | 13.9 | 13.5 | 14.1855 | 14.7416 | 13.1863 |  |  | 13.624
| 1 | 5 | 13.6 | 12.6783 | 15 | 13 | 11.064 | 14 | 11 | 13 | 13.2 | 14 | 14.1 | 14.2 | 13.8 | 12.1 | 13.3458 | 14.1 | 12 | 12.3 | 13.3 | 13.6 | 13.3 | 13.734 | 14.0274 | 13.0046 |  |  | 54
| 1 | 6 | 13.4 | 12.757 | 14 | 14 | 11.5827 | 13 | 11 | 13 | 13.6 | 14.3 | 15 | 13.5 | 13.7 | 11.8 | 13.3797 | 13.5 | 12.2 | 13.1 | 13.4 | 12.8 | 13.5 | 13.881 | 14.0274 | 13.0729
| Mean |  | 13.5833 | 12.6814 | 14.3333 | 13.6667 | 10.7573 | 13.3333 | 9.8333 | 13 | 13.4833 | 14.2667 | 14.65 | 13.8 | 13.8333 | 11.9167 | 13.4204 | 13.9167 | 12.12 | 12.5333 | 13.05 | 13.4667 | 13.4667 | 13.867 | 14.5083 | 13.0808
| Median |  | 13.6 | 12.6757 | 14 | 14 | 10.735 | 13 | 9.5 | 13 | 13.5 | 14.15 | 14.7 | 13.65 | 13.8 | 11.85 | 13.3773 | 13.85 | 11.915 | 12.55 | 13.2 | 13.4 | 13.45 | 13.8863 | 14.4416 | 13.0428
| Std Dev. |  | 0.098319 | 0.046056 | 0.5164 | 0.5164 | 0.5264 | 0.5164 | 0.98319 | 0 | 0.32506 | 0.40332 | 0.47223 | 0.40497 | 0.20656 | 0.19408 | 0.11219 | 0.47924 | 0.78442 | 0.56804 | 0.49295 | 0.48854 | 0.18619 | 0.24729 | 0.51598 | 0.091588
| Rel.Std.Dev. |  | 0.7238% | 0.3632% | 3.603% | 3.779% | 4.893% | 3.873% | 9.999% | 0% | 2.411% | 2.827% | 3.223% | 2.935% | 1.493% | 1.629% | 0.8359% | 3.444% | 6.472% | 4.532% | 3.777% | 3.628% | 1.383% | 1.783% | 3.556% | 0.7002%
| PDM3 |  | -0.2987% | -6.919% | 5.206% | 0.3129% | -21.04% | -2.134% | -27.82% | -4.58% | -1.033% | 4.717% | 7.531% | 1.292% | 1.536% | -12.53% | -1.494% | 2.148% | -11.04% | -8.006% | -4.213% | -1.155% | -1.155% | 1.783% | 6.49% | -3.988%
| Z-Score (Absolute) |  | 0.19 | 1.27 | NA | NA | 4.38 | NA | NA | NA | 0.03 | 1.29 | 1.92 | 0.54 | 0.59 | 2.51 | 0.07 | 0.73 | 2.18 | 1.51 | 0.67 | 0 | 0 | 0.65 | 1.69 | 0.62
Results from laboratories 1.05, 1.06, 1.08, 1.09 and 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 232b (Indicative Value 0.67 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | 1 | 1 | < 1 | < 1 | 1 | < 0.1 | 5 | 0.6 | < 1 | < 0.3 | 0.2069 | 3.4154 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | < 1 | 1 | < 1 | < 1 | 1 | < 0.1 | 5 | < 0.5 | < 1 | < 0.3 | 0.239 | 3.0171 | [7 empty] | 41
| 1 | 3 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | 5 | 1 | < 1 | < 0.3 | 0.2307 | 3.9465 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 1 | < 0.1 | 5 | < 0.5 | < 1 | < 0.3 | 0.2247 | 3.3068 | [7 empty] | 0.66521
| 1 | 5 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 1 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | 5 | 1 | < 1 | 0.4 | 0.1819 | 3.3912 | [7 empty] | 11
| 1 | 6 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | 5 | 0.6 | < 1 | 0.5 | 0.2557 | 3.1257
| Mean | [7 empty] | 1.6667 | 1 | 1 |  |  | 0.95 |  | 5 | 0.8 |  | 0.45 | 0.22315 | 3.3671
| Median | [7 empty] | 2 | 1 | 1 |  |  | 0.95 |  | 5 | 0.8 |  | 0.45 | 0.2277 | 3.349
| Std Dev. | [7 empty] | 0.57735 |  | 0 |  |  | 0.054772 |  | 0 | 0.23094 |  | 0.070711 | 0.025829 | 0.32343
| Rel.Std.Dev. | [7 empty] | 34.64% |  | 0% |  |  | 5.766% |  | 0% | 28.87% |  | 15.71% | 11.57% | 9.605%
| PDM3 | [7 empty] | 150.5% | 50.33% | 50.33% |  |  | 42.81% |  | 651.6% | 20.26% |  | -32.35% | -66.45% | 406.2%
| Z-Score (Absolute) |  | 0.61 | 3.52 | 0.98 | 3.52 | 0.77 | 0.15 | 0.92 | 0 | 0.77 | 0.15 | 0.15 | 0.67 | 0.98 | 8.12 | 0.06 | 0.15 | 0.61 | 0.66 | 5.12
| Analytical results for Sm in OREAS 232b (Certified Value 6.44 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.2098 | 6.4 | 6.8 | 5.6 | 6.8 | 5.389 | 7.05 | 6.2 | 7.1534 | 6.7111 | [16 empty] | 1
| 1 | 2 | 6.4432 | 6.55 | 6 | 6.1 | 6.9 | 5.879 | 6.38 | 6.3 | 7.1722 | 6.3914 | [16 empty] | 10
| 1 | 3 | 6.2923 | 6.7 | 5.8 | 5.7 | 7.2 | 5.881 | 6.03 | 6.4 | 7.285 | 6.6801 | [16 empty] | 16
| 1 | 4 | 6.2042 | 6.75 | 6.9 | 5.7 | 6.9 | 5.978 | 6.2 | 6.2 | 7.3414 | 6.6589 | [16 empty] | 6.438
| 1 | 5 | 6.2631 | 6.85 | 6 | 6 | 7 | 5.854 | 6.3 | 6.3 | 7.1628 | 6.8835 | [16 empty] | 55
| 1 | 6 | 5.9395 | 6.55 | 6.7 | 6.2 | 6.9 | 5.96 | 6.47 | 6.2 | 7.1158 | 6.6514
| Mean |  | 6.2253 | 6.6333 | 6.3667 | 5.8833 | 6.95 | 5.8235 | 6.405 | 6.2667 | 7.2051 | 6.6627
| Median |  | 6.2364 | 6.625 | 6.35 | 5.85 | 6.9 | 5.88 | 6.34 | 6.25 | 7.1675 | 6.6695
| Std Dev. |  | 0.16478 | 0.1633 | 0.48442 | 0.24833 | 0.13784 | 0.21845 | 0.35059 | 0.08165 | 0.087728 | 0.15823
| Rel.Std.Dev. |  | 2.647% | 2.462% | 7.609% | 4.221% | 1.983% | 3.751% | 5.474% | 1.303% | 1.218% | 2.375%
| PDM3 |  | -3.303% | 3.035% | -1.107% | -8.615% | 7.953% | -9.544% | -0.5119% | -2.661% | 11.92% | 3.491%
| Z-Score (Absolute) |  | 0.41 | 0.64 | 0.05 | 1.29 | 1.45 | 1.45 | 0.05 | 0.31 | 2.11 | 0.71
| Analytical results for Sn in OREAS 232b (Certified Value 3.44 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.7 | 3.7 | 2 | < 10 | 3.6 | 3.5 | 2.7 | 3.5 | 3.3 | 3.5 | 3.3 | 3.3 | 3.4 | 3.4829 | 3.48 | 3.204 | 4 | < 5 | 3.4 | 3.7 | 3.2 | 2.8129 | < 5 | 3.6302 | 3.5631 |  | 1
| 1 | 2 | 3.6 | 3.4 | 1 | < 10 | 3.7 | 3.3 | 2.6 | 3.7 | 3.1 | 3.4 | 3.7 | 3.3 | 3.5 | 3.3534 | 3.69 | 3.158 | 4 | < 5 | 3.2 | 3.4 | 3.3 | 2.7032 | < 5 | 3.7501 | 3.6034 |  | 26
| 1 | 3 | 3.6 | 3.3 | 2 | < 10 | 3.5 | 3.2 | 3.4 | 3.6 | 3.1 | 3.6 | 3.4 | 3.2 | 3.2 | 3.4726 | 3.59 | 3.233 | 3.9 | < 5 | 3.1 | 3.4 | 3.3 | 2.8759 | < 5 | 3.3499 | 3.571 |  | 16
| 1 | 4 | 3.6 | 3 | 2 | < 10 | 3.5 | 3.3 | 4.6 | 3.7 | 3.2 | 3.8 | 3.4 | 3.2 | 3.6 | 3.5045 | 3.48 | 3.238 | 4 | < 5 | 3.4 | 3.5 | 3.2 | 2.8532 | < 5 | 3.7945 | 3.5966 |  | 3.436
| 1 | 5 | 3.5 | 3.1 | 2 | < 10 | 3.5 | 3.4 | 3.1 | 3.7 | 3.1 | 3.5 | 3.6 | 3.3 | 3.4 | 3.4835 | 3.55 | 3.153 | 4 | < 5 | 3.2 | 3.5 | 3.1 | 2.7263 | < 5 | 3.8799 | 3.5304 |  | 56
| 1 | 6 | 3.5 | 3.7 | 2 | < 10 | 3.5 | 3.7 | 4 | 3.9 | 3.2 | 3.7 | 3.4 | 3.1 | 3.5 | 3.4027 | 3.75 | 3.255 | 4 | < 5 | 3.6 | 3.3 | 3.1 | 2.7702 | < 5 | 3.6801 | 3.6557
| Mean |  | 3.5833 | 3.3667 | 1.8333 |  | 3.55 | 3.4 | 3.4 | 3.6833 | 3.1667 | 3.5833 | 3.4667 | 3.2333 | 3.4333 | 3.4499 | 3.59 | 3.2068 | 3.9833 |  | 3.3167 | 3.4667 | 3.2 | 2.7903 |  | 3.6808 | 3.5867
| Median |  | 3.6 | 3.35 | 2 |  | 3.5 | 3.35 | 3.25 | 3.7 | 3.15 | 3.55 | 3.4 | 3.25 | 3.45 | 3.4777 | 3.57 | 3.2185 | 4 |  | 3.3 | 3.45 | 3.2 | 2.7915 |  | 3.7151 | 3.5838
| Std Dev. |  | 0.075277 | 0.29439 | 0.40825 |  | 0.083666 | 0.17889 | 0.77717 | 0.13292 | 0.08165 | 0.1472 | 0.15055 | 0.08165 | 0.13663 | 0.058729 | 0.11082 | 0.043051 | 0.040825 |  | 0.18348 | 0.13663 | 0.089443 | 0.06911 |  | 0.18408 | 0.042712
| Rel.Std.Dev. |  | 2.101% | 8.744% | 22.27% |  | 2.357% | 5.261% | 22.86% | 3.609% | 2.578% | 4.108% | 4.343% | 2.525% | 3.979% | 1.702% | 3.087% | 1.342% | 1.025% |  | 5.532% | 3.941% | 2.795% | 2.477% |  | 5.001% | 1.191%
| PDM3 |  | 4.289% | -2.017% | -46.64% |  | 3.319% | -1.047% | -1.047% | 7.199% | -7.838% | 4.289% | 0.8932% | -5.898% | -0.07691% | 0.4058% | 4.483% | -6.669% | 15.93% |  | -3.472% | 0.8932% | -6.868% | -18.79% |  | 7.125% | 4.387%
| Z-Score (Absolute) |  | 0.65 | 0.34 | NA | 7.16 | 0.5 | 0.19 | 0.19 | 1.11 | 1.26 | 0.65 | 0.12 | 0.96 | 0.04 | 0.04 | 0.68 | 1.08 | 2.49 | 4.33 | 0.57 | 0.12 | 1.11 | 2.99 | 4.33 | 1.1 | 0.67
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 232b (Certified Value 134 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 137.29 | 137.8762 | 137 | 120 | 137.6778 | 136.1 | 122.1 | 131 | 134 | 138 | 132.5 | 134 | 136.5 | 131 | 131.1119 | 135 | 117.7 | 128 | 130.3 | 129.3 | 135.9 | 131 | 158.2 | 135 | 134.9697 | 127.5935 | 1
| 1 | 2 | 136.24 | 139.4344 | 130 | 121 | 135.6314 | 135.5 | 118.1 | 129 | 135 | 132.5 | 136 | 142 | 134 | 130 | 129.6985 | 132 | 119.3 | 129 | 130.8 | 124.9 | 136.5 | 132 | 152.9 | 139 | 150.8929 | 130.5729 | 27
| 1 | 3 | 137.79 | 139.1012 | 133 | 118 | 137.7794 | 136.1 | 117.1 | 140 | 134 | 132 | 134.5 | 135 | 133.5 | 131 | 127.5451 | 134 | 121.8 | 127 | 131.2 | 121.5 | 135.4 | 137 | 160.9 | 137 | 137.5941 | 127.8791 | 16
| 1 | 4 | 135.97 | 137.4744 | 131 | 119 | 136.7554 | 136.2 | 127.1 | 135 | 130 | 135.5 | 145 | 133 | 132 | 134 | 128.8955 | 136 | 120.7 | 134 | 130.8 | 128.6 | 135.5 | 134 | 161.6 | 140 | 141.1775 | 128.9097 | 133.574
| 1 | 5 | 138.15 | 139.8166 | 136 | 118 | 136.5676 | 136.6 | 127.5 | 136 | 131 | 133.5 | 130 | 139.5 | 133 | 135 | 132.844 | 131 | 121.5 | 134 | 129 | 129.1 | 135.1 | 131 | 157.5 | 135 | 140.5214 | 127.9607 | 57
| 1 | 6 | 136.49 | 139.7382 | 130 | 126 | 138.3468 | 135.5 | 125 | 146 | 132 | 134 | 141.5 | 132 | 132.5 | 132 | 130.6065 | 138 | 121.5 | 130 | 130.6 | 132.5 | 133.6 | 131 | 157.1 | 136 | 135.3734 | 128.9557
| Mean |  | 136.9883 | 138.9068 | 132.8333 | 120.3333 | 137.1264 | 136 | 122.8167 | 136.1667 | 132.6667 | 134.25 | 136.5833 | 135.9167 | 133.5833 | 132.1667 | 130.1169 | 134.3333 | 120.4167 | 130.3333 | 130.45 | 127.65 | 135.3333 | 132.6667 | 158.0333 | 137 | 140.0882 | 128.6453
| Median |  | 136.89 | 139.2678 | 132 | 119.5 | 137.2166 | 136.1 | 123.55 | 135.5 | 133 | 133.75 | 135.25 | 134.5 | 133.25 | 131.5 | 130.1525 | 134.5 | 121.1 | 129.5 | 130.7 | 128.85 | 135.45 | 131.5 | 157.85 | 136.5 | 139.0577 | 128.4352
| Std Dev. |  | 0.8864 | 0.99489 | 3.0605 | 3.0111 | 0.99043 | 0.42895 | 4.483 | 6.1779 | 1.9664 | 2.2079 | 5.6517 | 3.955 | 1.5943 | 1.9408 | 1.8411 | 2.582 | 1.6105 | 3.0111 | 0.76877 | 3.8635 | 0.97707 | 2.4221 | 3.1136 | 2.0976 | 5.8782 | 1.0994
| Rel.Std.Dev. |  | 0.6471% | 0.7162% | 2.304% | 2.502% | 0.7223% | 0.3154% | 3.65% | 4.537% | 1.482% | 1.645% | 4.138% | 2.91% | 1.193% | 1.468% | 1.415% | 1.922% | 1.337% | 2.31% | 0.5893% | 3.027% | 0.722% | 1.826% | 1.97% | 1.531% | 4.196% | 0.8546%
| PDM3 |  | 2.556% | 3.992% | -0.5545% | -9.913% | 2.659% | 1.816% | -8.053% | 1.941% | -0.6793% | 0.5061% | 2.253% | 1.754% | 0.006967% | -1.054% | -2.588% | 0.5685% | -9.85% | -2.426% | -2.339% | -4.435% | 1.317% | -0.6793% | 18.31% | 2.565% | 4.877% | -3.69%
| Z-Score (Absolute) |  | 0.67 | 1.09 | 0.24 | 2.98 | 0.7 | 0.46 | 2.43 | 0.49 | 0.27 | 0.07 | 0.58 | 0.44 | 0.07 | 0.38 | 0.83 | 0.09 | 2.96 | 0.79 | 0.76 | 1.37 | 0.31 | 0.27 | 5.28 | 0.68 | 1.35 | 1.16
| Analytical results for Ta in OREAS 232b (Certified Value 1.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.1 | 1.0959 | < 0.1 | < 5 | 1.02 | 0.79 | 1.06 | 1 | 1.01 | 1.01 | 0.95 | 1 | 1.1 | 1.4917 | 1 | 0.92 | 1.15 | 1.28 | 0.8 | 0.7925 | 1.202 | [5 empty] | 1
| 1 | 2 | 1.04 | 0.98875 | < 0.1 | < 5 | 1 | 0.84 | 0.91 | 1.1 | 1 | 0.99 | 1 | 0.97 | 1 | 1.244 | 1 | 1 | 1.13 | 1.07 | 0.8 | 0.7349 | 1.1811 | [5 empty] | 28
| 1 | 3 | 1.04 | 1.0902 | < 0.1 | < 5 | 0.96 | 0.85 | 0.87 | 1 | 1 | 1 | 1.04 | 0.97 | 1 | 1.1768 | 1.03 | 0.95 | 1.09 | 1.08 | 0.9 | 0.7262 | 1.1668 | [5 empty] | 16
| 1 | 4 | 1 | 1.0158 | < 0.1 | < 5 | 0.96 | 0.74 | 0.87 | 1 | 1.04 | 1.06 | 1.02 | 0.95 | 1 | 1.3609 | 1.02 | 0.98 | 1.12 | 1.08 | 0.8 | 0.731 | 1.2078 | [5 empty] | 1.0078
| 1 | 5 | 1.07 | 0.93086 | < 0.1 | < 5 | 0.93 | 0.76 | 0.92 | 0.9 | 0.98 | 0.94 | 1 | 0.97 | 1 | 1.3654 | 1.02 | 0.94 | 1.11 | 1.08 | 0.8 | 0.7742 | 1.2215 | [5 empty] | 58
| 1 | 6 | 1.08 | 0.99943 | < 0.1 | < 5 | 0.92 | 0.96 | 0.95 | 0.9 | 1.01 | 1.06 | 0.96 | 0.97 | 1 | 1.18 | 1 | 0.94 | 1.16 | 1.03 | 0.8 | 0.6858 | 1.207
| Mean |  | 1.055 | 1.0202 |  |  | 0.965 | 0.82333 | 0.93 | 0.98333 | 1.0067 | 1.01 | 0.995 | 0.97167 | 1.0167 | 1.3031 | 1.0117 | 0.955 | 1.1267 | 1.1033 | 0.81667 | 0.74077 | 1.1977
| Median |  | 1.055 | 1.0076 |  |  | 0.96 | 0.815 | 0.915 | 1 | 1.005 | 1.005 | 1 | 0.97 | 1 | 1.3024 | 1.01 | 0.945 | 1.125 | 1.08 | 0.8 | 0.73295 | 1.2045
| Std Dev. |  | 0.035637 | 0.063337 |  |  | 0.038859 | 0.079666 | 0.070711 | 0.075277 | 0.019664 | 0.045607 | 0.034496 | 0.016021 | 0.040825 | 0.12442 | 0.013292 | 0.029496 | 0.02582 | 0.088694 | 0.040825 | 0.037835 | 0.020031
| Rel.Std.Dev. |  | 3.378% | 6.208% |  |  | 4.027% | 9.676% | 7.603% | 7.655% | 1.953% | 4.516% | 3.467% | 1.649% | 4.016% | 9.548% | 1.314% | 3.089% | 2.292% | 8.039% | 4.999% | 5.108% | 1.672%
| PDM3 |  | 4.684% | 1.228% |  |  | -4.247% | -18.3% | -7.72% | -2.428% | -0.1122% | 0.2185% | -1.27% | -3.585% | 0.88% | 29.3% | 0.3839% | -5.239% | 11.79% | 9.48% | -18.97% | -26.5% | 18.84%
| Z-Score (Absolute) |  | 0.63 | 0.18 | 12.49 | 19.49 | 0.54 | 2.39 | 1 | 0.3 | 0 | 0.04 | 0.15 | 0.46 | 0.13 | 3.87 | 0.07 | 0.67 | 1.57 | 1.26 | 2.48 | 3.47 | 2.49
| Analytical results for Tb in OREAS 232b (Certified Value 0.67 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.69622 | 0.66 | 0.7 | 0.7 | 0.64 | 0.592 | 0.68 | 0.67 | 0.6 | 2.08 | 5.8581 | [15 empty] | 1
| 1 | 2 | 0.68607 | 0.66 | 0.7 | 0.7 | 0.66 | 0.642 | 0.7 | 0.71 | 0.6 | 1.68 | 5.612 | [15 empty] | 11
| 1 | 3 | 0.68358 | 0.7 | 0.7 | 0.6 | 0.67 | 0.67 | 0.7 | 1.08 | 0.6 | 2.34 | 5.7481 | [15 empty] | 16
| 1 | 4 | 0.70496 | 0.66 | 0.7 | 0.6 | 0.65 | 0.646 | 0.71 | 0.68 | 0.6 | 1.75 | 5.9809 | [15 empty] | 0.67386
| 1 | 5 | 0.68954 | 0.7 | 0.7 | 0.6 | 0.67 | 0.622 | 0.72 | 0.68 | 0.6 | 1.87 | 5.9773 | [15 empty] | 59
| 1 | 6 | 0.6726 | 0.66 | 0.7 | 0.7 | 0.67 | 0.644 | 0.69 | 1.08 | 0.6 | 2.38 | 5.9108
| Mean |  | 0.68883 | 0.67333 | 0.7 | 0.65 | 0.66 | 0.636 | 0.7 | 0.81667 | 0.6 | 2.0167 | 5.8479
| Median |  | 0.6878 | 0.66 | 0.7 | 0.65 | 0.665 | 0.643 | 0.7 | 0.695 | 0.6 | 1.975 | 5.8844
| Std Dev. |  | 0.011081 | 0.020656 | 1.2162e-16 | 0.054772 | 0.012649 | 0.026412 | 0.014142 | 0.20442 | 0 | 0.29884 | 0.14428
| Rel.Std.Dev. |  | 1.609% | 3.068% | 1.737e-14% | 8.427% | 1.917% | 4.153% | 2.02% | 25.03% | 0% | 14.82% | 2.467%
| PDM3 |  | 2.221% | -0.0782% | 3.879% | -3.541% | -2.057% | -5.618% | 3.879% | 21.19% | -10.96% | 199.3% | 767.8%
| Z-Score (Absolute) |  | 0.08 | 0.31 | NA | NA | 0.5 | 0.85 | 0.08 | 1.78 | NA | 19.21 | 74.87
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 232b (Indicative Value 0.06 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | < 0.1 | 9.8027 | < 0.1 | 0.11 | 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.03091 | [8 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 10.0787 | < 0.1 | 0.09 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.02695 | [8 empty] | 6
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 9.288 | < 0.1 | 0.09 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.01507 | [8 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 9.2967 | < 0.1 | 0.08 | 0.4 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.08 | < 0.05 | 0.01496 | [8 empty] | 0.060429
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 8.8553 | < 0.1 | 0.08 | 0.3 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.02992 | [8 empty] | 12
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 7.9487 | < 0.1 | 0.1 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.01507
| Mean |  |  |  |  | 9.2117 |  | 0.091667 | 0.25 | 0.1 |  |  |  | 0.055 |  |  |  | 0.066667 |  | 0.022147
| Median |  |  |  |  | 9.2923 |  | 0.09 | 0.25 | 0.1 |  |  |  | 0.055 |  |  |  | 0.065 |  | 0.02101
| Std Dev. |  |  |  |  | 0.75315 |  | 0.01169 | 0.1291 | 0 |  |  |  | 0.0054772 |  |  |  | 0.008165 |  | 0.0079006
| Rel.Std.Dev. |  |  |  |  | 8.176% |  | 12.75% | 51.64% | 0% |  |  |  | 9.959% |  |  |  | 12.25% |  | 35.67%
| PDM3 |  |  |  |  | 15144% |  | 51.69% | 313.7% | 65.48% |  |  |  | -8.985% |  |  |  | 10.32% |  | -63.35%
| Z-Score (Absolute) |  | 1.35 | 1.35 | 0 | 247.11 | 0 | 1.12 | 3.6 | 0.45 | 0.67 | 0.67 | 0.67 | 0.13 | 0.67 | 0.67 | 0.94 | 0.45 | 0.67 | 0.75
| Analytical results for Th in OREAS 232b (Certified Value 14.3 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.8 | 14.9832 | 16.3 | 15.9 | 14.01 | 13 | 14.9 | 14.7 | 14.2 | 13.35 | 13.6 | 15.1 | 15.6 | 13.981 | 13.3 | 12.22 | 11.6 | 13.5 | 16.7 | 13.9 | 13.7776 | 14.0816 | 14.1731 |  |  |  | 1
| 1 | 2 | 14.8 | 15.0428 | 15.9 | 15.5 | 14.39 | 15 | 14.5 | 14.6 | 14.25 | 13.45 | 14.85 | 14.85 | 17.6 | 14.5329 | 13.4 | 13.05 | 11.8 | 13.2 | 15.2 | 13.9 | 13.5246 | 14.0816 | 13.5926 |  |  |  | 31
| 1 | 3 | 14.16 | 15.312 | 15.8 | 15 | 14.2 | 12 | 14.5 | 14.2 | 13.95 | 13.1 | 13.8 | 14.5 | 15.2 | 14.0478 | 13.4 | 13.39 | 11.5 | 12.6 | 15.5 | 14.1 | 13.8412 | 14.3772 | 14.1208 |  |  |  | 16
| 1 | 4 | 14.39 | 15.1798 | 16.9 | 15.5 | 14.34 | 12 | 14.3 | 14 | 14.7 | 14.95 | 13.4 | 14.5 | 16.2 | 14.0649 | 13.4 | 13.53 | 11.8 | 13.2 | 15.6 | 13.9 | 13.819 | 14.7319 | 14.2466 |  |  |  | 14.2812
| 1 | 5 | 14.64 | 14.9865 | 16.8 | 15.1 | 14.29 | 13 | 14.8 | 14.3 | 14.25 | 12.8 | 14.05 | 14.95 | 15 | 14.5291 | 13.4 | 13.1 | 11.8 | 13 | 15.4 | 13.9 | 13.6744 | 14.1407 | 14.1803 |  |  |  | 60
| 1 | 6 | 14.55 | 15.2236 | 17.1 | 15.9 | 13.97 | 16 | 15.5 | 14.4 | 14.75 | 14.4 | 13.65 | 14.8 | 17.2 | 14.1959 | 13.4 | 13.04 | 11.9 | 13.6 | 14.9 | 13.8 | 13.552 | 14.2196 | 14.1569
| Mean |  | 14.5567 | 15.1213 | 16.4667 | 15.4833 | 14.2 | 13.5 | 14.75 | 14.3667 | 14.35 | 13.675 | 13.8917 | 14.7833 | 16.1333 | 14.2253 | 13.3833 | 13.055 | 11.7333 | 13.1833 | 15.55 | 13.9167 | 13.6981 | 14.2721 | 14.0784
| Median |  | 14.595 | 15.1113 | 16.55 | 15.5 | 14.245 | 13 | 14.65 | 14.35 | 14.25 | 13.4 | 13.725 | 14.825 | 15.9 | 14.1304 | 13.4 | 13.075 | 11.8 | 13.2 | 15.45 | 13.9 | 13.726 | 14.1802 | 14.165
| Std Dev. |  | 0.24921 | 0.13686 | 0.5465 | 0.38166 | 0.17481 | 1.6432 | 0.42778 | 0.2582 | 0.31145 | 0.82508 | 0.51712 | 0.24221 | 1.0708 | 0.24685 | 0.040825 | 0.45549 | 0.15055 | 0.36009 | 0.61563 | 0.098319 | 0.13669 | 0.25102 | 0.24151
| Rel.Std.Dev. |  | 1.712% | 0.9051% | 3.319% | 2.465% | 1.231% | 12.17% | 2.9% | 1.797% | 2.17% | 6.033% | 3.723% | 1.638% | 6.637% | 1.735% | 0.305% | 3.489% | 1.283% | 2.731% | 3.959% | 0.7065% | 0.9978% | 1.759% | 1.715%
| PDM3 |  | 1.929% | 5.883% | 15.3% | 8.418% | -0.5685% | -5.47% | 3.283% | 0.5986% | 0.4819% | -4.245% | -2.727% | 3.516% | 12.97% | -0.3916% | -6.287% | -8.586% | -17.84% | -7.687% | 8.885% | -2.552% | -4.083% | -0.06361% | -1.42%
| Z-Score (Absolute) |  | 0.38 | 1.08 | 2.76 | 1.53 | 0.06 | NA | 0.62 | 0.15 | 0.13 | 0.71 | 0.44 | 0.66 | 2.34 | 0.03 | 1.08 | 1.48 | 3.13 | 1.32 | 1.62 | 0.41 | 0.68 | 0.03 | 0.21
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 232b (Certified Value 0.443 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.4444 | 0.42328 | 0.42 | 0.371 | 0.355 | 0.4892 | 0.433 | 0.442 | 0.464 | 0.449 | 0.43 | 0.438 | 0.45 | 0.4284 | 0.44437 | 0.512 | 0.44 | 0.45 | 0.44 | 0.48 | 0.446 | 0.38751 | 0.4495 | 0.4306 | 0.47651 |  | 1
| 1 | 2 | 0.4493 | 0.43513 | 0.425 | 0.275 | 0.36333 | 0.489 | 0.445 | 0.437 | 0.456 | 0.436 | 0.451 | 0.461 | 0.442 | 0.4447 | 0.45475 | 0.5006 | 0.44 | 0.46 | 0.44 | 0.48 | 0.445 | 0.37697 | 0.4657 | 0.4248 | 0.48465 |  | 32
| 1 | 3 | 0.4426 | 0.43127 | 0.425 | 0.209 | 0.35333 | 0.477 | 0.441 | 0.434 | 0.461 | 0.437 | 0.429 | 0.44 | 0.444 | 0.4325 | 0.44232 | 0.509 | 0.43 | 0.45 | 0.44 | 0.48 | 0.465 | 0.38785 | 0.4557 | 0.4354 | 0.47682 |  | 16
| 1 | 4 | 0.4426 | 0.42289 | 0.435 | 0.234 | 0.36667 | 0.478 | 0.445 | 0.424 | 0.448 | 0.436 | 0.466 | 0.441 | 0.437 | 0.4417 | 0.44788 | 0.5126 | 0.45 | 0.45 | 0.46 | 0.48 | 0.452 | 0.39525 | 0.4659 | 0.4318 | 0.47628 |  | 0.44255
| 1 | 5 | 0.4401 | 0.43522 | 0.425 | 0.198 | 0.365 | 0.484 | 0.44 | 0.435 | 0.452 | 0.439 | 0.427 | 0.455 | 0.457 | 0.4457 | 0.4437 | 0.5006 | 0.45 | 0.45 | 0.46 | 0.48 | 0.448 | 0.38666 | 0.4445 | 0.4299 | 0.47313 |  | 61
| 1 | 6 | 0.4446 | 0.43205 | 0.42 | 0.37 | 0.34667 | 0.477 | 0.457 | 0.438 | 0.455 | 0.439 | 0.458 | 0.428 | 0.443 | 0.4437 | 0.44881 | 0.5186 | 0.44 | 0.45 | 0.45 | 0.48 | 0.439 | 0.39031 | 0.4512 | 0.432 | 0.4728
| Mean |  | 0.44393 | 0.42997 | 0.425 | 0.27617 | 0.35833 | 0.48237 | 0.4435 | 0.435 | 0.456 | 0.43933 | 0.4435 | 0.44383 | 0.4455 | 0.43945 | 0.44697 | 0.5089 | 0.44167 | 0.45167 | 0.44833 | 0.48 | 0.44917 | 0.38743 | 0.45542 | 0.43075 | 0.4767
| Median |  | 0.4435 | 0.43166 | 0.425 | 0.2545 | 0.35917 | 0.481 | 0.443 | 0.436 | 0.4555 | 0.438 | 0.4405 | 0.4405 | 0.4435 | 0.4427 | 0.44612 | 0.5105 | 0.44 | 0.45 | 0.445 | 0.48 | 0.447 | 0.38768 | 0.45345 | 0.4312 | 0.47639
| Std Dev. |  | 0.0030891 | 0.0055693 | 0.0054772 | 0.077716 | 0.0078881 | 0.0058315 | 0.0079436 | 0.0060663 | 0.005831 | 0.0049261 | 0.016956 | 0.012057 | 0.0070071 | 0.0072133 | 0.0045555 | 0.0071439 | 0.0075277 | 0.0040825 | 0.0098319 | 0 | 0.0088412 | 0.0059947 | 0.0088055 | 0.0034766 | 0.0042741
| Rel.Std.Dev. |  | 0.6959% | 1.295% | 1.289% | 28.14% | 2.201% | 1.209% | 1.791% | 1.395% | 1.279% | 1.121% | 3.823% | 2.717% | 1.573% | 1.641% | 1.019% | 1.404% | 1.704% | 0.9039% | 2.193% | 0% | 1.968% | 1.547% | 1.934% | 0.8071% | 0.8966%
| PDM3 |  | 0.3135% | -2.841% | -3.965% | -37.6% | -19.03% | 8.998% | 0.2156% | -1.705% | 3.04% | -0.726% | 0.2156% | 0.2909% | 0.6675% | -0.6996% | 1% | 14.99% | -0.1987% | 2.061% | 1.308% | 8.463% | 1.496% | -12.46% | 2.908% | -2.666% | 7.717%
| Z-Score (Absolute) |  | 0.01 | 1.06 | 1.44 | 12.8 | 6.53 | 2.94 | 0.03 | 0.67 | 0.93 | 0.34 | 0.03 | 0 | 0.13 | 0.33 | 0.24 | 4.96 | 0.17 | 0.6 | 0.34 | 2.76 | 0.41 | 4.31 | 0.88 | 1 | 2.51
| Analytical results for Tl in OREAS 232b (Certified Value 0.77 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.82 | 0.70701 | 0.8 | 0.79 | 0.78 | 0.76 | 0.69 | 0.9 | 0.81 | 0.72 | 0.73 | 0.75 | 0.74 | 0.83215 | 0.738 | 0.7 | 0.76 | 0.79 | 0.77 | 0.91384 | 0.9004 | 0.80733 |  |  |  |  | 1
| 1 | 2 | 0.81 | 0.78799 | 0.8 | 0.79 | 0.79 | 0.7 | 0.77 | 0.9 | 0.78 | 0.78 | 0.8 | 0.75 | 0.79 | 0.79444 | 0.721 | 0.69 | 0.77 | 0.75 | 0.8 | 0.91168 | 0.9004 | 0.80707 |  |  |  |  | 33
| 1 | 3 | 0.81 | 0.72346 | 0.8 | 0.79 | 0.79 | 0.72 | 0.75 | 0.9 | 0.79 | 0.73 | 0.74 | 0.75 | 0.72 | 0.81309 | 0.743 | 0.69 | 0.74 | 0.76 | 0.8 | 0.95106 | 0.8791 | 0.82304 |  |  |  |  | 16
| 1 | 4 | 0.81 | 0.75725 | 0.8 | 0.82 | 0.79 | 0.71 | 0.79 | 0.8 | 0.81 | 0.82 | 0.77 | 0.74 | 0.79 | 0.81759 | 0.73 | 0.68 | 0.77 | 0.8 | 0.77 | 0.95326 | 0.8992 | 0.82864 |  |  |  |  | 0.76917
| 1 | 5 | 0.84 | 0.77889 | 0.8 | 0.79 | 0.78 | 0.66 | 0.81 | 0.9 | 0.79 | 0.73 | 0.75 | 0.76 | 0.75 | 0.82558 | 0.746 | 0.67 | 0.78 | 0.76 | 0.79 | 0.90662 | 0.8638 | 0.83517 |  |  |  |  | 62
| 1 | 6 | 0.84 | 0.73361 | 0.8 | 0.8 | 0.8 | 0.77 | 0.76 | 0.8 | 0.83 | 0.79 | 0.75 | 0.75 | 0.79 | 0.82257 | 0.738 | 0.7 | 0.81 | 0.74 | 0.75 | 0.92125 | 0.8773 | 0.81735
| Mean |  | 0.82167 | 0.74803 | 0.8 | 0.79667 | 0.78833 | 0.72 | 0.76167 | 0.86667 | 0.80167 | 0.76167 | 0.75667 | 0.75 | 0.76333 | 0.81757 | 0.736 | 0.68833 | 0.77167 | 0.76667 | 0.78 | 0.92629 | 0.8867 | 0.81977
| Median |  | 0.815 | 0.74543 | 0.8 | 0.79 | 0.79 | 0.715 | 0.765 | 0.9 | 0.8 | 0.755 | 0.75 | 0.75 | 0.77 | 0.82008 | 0.738 | 0.69 | 0.77 | 0.76 | 0.78 | 0.91755 | 0.88915 | 0.8202
| Std Dev. |  | 0.01472 | 0.032026 | 1.2162e-16 | 0.012111 | 0.0075277 | 0.040497 | 0.041191 | 0.05164 | 0.018348 | 0.040702 | 0.025033 | 0.0063246 | 0.030768 | 0.013092 | 0.0091433 | 0.01169 | 0.023166 | 0.023381 | 0.02 | 0.0206 | 0.015506 | 0.011388
| Rel.Std.Dev. |  | 1.791% | 4.281% | 1.52e-14% | 1.52% | 0.9549% | 5.625% | 5.408% | 5.958% | 2.289% | 5.344% | 3.308% | 0.8433% | 4.031% | 1.601% | 1.242% | 1.698% | 3.002% | 3.05% | 2.564% | 2.224% | 1.749% | 1.389%
| PDM3 |  | 6.825% | -2.748% | 4.008% | 3.575% | 2.492% | -6.392% | -0.9753% | 12.68% | 4.225% | -0.9753% | -1.625% | -2.492% | -0.7587% | 6.293% | -4.312% | -10.51% | 0.3248% | -0.3253% | 1.408% | 20.43% | 15.28% | 6.578%
| Z-Score (Absolute) |  | 1.46 | 0.59 | NA | 0.76 | 0.53 | 1.36 | 0.21 | NA | 0.9 | 0.21 | 0.35 | 0.53 | 0.16 | 1.34 | 0.92 | 2.24 | 0.07 | 0.07 | 0.3 | 4.36 | 3.26 | 1.4
Results from laboratories 1.05 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 232b (Certified Value 0.25 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.23166 | 0.3 | 0.3 | 0.3 | 0.216 | 0.27 | 0.2 | 0.2119 | 0.46418 | [17 empty] | 1
| 1 | 2 | 0.23397 | 0.3 | 0.3 | 0.2 | 0.218 | 0.27 | 0.2 | 0.2033 | 0.4343 | [17 empty] | 13
| 1 | 3 | 0.23723 | 0.3 | 0.3 | 0.3 | 0.211 | 0.25 | 0.2 | 0.2122 | 0.44589 | [17 empty] | 16
| 1 | 4 | 0.24105 | 0.35 | 0.3 | 0.3 | 0.216 | 0.27 | 0.2 | 0.2118 | 0.46132 | [17 empty] | 0.24925
| 1 | 5 | 0.23074 | 0.3 | 0.3 | 0.2 | 0.206 | 0.27 | 0.2 | 0.2125 | 0.46689 | [17 empty] | 63
| 1 | 6 | 0.23788 | 0.3 | 0.3 | 0.3 | 0.241 | 0.28 | 0.2 | 0.2096 | 0.46483
| Mean |  | 0.23542 | 0.30833 | 0.3 | 0.26667 | 0.218 | 0.26833 | 0.2 | 0.21022 | 0.45624
| Median |  | 0.2356 | 0.3 | 0.3 | 0.3 | 0.216 | 0.27 | 0.2 | 0.21185 | 0.46275
| Std Dev. |  | 0.0039808 | 0.020412 | 0 | 0.05164 | 0.012083 | 0.0098319 | 3.0405e-17 | 0.0035414 | 0.013149
| Rel.Std.Dev. |  | 1.691% | 6.62% | 0% | 19.36% | 5.543% | 3.664% | 1.52e-14% | 1.685% | 2.882%
| PDM3 |  | -5.55% | 23.7% | 20.36% | 6.986% | -12.54% | 7.654% | -19.76% | -15.66% | 83.04%
| Z-Score (Absolute) |  | 0.51 | 0.67 | 0.54 | 0 | 0.79 | 0.03 | 1.08 | 0.91 | 3.07
| Analytical results for U in OREAS 232b (Certified Value 2.71 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.78 | 2.7543 | 2.5 | 2.9 | 2.79 | 2.5 | 2.92 | 3 | 2.5 | 2.6 | 2.7 | 2.6 | 2.9 | 2.6452 | 2.65 | 2.481 | 2.14 | 2.74 | 2.9 | 2.6 | 2.6436 | 3.7383 | 2.9413 |  |  |  | 1
| 1 | 2 | 2.78 | 2.8316 | 2.5 | 2.8 | 2.82 | 2.5 | 2.77 | 2.9 | 2.5 | 2.6 | 2.8 | 2.6 | 2.8 | 2.5574 | 2.7 | 2.407 | 2.26 | 2.66 | 2.8 | 2.6 | 2.5884 | 4.0754 | 2.8326 |  |  |  | 35
| 1 | 3 | 2.71 | 2.9035 | 2.6 | 2.7 | 2.72 | 2.4 | 2.85 | 2.8 | 2.5 | 2.6 | 2.7 | 2.5 | 2.7 | 2.7686 | 2.71 | 2.537 | 2.27 | 2.53 | 2.8 | 2.6 | 2.6953 | 4.1321 | 2.9447 |  |  |  | 16
| 1 | 4 | 2.72 | 2.7608 | 2.7 | 2.8 | 2.78 | 2.5 | 2.83 | 2.8 | 2.5 | 2.9 | 2.6 | 2.5 | 2.9 | 2.5652 | 2.77 | 2.543 | 2.27 | 2.69 | 2.9 | 2.6 | 2.7539 | 3.7651 | 3.0014 |  |  |  | 2.7067
| 1 | 5 | 2.77 | 2.7113 | 2.7 | 2.7 | 2.75 | 2.6 | 3.01 | 2.9 | 2.5 | 2.5 | 2.8 | 2.6 | 2.6 | 2.6576 | 2.69 | 2.465 | 2.24 | 2.61 | 2.8 | 2.6 | 2.7028 | 3.7966 | 3.0007 |  |  |  | 64
| 1 | 6 | 2.78 | 2.733 | 2.6 | 3.1 | 2.7 | 2.7 | 3.05 | 2.8 | 2.6 | 2.8 | 2.6 | 2.6 | 2.8 | 2.6944 | 2.82 | 2.515 | 2.2 | 2.81 | 2.7 | 2.5 | 2.7083 | 3.8202 | 2.9903
| Mean |  | 2.7567 | 2.7824 | 2.6 | 2.8333 | 2.76 | 2.5333 | 2.905 | 2.8667 | 2.5167 | 2.6667 | 2.7 | 2.5667 | 2.7833 | 2.6481 | 2.7233 | 2.4913 | 2.23 | 2.6733 | 2.8167 | 2.5833 | 2.6821 | 3.888 | 2.9518
| Median |  | 2.775 | 2.7576 | 2.6 | 2.8 | 2.765 | 2.5 | 2.885 | 2.85 | 2.5 | 2.6 | 2.7 | 2.6 | 2.8 | 2.6514 | 2.705 | 2.498 | 2.25 | 2.675 | 2.8 | 2.6 | 2.699 | 3.8084 | 2.9675
| Std Dev. |  | 0.03266 | 0.0719 | 0.089443 | 0.15055 | 0.045166 | 0.10328 | 0.10877 | 0.08165 | 0.040825 | 0.15055 | 0.089443 | 0.05164 | 0.1169 | 0.079812 | 0.06121 | 0.051434 | 0.051381 | 0.098116 | 0.075277 | 0.040825 | 0.057804 | 0.17039 | 0.064335
| Rel.Std.Dev. |  | 1.185% | 2.584% | 3.44% | 5.314% | 1.636% | 4.077% | 3.744% | 2.848% | 1.622% | 5.646% | 3.313% | 2.012% | 4.2% | 3.014% | 2.248% | 2.065% | 2.304% | 3.67% | 2.673% | 1.58% | 2.155% | 4.383% | 2.179%
| PDM3 |  | 1.846% | 2.797% | -3.942% | 4.678% | 1.969% | -6.405% | 7.326% | 5.91% | -7.021% | -1.479% | -0.2475% | -5.174% | 2.831% | -2.166% | 0.6145% | -7.957% | -17.61% | -1.233% | 4.063% | -4.558% | -0.9107% | 43.64% | 9.057%
| Z-Score (Absolute) |  | 0.33 | 0.48 | 0.58 | 0.77 | 0.35 | 0.96 | 1.18 | 0.96 | 1.06 | 0.19 | 0 | 0.77 | 0.48 | 0.3 | 0.13 | 1.21 | 2.72 | 0.15 | 0.67 | 0.67 | 0.1 | 6.87 | 1.46
| Analytical results for V in OREAS 232b (Certified Value 98 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 102 | 87.61 | 90 | 85 | 107.7643 | 102 | 97 | 88 | 97 | 99 | 98 | 96 | 98 | 96 | 99.0078 | 97.8 | 100 | 98.06 | 101 | 100 | 89 | 93.5 | 96 | 109.7743 | 103.0544 |  | 1
| 1 | 2 | 102 | 85.66 | 100 | 76 | 106.3151 | 98 | 98 | 95 | 98 | 97 | 104 | 102 | 96 | 97 | 100.7823 | 97.9 | 101 | 98.44 | 100 | 101 | 88 | 92.64 | 98 | 99.921 | 100.829 |  | 36
| 1 | 3 | 102 | 85.68 | 100 | 65 | 108.2461 | 99 | 97 | 97 | 95 | 96 | 100 | 96 | 98 | 96 | 100.8757 | 98.4 | 97 | 97.01 | 103 | 102 | 89 | 94.49 | 98 | 109.8748 | 103.061 |  | 16
| 1 | 4 | 102 | 86.01 | 95 | 66 | 109.3407 | 98 | 96 | 92 | 97 | 97 | 106 | 96 | 96 | 98 | 99.9959 | 98.8 | 101 | 97.54 | 103 | 102 | 89 | 95.48 | 97 | 100.8862 | 104.5538 |  | 98.0314
| 1 | 5 | 102 | 83.79 | 95 | 60 | 108.889 | 100 | 97 | 92 | 96 | 97 | 98 | 97 | 98 | 97 | 98.7804 | 96.8 | 101 | 97.48 | 102 | 101 | 84 | 93.06 | 95 | 106.1145 | 101.4385 |  | 65
| 1 | 6 | 105 | 85.61 | 95 | 90 | 108.654 | 102 | 97 | 97 | 94 | 98 | 104 | 94 | 97 | 97 | 100.0936 | 99.6 | 101 | 97.93 | 101 | 102 | 98 | 93.83 | 95 | 103.2188 | 101.528
| Mean |  | 102.5 | 85.7267 | 95.8333 | 73.6667 | 108.2015 | 99.8333 | 97 | 93.5 | 96.1667 | 97.3333 | 101.6667 | 96.8333 | 97.1667 | 96.8333 | 99.9226 | 98.2167 | 100.1667 | 97.7433 | 101.6667 | 101.3333 | 89.5 | 93.8333 | 96.5 | 104.9649 | 102.4108
| Median |  | 102 | 85.67 | 95 | 71 | 108.4501 | 99.5 | 97 | 93.5 | 96.5 | 97 | 102 | 96 | 97.5 | 97 | 100.0447 | 98.15 | 101 | 97.735 | 101.5 | 101.5 | 89 | 93.665 | 96.5 | 104.6667 | 102.2912
| Std Dev. |  | 1.2247 | 1.2164 | 3.7639 | 12.0111 | 1.0702 | 1.8348 | 0.63246 | 3.5071 | 1.472 | 1.0328 | 3.4448 | 2.7142 | 0.98319 | 0.75277 | 0.87449 | 0.95586 | 1.6021 | 0.50377 | 1.2111 | 0.8165 | 4.5935 | 1.0272 | 1.3784 | 4.3286 | 1.3897
| Rel.Std.Dev. |  | 1.195% | 1.419% | 3.928% | 16.3% | 0.9891% | 1.838% | 0.652% | 3.751% | 1.531% | 1.061% | 3.388% | 2.803% | 1.012% | 0.7774% | 0.8752% | 0.9732% | 1.599% | 0.5154% | 1.191% | 0.8058% | 5.132% | 1.095% | 1.428% | 4.124% | 1.357%
| PDM3 |  | 4.558% | -12.55% | -2.242% | -24.85% | 10.37% | 1.838% | -1.052% | -4.622% | -1.902% | -0.7121% | 3.708% | -1.222% | -0.8821% | -1.222% | 1.929% | 0.189% | 2.178% | -0.2938% | 3.708% | 3.368% | -8.703% | -4.282% | -1.562% | 7.073% | 4.467%
| Z-Score (Absolute) |  | 1.23 | 2.76 | 0.36 | 5.63 | 2.59 | 0.59 | 0.08 | 0.91 | 0.28 | 0 | 1.03 | 0.12 | 0.04 | 0.12 | 0.62 | 0.21 | 0.67 | 0.1 | 1.03 | 0.95 | 1.86 | 0.83 | 0.2 | 1.82 | 1.21
| Analytical results for W in OREAS 232b (Certified Value 1.9 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 2.1 | 2.0939 | 1.5 | < 0.1 | 19.36 | 2.1 | 2.1 | 1.3 | 1.9 | 2 | 1.8 | 1.8 | 1.8 | 2.2 | 1.9222 | 2.9 | 1.6 | < 5 | 2.39 | 2.5 | 1.5 | 1.2337 | < 5 | 2.4806 |  |  | 1
| 1 | 2 | 2 | 2.0586 | 2.5 | < 0.1 | 20.251 | 2.1 | 2 | 0.7 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 2.2 | 1.8328 | 3 | 1.7 | < 5 | 2.33 | 2.3 | 1.5 | 1.3068 | < 5 | 2.3855 |  |  | 37
| 1 | 3 | 1.9 | 1.9843 | 2 | < 0.1 | 18.8914 | 1.8 | 2 | 1.3 | 1.7 | 2 | 1.8 | 1.9 | 1.7 | 2 | 1.8779 | 2.9 | 1.7 | < 5 | 2.16 | 2.6 | 1.5 | 1.2816 | < 5 | 2.3613 |  |  | 16
| 1 | 4 | 1.9 | 1.9817 | 2 | < 0.1 | 20.9481 | 2.1 | 1.7 | 3 | 1.8 | 2 | 1.9 | 1.9 | 1.7 | 2.2 | 1.9475 | 2.9 | 1.6 | < 5 | 2.28 | 2.3 | 1.5 | 1.2454 | < 5 | 2.429 |  |  | 1.8998
| 1 | 5 | 1.9 | 2.0605 | 2 | < 0.1 | 19.59 | 1.8 | 1.8 | 1.6 | 1.7 | 1.9 | 1.6 | 1.9 | 1.9 | 2.1 | 1.9141 | 2.9 | 1.7 | < 5 | 2.37 | 2.3 | 1.5 | 1.2966 | < 5 | 2.4989 |  |  | 66
| 1 | 6 | 2.2 | 2.0195 | 2 | < 0.1 | 19.08 | 1.9 | 1.7 | 1.6 | 1.7 | 2 | 1.9 | 1.9 | 1.7 | 2.1 | 1.9449 | 2.8 | 1.6 | < 5 | 2.45 | 2.2 | 1.5 | 1.2122 | < 5 | 2.4237
| Mean |  | 2 | 2.0331 | 2 |  | 19.6867 | 1.9667 | 1.8833 | 1.5833 | 1.7667 | 1.9667 | 1.8 | 1.8833 | 1.7667 | 2.1333 | 1.9066 | 2.9 | 1.65 |  | 2.33 | 2.3667 | 1.5 | 1.2627 |  | 2.4298
| Median |  | 1.95 | 2.0391 | 2 |  | 19.475 | 2 | 1.9 | 1.45 | 1.75 | 2 | 1.8 | 1.9 | 1.75 | 2.15 | 1.9182 | 2.9 | 1.65 |  | 2.35 | 2.3 | 1.5 | 1.2635 |  | 2.4263
| Std Dev. |  | 0.12649 | 0.045418 | 0.31623 |  | 0.77799 | 0.15055 | 0.17224 | 0.7679 | 0.08165 | 0.05164 | 0.10954 | 0.040825 | 0.08165 | 0.08165 | 0.044065 | 0.063246 | 0.054772 |  | 0.101 | 0.15055 | 0 | 0.037794 |  | 0.053019
| Rel.Std.Dev. |  | 6.325% | 2.234% | 15.81% |  | 3.952% | 7.655% | 9.145% | 48.5% | 4.622% | 2.626% | 6.086% | 2.168% | 4.622% | 3.827% | 2.311% | 2.181% | 3.32% |  | 4.335% | 6.361% | 0% | 2.993% |  | 2.182%
| PDM3 |  | 5.276% | 7.018% | 5.276% |  | 936.3% | 3.522% | -0.8649% | -16.66% | -7.006% | 3.522% | -5.251% | -0.8649% | -7.006% | 12.29% | 0.3587% | 52.65% | -13.15% |  | 22.65% | 24.58% | -21.04% | -33.53% |  | 27.9%
| Z-Score (Absolute) |  | 0.09 | 0.17 | 0.09 | 5 | 46.25 | 0 | 0.22 | 1 | 0.52 | 0 | 0.44 | 0.22 | 0.52 | 0.44 | 0.16 | 2.44 | 0.83 | 1.39 | 0.95 | 1.04 | 1.22 | 1.84 | 1.39 | 1.21
| Analytical results for Y in OREAS 232b (Certified Value 15.9 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 16.57 | 12.6765 | 17.1 | 18.7 | 9.715 | 16.7 | 11 | 18.8 | 15.3 | 17.8 | 17.5 | 15.6 | 16.9 | 14.1 | 15.2053 | 19 | 12.31 | 13.6 | 17.66 | 18.7 | 14.5 | 12.63 | 15.1913 | 27.6986 |  |  | 1
| 1 | 2 | 15.97 | 12.3585 | 18 | 18.2 | 11.05 | 16.6 | 10 | 16.8 | 14.2 | 16.9 | 15.8 | 17.5 | 16.2 | 14.1 | 16.2915 | 19.7 | 12.73 | 13.8 | 17.39 | 16.9 | 14.8 | 12.35 | 15.4479 | 28.307 |  |  | 38
| 1 | 3 | 16.07 | 12.6867 | 17.9 | 18.8 | 8.7667 | 17.1 | 10 | 17 | 13.6 | 17.3 | 16 | 21.6 | 15.7 | 14.2 | 15.7149 | 19.6 | 12.55 | 12.8 | 20.3 | 16.9 | 14.8 | 12.75 | 15.4992 | 27.9978 |  |  | 16
| 1 | 4 | 16.8 | 12.6396 | 17.8 | 18.7 | 8.7867 | 17 | 11 | 17 | 13.7 | 18.3 | 17.4 | 15.5 | 16.2 | 14.8 | 16.0091 | 19.2 | 12.94 | 14.2 | 17.54 | 17.2 | 14.5 | 12.86 | 16.0636 | 29.2376 |  |  | 15.8747
| 1 | 5 | 15.82 | 12.8718 | 17.5 | 18.4 | 10.565 | 16.5 | 13 | 16.7 | 14.1 | 17.2 | 15.2 | 17 | 16.4 | 14.6 | 15.7239 | 18.6 | 12.24 | 13.6 | 17.83 | 16.8 | 14.1 | 12.54 | 15.0886 | 28.7909 |  |  | 67
| 1 | 6 | 15.75 | 12.6644 | 17.9 | 20 | 9.11 | 16.9 | 12 | 18 | 15.2 | 17.3 | 16.5 | 29.7 | 15.5 | 14.3 | 15.3732 | 18.9 | 12.38 | 13.4 | 18.46 | 16.6 | 14.1 | 12.6 | 15.8071 | 29.0525
| Mean |  | 16.1633 | 12.6496 | 17.7 | 18.8 | 9.6656 | 16.8 | 11.1667 | 17.3833 | 14.35 | 17.4667 | 16.4 | 19.4833 | 16.15 | 14.35 | 15.7197 | 19.1667 | 12.525 | 13.5667 | 18.1967 | 17.1833 | 14.4667 | 12.6217 | 15.5163 | 28.514
| Median |  | 16.02 | 12.6705 | 17.85 | 18.7 | 9.4125 | 16.8 | 11 | 17 | 14.15 | 17.3 | 16.25 | 17.25 | 16.2 | 14.25 | 15.7194 | 19.1 | 12.465 | 13.6 | 17.745 | 16.9 | 14.5 | 12.615 | 15.4735 | 28.5489
| Std Dev. |  | 0.42557 | 0.16525 | 0.34059 | 0.62929 | 0.96096 | 0.23664 | 1.169 | 0.83526 | 0.73417 | 0.50067 | 0.91433 | 5.4756 | 0.501 | 0.2881 | 0.39886 | 0.42269 | 0.26957 | 0.46332 | 1.0952 | 0.7679 | 0.31411 | 0.17566 | 0.36814 | 0.61066
| Rel.Std.Dev. |  | 2.633% | 1.306% | 1.924% | 3.347% | 9.942% | 1.409% | 10.47% | 4.805% | 5.116% | 2.866% | 5.575% | 28.1% | 3.102% | 2.008% | 2.537% | 2.205% | 2.152% | 3.415% | 6.019% | 4.469% | 2.171% | 1.392% | 2.373% | 2.142%
| PDM3 |  | 1.818% | -20.32% | 11.5% | 18.43% | -39.11% | 5.829% | -29.66% | 9.504% | -9.604% | 10.03% | 3.309% | 22.73% | 1.734% | -9.604% | -0.9764% | 20.74% | -21.1% | -14.54% | 14.63% | 8.244% | -8.87% | -20.49% | -2.258% | 79.62%
| Z-Score (Absolute) |  | 0 | 1.31 | 0.57 | 0.98 | 2.42 | 0.24 | NA | 0.45 | 0.67 | 0.48 | 0.09 | 1.23 | 0 | 0.67 | 0.16 | 1.12 | 1.35 | 0.97 | 0.76 | 0.38 | 0.63 | 1.32 | 0.24 | 4.59
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 232b (Certified Value 1.7 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.32 | 1.34 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 1.6138 | 1.85 | 2.2 | 1.9 | 1.89 | 1.424 | 1.6 | 1.96 | 1.5 | 1.11 | 1.5998 | 3.0528 | [14 empty] | 1
| 1 | 2 | 1.5504 | 1.9 | 2.1 | 1.8 | 1.96 | 1.715 | 1.6 | 1.97 | 1.5 | 1.08 | 1.6848 | 2.9346 | [14 empty] | 39
| 1 | 3 | 1.5929 | 1.95 | 2 | 1.5 | 1.96 | 1.485 | 1.5 | 1.87 | 1.5 | 1.12 | 1.6423 | 2.9815 | [14 empty] | 16
| 1 | 4 | 1.5852 | 1.95 | 2.1 | 3.1 | 1.9 | 1.651 | 1.6 | 1.94 | 1.5 | 1.38 | 1.7273 | 3.0484 | [14 empty] | 1.7021
| 1 | 5 | 1.4879 | 1.85 | 2.1 | 1.5 | 1.91 | 1.471 | 1.6 | 1.95 | 1.5 | 1.28 | 1.5998 | 3.1119 | [14 empty] | 68
| 1 | 6 | 1.4854 | 1.95 | 2.1 | 2.1 | 1.9 | 1.579 | 1.6 | 2.07 | 1.5 | 1.4 | 1.6848 | 3.1421
| Mean |  | 1.5526 | 1.9083 | 2.1 | 1.9833 | 1.92 | 1.5542 | 1.5833 | 1.96 | 1.5 | 1.2283 | 1.6565 | 3.0452
| Median |  | 1.5678 | 1.925 | 2.1 | 1.85 | 1.905 | 1.532 | 1.6 | 1.955 | 1.5 | 1.2 | 1.6636 | 3.0506
| Std Dev. |  | 0.055022 | 0.04916 | 0.063246 | 0.5947 | 0.031623 | 0.11358 | 0.040825 | 0.064498 | 0 | 0.14345 | 0.05147 | 0.077705
| Rel.Std.Dev. |  | 3.544% | 2.576% | 3.012% | 29.98% | 1.647% | 7.308% | 2.578% | 3.291% | 0% | 11.68% | 3.107% | 2.552%
| PDM3 |  | -8.784% | 12.12% | 23.38% | 16.52% | 12.8% | -8.692% | -6.978% | 15.15% | -11.87% | -27.83% | -2.682% | 78.91%
| Z-Score (Absolute) |  | 0.72 | 0.4 | 1 | 0.63 | 0.43 | 0.72 | 0.63 | 0.56 | 0.89 | 1.74 | 0.4 | 3.97
| Analytical results for Zn in OREAS 232b (Certified Value 91 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 93 | 86.9952 | 88 | 97 | 94.6444 | 92 | 93 | 88 | 94 | 93 | 90 | 89 | 92 | 88.4 | 90.8611 | 99 | 91.6 | 90 | 111 | 88 | 89 | 90.4 | 94.061 | 90 | 84.6392 | 86.6865 | 1
| 1 | 2 | 94 | 87.8792 | 90 | 100 | 92.625 | 93 | 93 | 89 | 92 | 90 | 94 | 93 | 90 | 91.1 | 90.0659 | 98 | 93.03 | 92 | 115.4 | 87 | 89 | 91 | 93.31 | 93 | 85.9428 | 88.2252 | 40
| 1 | 3 | 95 | 86.652 | 88 | 95 | 91.4081 | 92 | 89 | 90 | 91 | 90 | 90 | 89 | 90 | 89.1 | 89.0198 | 99 | 92.93 | 88 | 106.9 | 86 | 88 | 92.4 | 94.787 | 92 | 88.8978 | 90.0054 | 16
| 1 | 4 | 109 | 84.2912 | 90 | 97 | 92.1982 | 92 | 91 | 89 | 92 | 90 | 96 | 87 | 90 | 90.3 | 91.1129 | 99 | 91.87 | 89 | 110.2 | 89 | 87 | 90.4 | 95.898 | 93 | 93.6199 | 90.9869 | 90.5943
| 1 | 5 | 93 | 85.7576 | 88 | 97 | 93.3213 | 93 | 97 | 91 | 92 | 90 | 88 | 92 | 92 | 89.9 | 89.3699 | 101 | 91.53 | 91 | 109.9 | 91 | 88 | 88.8 | 93.434 | 89 | 84.0743 | 86.7189 | 69
| 1 | 6 | 93 | 84.0728 | 88 | 96 | 92.6468 | 93 | 90 | 91 | 92 | 90 | 95 | 85 | 91 | 89.7 | 90.0215 | 97 | 93.04 | 89 | 109.6 | 88 | 88 | 86.5 | 94.831 | 90 | 86.4498 | 87.6935
| Mean |  | 96.1667 | 85.9413 | 88.6667 | 97 | 92.8073 | 92.5 | 92.1667 | 89.6667 | 92.1667 | 90.5 | 92.1667 | 89.1667 | 90.8333 | 89.75 | 90.0752 | 98.8333 | 92.3333 | 89.8333 | 110.5 | 88.1667 | 88.1667 | 89.9167 | 94.3868 | 91.1667 | 87.2706 | 88.3861
| Median |  | 93.5 | 86.2048 | 88 | 97 | 92.6359 | 92.5 | 92 | 89.5 | 92 | 90 | 92 | 89 | 90.5 | 89.8 | 90.0437 | 99 | 92.4 | 89.5 | 110.05 | 88 | 88 | 90.4 | 94.424 | 91 | 86.1963 | 87.9594
| Std Dev. |  | 6.3377 | 1.5244 | 1.0328 | 1.6733 | 1.0979 | 0.54772 | 2.8577 | 1.2111 | 0.98319 | 1.2247 | 3.2506 | 2.9944 | 0.98319 | 0.93755 | 0.81348 | 1.3292 | 0.74007 | 1.472 | 2.7742 | 1.7224 | 0.75277 | 2.0361 | 0.98129 | 1.7224 | 3.5366 | 1.7641
| Rel.Std.Dev. |  | 6.59% | 1.774% | 1.165% | 1.725% | 1.183% | 0.5921% | 3.101% | 1.351% | 1.067% | 1.353% | 3.527% | 3.358% | 1.082% | 1.045% | 0.9031% | 1.345% | 0.8015% | 1.639% | 2.511% | 1.954% | 0.8538% | 2.264% | 1.04% | 1.889% | 4.052% | 1.996%
| PDM3 |  | 6.151% | -5.136% | -2.128% | 7.071% | 2.443% | 2.104% | 1.736% | -1.024% | 1.736% | -0.1041% | 1.736% | -1.576% | 0.2638% | -0.932% | -0.573% | 9.094% | 1.92% | -0.84% | 21.97% | -2.68% | -2.68% | -0.748% | 4.186% | 0.6317% | -3.669% | -2.438%
| Z-Score (Absolute) |  | 2.12 | 1.82 | 0.77 | 2.44 | 0.82 | 0.71 | 0.58 | 0.39 | 0.58 | 0.06 | 0.58 | 0.58 | 0.06 | 0.35 | 0.23 | 3.15 | 0.64 | 0.32 | 7.64 | 0.96 | 0.96 | 0.29 | 1.43 | 0.19 | 1.31 | 0.88
| Analytical results for Zr in OREAS 232b (Certified Value 140 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 152.5 | 136.8472 | 140 | 130 | 115.2 | 140.4 | 140 | 147 | 149.1 | 146.5 | 137.5 | 140 | 146 | 127.9 | 137.5766 | 153 | 134 | 144.4 | 141.4 | 132.4 | 123.9 | 128.634 | 142 | 137.6285 | 238.2081 |  | 1
| 1 | 2 | 149.7 | 138.7484 | 147 | 115 | 113.35 | 141 | 141.8 | 143 | 150.2 | 139 | 141 | 145.5 | 145 | 129.5 | 133.6522 | 148 | 135 | 148.5 | 133.8 | 137.3 | 125.3 | 123.882 | 147 | 167.7408 | 241.4897 |  | 41
| 1 | 3 | 151 | 138.9248 | 148 | 96 | 112.54 | 137.5 | 135 | 144 | 145.2 | 138.5 | 141.5 | 139.5 | 136 | 128.7 | 134.9506 | 151 | 128 | 143.8 | 132.2 | 136 | 128 | 126.379 | 144 | 139.7938 | 240.9101 |  | 16
| 1 | 4 | 144.2 | 137.1314 | 150 | 101 | 116.52 | 139.1 | 142.2 | 141 | 150 | 145 | 154 | 138 | 144 | 132.8 | 136.4451 | 150 | 135 | 149.8 | 137.2 | 134.5 | 126.6 | 129.217 | 145 | 136.848 | 238.3542 |  | 139.9245
| 1 | 5 | 149 | 138.915 | 144 | 89 | 116.65 | 140 | 147.6 | 145 | 148.4 | 138 | 136.5 | 144.5 | 145 | 130.8 | 134.8982 | 145 | 135 | 146.7 | 137.1 | 133.8 | 126.2 | 124.179 | 140 | 141.6569 | 237.1391 |  | 70
| 1 | 6 | 145.1 | 140.3948 | 149 | 94 | 112.84 | 137.8 | 143.1 | 151 | 145.1 | 141.5 | 146.5 | 136.5 | 138.5 | 130.9 | 135.1389 | 153 | 133 | 145 | 141 | 134.3 | 131 | 126.39 | 139 | 140.3477 | 243.8385
| Mean |  | 148.5833 | 138.4936 | 146.3333 | 104.1667 | 114.5167 | 139.3 | 141.6167 | 145.1667 | 148 | 141.4167 | 142.8333 | 140.6667 | 142.4167 | 130.1 | 135.4436 | 150 | 133.3333 | 146.3667 | 137.1167 | 134.7167 | 126.8333 | 126.4468 | 142.8333 | 144.0026 | 239.99
| Median |  | 149.35 | 138.8317 | 147.5 | 98.5 | 114.275 | 139.55 | 142 | 144.5 | 148.75 | 140.25 | 141.25 | 139.75 | 144.5 | 130.15 | 135.0448 | 150.5 | 134.5 | 145.85 | 137.15 | 134.4 | 126.4 | 126.3845 | 143 | 140.0708 | 239.6321
| Std Dev. |  | 3.2848 | 1.3122 | 3.7238 | 15.4585 | 1.8499 | 1.4227 | 4.1175 | 3.4881 | 2.3004 | 3.5975 | 6.509 | 3.587 | 4.1282 | 1.7652 | 1.3708 | 3.0984 | 2.7325 | 2.3989 | 3.7043 | 1.7175 | 2.4549 | 2.1994 | 3.0605 | 11.7635 | 2.5253
| Rel.Std.Dev. |  | 2.211% | 0.9475% | 2.545% | 14.84% | 1.615% | 1.021% | 2.907% | 2.403% | 1.554% | 2.544% | 4.557% | 2.55% | 2.899% | 1.357% | 1.012% | 2.066% | 2.049% | 1.639% | 2.702% | 1.275% | 1.936% | 1.739% | 2.143% | 8.169% | 1.052%
| PDM3 |  | 6.188% | -1.023% | 4.58% | -25.56% | -18.16% | -0.4463% | 1.209% | 3.746% | 5.771% | 1.066% | 2.079% | 0.5304% | 1.781% | -7.021% | -3.202% | 7.201% | -4.711% | 4.604% | -2.007% | -3.722% | -9.356% | -9.632% | 2.079% | 2.915% | 71.51%
| Z-Score (Absolute) |  | 0.98 | 0.4 | 0.67 | 5.07 | 3.66 | 0.29 | 0.03 | 0.51 | 0.9 | 0 | 0.19 | 0.1 | 0.14 | 1.54 | 0.81 | 1.17 | 1.1 | 0.67 | 0.59 | 0.91 | 1.99 | 2.04 | 0.19 | 0.35 | 13.43
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 232b (Certified Value 0.102 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.09 | 0.101 | < 0.5 | 0.1 | 0.08 | 0.12 | 0.1 | 0.18 | 0.1 | 0.1 | 0.1 | 0.088 | < 1 | < 0.2 | 0.1 | 0.11 | < 0.5 | 0.11 | 0.13 | < 1 | 1.7009 |  |  |  | 1
| 1 | 2 | 0.1 | 0.113 | < 0.5 | 0.11 | 0.1 | 0.11 | 0.1 | 0.12 | 0.1 | 0.1 | 0.1 | 0.09 | < 1 | 0.3 | 0.1 | 0.1 | < 0.5 | 0.11 | 0.1 | < 1 | 1.2072 |  |  |  | 25
| 1 | 3 | 0.1 | 0.106 | < 0.5 | 0.11 | 0.08 | 0.11 | < 0.1 | 0.12 | 0.1 | 0.1 | 0.1 | 0.092 | < 1 | < 0.2 | 0.09 | 0.09 | < 0.5 | 0.11 | 0.1 | < 1 | 1.7867 |  |  |  | 16
| 1 | 4 | 0.09 | 0.104 | < 0.5 | 0.11 | 0.09 | 0.11 | < 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.089 | < 1 | < 0.2 | 0.11 | 0.08 | < 0.5 | 0.1 | 0.13 | < 1 | 1.9613 |  |  |  | 0.10232
| 1 | 5 | 0.09 | 0.113 | < 0.5 | 0.09 | 0.1 | 0.13 | < 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.08 | < 1 | < 0.2 | 0.1 | 0.12 | < 0.5 | 0.11 | 0.11 | < 1 | 2.3661 |  |  |  | 72
| 1 | 6 | 0.1 | 0.115 | < 0.5 | 0.11 | 0.07 | 0.18 | < 0.1 | 0.1 | 0.11 | 0.1 | 0.09 | 0.093 | < 1 | 0.3 | 0.11 | 0.11 | < 0.5 | 0.11 | 0.09 | < 1 | 1.2072
| Mean |  | 0.095 | 0.10867 |  | 0.105 | 0.086667 | 0.12667 | 0.1 | 0.12 | 0.105 | 0.1 | 0.098333 | 0.088667 |  | 0.3 | 0.10167 | 0.10167 |  | 0.10833 | 0.11 |  | 1.7049
| Median |  | 0.095 | 0.1095 |  | 0.11 | 0.085 | 0.115 | 0.1 | 0.11 | 0.105 | 0.1 | 0.1 | 0.0895 |  | 0.3 | 0.1 | 0.105 |  | 0.11 | 0.105 |  | 1.7438
| Std Dev. |  | 0.0054772 | 0.0057504 |  | 0.0083666 | 0.012111 | 0.027325 | 0 | 0.030984 | 0.0054772 | 1.5202e-17 | 0.0040825 | 0.0046332 |  | 0 | 0.0075277 | 0.01472 |  | 0.0040825 | 0.016733 |  | 0.44834
| Rel.Std.Dev. |  | 5.766% | 5.292% |  | 7.968% | 13.97% | 21.57% | 0% | 25.82% | 5.216% | 1.52e-14% | 4.152% | 5.225% |  | 0% | 7.404% | 14.48% |  | 3.768% | 15.21% |  | 26.3%
| PDM3 |  | -7.15% | 6.207% |  | 2.624% | -15.29% | 23.8% | -2.263% | 17.28% | 2.624% | -2.263% | -3.892% | -13.34% |  | 193.2% | -0.6342% | -0.6342% |  | 5.882% | 7.511% |  | 1566%
| Z-Score (Absolute) |  | 0.67 | 0.02 | 7.16 | 0.17 | 1.1 | 0.93 | 2.11 | 0.59 | 0.17 | 0.42 | 0.51 | 0.99 | 19.81 | 2.95 | 0.34 | 0.34 | 7.16 | 0 | 0.08 | 19.81 | 80.74
| Analytical results for Al in OREAS 232b (Certified Value 2.72 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 2.477 | 3.0798 | 2.93 | 2.8733 | 2.8 | 2.86 | 3.27 | 2.57 | 2.86 | 2.72 | 2.66 | 2.63 | 2.71 | 2.6815 | 2.472 | 2.57 | 2.53 | 2.51 | 2.4568 | 4.9874 |  |  |  |  | 1
| 1 | 2 | 2.431 | 3.0877 | 2.92 | 2.8733 | 2.814 | 2.7 | 3.3 | 2.56 | 2.8 | 2.76 | 2.68 | 2.66 | 2.71 | 2.6473 | 2.514 | 2.59 | 2.57 | 2.51 | 2.4709 | 4.9817 |  |  |  |  | #N/A
| 1 | 3 | 2.48 | 3.0223 | 2.8 | 2.87 | 2.824 | 2.68 | 3.23 | 2.58 | 2.96 | 2.72 | 2.69 | 2.64 | 2.72 | 2.7836 | 2.461 | 2.59 | 2.53 | 2.57 | 2.479 | 4.9881 |  |  |  |  | 16
| 1 | 4 | 2.496 | 3.0127 | 2.75 | 2.91 | 2.786 | 2.71 | 3.2 | 2.55 | 2.83 | 2.76 | 2.6 | 2.64 | 2.66 | 2.7052 | 2.482 | 2.62 | 2.57 | 2.53 | 2.5124 | 4.9684 |  |  |  |  | 2.7163
| 1 | 5 | 2.386 | 3.0412 | 2.94 | 2.8933 | 2.761 | 2.8 | 3.18 | 2.59 | 2.76 | 2.78 | 2.7 | 2.66 | 2.66 | 2.828 | 2.45 | 2.6 | 2.51 | 2.55 | 2.4717 | 4.9878 |  |  |  |  | 73
| 1 | 6 | 2.436 | 3.0358 | 3.14 | 2.8967 | 2.764 | 2.86 | 3.21 | 2.6 | 2.84 | 2.8 | 2.66 | 2.66 | 2.61 | 2.6923 | 2.472 | 2.61 | 2.56 | 2.57 | 2.4675 | 4.9893
| Mean |  | 2.451 | 3.0466 | 2.9133 | 2.8861 | 2.7915 | 2.7683 | 3.2317 | 2.575 | 2.8417 | 2.7567 | 2.665 | 2.6483 | 2.6783 | 2.723 | 2.4752 | 2.5967 | 2.545 | 2.54 | 2.4764 | 4.9838
| Median |  | 2.4565 | 3.0385 | 2.925 | 2.8833 | 2.793 | 2.755 | 3.22 | 2.575 | 2.835 | 2.76 | 2.67 | 2.65 | 2.685 | 2.6988 | 2.472 | 2.595 | 2.545 | 2.54 | 2.4713 | 4.9876
| Std Dev. |  | 0.040968 | 0.030584 | 0.1356 | 0.01625 | 0.02589 | 0.082077 | 0.04535 | 0.018708 | 0.067651 | 0.032042 | 0.035637 | 0.013292 | 0.042622 | 0.068413 | 0.021949 | 0.017512 | 0.0251 | 0.027568 | 0.019071 | 0.0079908
| Rel.Std.Dev. |  | 1.671% | 1.004% | 4.654% | 0.5631% | 0.9275% | 2.965% | 1.403% | 0.7265% | 2.381% | 1.162% | 1.337% | 0.5019% | 1.591% | 2.512% | 0.8868% | 0.6744% | 0.9862% | 1.085% | 0.7701% | 0.1603%
| PDM3 |  | -9.767% | 12.16% | 7.254% | 6.252% | 2.769% | 1.916% | 18.97% | -5.202% | 4.615% | 1.486% | -1.889% | -2.502% | -1.398% | 0.2468% | -8.877% | -4.404% | -6.306% | -6.49% | -8.832% | 83.48%
| Z-Score (Absolute) |  | 1.13 | 1.57 | 0.97 | 0.84 | 0.41 | 0.31 | 2.41 | 0.57 | 0.64 | 0.25 | 0.16 | 0.24 | 0.1 | 0.1 | 1.03 | 0.47 | 0.71 | 0.73 | 1.02 | 10.38
| Analytical results for As in OREAS 232b (Certified Value 455 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 464.42 | 475.89 | 436 | 441.3088 | 453 | 487 | 439 | 443.3 | 484 | 472 | 465 | 459 | 466.6 | 467.8645 | 441 | 401 | 522.4 | 485.2 | 421.1 | 425.8 | 426 | 691.9082 | 453.2066 |  | 1
| 1 | 2 | 459.52 | 478.06 | 436 | 436.01 | 454 | 502 | 452 | 453 | 467 | 465 | 483 | 455 | 480.7 | 473.389 | 442 | 401 | 487.45 | 490.3 | 438.9 | 416.2 | 442 | 665.7504 | 439.995 |  | 26
| 1 | 3 | 459.42 | 479.48 | 419 | 438.8029 | 452 | 494 | 442 | 446.4 | 501 | 470 | 477 | 456 | 460.8 | 440.8748 | 424 | 426 | 465.06 | 483.4 | 427.8 | 427.3 | 436 | 640.9717 | 432.048 |  | 16
| 1 | 4 | 459.37 | 472.4 | 416 | 440.8085 | 454 | 485 | 423 | 460.3 | 484 | 477 | 442 | 471 | 450 | 456.0115 | 438 | 418 | 505.7 | 494.1 | 438.6 | 430.6 | 426 | 640.6784 | 426.8099 |  | 455.0471
| 1 | 5 | 452.75 | 477.13 | 432 | 437.42 | 451 | 506 | 437 | 450.6 | 480 | 484 | 467 | 471 | 446.6 | 464.2715 | 446 | 418 | 464.38 | 487.7 | 427.9 | 435.3 | 429 | 667.8692 | 424.8522 |  | 74
| 1 | 6 | 468.75 | 482.92 | 442 | 442.81 | 450 | 493 | 447 | 455.2 | 478 | 475 | 452 | 468 | 459.4 | 466.199 | 434 | 409 | 486.35 | 486.5 | 437.7 | 415.7 | 443 | 661.8658 | 449.417
| Mean |  | 460.705 | 477.6467 | 430.1667 | 439.5267 | 452.3333 | 494.5 | 440 | 451.4667 | 482.3333 | 473.8333 | 464.3333 | 463.3333 | 460.6833 | 461.4351 | 437.5 | 412.1667 | 488.5567 | 487.8667 | 432 | 425.15 | 433.6667 | 661.5073 | 437.7215
| Median |  | 459.47 | 477.595 | 434 | 439.8057 | 452.5 | 493.5 | 440.5 | 451.8 | 482 | 473.5 | 466 | 463.5 | 460.1 | 465.2353 | 439.5 | 413.5 | 486.9 | 487.1 | 432.8 | 426.55 | 432.5 | 663.8081 | 436.0215
| Std Dev. |  | 5.4156 | 3.5276 | 10.3618 | 2.5663 | 1.633 | 8.2158 | 9.9599 | 6.1226 | 11.0755 | 6.4936 | 15.2796 | 7.5011 | 12.2459 | 11.5509 | 7.7395 | 10.1866 | 22.7243 | 3.8401 | 7.442 | 7.838 | 7.7632 | 19.1749 | 11.8204
| Rel.Std.Dev. |  | 1.176% | 0.7385% | 2.409% | 0.5839% | 0.361% | 1.661% | 2.264% | 1.356% | 2.296% | 1.37% | 3.291% | 1.619% | 2.658% | 2.503% | 1.769% | 2.471% | 4.651% | 0.7871% | 1.723% | 1.844% | 1.79% | 2.899% | 2.7%
| PDM3 |  | 1.243% | 4.966% | -5.468% | -3.411% | -0.5964% | 8.67% | -3.307% | -0.7868% | 5.996% | 4.128% | 2.041% | 1.821% | 1.239% | 1.404% | -3.856% | -9.423% | 7.364% | 7.212% | -5.065% | -6.57% | -4.699% | 45.37% | -3.807%
| Z-Score (Absolute) |  | 0 | 0.53 | 0.95 | 0.66 | 0.26 | 1.05 | 0.64 | 0.29 | 0.67 | 0.41 | 0.11 | 0.08 | 0 | 0.02 | 0.72 | 1.51 | 0.87 | 0.85 | 0.89 | 1.11 | 0.84 | 6.25 | 0.72
| Analytical results for B in OREAS 232b (Indicative Value 14.3 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [13 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [13 empty] | 1
| 1 | 1 | < 5 | 5 | 25.44 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 19 | [13 empty] | 1
| 1 | 2 | < 5 | 4 | 24.935 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 18 | [13 empty] | 4
| 1 | 3 | < 5 | 4 | 24.93 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 19 | [13 empty] | 16
| 1 | 4 | < 5 | 4 | 24.035 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 18 | [13 empty] | 14.3392
| 1 | 5 | < 5 | 5 | 23.622 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 18 | [13 empty] | 14
| 1 | 6 | < 5 | 4 | 25.18 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 18
| Mean |  |  | 4.3333 | 24.6903 |  |  | 10 |  |  |  |  | 18.3333
| Median |  |  | 4 | 24.9325 |  |  | 10 |  |  |  |  | 18
| Std Dev. |  |  | 0.5164 | 0.70567 |  |  | 0 |  |  |  |  | 0.5164
| Rel.Std.Dev. |  |  | 11.92% | 2.858% |  |  | 0% |  |  |  |  | 2.817%
| PDM3 |  |  | -69.78% | 72.19% |  |  | -30.26% |  |  |  |  | 27.85%
| Z-Score (Absolute) |  | 2.53 | 0.67 | 19.92 | 0 | 0 | 5.06 | 0 | 0 | 5.06 | 0 | 13.49
| Analytical results for Ba in OREAS 232b (Certified Value 118 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 102.01 | 129.6966 | 131 | 118.615 | 114 | 124 | 131 | 107 | 120 | 100 | 120 | 110 | 114.8 | 113.3249 | 111 | 123.3 | 91 | 120.34 | 127.1 | 115 | 116.01 | 604.132 |  |  | 1
| 1 | 2 | 103.4 | 129.3233 | 130 | 118.925 | 117 | 120 | 134 | 109 | 120 | 110 | 120 | 110 | 114.1 | 113.8503 | 111 | 484.5 | 90 | 128.17 | 121.9 | 115 | 116.23 | 613.1058 |  |  | 27
| 1 | 3 | 104.85 | 128.5437 | 129 | 119.335 | 115 | 116 | 135 | 111 | 120 | 100 | 120 | 110 | 113.6 | 111.1316 | 111 | 122 | 91 | 122.35 | 124.2 | 118 | 116.06 | 636.106 |  |  | 16
| 1 | 4 | 104.23 | 127.9664 | 128 | 118.875 | 113 | 117 | 130 | 109 | 120 | 110 | 110 | 110 | 112.5 | 111.1751 | 111 | 114.6 | 92 | 126.61 | 121.5 | 116 | 118.09 | 622.0797 |  |  | 117.8455
| 1 | 5 | 100.89 | 128.6622 | 124 | 119.795 | 116 | 121 | 129 | 111 | 120 | 110 | 120 | 110 | 111.9 | 115.2828 | 110 | 141.6 | 91 | 123.32 | 123.9 | 117 | 115.04 | 598.8708 |  |  | 75
| 1 | 6 | 105.07 | 131.456 | 138 | 119.51 | 114 | 124 | 131 | 113 | 120 | 110 | 120 | 110 | 110.8 | 106.6932 | 110 | 161.2 | 92 | 134.2 | 124.9 | 117 | 115.98 | 608.0535
| Mean |  | 103.4083 | 129.2747 | 130 | 119.1758 | 114.8333 | 120.3333 | 131.6667 | 110 | 120 | 106.6667 | 118.3333 | 110 | 112.95 | 111.9097 | 110.6667 | 191.2 | 91.1667 | 125.8317 | 123.9167 | 116.3333 | 116.235 | 613.7246
| Median |  | 103.815 | 128.9928 | 129.5 | 119.13 | 114.5 | 120.5 | 131 | 110 | 120 | 110 | 120 | 110 | 113.05 | 112.25 | 111 | 132.45 | 91 | 124.965 | 124.05 | 116.5 | 116.035 | 610.5797
| Std Dev. |  | 1.6621 | 1.2296 | 4.6043 | 0.44459 | 1.472 | 3.3862 | 2.3381 | 2.0976 | 0 | 5.164 | 4.0825 | 0 | 1.4896 | 3.0167 | 0.5164 | 144.6732 | 0.75277 | 4.9938 | 2.0537 | 1.2111 | 1.0015 | 13.5346
| Rel.Std.Dev. |  | 1.607% | 0.9512% | 3.542% | 0.3731% | 1.282% | 2.814% | 1.776% | 1.907% | 0% | 4.841% | 3.45% | 0% | 1.319% | 2.696% | 0.4666% | 75.67% | 0.8257% | 3.969% | 1.657% | 1.041% | 0.8616% | 2.205%
| PDM3 |  | -12.25% | 9.698% | 10.31% | 1.129% | -2.556% | 2.111% | 11.73% | -6.657% | 1.828% | -9.486% | 0.414% | -6.657% | -4.154% | -5.037% | -6.092% | 62.25% | -22.64% | 6.777% | 5.152% | -1.283% | -1.367% | 420.8%
| Z-Score (Absolute) |  | 1.34 | 1.15 | 1.22 | 0.18 | 0.24 | 0.29 | 1.38 | 0.71 | 0.26 | 1.03 | 0.1 | 0.71 | 0.42 | 0.52 | 0.64 | 7.12 | 2.52 | 0.82 | 0.63 | 0.1 | 0.11 | 47.82
| Analytical results for Be in OREAS 232b (Certified Value 1.26 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.24 | 1.609 | 1.3 | < 5 | 1.2 | 1.42 | 1.42 | 1 | 1.33 | 1.25 | 1.35 | 1.24 | 1.4542 | 1.3 | 1.323 | 1.2 | 1.1 | < 2 | 1.26 | 0.999 | 1.0797 | 1.1357 |  |  | 1
| 1 | 2 | 1.26 | 1.5 | 1.4 | < 5 | 1.2 | 1.38 | 1.39 | 1 | 1.31 | 1.24 | 1.32 | 1.23 | 1.3188 | 1.3 | 1.264 | 1.4 | 1.11 | < 2 | 1.18 | 1.006 | 1.1015 | 1.2116 |  |  | 28
| 1 | 3 | 1.28 | 1.6644 | 1.3 | < 5 | 1.2 | 1.32 | 1.41 | 1 | 1.4 | 1.26 | 1.32 | 1.31 | 1.3446 | 1.3 | 1.244 | 1.4 | 1.13 | < 2 | 1.19 | 1.011 | 1.0799 | 1.1807 |  |  | 16
| 1 | 4 | 1.21 | 1.5181 | 1.3 | < 5 | 1.2 | 1.35 | 1.39 | 1 | 1.29 | 1.24 | 1.31 | 1.29 | 1.3072 | 1.4 | 1.234 | 1.4 | 1.17 | < 2 | 1.16 | 1.026 | 0.9502 | 1.1649 |  |  | 1.2557
| 1 | 5 | 1.2 | 1.5375 | 1.3 | < 5 | 1.2 | 1.35 | 1.5 | 1 | 1.35 | 1.27 | 1.33 | 1.34 | 1.3143 | 1.3 | 1.273 | 1.4 | 1.17 | < 2 | 1.18 | 1.01 | 1.0802 | 1.2716 |  |  | 76
| 1 | 6 | 1.22 | 1.8351 | 1.2 | < 5 | 1.2 | 1.38 | 1.49 | 1 | 1.4 | 1.28 | 1.38 | 1.27 | 1.2978 | 1.3 | 1.266 | 1.5 | 1.16 | < 2 | 1.15 | 1.005 | 0.9287 | 1.279
| Mean |  | 1.235 | 1.6107 | 1.3 |  | 1.2 | 1.3667 | 1.4333 | 1 | 1.3467 | 1.2567 | 1.335 | 1.28 | 1.3395 | 1.3167 | 1.2673 | 1.3833 | 1.14 |  | 1.1867 | 1.0095 | 1.0367 | 1.2072
| Median |  | 1.23 | 1.5732 | 1.3 |  | 1.2 | 1.365 | 1.415 | 1 | 1.34 | 1.255 | 1.325 | 1.28 | 1.3166 | 1.3 | 1.265 | 1.4 | 1.145 |  | 1.18 | 1.008 | 1.0798 | 1.1961
| Std Dev. |  | 0.030822 | 0.12603 | 0.063246 |  | 0 | 0.034448 | 0.049261 | 0 | 0.045898 | 0.01633 | 0.025884 | 0.041952 | 0.058357 | 0.040825 | 0.030982 | 0.098319 | 0.030984 |  | 0.038816 | 0.0091378 | 0.076096 | 0.058204
| Rel.Std.Dev. |  | 2.496% | 7.824% | 4.865% |  | 0% | 2.521% | 3.437% | 0% | 3.408% | 1.299% | 1.939% | 3.278% | 4.357% | 3.101% | 2.445% | 7.107% | 2.718% |  | 3.271% | 0.9052% | 7.34% | 4.821%
| PDM3 |  | -1.649% | 28.27% | 3.528% |  | -4.436% | 8.837% | 14.15% | -20.36% | 7.244% | 0.07671% | 6.315% | 1.935% | 6.673% | 4.855% | 0.9262% | 10.16% | -9.214% |  | -5.498% | -19.61% | -17.44% | -3.859%
| Z-Score (Absolute) |  | 0.32 | 2.83 | 0.22 | 10.29 | 0.62 | 0.78 | 1.34 | 2.3 | 0.61 | 0.14 | 0.51 | 0.05 | 0.55 | 0.36 | 0.05 | 0.92 | 1.12 | 2.3 | 0.73 | 2.22 | 1.99 | 0.56
| Analytical results for Bi in OREAS 232b (Certified Value 0.3 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.29 | 0.33251 | 0.72 | < 2 | 0.38 | 0.28 | 0.32 | 0.3 | 0.3 | 0.32 | 0.26 | 0.34 | 0.29 | 0.3313 | 5 | 0.4 | 0.3 | 49.75 | < 5 | 0.34 | 0.4884 | 1.4103 |  |  | 1
| 1 | 2 | 0.27 | 0.33739 | 0.31 | < 2 | 0.41 | 0.31 | 0.32 | 0.3 | 0.29 | 0.28 | 0.27 | 0.31 | 0.27 | 0.29269 | 4 | 0.345 | 0.32 | 28.9 | < 5 | 0.29 | 0.6287 | 1.1198 |  |  | 29
| 1 | 3 | 0.28 | 0.35939 | 0.29 | < 2 | 0.37 | 0.27 | 0.31 | 0.3 | 0.3 | 0.42 | 0.3 | 0.29 | 0.28 | 0.31533 | 3 | 0.371 | 0.34 | 35.62 | < 5 | 0.31 | 0.5427 | 1.8797 |  |  | 16
| 1 | 4 | 0.31 | 0.34634 | 0.28 | < 2 | 0.41 | 0.3 | 0.3 | 0.3 | 0.3 | 0.29 | 0.27 | 0.29 | 0.26 | 0.31909 | 5 | 0.313 | 0.29 | 38.22 | < 5 | 0.31 | 0.4342 | 1.9441 |  |  | 0.29887
| 1 | 5 | 0.27 | 0.32501 | 0.29 | < 2 | 0.39 | 0.31 | 0.3 | 0.3 | 0.28 | 0.29 | 0.29 | 0.3 | 0.27 | 0.2906 | 4 | 0.55 | 0.29 | 43.43 | < 5 | 0.3 | 0.6105 | 1.6696 |  |  | 77
| 1 | 6 | 0.29 | 0.3345 | 0.32 | < 2 | 0.36 | 0.27 | 0.32 | 0.3 | 0.27 | 0.29 | 0.29 | 0.29 | 0.28 | 0.32966 | 3 | 0.419 | 0.29 | 31.04 | < 5 | 0.3 | 0.5112 | 1.9385
| Mean |  | 0.285 | 0.33919 | 0.36833 |  | 0.38667 | 0.29 | 0.31167 | 0.3 | 0.29 | 0.315 | 0.28 | 0.30333 | 0.275 | 0.31311 | 4 | 0.39967 | 0.305 | 37.8267 |  | 0.30833 | 0.53595 | 1.6603
| Median |  | 0.285 | 0.33594 | 0.3 |  | 0.385 | 0.29 | 0.315 | 0.3 | 0.295 | 0.29 | 0.28 | 0.295 | 0.275 | 0.31721 | 4 | 0.3855 | 0.295 | 36.92 |  | 0.305 | 0.52695 | 1.7746
| Std Dev. |  | 0.015166 | 0.012085 | 0.17291 |  | 0.020656 | 0.018974 | 0.0098319 | 0 | 0.012649 | 0.053198 | 0.015492 | 0.019664 | 0.010488 | 0.017712 | 0.89443 | 0.082819 | 0.020736 | 7.8017 |  | 0.017224 | 0.074071 | 0.3349
| Rel.Std.Dev. |  | 5.321% | 3.563% | 46.94% |  | 5.342% | 6.543% | 3.155% | 0% | 4.362% | 16.89% | 5.533% | 6.483% | 3.814% | 5.657% | 22.36% | 20.72% | 6.799% | 20.62% |  | 5.586% | 13.82% | 20.17%
| PDM3 |  | -4.641% | 13.49% | 23.24% |  | 29.38% | -2.968% | 4.282% | 0.3783% | -2.968% | 5.397% | -6.314% | 1.494% | -7.987% | 4.765% | 1238% | 33.73% | 2.051% | 12557% |  | 3.167% | 79.33% | 455.5%
| Z-Score (Absolute) |  | 0.62 | 0.54 | 1.16 | 14.66 | 1.55 | 0.51 | 0.05 | NA | 0.51 | 0.02 | 0.73 | 0.23 | 0.83 | 0.02 | NA | 1.83 | 0.19 | 801.61 | 46.71 | 0.12 | 4.74 | 28.77
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 232b (Certified Value 0.337 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.317 | 0.34082 | 0.4 | 0.375 | 0.34 | 0.35 | 0.36 | 0.35 | 0.36 | 0.33 | 0.34 | 0.32 | 0.33 | 0.36144 | 0.314 | 0.3 | 0.32557 | 0.32 | 0.33 | 0.2903 |  |  |  |  | 1
| 1 | 2 | 0.314 | 0.33865 | 0.4 | 0.376 | 0.35 | 0.33 | 0.35 | 0.34 | 0.35 | 0.34 | 0.34 | 0.33 | 0.34 | 0.34684 | 0.314 | 0.3 | 0.33408 | 0.31 | 0.33 | 0.287 |  |  |  |  | #N/A
| 1 | 3 | 0.323 | 0.34129 | 0.39 | 0.374 | 0.35 | 0.34 | 0.35 | 0.35 | 0.37 | 0.34 | 0.35 | 0.32 | 0.33 | 0.35608 | 0.307 | 0.3 | 0.33278 | 0.32 | 0.34 | 0.2875 |  |  |  |  | 16
| 1 | 4 | 0.316 | 0.33711 | 0.39 | 0.374 | 0.34 | 0.34 | 0.34 | 0.34 | 0.36 | 0.34 | 0.34 | 0.33 | 0.33 | 0.34298 | 0.314 | 0.3 | 0.33659 | 0.32 | 0.33 | 0.2903 |  |  |  |  | 0.33697
| 1 | 5 | 0.313 | 0.33939 | 0.42 | 0.37067 | 0.34 | 0.35 | 0.34 | 0.35 | 0.35 | 0.34 | 0.35 | 0.33 | 0.33 | 0.36291 | 0.314 | 0.3 | 0.34203 | 0.33 | 0.33 | 0.2887 |  |  |  |  | 78
| 1 | 6 | 0.325 | 0.33829 | 0.4 | 0.375 | 0.35 | 0.35 | 0.35 | 0.34 | 0.36 | 0.34 | 0.34 | 0.33 | 0.33 | 0.35944 | 0.314 | 0.3 | 0.33372 | 0.32 | 0.33 | 0.2896
| Mean |  | 0.318 | 0.33926 | 0.4 | 0.37411 | 0.345 | 0.34333 | 0.34833 | 0.345 | 0.35833 | 0.33833 | 0.34333 | 0.32667 | 0.33167 | 0.35495 | 0.31283 | 0.3 | 0.33413 | 0.32 | 0.33167 | 0.2889
| Median |  | 0.3165 | 0.33902 | 0.4 | 0.3745 | 0.345 | 0.345 | 0.35 | 0.345 | 0.36 | 0.34 | 0.34 | 0.33 | 0.33 | 0.35776 | 0.314 | 0.3 | 0.3339 | 0.32 | 0.33 | 0.28915
| Std Dev. |  | 0.004899 | 0.0015831 | 0.010954 | 0.0018459 | 0.0054772 | 0.008165 | 0.0075277 | 0.0054772 | 0.0075277 | 0.0040825 | 0.005164 | 0.005164 | 0.0040825 | 0.0081977 | 0.0028577 | 0 | 0.0053627 | 0.0063246 | 0.0040825 | 0.0014156
| Rel.Std.Dev. |  | 1.541% | 0.4666% | 2.739% | 0.4934% | 1.588% | 2.378% | 2.161% | 1.588% | 2.101% | 1.207% | 1.504% | 1.581% | 1.231% | 2.31% | 0.9135% | 0% | 1.605% | 1.976% | 1.231% | 0.49%
| PDM3 |  | -5.628% | 0.6801% | 18.71% | 11.02% | 2.384% | 1.89% | 3.373% | 2.384% | 6.341% | 0.4058% | 1.89% | -3.056% | -1.573% | 5.336% | -7.162% | -10.97% | -0.8419% | -5.035% | -1.573% | -14.26%
| Z-Score (Absolute) |  | 1.29 | 0.03 | 3.81 | 2.2 | 0.39 | 0.28 | 0.59 | 0.39 | 1.22 | 0.03 | 0.28 | 0.75 | 0.44 | 1.01 | 1.59 | 2.41 | 0.29 | 1.17 | 0.44 | 3.11
| Analytical results for Cd in OREAS 232b (Certified Value 0.045 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.03 | 1.13 | 0.05 | 0.58917 | 0.07 | 0.09 | 0.06 | < 0.1 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | 0.055048 | < 1 | 0.142 | 0.04 | 18.35 | 8.4 | 0.04 | 0.81 | 0.3402 |  |  | 1
| 1 | 2 | 0.033 | 1.12 | 0.02 | 0.59785 | 0.05 | 0.07 | 0.07 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.05 | 0.060332 | < 1 | < 0.1 | 0.03 | 16.38 | 8.3 | 0.03 | 0.76 | 0.3453 |  |  | 5
| 1 | 3 | 0.033 | 1.09 | 0.06 | 0.57 | 0.07 | 0.08 | 0.07 | < 0.1 | 0.04 | 0.03 | 0.03 | 0.03 | 0.05 | 0.054634 | < 1 | < 0.1 | 0.04 | 16.6 | 9 | 0.04 | 0.67 | 0.3402 |  |  | 16
| 1 | 4 | 0.034 | 1.13 | 0.03 | 0.58 | 0.06 | 0.08 | 0.06 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.056754 | < 1 | < 0.1 | 0.04 | 17.44 | 8.3 | 0.04 | 0.74 | 0.3936 |  |  | 0.044859
| 1 | 5 | 0.03 | 1.12 | 0.01 | 0.60208 | 0.06 | 0.08 | 0.07 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.03 | 0.04 | 0.04872 | < 1 | < 0.1 | 0.05 | 16.48 | 8.4 | 0.04 | 0.83 | 0.3302 |  |  | 79
| 1 | 6 | 0.034 | 1.14 | 0.04 | 0.59313 | 0.06 | 0.08 | 0.06 | < 0.1 | 0.04 | 0.03 | 0.03 | 0.03 | 0.05 | 0.049513 | < 1 | < 0.1 | 0.03 | 16.97 | 8.2 | 0.04 | 0.86 | 0.348
| Mean |  | 0.032333 | 1.1217 | 0.035 | 0.58871 | 0.061667 | 0.08 | 0.065 |  | 0.035 | 0.033333 | 0.03 | 0.033333 | 0.046667 | 0.054167 |  | 0.142 | 0.038333 | 17.0367 | 8.4333 | 0.038333 | 0.77833 | 0.34958
| Median |  | 0.033 | 1.125 | 0.035 | 0.59115 | 0.06 | 0.08 | 0.065 |  | 0.035 | 0.03 | 0.03 | 0.03 | 0.05 | 0.054841 |  | 0.142 | 0.04 | 16.785 | 8.35 | 0.04 | 0.785 | 0.34275
| Std Dev. |  | 0.0018619 | 0.017224 | 0.018708 | 0.0119 | 0.0075277 | 0.0063246 | 0.0054772 |  | 0.0054772 | 0.005164 | 0 | 0.005164 | 0.005164 | 0.0044052 |  |  | 0.0075277 | 0.7515 | 0.28752 | 0.0040825 | 0.069113 | 0.022406
| Rel.Std.Dev. |  | 5.758% | 1.536% | 53.45% | 2.021% | 12.21% | 7.906% | 8.427% |  | 15.65% | 15.49% | 0% | 15.49% | 11.07% | 8.133% |  |  | 19.64% | 4.411% | 3.409% | 10.65% | 8.88% | 6.409%
| PDM3 |  | -27.92% | 2400% | -21.98% | 1212% | 37.47% | 78.34% | 44.9% |  | -21.98% | -25.69% | -33.12% | -25.69% | 4.03% | 20.75% |  | 216.5% | -14.55% | 37878% | 18700% | -14.55% | 1635% | 679.3%
| Z-Score (Absolute) |  | 0.73 | 30.2 | 0.65 | 15.07 | 0.11 | 0.63 | 0.2 | 0.22 | 0.65 | 0.7 | 0.79 | 0.7 | 0.32 | 0.11 | 12.55 | 0.21 | 0.56 | 482.06 | 237.79 | 0.56 | 20.45 | 8.28
| Analytical results for Ce in OREAS 232b (Certified Value 52 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 51.721 | 82.017 | 54.1 | 64.05 | 44.68 | 67.5 | 51 | 54 | 50.5 | 52.7 | 52.6 | 52.9594 | 52.15 | 52.53 | 50.4 | 38.2935 | 0.2196 | [7 empty] | 1
| 1 | 2 | 51.37 | 82.06 | 54.8 | 66.4 | 48.59 | 71.6 | 51 | 52.7 | 49.8 | 54.4 | 53.2 | 51.7322 | 54.45 | 53.15 | 48.28 | 46.0269 | 0.2308 | [7 empty] | 31
| 1 | 3 | 51.908 | 82.2859 | 53.5 | 64.75 | 43.55 | 71 | 53 | 54.2 | 52.1 | 54 | 53.9 | 50.3552 | 56.48 | 51.65 | 48.85 | 45.5314 | 0.2278 | [7 empty] | 16
| 1 | 4 | 52.527 | 82.3129 | 54.2 | 65.38 | 46.52 | 66 | 51 | 52.7 | 51.3 | 51.7 | 53.9 | 52.5373 | 53.41 | 53.61 | 48.03 | 38.1927 | 0.2355 | [7 empty] | 52.2172
| 1 | 5 | 51.706 | 82.6972 | 52.5 | 66.49 | 47.37 | 71.9 | 53 | 51.7 | 52.8 | 54.9 | 55.2 | 54.3991 | 58.2 | 55.38 | 49.91 | 45.0034 | 0.2257 | [7 empty] | 80
| 1 | 6 | 51.279 | 82.3904 | 52.7 | 63.95 | 49.91 | 69.8 | 52 | 52.8 | 49 | 54.6 | 53.9 | 52.1213 | 57.98 | 55.53 | 49.15 | 36.9463 | 0.2321
| Mean |  | 51.7518 | 82.2939 | 53.6333 | 65.17 | 46.77 | 69.6333 | 51.8333 | 53.0167 | 50.9167 | 53.7167 | 53.7833 | 52.3508 | 55.445 | 53.6417 | 49.1033 | 41.6657 | 0.22858
| Median |  | 51.7135 | 82.2994 | 53.8 | 65.065 | 46.945 | 70.4 | 51.5 | 52.75 | 50.9 | 54.2 | 53.9 | 52.3293 | 55.465 | 53.38 | 49 | 41.6484 | 0.2293
| Std Dev. |  | 0.4467 | 0.2465 | 0.90259 | 1.1152 | 2.3804 | 2.3939 | 0.98319 | 0.93256 | 1.4275 | 1.2513 | 0.87044 | 1.343 | 2.493 | 1.5513 | 0.9196 | 4.2617 | 0.0055632
| Rel.Std.Dev. |  | 0.8632% | 0.2995% | 1.683% | 1.711% | 5.09% | 3.438% | 1.897% | 1.759% | 2.804% | 2.329% | 1.618% | 2.565% | 4.496% | 2.892% | 1.873% | 10.23% | 2.434%
| PDM3 |  | -0.8912% | 57.6% | 2.712% | 24.81% | -10.43% | 33.35% | -0.7352% | 1.531% | -2.491% | 2.872% | 2.999% | 0.2558% | 6.181% | 2.728% | -5.963% | -20.21% | -99.56%
| Z-Score (Absolute) |  | 0.41 | 9.4 | 0.2 | 3.9 | 2.01 | 5.34 | 0.38 | 0 | 0.67 | 0.22 | 0.25 | 0.21 | 0.78 | 0.2 | 1.26 | 3.64 | 16.95
| Analytical results for Co in OREAS 232b (Certified Value 15.4 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 15 | 16.245 | 15.5 | 16.9628 | 15 | 16 | 15.7 | 13.8 | 15.6 | 15 | 15.4 | 14.7 | 15.3 | 16.2223 | 13 | 15.66 | 14.2 | 20.53 | 19.7 | 15.7 | 16.815 | 13.8192 | 14.9654 |  | 1
| 1 | 2 | 15.14 | 16.281 | 16.4 | 16.93 | 15.5 | 15 | 16.1 | 13.5 | 14.9 | 15.1 | 15.6 | 14.4 | 15.8 | 15.6227 | 13 | 14.15 | 14.5 | 24.58 | 19.6 | 14.8 | 16.53 | 14.5565 | 15.0861 |  | 32
| 1 | 3 | 14.98 | 16.406 | 15.7 | 16.9269 | 15.2 | 16 | 16.3 | 13.7 | 15.8 | 15.1 | 15.2 | 14.8 | 15 | 16.5333 | 14 | 13.53 | 15.5 | 20.14 | 20.4 | 14.8 | 17.005 | 13.638 | 14.9584 |  | 16
| 1 | 4 | 15.33 | 15.903 | 15.8 | 17.0312 | 15.2 | 15 | 14.8 | 13.3 | 15.3 | 14.8 | 14.9 | 14.3 | 16 | 16.8374 | 13 | 13.24 | 15.6 | 23.51 | 19.4 | 14.5 | 17.005 | 13.1794 | 14.8984 |  | 15.3523
| 1 | 5 | 14.9 | 16.542 | 16.5 | 16.9467 | 15.5 | 15 | 16.3 | 13.7 | 15.5 | 15.5 | 15.5 | 15 | 15.4 | 16.5258 | 13 | 14 | 15.4 | 18.61 | 19.4 | 14.8 | 16.72 | 14.1008 | 15.0894 |  | 81
| 1 | 6 | 15.58 | 16.659 | 16.5 | 17.0375 | 15.2 | 15 | 16.2 | 13.9 | 15.2 | 15.3 | 15.4 | 14.4 | 15.8 | 15.6871 | 13 | 14.25 | 16 | 20.91 | 19.5 | 14.7 | 17.29 | 14.5623 | 14.749
| Mean |  | 15.155 | 16.3393 | 16.0667 | 16.9725 | 15.2667 | 15.3333 | 15.9 | 13.65 | 15.3833 | 15.1333 | 15.3333 | 14.6 | 15.55 | 16.2381 | 13.1667 | 14.1383 | 15.2 | 21.38 | 19.6667 | 14.8833 | 16.8942 | 13.976 | 14.9578
| Median |  | 15.07 | 16.3435 | 16.1 | 16.9547 | 15.2 | 15 | 16.15 | 13.7 | 15.4 | 15.1 | 15.4 | 14.55 | 15.6 | 16.374 | 13 | 14.075 | 15.45 | 20.72 | 19.55 | 14.8 | 16.91 | 13.96 | 14.9619
| Std Dev. |  | 0.25735 | 0.26472 | 0.45019 | 0.049619 | 0.19664 | 0.5164 | 0.5831 | 0.21679 | 0.31885 | 0.24221 | 0.25033 | 0.27568 | 0.37815 | 0.49227 | 0.40825 | 0.83956 | 0.6957 | 2.233 | 0.37771 | 0.41673 | 0.26475 | 0.54216 | 0.12723
| Rel.Std.Dev. |  | 1.698% | 1.62% | 2.802% | 0.2924% | 1.288% | 3.368% | 3.667% | 1.588% | 2.073% | 1.601% | 1.633% | 1.888% | 2.432% | 3.032% | 3.101% | 5.938% | 4.577% | 10.44% | 1.921% | 2.8% | 1.567% | 3.879% | 0.8506%
| PDM3 |  | -1.285% | 6.429% | 4.653% | 10.55% | -0.5581% | -0.1239% | 3.567% | -11.09% | 0.2018% | -1.427% | -0.1239% | -4.901% | 1.287% | 5.77% | -14.24% | -7.908% | -0.9923% | 39.26% | 28.1% | -3.055% | 10.04% | -8.965% | -2.57%
| Z-Score (Absolute) |  | 0.16 | 0.93 | 0.67 | 1.51 | 0.06 | NA | 0.52 | 1.55 | 0.05 | 0.18 | 0 | 0.67 | 0.2 | 0.83 | NA | 1.1 | 0.12 | 5.56 | 3.98 | 0.41 | 1.44 | 1.25 | 0.35
Results from laboratories 1.09 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 232b (Certified Value 109 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 114.7 | 104.88 | 97 | 114.3533 | 100 | 114 | 113 | 105 | 118 | 112 | 110 | 104 | 111.6 | 101.7708 | 107 | 132.5 | 103 | 101.93 | 110 | 104.4 | 82.76 | 110.8002 |  |  | 1
| 1 | 2 | 114.8 | 104.6615 | 102 | 113.9233 | 105 | 110 | 115 | 106 | 115 | 114 | 112 | 105 | 108.7 | 102.6787 | 108 | 122.6 | 103 | 110.16 | 111 | 103.9 | 77.181 | 111.4168 |  |  | 33
| 1 | 3 | 115.2 | 105.336 | 97 | 113.6933 | 103 | 111 | 114 | 111 | 122 | 114 | 112 | 105 | 108.4 | 110.2923 | 106 | 123 | 102 | 103.55 | 112 | 104.8 | 82.2923 | 114.1715 |  |  | 16
| 1 | 4 | 112.4 | 102.98 | 97 | 114.5033 | 102 | 110 | 113 | 105 | 117 | 115 | 109 | 106 | 107.1 | 106.7866 | 108 | 117.3 | 103 | 107.56 | 111 | 106.5 | 87.4035 | 110.4758 |  |  | 108.6305
| 1 | 5 | 112.2 | 106.875 | 103 | 114.38 | 106 | 113 | 112 | 110 | 116 | 116 | 111 | 107 | 106.9 | 105.6528 | 106 | 125.1 | 103 | 103.06 | 111 | 104.2 | 86.737 | 112.6382 |  |  | 82
| 1 | 6 | 115.8 | 105.6115 | 102 | 115.3367 | 102 | 114 | 113 | 112 | 116 | 115 | 110 | 105 | 107.2 | 101.9739 | 106 | 120.2 | 103 | 112.6 | 111 | 104.7 | 86.0705 | 110.65
| Mean |  | 114.1833 | 105.0573 | 99.6667 | 114.365 | 103 | 112 | 113.3333 | 108.1667 | 117.3333 | 114.3333 | 110.6667 | 105.3333 | 108.3167 | 104.8592 | 106.8333 | 123.45 | 102.8333 | 106.4767 | 111 | 104.75 | 83.7407 | 111.6921
| Median |  | 114.75 | 105.108 | 99.5 | 114.3667 | 102.5 | 112 | 113 | 108 | 116.5 | 114.5 | 110.5 | 105 | 107.8 | 104.1657 | 106.5 | 122.8 | 103 | 105.555 | 111 | 104.55 | 84.4153 | 111.1085
| Std Dev. |  | 1.5105 | 1.28 | 2.9439 | 0.56742 | 2.1909 | 1.8974 | 1.0328 | 3.1885 | 2.5033 | 1.3663 | 1.2111 | 1.0328 | 1.7702 | 3.361 | 0.98319 | 5.1756 | 0.40825 | 4.3162 | 0.63246 | 0.91815 | 3.8448 | 1.4477
| Rel.Std.Dev. |  | 1.323% | 1.218% | 2.954% | 0.4962% | 2.127% | 1.694% | 0.9113% | 2.948% | 2.134% | 1.195% | 1.094% | 0.9805% | 1.634% | 3.205% | 0.9203% | 4.192% | 0.397% | 4.054% | 0.5698% | 0.8765% | 4.591% | 1.296%
| PDM3 |  | 5.112% | -3.289% | -8.252% | 5.279% | -5.183% | 3.102% | 4.329% | -0.427% | 8.011% | 5.25% | 1.874% | -3.035% | -0.2889% | -3.472% | -1.654% | 13.64% | -5.337% | -1.983% | 2.181% | -3.572% | -22.91% | 2.818%
| Z-Score (Absolute) |  | 1.11 | 0.59 | 1.6 | 1.14 | 0.98 | 0.7 | 0.95 | 0.01 | 1.69 | 1.13 | 0.45 | 0.54 | 0.01 | 0.63 | 0.26 | 2.83 | 1.01 | 0.33 | 0.51 | 0.65 | 4.56 | 0.64
| Analytical results for Cs in OREAS 232b (Certified Value 6.75 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.20 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 6.28 | 5.8374 | 6.6 | 6.51 | 6.79 | 7.68 | 6.1 | 7.53 | 6.95 | 6.65 | 6.61 | 6.749 | 6.36 | 7.29 | 6.7882 | [9 empty] | 1
| 1 | 2 | 6.29 | 5.6056 | 6.78 | 6.82 | 6.76 | 7.78 | 6.2 | 7.24 | 6.69 | 6.66 | 6.69 | 6.484 | 6.42 | 6.92 | 7.8563 | [9 empty] | 34
| 1 | 3 | 6.46 | 5.6861 | 6.44 | 6.61 | 6.69 | 7.96 | 6.3 | 7.62 | 6.94 | 6.6 | 6.86 | 6.567 | 6.65 | 7.05 | 7.8251 | [9 empty] | 16
| 1 | 4 | 6.46 | 5.533 | 6.11 | 6.69 | 6.92 | 7.32 | 6.2 | 7.24 | 6.93 | 6.52 | 6.83 | 6.234 | 6.82 | 6.88 | 6.6874 | [9 empty] | 6.7465
| 1 | 5 | 6.28 | 5.5721 | 6.61 | 6.79 | 6.96 | 7.95 | 6.4 | 7.26 | 7 | 6.73 | 6.8 | 6.539 | 6.48 | 6.97 | 7.7817 | [9 empty] | 83
| 1 | 6 | 6.48 | 5.4563 | 7.04 | 6.61 | 6.9 | 7.81 | 6.4 | 7.4 | 6.85 | 6.73 | 6.62 | 6.343 | 6.89 | 6.91 | 6.4059
| Mean |  | 6.375 | 5.6151 | 6.5967 | 6.6717 | 6.8367 | 7.75 | 6.2667 | 7.3817 | 6.8933 | 6.6483 | 6.735 | 6.486 | 6.6033 | 7.0033 | 7.2241
| Median |  | 6.375 | 5.5888 | 6.605 | 6.65 | 6.845 | 7.795 | 6.25 | 7.33 | 6.935 | 6.655 | 6.745 | 6.5115 | 6.565 | 6.945 | 7.285
| Std Dev. |  | 0.10075 | 0.13291 | 0.3136 | 0.11839 | 0.10558 | 0.23597 | 0.12111 | 0.16376 | 0.11075 | 0.080353 | 0.10932 | 0.18028 | 0.21878 | 0.15253 | 0.66623
| Rel.Std.Dev. |  | 1.58% | 2.367% | 4.754% | 1.775% | 1.544% | 3.045% | 1.933% | 2.218% | 1.607% | 1.209% | 1.623% | 2.779% | 3.313% | 2.178% | 9.222%
| PDM3 |  | -5.507% | -16.77% | -2.222% | -1.11% | 1.336% | 14.87% | -7.113% | 9.414% | 2.176% | -1.456% | -0.1711% | -3.862% | -2.123% | 3.806% | 7.078%
| Z-Score (Absolute) |  | 0.9 | 3.21 | 0.23 | 0 | 0.5 | 3.28 | 1.23 | 2.16 | 0.67 | 0.07 | 0.19 | 0.56 | 0.21 | 1.01 | 1.68
| Analytical results for Cu in OREAS 232b (Certified Value 25.4 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 25.76 | 22.7312 | 26.1 | 27.4 | 21 | 24.5 | 27 | 23.7 | 298 | 24.8 | 25.2 | 25.9 | 27.41 | 27.5534 | 27 | 31.5 | 23.8 | 22.6 | 31 | 27.2 | 26.6665 | 25 | 1 | 25.7235 | 1
| 1 | 2 | 25.49 | 23.1374 | 26.4 | 26.46 | 22 | 26.3 | 25 | 24.4 | 61.2 | 24.5 | 25.5 | 25.7 | 26.79 | 27.8991 | 27 | 27.72 | 23.5 | 23.33 | 30.6 | 25.5 | 26.353 | 24 | 1 | 25.8131 | 35
| 1 | 3 | 26 | 22.5457 | 25.6 | 25.84 | 21 | 25.4 | 25 | 26.5 | 56.7 | 25.4 | 25.5 | 26.3 | 26.12 | 27.5939 | 26 | 24.95 | 23.6 | 22.37 | 31.4 | 25.5 | 26.448 | 24 | 1 | 25.7052 | 16
| 1 | 4 | 24.88 | 22.5051 | 24.9 | 24.8367 | 22 | 25.7 | 23 | 26.8 | 33.7 | 26.1 | 25 | 26.2 | 25.01 | 28.4899 | 27 | 28.66 | 24.2 | 24 | 31.5 | 25.3 | 26.41 | 24 | 1 | 25.5678 | 25.4034
| 1 | 5 | 25.01 | 21.9093 | 26.1 | 24.02 | 22 | 25.2 | 24 | 23.3 | 30.4 | 25.7 | 25.8 | 26.6 | 24.71 | 26.7628 | 26 | 26.09 | 23.5 | 23.59 | 31.5 | 25.8 | 26.125 | 25 |  | 26.1899 | 84
| 1 | 6 | 25.86 | 22.4708 | 26.3 | 24.5967 | 23 | 25.3 | 23 | 24.2 | 31.4 | 25.2 | 26.1 | 26 | 24.85 | 28.7257 | 26 | 27.5 | 24.1 | 25.25 | 29.2 | 25.9 | 26.809 | 25 | 1 | 25.4803
| Mean |  | 25.5 | 22.5499 | 25.9 | 25.5256 | 21.8333 | 25.4 | 24.5 | 24.8167 | 85.2333 | 25.2833 | 25.5167 | 26.1167 | 25.815 | 27.8375 | 26.5 | 27.7367 | 23.7833 | 23.5233 | 30.8667 | 25.8667 | 26.4686 | 24.5 | 1 | 25.7466
| Median |  | 25.625 | 22.5254 | 26.1 | 25.3383 | 22 | 25.35 | 24.5 | 24.3 | 45.2 | 25.3 | 25.5 | 26.1 | 25.565 | 27.7465 | 26.5 | 27.61 | 23.7 | 23.46 | 31.2 | 25.65 | 26.429 | 24.5 | 1 | 25.7144
| Std Dev. |  | 0.46299 | 0.39914 | 0.56214 | 1.272 | 0.75277 | 0.5933 | 1.5166 | 1.4743 | 105.0922 | 0.58452 | 0.39707 | 0.31885 | 1.1303 | 0.709 | 0.54772 | 2.2605 | 0.30605 | 1.0424 | 0.88919 | 0.68896 | 0.24096 | 0.54772 | 0 | 0.24743
| Rel.Std.Dev. |  | 1.816% | 1.77% | 2.17% | 4.983% | 3.448% | 2.336% | 6.19% | 5.941% | 123.3% | 2.312% | 1.556% | 1.221% | 4.378% | 2.547% | 2.067% | 8.15% | 1.287% | 4.431% | 2.881% | 2.664% | 0.9104% | 2.236% | 0% | 0.961%
| PDM3 |  | 0.3804% | -11.23% | 1.955% | 0.481% | -14.05% | -0.01321% | -3.556% | -2.309% | 235.5% | -0.4725% | 0.446% | 2.808% | 1.62% | 9.582% | 4.317% | 9.185% | -6.377% | -7.401% | 21.51% | 1.824% | 4.193% | -3.556% | -96.06% | 1.351%
| Z-Score (Absolute) |  | 0.01 | 2.08 | 0.27 | 0 | 2.58 | 0.08 | 0.71 | 0.49 | 41.8 | 0.17 | 0 | 0.42 | 0.21 | 1.62 | 0.69 | 1.55 | 1.22 | 1.4 | 3.74 | 0.24 | 0.66 | 0.71 | 17.17 | 0.16
| Analytical results for Dy in OREAS 232b (Indicative Value 2.09 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 2.2572 | 2.1 | 1.971 | 1.8287 | [20 empty] | 1
| 1 | 2 | 2.2672 | 2.1 | 2.058 | 2.131 | [20 empty] | 9
| 1 | 3 | 2.2492 | 2 | 2.085 | 2.155 | [20 empty] | 16
| 1 | 4 | 2.273 | 2 | 2.019 | 1.8081 | [20 empty] | 2.0888
| 1 | 5 | 2.2245 | 2.1 | 2.15 | 2.1319 | [20 empty] | 15
| 1 | 6 | 2.2923 | 2.1 | 2.09 | 1.7409
| Mean |  | 2.2606 | 2.0667 | 2.0622 | 1.9659
| Median |  | 2.2622 | 2.1 | 2.0715 | 1.9798
| Std Dev. |  | 0.023 | 0.05164 | 0.061972 | 0.19231
| Rel.Std.Dev. |  | 1.017% | 2.499% | 3.005% | 9.782%
| PDM3 |  | 8.222% | -1.061% | -1.277% | -5.884%
| Z-Score (Absolute) |  | 2.63 | 0.03 | 0.03 | 1.32
| Analytical results for Er in OREAS 232b (Indicative Value 0.91 ppm) | [25 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.97658 | 0.8 | 0.89 | 0.7836 | [20 empty] | 1
| 1 | 2 | 1.0538 | 0.9 | 0.92 | 0.9256 | [20 empty] | 10
| 1 | 3 | 1.0212 | 0.8 | 0.926 | 0.9419 | [20 empty] | 16
| 1 | 4 | 0.98091 | 0.8 | 0.904 | 0.7892 | [20 empty] | 0.90798
| 1 | 5 | 1.0257 | 0.9 | 0.958 | 0.9281 | [20 empty] | 16
| 1 | 6 | 0.98202 | 0.9 | 0.948 | 0.7368
| Mean |  | 1.0067 | 0.85 | 0.92433 | 0.85087
| Median |  | 1.0016 | 0.85 | 0.923 | 0.8574
| Std Dev. |  | 0.031551 | 0.054772 | 0.025719 | 0.090748
| Rel.Std.Dev. |  | 3.134% | 6.444% | 2.782% | 10.67%
| PDM3 |  | 10.87% | -6.386% | 1.801% | -6.29%
| Z-Score (Absolute) |  | 2.16 | 0.68 | 0.67 | 0.67
| Analytical results for Eu in OREAS 232b (Indicative Value 0.58 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.72794 | 0.5 | 0.543 | 0.4678 | [20 empty] | 1
| 1 | 2 | 0.74041 | 0.5 | 0.614 | 0.5518 | [20 empty] | 11
| 1 | 3 | 0.69311 | 0.5 | 0.579 | 0.5562 | [20 empty] | 16
| 1 | 4 | 0.70352 | 0.5 | 0.561 | 0.4604 | [20 empty] | 0.57985
| 1 | 5 | 0.70145 | 0.5 | 0.612 | 0.552 | [20 empty] | 17
| 1 | 6 | 0.6938 | 0.6 | 0.609 | 0.4499
| Mean |  | 0.71004 | 0.51667 | 0.58633 | 0.50635
| Median |  | 0.70248 | 0.5 | 0.594 | 0.5098
| Std Dev. |  | 0.019541 | 0.040825 | 0.030038 | 0.051805
| Rel.Std.Dev. |  | 2.752% | 7.902% | 5.123% | 10.23%
| PDM3 |  | 22.45% | -10.9% | 1.119% | -12.68%
| Z-Score (Absolute) |  | 2.67 | 0.59 | 0.59 | 0.76
| Analytical results for Fe in OREAS 232b (Certified Value 3.3 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.199 | 3.339 | 3.27 | 3.3867 | 3.3 | 3.31 | 3.7 | 3.34 | 3.58 | 3.32 | 3.28 | 3.09 | 3.44 | 3.4584 | 3.231 | 3.11 | 3.1206 | 3.43 | 3.32 | 3.1495 | 4.0985 |  |  |  | 1
| 1 | 2 | 3.239 | 3.3277 | 3.46 | 3.4033 | 3.35 | 3.21 | 3.71 | 3.32 | 3.49 | 3.35 | 3.29 | 3.12 | 3.37 | 3.4566 | 3.273 | 3.11 | 3.1913 | 3.41 | 3.33 | 3.1101 | 3.844 |  |  |  | #N/A
| 1 | 3 | 3.207 | 3.3576 | 3.25 | 3.39 | 3.29 | 3.29 | 3.69 | 3.34 | 3.69 | 3.33 | 3.31 | 3.1 | 3.34 | 3.4642 | 3.168 | 3.1 | 3.1894 | 3.43 | 3.38 | 3.1366 | 3.9796 |  |  |  | 16
| 1 | 4 | 3.241 | 3.2724 | 3.18 | 3.4 | 3.29 | 3.22 | 3.67 | 3.28 | 3.54 | 3.36 | 3.21 | 3.11 | 3.31 | 3.4512 | 3.245 | 3.15 | 3.236 | 3.42 | 3.35 | 3.1671 | 3.8688 |  |  |  | 3.3021
| 1 | 5 | 3.181 | 3.3174 | 3.42 | 3.4 | 3.29 | 3.33 | 3.59 | 3.35 | 3.49 | 3.38 | 3.31 | 3.14 | 3.31 | 3.4602 | 3.231 | 3.11 | 3.2664 | 3.4 | 3.34 | 3.1349 | 3.9363 |  |  |  | 85
| 1 | 6 | 3.322 | 3.3378 | 3.45 | 3.42 | 3.34 | 3.31 | 3.68 | 3.35 | 3.54 | 3.41 | 3.24 | 3.12 | 3.29 | 3.454 | 3.217 | 3.13 | 3.2121 | 3.4 | 3.36 | 3.1523 | 3.9211
| Mean |  | 3.2315 | 3.3253 | 3.3383 | 3.4 | 3.31 | 3.2783 | 3.6733 | 3.33 | 3.555 | 3.3583 | 3.2733 | 3.1133 | 3.3433 | 3.4574 | 3.2275 | 3.1183 | 3.2026 | 3.415 | 3.3467 | 3.1418 | 3.9414
| Median |  | 3.223 | 3.3327 | 3.345 | 3.4 | 3.295 | 3.3 | 3.685 | 3.34 | 3.54 | 3.355 | 3.285 | 3.115 | 3.325 | 3.4575 | 3.231 | 3.11 | 3.2017 | 3.415 | 3.345 | 3.143 | 3.9287
| Std Dev. |  | 0.050087 | 0.029145 | 0.11957 | 0.011738 | 0.027568 | 0.050761 | 0.043205 | 0.026833 | 0.074498 | 0.033116 | 0.040332 | 0.017512 | 0.055015 | 0.0045927 | 0.034789 | 0.018348 | 0.049599 | 0.013784 | 0.021602 | 0.019438 | 0.090923
| Rel.Std.Dev. |  | 1.55% | 0.8765% | 3.582% | 0.3452% | 0.8329% | 1.548% | 1.176% | 0.8058% | 2.096% | 0.9861% | 1.232% | 0.5625% | 1.646% | 0.1328% | 1.078% | 0.5884% | 1.549% | 0.4036% | 0.6455% | 0.6187% | 2.307%
| PDM3 |  | -2.138% | 0.7037% | 1.098% | 2.965% | 0.2397% | -0.7192% | 11.24% | 0.8454% | 7.659% | 1.703% | -0.8707% | -5.716% | 1.249% | 4.704% | -2.259% | -5.565% | -3.012% | 3.42% | 1.35% | -4.856% | 19.36%
| Z-Score (Absolute) |  | 0.78 | 0.04 | 0.07 | 0.56 | 0.16 | 0.41 | 2.72 | 0 | 1.78 | 0.22 | 0.45 | 1.72 | 0.11 | 1.01 | 0.81 | 1.68 | 1.01 | 0.67 | 0.13 | 1.49 | 4.85
| Analytical results for Ga in OREAS 232b (Certified Value 8.58 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.27 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 8.8 | 8.0122 | 8.53 | < 2 | 8.9 | 11 | 9 | 9 | 8.5 | 8.42 | 8.63 | 8.3 | 9 | 8.6569 | 8.1 | 10.5 | 9.75 | 7.1076 | 7.4559 | [5 empty] | 1
| 1 | 2 | 8.6 | 8.2191 | 8.91 | < 2 | 8.9 | 9 | 8.6 | 9 | 8.16 | 8.18 | 8.72 | 8.15 | 8.8 | 8.6977 | 8.2 | 10.7 | 9.22 | 8.557 | 7.9649 | [5 empty] | 39
| 1 | 3 | 8.7 | 8.1135 | 8.57 | < 2 | 8.8 | 9 | 8.8 | 9 | 8.62 | 8.5 | 8.78 | 8.33 | 8.5 | 9.0832 | 8.6 | 10.6 | 9.36 | 8.9785 | 7.9743 | [5 empty] | 16
| 1 | 4 | 8.6 | 8.0187 | 8.27 | < 2 | 8.8 | 11 | 8.9 | 9 | 8.07 | 8.5 | 8.53 | 8.24 | 8.6 | 8.9004 | 8.7 | 10.5 | 9.32 | 7.8433 | 7.7653 | [5 empty] | 8.5777
| 1 | 5 | 8.5 | 7.7881 | 9.08 | < 2 | 8.9 | 12 | 8.9 | 9 | 8.24 | 8.73 | 8.81 | 8.26 | 8.7 | 9.1469 | 8.9 | 10 | 9.33 | 8.6864 | 8.2876 | [5 empty] | 86
| 1 | 6 | 8.8 | 7.9013 | 8.65 | < 2 | 8.9 | 12 | 8.6 | 9 | 8.38 | 8.37 | 8.81 | 8.19 | 8.3 | 9.0348 | 8.8 | 11 | 9.33 | 7.0858 | 8.3689
| Mean |  | 8.6667 | 8.0088 | 8.6683 |  | 8.8667 | 10.6667 | 8.8 | 9 | 8.3283 | 8.45 | 8.7133 | 8.245 | 8.65 | 8.92 | 8.55 | 10.55 | 9.385 | 8.0431 | 7.9695
| Median |  | 8.65 | 8.0155 | 8.61 |  | 8.9 | 11 | 8.85 | 9 | 8.31 | 8.46 | 8.75 | 8.25 | 8.65 | 8.9676 | 8.65 | 10.55 | 9.33 | 8.2002 | 7.9696
| Std Dev. |  | 0.12111 | 0.15204 | 0.28833 |  | 0.05164 | 1.3663 | 0.16733 | 0 | 0.2098 | 0.18089 | 0.1129 | 0.067157 | 0.2429 | 0.2051 | 0.32711 | 0.32711 | 0.18512 | 0.82273 | 0.33651
| Rel.Std.Dev. |  | 1.397% | 1.898% | 3.326% |  | 0.5824% | 12.81% | 1.902% | 0% | 2.519% | 2.141% | 1.296% | 0.8145% | 2.808% | 2.299% | 3.826% | 3.101% | 1.973% | 10.23% | 4.222%
| PDM3 |  | 1.038% | -6.632% | 1.057% |  | 3.369% | 24.35% | 2.592% | 4.924% | -2.907% | -1.488% | 1.582% | -3.878% | 0.8433% | 3.991% | -0.3225% | 22.99% | 9.412% | -6.232% | -7.09%
| Z-Score (Absolute) |  | 0.04 | 1.6 | 0.05 | 19.11 | 0.54 | NA | 0.37 | NA | 0.8 | 0.5 | 0.16 | 1.01 | 0 | 0.67 | 0.25 | 4.75 | 1.84 | 1.52 | 1.7
Results from laboratories 1.09 and 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 232b (Indicative Value 4.05 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 5.583 | 3.3 | 3.55 | 3.4171 | [20 empty] | 1
| 1 | 2 | 5.5136 | 3.2 | 3.691 | 4.0306 | [20 empty] | 12
| 1 | 3 | 5.5825 | 3.3 | 3.775 | 3.9895 | [20 empty] | 16
| 1 | 4 | 5.5234 | 3.4 | 3.622 | 3.4062 | [20 empty] | 4.0497
| 1 | 5 | 5.495 | 3.3 | 3.952 | 4.0015 | [20 empty] | 18
| 1 | 6 | 5.28 | 3.2 | 3.849 | 3.2314
| Mean |  | 5.4962 | 3.2833 | 3.7398 | 3.6794
| Median |  | 5.5185 | 3.3 | 3.733 | 3.7033
| Std Dev. |  | 0.11206 | 0.075277 | 0.14864 | 0.36535
| Rel.Std.Dev. |  | 2.039% | 2.293% | 3.975% | 9.93%
| PDM3 |  | 35.72% | -18.92% | -7.652% | -9.144%
| Z-Score (Absolute) |  | 5.28 | 1.26 | 0.09 | 0.09
| Analytical results for Ge in OREAS 232b (Certified Value 0.12 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.12 | 0.11 | 0.09 | 0.11 | 0.2 | 0.14 | 0.085 | [13 empty] | 1
| 1 | 2 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.12 | 0.12 | 0.07 | 0.12 | 0.2 | 0.11 | 0.0842 | [13 empty] | 6
| 1 | 3 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.12 | 0.11 | 0.08 | 0.13 | 0.2 | 0.12 | 0.0737 | [13 empty] | 16
| 1 | 4 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.11 | 0.11 | 0.08 | 0.12 | 0.2 | 0.12 | 0.0733 | [13 empty] | 0.123
| 1 | 5 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.11 | 0.12 | 0.08 | 0.13 | 0.2 | 0.12 | 0.0842 | [13 empty] | 87
| 1 | 6 | < 0.05 | < 0.1 | 0.14 | 1.6 | 0.11 | 0.12 | 0.08 | 0.12 | 0.2 | 0.13 | 0.0842
| Mean |  |  |  | 0.14 | 1.6 | 0.115 | 0.115 | 0.08 | 0.12167 | 0.2 | 0.12333 | 0.080767
| Median |  |  |  | 0.14 | 1.6 | 0.115 | 0.115 | 0.08 | 0.12 | 0.2 | 0.12 | 0.0842
| Std Dev. |  |  |  | 0 | 2.4324e-16 | 0.0054772 | 0.0054772 | 0.0063246 | 0.0075277 | 3.0405e-17 | 0.010328 | 0.0056387
| Rel.Std.Dev. |  |  |  | 0% | 1.52e-14% | 4.763% | 4.763% | 7.906% | 6.187% | 1.52e-14% | 8.374% | 6.981%
| PDM3 |  |  |  | 13.82% | 1201% | -6.504% | -6.504% | -34.96% | -1.084% | 62.6% | 0.271% | -34.34%
| Z-Score (Absolute) |  | 2.43 | 1.75 | 0.67 | NA | 0 | 0 | 0.94 | 0.18 | NA | 0.22 | 0.92
Results from laboratories 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 232b (Certified Value 0.5 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.47 | 0.7911 | 0.2 | 0.66 | 0.55 | 0.74 | 0.5 | 0.51 | 0.47 | 0.48 | 0.54 | 0.51 | 0.37 | 0.2969 | 0.4274 | [9 empty] | 1
| 1 | 2 | 0.49 | 0.84094 | 0.2 | 0.71 | 0.48 | 0.77 | 0.5 | 0.5 | 0.46 | 0.48 | 0.54 | 0.53 | 0.32 | 0.3397 | 0.4247 | [9 empty] | 25
| 1 | 3 | 0.48 | 0.82758 | 0.2 | 0.66 | 0.49 | 0.71 | 0.5 | 0.5 | 0.49 | 0.5 | 0.55 | 0.53 | 0.33 | 0.336 | 0.422 | [9 empty] | 16
| 1 | 4 | 0.48 | 0.77675 | 0.3 | 0.69 | 0.5 | 0.75 | 0.5 | 0.51 | 0.49 | 0.48 | 0.55 | 0.54 | 0.31 | 0.278 | 0.433 | [9 empty] | 0.4998
| 1 | 5 | 0.5 | 0.84488 | 0.1 | 0.68 | 0.53 | 0.77 | 0.5 | 0.49 | 0.46 | 0.49 | 0.55 | 0.57 | 0.34 | 0.3369 | 0.444 | [9 empty] | 88
| 1 | 6 | 0.52 | 0.75569 | 0.2 | 0.69 | 0.55 | 0.73 | 0.5 | 0.5 | 0.48 | 0.5 | 0.54 | 0.55 | 0.34 | 0.2806 | 0.4248
| Mean |  | 0.49 | 0.80616 | 0.2 | 0.68167 | 0.51667 | 0.745 | 0.5 | 0.50167 | 0.475 | 0.48833 | 0.545 | 0.53833 | 0.335 | 0.31135 | 0.42932
| Median |  | 0.485 | 0.80934 | 0.2 | 0.685 | 0.515 | 0.745 | 0.5 | 0.5 | 0.475 | 0.485 | 0.545 | 0.535 | 0.335 | 0.31645 | 0.4261
| Std Dev. |  | 0.017889 | 0.036895 | 0.063246 | 0.019408 | 0.030768 | 0.023452 | 0 | 0.0075277 | 0.013784 | 0.0098319 | 0.0054772 | 0.020412 | 0.020736 | 0.02943 | 0.0081005
| Rel.Std.Dev. |  | 3.651% | 4.577% | 31.62% | 2.847% | 5.955% | 3.148% | 0% | 1.501% | 2.902% | 2.013% | 1.005% | 3.792% | 6.19% | 9.453% | 1.887%
| PDM3 |  | -1.962% | 61.29% | -59.98% | 36.39% | 3.374% | 49.06% | 0.03908% | 0.3725% | -4.963% | -2.295% | 9.043% | 7.709% | -32.97% | -37.71% | -14.1%
| Z-Score (Absolute) |  | 0.18 | 4.74 | NA | 2.8 | 0.23 | 3.79 | NA | 0 | 0.41 | 0.21 | 0.67 | 0.57 | 2.59 | 2.96 | 1.13
Results from laboratories 1.06 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 232b (Indicative Value 0.01 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.102 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.028 | < 1 | < 0.01 | < 5 | 0.009 | 0.0091 | [9 empty] | 1
| 1 | 2 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.094 | 0.01 | 0.01 | < 0.01 | < 0.01 | 0.019 | < 1 | < 0.01 | < 5 | < 0.005 | 0.0102 | [9 empty] | 13
| 1 | 3 | < 0.01 | 0.04 | < 0.05 | < 2 | 0.102 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.021 | < 1 | < 0.01 | < 5 | < 0.005 | 0.0094 | [9 empty] | 16
| 1 | 4 | < 0.01 | 0.06 | < 0.05 | < 2 | 0.106 | 0.01 | 0.02 | < 0.01 | < 0.01 | 0.016 | < 1 | < 0.01 | < 5 | 0.006 | 0.0079 | [9 empty] | 0.0095611
| 1 | 5 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.107 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.013 | < 1 | < 0.01 | < 5 | 0.006 | 0.0027 | [9 empty] | 19
| 1 | 6 | < 0.01 | 0.04 | < 0.05 | < 2 | 0.091 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.015 | < 1 | < 0.01 | < 5 | < 0.005 | 0.0056
| Mean |  |  | 0.048333 |  |  | 0.10033 | 0.01 | 0.011667 | 0.01 |  | 0.018667 |  |  |  | 0.007 | 0.0074833
| Median |  |  | 0.05 |  |  | 0.102 | 0.01 | 0.01 | 0.01 |  | 0.0175 |  |  |  | 0.006 | 0.0085
| Std Dev. |  |  | 0.0075277 |  |  | 0.0064704 | 0 | 0.0040825 | 0 |  | 0.0053914 |  |  |  | 0.0017321 | 0.0028379
| Rel.Std.Dev. |  |  | 15.57% |  |  | 6.449% | 0% | 34.99% | 0% |  | 28.88% |  |  |  | 24.74% | 37.92%
| PDM3 |  |  | 405.5% |  |  | 949.4% | 4.59% | 22.02% | 4.59% |  | 95.24% |  |  |  | -26.79% | -21.73%
| Z-Score (Absolute) |  | 0.65 | 3.57 | 1.3 | 96.35 | 8.64 | 0.24 | 0 | 0.41 | 0.65 | 0.68 | 47.61 | 0.65 | 242.59 | 0.67 | 0.41
| Analytical results for Ho in OREAS 232b (Indicative Value 0.34 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.38477 | 0.3 | 0.335 | 0.3095 | [20 empty] | 1
| 1 | 2 | 0.39141 | 0.3 | 0.348 | 0.3609 | [20 empty] | 14
| 1 | 3 | 0.3878 | 0.3 | 0.35 | 0.3651 | [20 empty] | 16
| 1 | 4 | 0.38324 | 0.3 | 0.337 | 0.3063 | [20 empty] | 0.34103
| 1 | 5 | 0.38308 | 0.3 | 0.364 | 0.3573 | [20 empty] | 20
| 1 | 6 | 0.37695 | 0.3 | 0.353 | 0.2914
| Mean |  | 0.38454 | 0.3 | 0.34783 | 0.33175
| Median |  | 0.384 | 0.3 | 0.349 | 0.3334
| Std Dev. |  | 0.0048857 | 0 | 0.010722 | 0.03282
| Rel.Std.Dev. |  | 1.271% | 0% | 3.083% | 9.893%
| PDM3 |  | 12.76% | -12.03% | 1.995% | -2.722%
| Z-Score (Absolute) |  | 1.26 | 1.12 | 0.23 | 0.23
| Analytical results for In in OREAS 232b (Certified Value 0.034 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.03 | 0.04 | < 5 | < 0.05 | 0.03 | < 0.05 | 0.04 | 0.033 | 0.033 | 0.032 | 0.033 | 0.03 | 0.035 | 0.0337 | [10 empty] | 1
| 1 | 2 | 0.03 | 0.04 | < 5 | < 0.05 | 0.029 | < 0.05 | 0.04 | 0.033 | 0.03 | 0.032 | 0.03 | 0.03 | 0.031 | 0.0425 | [10 empty] | 27
| 1 | 3 | 0.03 | 0.03 | < 5 | < 0.05 | 0.038 | < 0.05 | 0.04 | 0.036 | 0.032 | 0.031 | 0.033 | 0.03 | 0.033 | 0.0415 | [10 empty] | 16
| 1 | 4 | 0.03 | 0.02 | < 5 | < 0.05 | 0.042 | < 0.05 | 0.04 | 0.034 | 0.032 | 0.032 | 0.034 | 0.03 | 0.031 | 0.0377 | [10 empty] | 0.033608
| 1 | 5 | 0.03 | 0.03 | < 5 | < 0.05 | 0.045 | < 0.05 | 0.04 | 0.035 | 0.034 | 0.029 | 0.034 | 0.03 | 0.031 | 0.0365 | [10 empty] | 89
| 1 | 6 | 0.03 | 0.03 | < 5 | < 0.05 | 0.033 | < 0.05 | 0.04 | 0.036 | 0.033 | 0.032 | 0.031 | 0.03 | 0.031 | 0.03
| Mean |  | 0.03 | 0.031667 |  |  | 0.036167 |  | 0.04 | 0.0345 | 0.032333 | 0.031333 | 0.0325 | 0.03 | 0.032 | 0.036983
| Median |  | 0.03 | 0.03 |  |  | 0.0355 |  | 0.04 | 0.0345 | 0.0325 | 0.032 | 0.033 | 0.03 | 0.031 | 0.0371
| Std Dev. |  | 0 | 0.0075277 |  |  | 0.0065549 |  | 0 | 0.0013784 | 0.0013663 | 0.0012111 | 0.0016432 | 0 | 0.0016733 | 0.0047144
| Rel.Std.Dev. |  | 0% | 23.77% |  |  | 18.12% |  | 0% | 3.995% | 4.226% | 3.865% | 5.056% | 0% | 5.229% | 12.75%
| PDM3 |  | -10.73% | -5.775% |  |  | 7.615% |  | 19.02% | 2.655% | -3.792% | -6.767% | -3.296% | -10.73% | -4.783% | 10.04%
| Z-Score (Absolute) |  | 0.65 | 0.15 | 739.59 | 2.15 | 1.2 | 2.15 | 2.35 | 0.7 | 0.05 | 0.25 | 0.1 | 0.65 | 0.05 | 1.44
| Analytical results for Ir in OREAS 232b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Ir | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.0026 | [23 empty] | 1
| 1 | 2 | 0.0008 | [23 empty] | 4
| 1 | 3 | < 0.0003 | [23 empty] | 16
| 1 | 4 | < 0.0012 | [23 empty] | 0.00088333
| 1 | 5 | < 0.0008 | [23 empty] | 14
| 1 | 6 | < 0.0015
| Mean |  | 0.0017
| Median |  | 0.0017
| Std Dev. |  | 0.0012728
| Rel.Std.Dev. |  | 74.87%
| PDM3 |  | 92.45%
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 232b (Certified Value 0.869 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.806 | 0.94103 | 0.78 | 0.8675 | 0.898 | 0.9 | 0.923 | 0.83 | 0.93 | 0.88 | 0.89 | 0.82 | 0.88 | 0.8662 | 0.863 | 0.81 | 0.82681 | 0.85 | 0.88 | 0.8957 | 1.0081 |  |  |  | 1
| 1 | 2 | 0.809 | 0.94362 | 0.76 | 0.872 | 0.914 | 0.86 | 0.939 | 0.83 | 0.91 | 0.92 | 0.9 | 0.83 | 0.86 | 0.86267 | 0.872 | 0.81 | 0.84408 | 0.85 | 0.88 | 0.8907 | 1.007 |  |  |  | #N/A
| 1 | 3 | 0.825 | 0.94308 | 0.79 | 0.8675 | 0.897 | 0.87 | 0.929 | 0.83 | 0.96 | 0.9 | 0.9 | 0.83 | 0.85 | 0.85771 | 0.855 | 0.81 | 0.84453 | 0.86 | 0.9 | 0.8978 | 1.0108 |  |  |  | 16
| 1 | 4 | 0.816 | 0.94013 | 0.77 | 0.8775 | 0.905 | 0.88 | 0.92 | 0.82 | 0.92 | 0.9 | 0.87 | 0.83 | 0.85 | 0.86523 | 0.863 | 0.82 | 0.85603 | 0.85 | 0.89 | 0.9021 | 1.0035 |  |  |  | 0.86889
| 1 | 5 | 0.8 | 0.94866 | 0.83 | 0.8675 | 0.914 | 0.9 | 0.904 | 0.83 | 0.92 | 0.91 | 0.9 | 0.83 | 0.85 | 0.84681 | 0.855 | 0.81 | 0.86024 | 0.85 | 0.89 | 0.8973 | 1.0091 |  |  |  | 90
| 1 | 6 | 0.812 | 0.94859 | 0.81 | 0.8865 | 0.89 | 0.91 | 0.923 | 0.83 | 0.93 | 0.91 | 0.89 | 0.83 | 0.84 | 0.83771 | 0.863 | 0.82 | 0.84867 | 0.86 | 0.89 | 0.8976 | 1.0086
| Mean |  | 0.81133 | 0.94419 | 0.79 | 0.87308 | 0.903 | 0.88667 | 0.923 | 0.82833 | 0.92833 | 0.90333 | 0.89167 | 0.82833 | 0.855 | 0.85605 | 0.86183 | 0.81333 | 0.84673 | 0.85333 | 0.88833 | 0.89687 | 1.0079
| Median |  | 0.8105 | 0.94335 | 0.785 | 0.86975 | 0.9015 | 0.89 | 0.923 | 0.83 | 0.925 | 0.905 | 0.895 | 0.83 | 0.85 | 0.86019 | 0.863 | 0.81 | 0.8466 | 0.85 | 0.89 | 0.89745 | 1.0084
| Std Dev. |  | 0.0086178 | 0.0036727 | 0.026077 | 0.0076708 | 0.009757 | 0.019664 | 0.011507 | 0.0040825 | 0.017224 | 0.013663 | 0.01169 | 0.0040825 | 0.013784 | 0.01145 | 0.0063377 | 0.005164 | 0.011674 | 0.005164 | 0.0075277 | 0.0036979 | 0.0024712
| Rel.Std.Dev. |  | 1.062% | 0.389% | 3.301% | 0.8786% | 1.081% | 2.218% | 1.247% | 0.4929% | 1.855% | 1.512% | 1.311% | 0.4929% | 1.612% | 1.338% | 0.7354% | 0.6349% | 1.379% | 0.6052% | 0.8474% | 0.4123% | 0.2452%
| PDM3 |  | -6.625% | 8.665% | -9.08% | 0.4821% | 3.925% | 2.045% | 6.227% | -4.668% | 6.841% | 3.964% | 2.621% | -4.668% | -1.599% | -1.478% | -0.8127% | -6.394% | -2.551% | -1.791% | 2.237% | 3.219% | 15.99%
| Z-Score (Absolute) |  | 1.39 | 1.6 | 1.87 | 0 | 0.67 | 0.31 | 1.13 | 1.01 | 1.25 | 0.68 | 0.42 | 1.01 | 0.41 | 0.38 | 0.25 | 1.35 | 0.59 | 0.45 | 0.34 | 0.54 | 3.04
| Analytical results for La in OREAS 232b (Certified Value 25.5 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 25.711 | 33.968 | 24.8 | 23.56 | 31.7 | 29.1 | 33.1 | 25 | 26.6 | 24.2 | 25.7 | 26 | 26 | 25.0831 | 14 | 25.58 | 25.6 | 31.7 | 21.7 | 21.4396 |  |  |  |  | 1
| 1 | 2 | 25.258 | 35.552 | 25.2 | 24.1667 | 31.73 | 24.1 | 34.3 | 25 | 26.1 | 24.1 | 26.6 | 25.8 | 26.3 | 26.2191 | 16 | 26.62 | 26.1 | 30.6 | 20.3 | 25.7634 |  |  |  |  | 28
| 1 | 3 | 25.659 | 35.453 | 24.3 | 24.01 | 32.55 | 21.9 | 34.2 | 25 | 26.5 | 24.3 | 26.1 | 26.7 | 26.6 | 26.3404 | 16 | 27.55 | 25.4 | 31.6 | 20.8 | 25.5676 |  |  |  |  | 16
| 1 | 4 | 25.722 | 34.463 | 25.7 | 24.0567 | 32.56 | 23.2 | 31.8 | 26 | 26.1 | 24.6 | 25 | 26.8 | 25.5 | 26.0033 | 16 | 26.12 | 26.7 | 31.2 | 21 | 21.4652 |  |  |  |  | 25.5034
| 1 | 5 | 25.457 | 35.321 | 24.7 | 24 | 31.41 | 27.1 | 35.3 | 26 | 26.2 | 24.7 | 27.3 | 26.8 | 25.6 | 25.6963 | 16 | 28.23 | 25.6 | 32.2 | 21.2 | 25.1106 |  |  |  |  | 91
| 1 | 6 | 26.257 | 35.167 | 24.6 | 24.27 | 31.67 | 26.2 | 34 | 25 | 26 | 23.3 | 26.7 | 26.4 | 24.9 | 26.7372 | 16 | 28.26 | 26.7 | 31.5 | 21.1 | 20.5486
| Mean |  | 25.6773 | 34.9873 | 24.8833 | 24.0106 | 31.9367 | 25.2667 | 33.7833 | 25.3333 | 26.25 | 24.2 | 26.2333 | 26.4167 | 25.8167 | 26.0132 | 15.6667 | 27.06 | 26.0167 | 31.4667 | 21.0167 | 23.3158
| Median |  | 25.685 | 35.244 | 24.75 | 24.0333 | 31.715 | 25.15 | 34.1 | 25 | 26.15 | 24.25 | 26.35 | 26.55 | 25.8 | 26.1112 | 16 | 27.085 | 25.85 | 31.55 | 21.05 | 23.2879
| Std Dev. |  | 0.33582 | 0.63139 | 0.49565 | 0.24369 | 0.49233 | 2.6793 | 1.199 | 0.5164 | 0.2429 | 0.498 | 0.81404 | 0.43089 | 0.61128 | 0.57259 | 0.8165 | 1.124 | 0.57764 | 0.53541 | 0.46224 | 2.4035
| Rel.Std.Dev. |  | 1.308% | 1.805% | 1.992% | 1.015% | 1.542% | 10.6% | 3.549% | 2.038% | 0.9253% | 2.058% | 3.103% | 1.631% | 2.368% | 2.201% | 5.212% | 4.154% | 2.22% | 1.702% | 2.199% | 10.31%
| PDM3 |  | 0.6822% | 37.19% | -2.431% | -5.853% | 25.23% | -0.9281% | 32.47% | -0.6667% | 2.928% | -5.111% | 2.862% | 3.581% | 1.228% | 1.999% | -38.57% | 6.104% | 2.013% | 23.38% | -17.59% | -8.577%
| Z-Score (Absolute) |  | 0.15 | 5.62 | 0.64 | 1.18 | 3.73 | 0.4 | 4.88 | 0.36 | 0.21 | 1.06 | 0.2 | 0.31 | 0.06 | 0.06 | 6.35 | 0.71 | 0.06 | 3.44 | 3.03 | 1.61
| Analytical results for Li in OREAS 232b (Certified Value 41.5 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [6 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 39.53 | 45.76 | 41.9 | 43.43 | 40.7 | 44 | 48.8 | 42.8 | 44.9 | 38.6 | 38.5 | 37.6 | 41.4945 | 46.2 | 38 | 46.8 | 35.9 | 37.5902 | [6 empty] | 1
| 1 | 2 | 39.1 | 46.11 | 44.3 | 43.59 | 43.7 | 40 | 50.5 | 42.6 | 43.5 | 38.6 | 39.9 | 39.4 | 40.2395 | 38.36 | 39 | 43.1 | 36.2 | 38.1983 | [6 empty] | #N/A
| 1 | 3 | 38.19 | 46.39 | 41.9 | 42.9 | 44.1 | 41 | 50.7 | 43 | 46.3 | 38.3 | 40.1 | 40.5 | 43.3708 | 38.4 | 38 | 43.2 | 36.3 | 37.7426 | [6 empty] | 16
| 1 | 4 | 38.13 | 45.95 | 41.9 | 43.78 | 40.8 | 42 | 47.4 | 40.1 | 44.8 | 38.4 | 39 | 40.6 | 40.382 | 37.17 | 39 | 42.5 | 37.1 | 36.1362 | [6 empty] | 41.4797
| 1 | 5 | 37.11 | 45.73 | 43.3 | 43.385 | 43.2 | 43 | 50.5 | 40 | 44.5 | 40.1 | 40.5 | 42.1 | 43.6932 | 39.2 | 39 | 42.6 | 36.2 | 40.0197 | [6 empty] | 92
| 1 | 6 | 38.27 | 45.92 | 43.7 | 43.78 | 41.2 | 44 | 49.3 | 41.4 | 45.7 | 39.6 | 40.7 | 39.7 | 42.6709 | 38.85 | 39 | 42.9 | 36.3 | 39.9778
| Mean |  | 38.3883 | 45.9767 | 42.8333 | 43.4775 | 42.2833 | 42.3333 | 49.5333 | 41.65 | 44.95 | 38.9333 | 39.7833 | 39.9833 | 41.9752 | 39.6967 | 38.6667 | 43.5167 | 36.3333 | 38.2775
| Median |  | 38.23 | 45.935 | 42.6 | 43.51 | 42.2 | 42.5 | 49.9 | 42 | 44.85 | 38.6 | 40 | 40.1 | 42.0827 | 38.625 | 39 | 43 | 36.25 | 37.9704
| Std Dev. |  | 0.84428 | 0.24509 | 1.0708 | 0.32863 | 1.551 | 1.633 | 1.2941 | 1.359 | 0.97108 | 0.73666 | 0.86352 | 1.4986 | 1.4943 | 3.2591 | 0.5164 | 1.6315 | 0.40332 | 1.5018
| Rel.Std.Dev. |  | 2.199% | 0.5331% | 2.5% | 0.7559% | 3.668% | 3.857% | 2.613% | 3.263% | 2.16% | 1.892% | 2.171% | 3.748% | 3.56% | 8.21% | 1.336% | 3.749% | 1.11% | 3.923%
| PDM3 |  | -7.453% | 10.84% | 3.263% | 4.816% | 1.937% | 2.058% | 19.42% | 0.4106% | 8.366% | -6.139% | -4.09% | -3.607% | 1.194% | -4.299% | -6.782% | 4.911% | -12.41% | -7.72%
| Z-Score (Absolute) |  | 1.11 | 1.35 | 0.33 | 0.54 | 0.15 | 0.17 | 2.51 | 0.05 | 1.02 | 0.94 | 0.66 | 0.6 | 0.05 | 0.69 | 1.02 | 0.55 | 1.78 | 1.15
| Analytical results for Lu in OREAS 232b (Certified Value 0.12 ppm) | [25 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.12929 | 0.1 | < 0.1 | 0.11 | 0.1262 | [19 empty] | 1
| 1 | 2 | 0.13399 | 0.1 | 0.104 | 0.11 | 0.1329 | [19 empty] | 29
| 1 | 3 | 0.12799 | 0.1 | 0.102 | 0.1 | 0.1711 | [19 empty] | 16
| 1 | 4 | 0.12018 | < 0.1 | 0.1 | 0.1 | 0.1655 | [19 empty] | 0.11561
| 1 | 5 | 0.12768 | 0.1 | 0.106 | 0.1 | 0.165 | [19 empty] | 93
| 1 | 6 | 0.11959 | 0.1 | 0.104 | 0.11 | 0.1498
| Mean |  | 0.12646 | 0.1 | 0.1032 | 0.105 | 0.15175
| Median |  | 0.12784 | 0.1 | 0.104 | 0.105 | 0.1574
| Std Dev. |  | 0.0055728 | 0 | 0.0022804 | 0.0054772 | 0.018713
| Rel.Std.Dev. |  | 4.407% | 0% | 2.21% | 5.216% | 12.33%
| PDM3 |  | 9.377% | -13.51% | -10.74% | -9.181% | 31.26%
| Z-Score (Absolute) |  | 1.09 | 0.67 | 0.54 | 0 | 2.36
| Analytical results for Mg in OREAS 232b (Certified Value 1.33 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.174 | 1.2533 | 1.42 | 1.3367 | 1.3 | 1.4 | 1.46 | 1.24 | 1.36 | 1.32 | 1.32 | 1.24 | 1.38 | 1.4513 | 1.369 | 1.19 | 1.2208 | 1.31 | 1.34 | 1.4352 | 1.4746 |  |  |  | 1
| 1 | 2 | 1.173 | 1.2474 | 1.45 | 1.3433 | 1.32 | 1.35 | 1.47 | 1.23 | 1.34 | 1.34 | 1.33 | 1.26 | 1.35 | 1.4294 | 1.381 | 1.2 | 1.2513 | 1.34 | 1.35 | 1.4338 | 1.4469 |  |  |  | #N/A
| 1 | 3 | 1.181 | 1.2688 | 1.38 | 1.33 | 1.3 | 1.37 | 1.45 | 1.24 | 1.42 | 1.32 | 1.35 | 1.25 | 1.35 | 1.4044 | 1.339 | 1.19 | 1.2472 | 1.36 | 1.37 | 1.4391 | 1.4161 |  |  |  | 16
| 1 | 4 | 1.164 | 1.2545 | 1.33 | 1.3467 | 1.29 | 1.37 | 1.45 | 1.22 | 1.36 | 1.32 | 1.3 | 1.26 | 1.33 | 1.4393 | 1.381 | 1.21 | 1.2594 | 1.35 | 1.35 | 1.4587 | 1.5669 |  |  |  | 1.3332
| 1 | 5 | 1.16 | 1.2918 | 1.42 | 1.34 | 1.3 | 1.4 | 1.42 | 1.24 | 1.35 | 1.34 | 1.34 | 1.27 | 1.34 | 1.4411 | 1.357 | 1.19 | 1.2751 | 1.34 | 1.35 | 1.4415 | 1.4315 |  |  |  | 94
| 1 | 6 | 1.174 | 1.2584 | 1.47 | 1.3433 | 1.32 | 1.41 | 1.45 | 1.24 | 1.37 | 1.36 | 1.31 | 1.26 | 1.32 | 1.4302 | 1.357 | 1.2 | 1.2567 | 1.37 | 1.35 | 1.4446 | 1.4232
| Mean |  | 1.171 | 1.2624 | 1.4117 | 1.34 | 1.305 | 1.3833 | 1.45 | 1.235 | 1.3667 | 1.3333 | 1.325 | 1.2567 | 1.345 | 1.4326 | 1.364 | 1.1967 | 1.2517 | 1.345 | 1.3517 | 1.4421 | 1.4599
| Median |  | 1.1735 | 1.2564 | 1.42 | 1.3417 | 1.3 | 1.385 | 1.45 | 1.24 | 1.36 | 1.33 | 1.325 | 1.26 | 1.345 | 1.4347 | 1.363 | 1.195 | 1.254 | 1.345 | 1.35 | 1.4403 | 1.4392
| Std Dev. |  | 0.007642 | 0.016052 | 0.050365 | 0.0059628 | 0.012247 | 0.023381 | 0.016733 | 0.0083666 | 0.028048 | 0.01633 | 0.018708 | 0.010328 | 0.020736 | 0.015997 | 0.016285 | 0.008165 | 0.017935 | 0.020736 | 0.0098319 | 0.0090295 | 0.056402
| Rel.Std.Dev. |  | 0.6526% | 1.272% | 3.568% | 0.445% | 0.9385% | 1.69% | 1.154% | 0.6775% | 2.052% | 1.225% | 1.412% | 0.8219% | 1.542% | 1.117% | 1.194% | 0.6823% | 1.433% | 1.542% | 0.7274% | 0.6261% | 3.864%
| PDM3 |  | -12.16% | -5.311% | 5.888% | 0.5122% | -2.113% | 3.763% | 8.763% | -7.364% | 2.512% | 0.01217% | -0.6129% | -5.739% | 0.8873% | 7.459% | 2.312% | -10.24% | -6.107% | 0.8873% | 1.387% | 8.174% | 9.503%
| Z-Score (Absolute) |  | 1.76 | 0.84 | 0.67 | 0.05 | 0.4 | 0.39 | 1.06 | 1.11 | 0.22 | 0.12 | 0.2 | 0.89 | 0 | 0.89 | 0.19 | 1.5 | 0.94 | 0 | 0.07 | 0.98 | 1.16
| Analytical results for Mn in OREAS 232b (Certified Value 0.03 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02945 | 0.028606 | 0.0315 | 0.030523 | 0.0306 | 0.03 | 0.0294 | 0.0317 | 0.0307 | 0.0288 | 0.0291 | 0.0264 | 0.0294 | 0.031353 | 0.0271 | 0.0286 | 0.026748 | 0.03274 | 0.0291 | 0.02753 | 0.0343 |  |  |  | 1
| 1 | 2 | 0.03002 | 0.029057 | 0.0326 | 0.030332 | 0.0312 | 0.0289 | 0.0296 | 0.0306 | 0.0297 | 0.0296 | 0.0296 | 0.0272 | 0.029 | 0.031139 | 0.0271 | 0.0287 | 0.027251 | 0.03247 | 0.0292 | 0.02774 | 0.0339 |  |  |  | 30
| 1 | 3 | 0.02957 | 0.029115 | 0.0309 | 0.03039 | 0.0307 | 0.0292 | 0.0293 | 0.0319 | 0.0315 | 0.0295 | 0.0295 | 0.0271 | 0.0288 | 0.031705 | 0.0263 | 0.0287 | 0.027134 | 0.03255 | 0.0297 | 0.02749 | 0.0347 |  |  |  | 16
| 1 | 4 | 0.02956 | 0.028672 | 0.0303 | 0.030726 | 0.0308 | 0.0287 | 0.029 | 0.0309 | 0.0303 | 0.0298 | 0.0283 | 0.0271 | 0.0286 | 0.030657 | 0.0263 | 0.029 | 0.027511 | 0.03237 | 0.0294 | 0.02783 | 0.0343 |  |  |  | 0.029529
| 1 | 5 | 0.02935 | 0.029632 | 0.0329 | 0.03056 | 0.0311 | 0.0295 | 0.0288 | 0.0319 | 0.03 | 0.0299 | 0.0294 | 0.0273 | 0.0286 | 0.031263 | 0.0263 | 0.0286 | 0.027935 | 0.03283 | 0.0294 | 0.02715 | 0.0343 |  |  |  | 95
| 1 | 6 | 0.02987 | 0.028995 | 0.0329 | 0.030726 | 0.0309 | 0.0299 | 0.0293 | 0.0324 | 0.0304 | 0.03 | 0.0289 | 0.0269 | 0.0284 | 0.031537 | 0.0263 | 0.0289 | 0.027533 | 0.03283 | 0.0295 | 0.0275 | 0.0343
| Mean |  | 0.029637 | 0.029013 | 0.03185 | 0.030543 | 0.030883 | 0.029367 | 0.029233 | 0.031567 | 0.030433 | 0.0296 | 0.029133 | 0.027 | 0.0288 | 0.031276 | 0.026567 | 0.02875 | 0.027352 | 0.032632 | 0.029383 | 0.02754 | 0.0343
| Median |  | 0.029565 | 0.029026 | 0.03205 | 0.030542 | 0.03085 | 0.02935 | 0.0293 | 0.0318 | 0.03035 | 0.0297 | 0.02925 | 0.0271 | 0.0287 | 0.031308 | 0.0263 | 0.0287 | 0.027381 | 0.032645 | 0.0294 | 0.027515 | 0.0343
| Std Dev. |  | 0.00025641 | 0.00036819 | 0.0011131 | 0.00016464 | 0.00023166 | 0.00052789 | 0.00028752 | 0.0006802 | 0.00062503 | 0.00043359 | 0.00048442 | 0.00032249 | 0.00035777 | 0.00036348 | 0.00041312 | 0.00016432 | 0.00040509 | 0.0001958 | 0.0002137 | 0.00023681 | 0.00025298
| Rel.Std.Dev. |  | 0.8652% | 1.269% | 3.495% | 0.539% | 0.7501% | 1.798% | 0.9835% | 2.155% | 2.054% | 1.465% | 1.663% | 1.194% | 1.242% | 1.162% | 1.555% | 0.5715% | 1.481% | 0.6% | 0.7273% | 0.8599% | 0.7376%
| PDM3 |  | 0.3644% | -1.748% | 7.86% | 3.433% | 4.586% | -0.55% | -1.002% | 6.9% | 3.062% | 0.2402% | -1.34% | -8.565% | -2.469% | 5.915% | -10.03% | -2.638% | -7.373% | 10.51% | -0.4935% | -6.736% | 16.16%
| Z-Score (Absolute) |  | 0.15 | 0.22 | 1.43 | 0.67 | 0.87 | 0.01 | 0.09 | 1.27 | 0.61 | 0.13 | 0.15 | 1.39 | 0.34 | 1.1 | 1.62 | 0.37 | 1.18 | 1.89 | 0 | 1.07 | 2.86
| Analytical results for Mo in OREAS 232b (Certified Value 0.94 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1 | 0.94 | < 2 | 1 | 1.25 | 1 | 0.9 | 1.01 | 1 | 0.86 | 0.91 | 1 | 0.99199 | < 1 | 1.443 | 0.87 |  | < 1 | 1.01 | 0.7159 | 2.8001 |  |  |  | 1
| 1 | 2 | 0.97 | 0.98 | < 2 | 1 | 1.31 | 1.1 | 0.9 | 0.86 | 0.95 | 0.85 | 0.86 | 0.96 | 1.0131 | < 1 | 1.265 | 0.88 |  | < 1 | 0.96 | 0.8793 | 2.5837 |  |  |  | 31
| 1 | 3 | 0.97 | 0.89 | < 2 | 1 | 1.26 | 1.1 | 1 | 0.97 | 0.97 | 0.86 | 0.89 | 0.94 | 1.0186 | < 1 | 1.213 | 0.96 |  | < 1 | 1 | 0.8754 | 2.8236 |  |  |  | 16
| 1 | 4 | 0.94 | 0.92 | < 2 | 0.9 | 1.34 | 1 | 1 | 0.85 | 1 | 0.9 | 0.89 | 0.99 | 1.0386 | < 1 | 1.109 | 1.05 |  | < 1 | 0.94 | 0.7742 | 7.3691 |  |  |  | 0.93558
| 1 | 5 | 0.95 | 0.93 | < 2 | 0.9 | 1.4 | 1.1 | 0.9 | 0.87 | 1 | 0.89 | 0.9 | 0.96 | 1.0222 | < 1 | 1.234 | 0.92 | 0.02 | < 1 | 0.97 | 0.7794 | 2.283 |  |  |  | 96
| 1 | 6 | 0.98 | 1.05 | < 2 | 1 | 1.32 | 1.3 | 1 | 0.92 | 0.92 | 0.9 | 0.92 | 1.01 | 0.98603 | < 1 | 1.11 | 0.94 |  | < 1 | 0.98 | 0.7483 | 2.2832
| Mean |  | 0.96833 | 0.95167 |  | 0.96667 | 1.3133 | 1.1 | 0.95 | 0.91333 | 0.97333 | 0.87667 | 0.895 | 0.97667 | 1.0117 |  | 1.229 | 0.93667 | 0.02 |  | 0.97667 | 0.79542 | 3.3571
| Median |  | 0.97 | 0.935 |  | 1 | 1.315 | 1.1 | 0.95 | 0.895 | 0.985 | 0.875 | 0.895 | 0.975 | 1.0158 |  | 1.2235 | 0.93 | 0.02 |  | 0.975 | 0.7768 | 2.6919
| Std Dev. |  | 0.02137 | 0.056362 |  | 0.05164 | 0.055015 | 0.10954 | 0.054772 | 0.06532 | 0.033267 | 0.022509 | 0.020736 | 0.027325 | 0.019637 |  | 0.12323 | 0.06532 |  |  | 0.02582 | 0.067355 | 1.9797
| Rel.Std.Dev. |  | 2.207% | 5.922% |  | 5.342% | 4.189% | 9.959% | 5.766% | 7.152% | 3.418% | 2.568% | 2.317% | 2.798% | 1.941% |  | 10.03% | 6.974% |  |  | 2.644% | 8.468% | 58.97%
| PDM3 |  | 3.501% | 1.719% |  | 3.323% | 40.38% | 17.57% | 1.541% | -2.378% | 4.035% | -6.297% | -4.338% | 4.391% | 8.14% |  | 31.36% | 0.116% | -97.86% |  | 4.391% | -14.98% | 258.8%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 0.46 | NA | 4.08 | NA | NA | 0.54 | 0.15 | 0.96 | 0.75 | 0.19 | 0.6 | 5.31 | 3.11 | 0.27 | 10.86 | 5.31 | 0.19 | 1.9 | 27.69
Results from laboratories 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 232b (Certified Value 0.1 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.099 | 0.13621 | 0.12 | 0.093333 | 0.11 | 0.1 | 0.13 | 0.15 | 0.1 | 0.1 | 0.1 | 0.09 | 0.107 | 0.10539 | 0.089 | 0.1 | 0.098954 | 0.09 | 0.11 | 0.091 | 0.0727 |  |  |  | 1
| 1 | 2 | 0.098 | 0.13653 | 0.126 | 0.093333 | 0.11 | 0.09 | 0.13 | 0.14 | 0.1 | 0.1 | 0.1 | 0.09 | 0.106 | 0.10403 | 0.089 | 0.1 | 0.10025 | 0.09 | 0.11 | 0.0915 | 0.0739 |  |  |  | #N/A
| 1 | 3 | 0.101 | 0.13935 | 0.122 | 0.093333 | 0.11 | 0.09 | 0.13 | 0.15 | 0.11 | 0.1 | 0.1 | 0.09 | 0.106 | 0.10519 | 0.089 | 0.1 | 0.099684 | 0.09 | 0.11 | 0.0917 | 0.0751 |  |  |  | 16
| 1 | 4 | 0.098 | 0.13772 | 0.111 | 0.093333 | 0.11 | 0.09 | 0.13 | 0.14 | 0.1 | 0.1 | 0.1 | 0.09 | 0.105 | 0.10261 | 0.089 | 0.1 | 0.10145 | 0.09 | 0.11 | 0.0924 | 0.213 |  |  |  | 0.10035
| 1 | 5 | 0.099 | 0.13535 | 0.126 | 0.09 | 0.11 | 0.1 | 0.13 | 0.14 | 0.1 | 0.1 | 0.11 | 0.09 | 0.104 | 0.10791 | 0.089 | 0.1 | 0.10326 | 0.09 | 0.11 | 0.0909 | 0.0797 |  |  |  | 97
| 1 | 6 | 0.101 | 0.13357 | 0.121 | 0.09 | 0.11 | 0.1 | 0.13 | 0.13 | 0.1 | 0.1 | 0.1 | 0.09 | 0.103 | 0.10097 | 0.089 | 0.1 | 0.10098 | 0.09 | 0.11 | 0.0916 | 0.0732
| Mean |  | 0.099333 | 0.13646 | 0.121 | 0.092222 | 0.11 | 0.095 | 0.13 | 0.14167 | 0.10167 | 0.1 | 0.10167 | 0.09 | 0.10517 | 0.10435 | 0.089 | 0.1 | 0.10076 | 0.09 | 0.11 | 0.091517 | 0.097933
| Median |  | 0.099 | 0.13637 | 0.1215 | 0.093333 | 0.11 | 0.095 | 0.13 | 0.14 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1055 | 0.10461 | 0.089 | 0.1 | 0.10061 | 0.09 | 0.11 | 0.09155 | 0.0745
| Std Dev. |  | 0.0013663 | 0.0019775 | 0.0055136 | 0.0017213 | 0 | 0.0054772 | 0 | 0.0075277 | 0.0040825 | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 0.001472 | 0.0024087 | 1.5202e-17 | 1.5202e-17 | 0.0015142 | 1.5202e-17 | 0 | 0.00054191 | 0.056427
| Rel.Std.Dev. |  | 1.375% | 1.449% | 4.557% | 1.866% | 0% | 5.766% | 0% | 5.314% | 4.016% | 1.52e-14% | 4.016% | 1.689e-14% | 1.4% | 2.308% | 1.708e-14% | 1.52e-14% | 1.503% | 1.689e-14% | 0% | 0.5921% | 57.62%
| PDM3 |  | -1.012% | 35.98% | 20.58% | -8.098% | 9.618% | -5.33% | 29.55% | 41.17% | 1.313% | -0.3475% | 1.313% | -10.31% | 4.801% | 3.987% | -11.31% | -0.3475% | 0.413% | -10.31% | 9.618% | -8.801% | -2.407%
| Z-Score (Absolute) |  | 0.11 | 2.82 | 1.6 | 0.67 | 0.73 | 0.46 | 2.31 | 3.23 | 0.07 | 0.06 | 0.07 | 0.85 | 0.35 | 0.28 | 0.93 | 0.06 | 0 | 0.85 | 0.73 | 0.73 | 0.22
| Analytical results for Nb in OREAS 232b (Certified Value 0.36 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.53 | 0.3 | < 5 | 0.8 | < 1 | 0.4 | 0.3 | 0.3 | 0.35 | 0.3 | 0.35 | 0.60478 | 0.56 | 0.35 | 0.3516 | [9 empty] | 1
| 1 | 2 | 0.53 | 0.4 | < 5 | 0.8 | < 1 | 0.5 | 0.3 | 0.26 | 0.34 | 0.32 | 0.35 | 0.59075 | 0.55 | 0.32 | 0.3947 | [9 empty] | 32
| 1 | 3 | 0.53 | 0.4 | < 5 | 0.7 | < 1 | 0.5 | 0.2 | 0.33 | 0.4 | 0.31 | 0.34 | 0.60726 | 0.62 | 0.34 | 0.3766 | [9 empty] | 16
| 1 | 4 | 0.52 | 0.3 | < 5 | 0.8 | < 1 | 0.4 | 0.3 | 0.29 | 0.38 | 0.27 | 0.4 | 0.57585 | 0.56 | 0.32 | 0.3167 | [9 empty] | 0.3622
| 1 | 5 | 0.49 | 0.6 | < 5 | 0.7 | < 1 | 0.5 | 0.3 | 0.32 | 0.37 | 0.28 | 0.37 | 0.62103 | 0.54 | 0.35 | 0.3828 | [9 empty] | 98
| 1 | 6 | 0.55 | 0.6 | < 5 | 0.8 | < 1 | 0.5 | 0.3 | 0.32 | 0.35 | 0.29 | 0.35 | 0.62224 | 0.66 | 0.33 | 0.338
| Mean |  | 0.525 | 0.43333 |  | 0.76667 |  | 0.46667 | 0.28333 | 0.30333 | 0.365 | 0.295 | 0.36 | 0.60365 | 0.58167 | 0.335 | 0.36007
| Median |  | 0.53 | 0.4 |  | 0.8 |  | 0.5 | 0.3 | 0.31 | 0.36 | 0.295 | 0.35 | 0.60602 | 0.56 | 0.335 | 0.3641
| Std Dev. |  | 0.019748 | 0.13663 |  | 0.05164 |  | 0.05164 | 0.040825 | 0.02582 | 0.022583 | 0.018708 | 0.021909 | 0.017894 | 0.047504 | 0.013784 | 0.029762
| Rel.Std.Dev. |  | 3.762% | 31.53% |  | 6.736% |  | 11.07% | 14.41% | 8.512% | 6.187% | 6.342% | 6.086% | 2.964% | 8.167% | 4.115% | 8.266%
| PDM3 |  | 44.95% | 19.64% |  | 111.7% |  | 28.84% | -21.77% | -16.25% | 0.7731% | -18.55% | -0.6074% | 66.66% | 60.59% | -7.51% | -0.589%
| Z-Score (Absolute) |  | 1.54 | NA | 20.57 | NA | 1.3 | NA | NA | 0.59 | 0 | 0.67 | 0.05 | 2.3 | 2.09 | 0.29 | 0.05
Results from laboratories 1.06, 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 232b (Indicative Value 22.9 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 1
| 1 | 1 | 26.3448 | 22.2 | 22.29 | 18.6569 | [20 empty] | 1
| 1 | 2 | 26.8795 | 21.5 | 23.23 | 22.3184 | [20 empty] | 9
| 1 | 3 | 26.6899 | 20.7 | 24.05 | 22.1743 | [20 empty] | 16
| 1 | 4 | 25.9964 | 22.3 | 22.85 | 18.6802 | [20 empty] | 22.942
| 1 | 5 | 26.0666 | 21.6 | 24.92 | 22.0501 | [20 empty] | 15
| 1 | 6 | 25.7023 | 20.5 | 24.88 | 18.0288
| Mean |  | 26.2799 | 21.4667 | 23.7033 | 20.3181
| Median |  | 26.2057 | 21.55 | 23.64 | 20.3651
| Std Dev. |  | 0.44527 | 0.74476 | 1.0895 | 2.0557
| Rel.Std.Dev. |  | 1.694% | 3.469% | 4.596% | 10.12%
| PDM3 |  | 14.55% | -6.431% | 3.318% | -11.44%
| Z-Score (Absolute) |  | 1.47 | 0.45 | 0.45 | 0.9
| Analytical results for Ni in OREAS 232b (Certified Value 59 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 60.4 | 51.3475 | 56.8 | 54.5667 | 58 | 62 | 64 | 62.8 | 59.2 | 58.4 | 60.2 | 55.4 | 62.3 | 64.0324 | 59 | 50.45 | 57 | 67.15 | 55.8 | 66.8 | 54.15 | 55.2429 | 60.058 |  | 1
| 1 | 2 | 58.1 | 51.338 | 59.6 | 53.9267 | 61 | 61 | 64 | 62 | 57.8 | 58.5 | 62 | 55 | 59.8 | 63.4912 | 59 | 45.57 | 58.4 | 71.54 | 55.94 | 63.4 | 53.295 | 58.8178 | 59.4485 |  | 34
| 1 | 3 | 57.4 | 51.4805 | 55.7 | 53.58 | 59 | 61 | 65 | 63.2 | 60.9 | 59.5 | 62 | 56.2 | 59.6 | 61.1197 | 58 | 43.91 | 57.8 | 69.08 | 56.51 | 63 | 54.34 | 59.6443 | 59.1716 |  | 16
| 1 | 4 | 57.1 | 51.9225 | 56.4 | 53.8533 | 58 | 61 | 64 | 61.8 | 58.7 | 58.8 | 59.5 | 55.8 | 58.6 | 62.7606 | 59 | 44.68 | 58.7 | 69.12 | 56.44 | 62.2 | 54.91 | 57.0303 | 59.3443 |  | 58.8159
| 1 | 5 | 57.4 | 52.2975 | 59.6 | 53.57 | 61 | 61 | 63 | 62.8 | 59.3 | 60.7 | 62.7 | 57.7 | 58.9 | 61.5686 | 59 | 49.68 | 58.6 | 68.79 | 56.14 | 63.1 | 54.055 | 59.1228 | 59.6738 |  | 99
| 1 | 6 | 58.6 | 51.699 | 58.5 | 54.18 | 58 | 62 | 64 | 63.7 | 59.3 | 60.6 | 61.7 | 55.9 | 59.3 | 62.3966 | 59 | 86.06 | 61.5 | 74.42 | 55.92 | 62.7 | 54.91 | 59.8675 | 59.7656
| Mean |  | 58.1667 | 51.6808 | 57.7667 | 53.9461 | 59.1667 | 61.3333 | 64 | 62.7167 | 59.2 | 59.4167 | 61.35 | 56 | 59.75 | 62.5615 | 58.8333 | 53.3917 | 58.6667 | 70.0167 | 56.125 | 63.5333 | 54.2767 | 58.2876 | 59.577
| Median |  | 57.75 | 51.5897 | 57.65 | 53.89 | 58.5 | 61 | 64 | 62.8 | 59.25 | 59.15 | 61.85 | 55.85 | 59.45 | 62.5786 | 59 | 47.625 | 58.5 | 69.1 | 56.04 | 63.05 | 54.245 | 58.9703 | 59.5611
| Std Dev. |  | 1.2242 | 0.37624 | 1.6931 | 0.3806 | 1.472 | 0.5164 | 0.63246 | 0.71671 | 1.0119 | 1.0304 | 1.2276 | 0.93167 | 1.3248 | 1.1099 | 0.40825 | 16.2273 | 1.5253 | 2.5735 | 0.2931 | 1.6513 | 0.60582 | 1.7981 | 0.3198
| Rel.Std.Dev. |  | 2.105% | 0.728% | 2.931% | 0.7055% | 2.488% | 0.842% | 0.9882% | 1.143% | 1.709% | 1.734% | 2.001% | 1.664% | 2.217% | 1.774% | 0.6939% | 30.39% | 2.6% | 3.676% | 0.5222% | 2.599% | 1.116% | 3.085% | 0.5368%
| PDM3 |  | -1.104% | -12.13% | -1.784% | -8.28% | 0.5963% | 4.28% | 8.814% | 6.632% | 0.653% | 1.021% | 4.308% | -4.788% | 1.588% | 6.368% | 0.02958% | -9.222% | -0.2538% | 19.04% | -4.575% | 8.021% | -7.718% | -0.8983% | 1.294%
| Z-Score (Absolute) |  | 0.31 | 2.31 | 0.43 | 1.61 | 0 | 0.67 | 1.49 | 1.1 | 0.01 | 0.08 | 0.67 | 0.98 | 0.18 | 1.05 | 0.1 | 1.78 | 0.15 | 3.35 | 0.94 | 1.35 | 1.51 | 0.27 | 0.13
| Analytical results for P in OREAS 232b (Certified Value 0.059 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.0552 | 0.056 | 0.05806 | 0.058 | 0.0597 | 0.0587 | 0.059 | 0.06 | 0.059 | 0.06 | 0.058 | 0.06 | 0.06345 | 0.0559 | 0.05 | 0.057295 | 0.06 | 0.0592 | 0.0657 | 0.0643 |  |  |  |  | 1
| 1 | 2 | 0.0551 | 0.056 | 0.058121 | 0.058 | 0.0584 | 0.0598 | 0.059 | 0.06 | 0.06 | 0.06 | 0.058 | 0.058 | 0.063856 | 0.0554 | 0.05 | 0.054917 | 0.06 | 0.0595 | 0.0649 | 0.0645 |  |  |  |  | #N/A
| 1 | 3 | 0.0544 | 0.055 | 0.058153 | 0.058 | 0.059 | 0.0592 | 0.06 | 0.063 | 0.059 | 0.061 | 0.057 | 0.058 | 0.06125 | 0.0576 | 0.05 | 0.055907 | 0.06 | 0.0603 | 0.0654 | 0.0653 |  |  |  |  | 16
| 1 | 4 | 0.054 | 0.055 | 0.058469 | 0.058 | 0.0594 | 0.0579 | 0.06 | 0.061 | 0.06 | 0.059 | 0.058 | 0.057 | 0.061812 | 0.0554 | 0.05 | 0.058173 | 0.06 | 0.0598 | 0.0659 | 0.0663 |  |  |  |  | 0.058827
| 1 | 5 | 0.0544 | 0.055 | 0.058609 | 0.058 | 0.0599 | 0.0576 | 0.06 | 0.06 | 0.06 | 0.06 | 0.059 | 0.057 | 0.063476 | 0.0554 | 0.05 | 0.055594 | 0.06 | 0.0594 | 0.0648 | 0.0646 |  |  |  |  | 100
| 1 | 6 | 0.0543 | 0.056 | 0.058619 | 0.058 | 0.0602 | 0.0583 | 0.061 | 0.061 | 0.062 | 0.059 | 0.058 | 0.057 | 0.061537 | 0.0576 | 0.05 | 0.058037 | 0.06 | 0.0594 | 0.0655 | 0.0647
| Mean |  | 0.054567 | 0.0555 | 0.058339 | 0.058 | 0.059433 | 0.058583 | 0.059833 | 0.060833 | 0.06 | 0.059833 | 0.058 | 0.057833 | 0.062563 | 0.056217 | 0.05 | 0.056654 | 0.06 | 0.0596 | 0.065367 | 0.06495
| Median |  | 0.0544 | 0.0555 | 0.058311 | 0.058 | 0.05955 | 0.0585 | 0.06 | 0.0605 | 0.06 | 0.06 | 0.058 | 0.0575 | 0.062631 | 0.05565 | 0.05 | 0.056601 | 0.06 | 0.05945 | 0.06545 | 0.06465
| Std Dev. |  | 0.0004761 | 0.00054772 | 0.00025621 | 0 | 0.0006532 | 0.00082321 | 0.00075277 | 0.001169 | 0.0010954 | 0.00075277 | 0.00063246 | 0.001169 | 0.0011518 | 0.0010889 | 7.6012e-18 | 0.001366 | 0 | 0.00039497 | 0.00043665 | 0.00074229
| Rel.Std.Dev. |  | 0.8725% | 0.9869% | 0.4392% | 0% | 1.099% | 1.405% | 1.258% | 1.922% | 1.826% | 1.258% | 1.09% | 2.021% | 1.841% | 1.937% | 1.52e-14% | 2.411% | 0% | 0.6627% | 0.668% | 1.143%
| PDM3 |  | -7.241% | -5.655% | -0.8297% | -1.405% | 1.031% | -0.4135% | 1.711% | 3.411% | 1.995% | 1.711% | -1.405% | -1.688% | 6.352% | -4.437% | -15% | -3.693% | 1.995% | 1.315% | 11.12% | 10.41%
| Z-Score (Absolute) |  | 2.74 | 2.17 | 0.41 | 0.62 | 0.26 | 0.26 | 0.51 | 1.13 | 0.61 | 0.51 | 0.62 | 0.73 | 2.2 | 1.72 | 5.56 | 1.45 | 0.61 | 0.37 | 3.93 | 3.67
| Analytical results for Pb in OREAS 232b (Certified Value 8.62 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  | 3
| 1 | 1 | 9.1 | 9.1 | 7.55 | 10.2 | 8 | 8.2 | 9.1 | 8.7 | 8.1 | 8.6 | 9 | 9.72 | 9.9245 | 7 | 20.66 | 8.4 | 0.69 | 9 | 7.9 | 11 | 10 | 7.5223 |  |  | 1
| 1 | 2 | 8.4 | 9.7 | 8.6133 | 9.7 | 8 | 8.2 | 8.9 | 8.5 | 8.1 | 8.6 | 8.6 | 8.34 | 9.6899 | 8 | 19.21 | 8.5 |  | 8.4 | 7.8 | 9.9 | 10 | 7.2327 |  |  | 35
| 1 | 3 | 8.4 | 8.9 | 8.16 | 9.6 | 6 | 8.3 | 8.4 | 8.6 | 8.1 | 9 | 8.7 | 8.84 | 9.0211 | 7 | 24.96 | 8.2 |  | 8.5 | 7.8 | 10.2 | 10 | 7.9469 |  |  | 16
| 1 | 4 | 8.4 | 8.7 | 8.1733 | 10.1 | 6 | 8 | 9.5 | 8.1 | 8.6 | 8.4 | 8.7 | 8.43 | 10.0868 | 8 | 16.56 | 8.4 | 2.75 | 8.7 | 7.8 | 10.1 | 10 | 8.6347 |  |  | 8.6164
| 1 | 5 | 8.3 | 9.2 | 7.8233 | 10.2 | 7 | 8.5 | 9.8 | 8.7 | 8.1 | 9 | 8.6 | 8.29 | 9.7159 | 7 | 15.45 | 8.4 | 0.57 | 8.6 | 7.9 | 10 | 10 | 7.9469 |  |  | 101
| 1 | 6 | 8.6 | 9.8 | 7.9267 | 9.9 | 6 | 8.3 | 9.2 | 8.1 | 8.3 | 8.6 | 8.7 | 8.62 | 9.0982 | 7 | 18.26 | 8.6 | 5.05 | 8.6 | 7.9 | 10.6 | 10 | 7.9602
| Mean |  | 8.5333 | 9.2333 | 8.0411 | 9.95 | 6.8333 | 8.25 | 9.15 | 8.45 | 8.2167 | 8.7 | 8.7167 | 8.7067 | 9.5894 | 7.3333 | 19.1833 | 8.4167 | 2.265 | 8.6333 | 7.85 | 10.3 | 10 | 7.874
| Median |  | 8.4 | 9.15 | 8.0433 | 10 | 6.5 | 8.25 | 9.15 | 8.55 | 8.1 | 8.6 | 8.7 | 8.525 | 9.7029 | 7 | 18.735 | 8.4 | 1.72 | 8.6 | 7.85 | 10.15 | 10 | 7.9469
| Std Dev. |  | 0.29439 | 0.43665 | 0.36353 | 0.25884 | 0.98319 | 0.16432 | 0.48477 | 0.28107 | 0.20412 | 0.24495 | 0.1472 | 0.53605 | 0.43589 | 0.5164 | 3.383 | 0.13292 | 2.1091 | 0.20656 | 0.054772 | 0.41952 | 0 | 0.47575
| Rel.Std.Dev. |  | 3.45% | 4.729% | 4.521% | 2.601% | 14.39% | 1.992% | 5.298% | 3.326% | 2.484% | 2.816% | 1.689% | 6.157% | 4.546% | 7.042% | 17.64% | 1.579% | 93.12% | 2.393% | 0.6977% | 4.073% | 0% | 6.042%
| PDM3 |  | -0.9636% | 7.16% | -6.676% | 15.48% | -20.69% | -4.252% | 6.193% | -1.931% | -4.639% | 0.9707% | 1.164% | 1.048% | 11.29% | -14.89% | 122.6% | -2.318% | -73.71% | 0.197% | -8.894% | 19.54% | 16.06% | -8.616%
| Z-Score (Absolute) |  | 0.13 | 0.78 | 0.77 | 1.72 | NA | 0.5 | 0.67 | 0.24 | 0.54 | 0.09 | 0.11 | 0.1 | 1.25 | NA | 13.77 | 0.28 | 8.31 | 0 | 1.02 | 2.18 | NA | 0.99
Results from laboratories 1.09, 1.19 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 232b (Indicative Value 15.9 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 1
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 1
| 1 | 1 | < 1 | < 10 | 27.8 | [21 empty] | 1
| 1 | 2 | < 1 | < 10 | 26.9 | [21 empty] | 1
| 1 | 3 | 1 | < 10 | 36.8 | [21 empty] | 16
| 1 | 4 | < 1 | < 10 | 33.2 | [21 empty] | 15.9167
| 1 | 5 | < 1 | < 10 | 44.9 | [21 empty] | 16
| 1 | 6 | < 1 | < 10 | 17.9
| Mean |  | 1 |  | 31.25
| Median |  | 1 |  | 30.5
| Std Dev. |  |  |  | 9.2733
| Rel.Std.Dev. |  |  |  | 29.67%
| PDM3 |  | -93.72% |  | 96.34%
| Z-Score (Absolute) |  | 0.67 | 0 | 4.01
| Analytical results for Pr in OREAS 232b (Indicative Value 6.24 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 6.9887 | 6.1 | 5.971 | 4.9987 | [20 empty] | 1
| 1 | 2 | 7.0389 | 6.4 | 6.151 | 6.0109 | [20 empty] | 11
| 1 | 3 | 6.9527 | 6.3 | 6.603 | 5.9791 | [20 empty] | 16
| 1 | 4 | 6.9556 | 6.2 | 6.07 | 4.9961 | [20 empty] | 6.2389
| 1 | 5 | 6.9653 | 6.2 | 6.694 | 5.8559 | [20 empty] | 17
| 1 | 6 | 6.9979 | 5.9 | 6.585 | 4.8191
| Mean |  | 6.9832 | 6.1833 | 6.3457 | 5.4433
| Median |  | 6.977 | 6.2 | 6.368 | 5.4273
| Std Dev. |  | 0.032752 | 0.17224 | 0.31594 | 0.55978
| Rel.Std.Dev. |  | 0.469% | 2.786% | 4.979% | 10.28%
| PDM3 |  | 11.93% | -0.8902% | 1.712% | -12.75%
| Z-Score (Absolute) |  | 1.21 | 0.14 | 0.14 | 1.38
| Analytical results for Pt in OREAS 232b (Indicative Value 2.22 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 2
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 2
| 1 | 1 | 2 | < 5 | 2.6 | [21 empty] | 1
| 1 | 2 | 2 | < 5 | 1.9 | [21 empty] | 1
| 1 | 3 | 2 | < 5 | 1.5 | [21 empty] | 16
| 1 | 4 | 2 | < 5 | 3.6 | [21 empty] | 2.2167
| 1 | 5 | 2 | < 5 | 2.6 | [21 empty] | 18
| 1 | 6 | 2 | < 5 | 2.4
| Mean |  | 2 |  | 2.4333
| Median |  | 2 |  | 2.5
| Std Dev. |  | 0 |  | 0.7174
| Rel.Std.Dev. |  | 0% |  | 29.48%
| PDM3 |  | -9.774% |  | 9.774%
| Z-Score (Absolute) |  | 4.38 | 0.67 | 0
| Analytical results for Rb in OREAS 232b (Certified Value 84 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 84.27 | 67.6111 | 68.1 | 82.3 | 84.1 | 92.5 | 77 | 87.7 | 84 | 85.6 | 81.1 | 82.0316 | 81.04 | 81.9 | 93.8 | 82.9856 | 81.6001 | [7 empty] | 1
| 1 | 2 | 84.36 | 69.2191 | 70.5 | 86.48 | 83.6 | 95.2 | 77.1 | 86 | 83.1 | 86 | 80.3 | 84.2854 | 75.31 | 80.5 | 88.6 | 96.2278 | 82.3574 | [7 empty] | 37
| 1 | 3 | 84.88 | 68.7413 | 65.8 | 84.25 | 83.1 | 96.6 | 79.2 | 89.8 | 84.3 | 86.3 | 81.2 | 85.9317 | 72 | 83.2 | 89 | 95.2678 | 81.7346 | [7 empty] | 16
| 1 | 4 | 82.94 | 68.475 | 65.7 | 84.11 | 84.9 | 89.1 | 77.3 | 83.9 | 84.8 | 82.3 | 81.9 | 83.3767 | 72.49 | 83.8 | 87.8 | 82.3911 | 80.0165 | [7 empty] | 83.6152
| 1 | 5 | 82.92 | 68.5029 | 69.4 | 86.2 | 83.7 | 97 | 79.2 | 86.5 | 85.4 | 86.2 | 80.9 | 84.1073 | 73.27 | 85.1 | 88.3 | 95.9374 | 84.3941 | [7 empty] | 102
| 1 | 6 | 84.81 | 67.3675 | 70.4 | 83.93 | 85 | 95.9 | 80.2 | 85.6 | 82.7 | 85.9 | 80.8 | 82.9156 | 76.32 | 88.7 | 88.3 | 78.287 | 83.9511
| Mean |  | 84.03 | 68.3195 | 68.3167 | 84.545 | 84.0667 | 94.3833 | 78.3333 | 86.5833 | 84.05 | 85.3833 | 81.0333 | 83.7747 | 75.0717 | 83.8667 | 89.3 | 88.5161 | 82.3423
| Median |  | 84.315 | 68.489 | 68.75 | 84.18 | 83.9 | 95.55 | 78.25 | 86.25 | 84.15 | 85.95 | 81 | 83.742 | 74.29 | 83.5 | 88.45 | 89.1267 | 82.046
| Std Dev. |  | 0.88517 | 0.70046 | 2.1683 | 1.5606 | 0.75542 | 3.0394 | 1.3677 | 2.0054 | 1.0173 | 1.5303 | 0.52789 | 1.3386 | 3.364 | 2.8472 | 2.2396 | 8.1593 | 1.6204
| Rel.Std.Dev. |  | 1.053% | 1.025% | 3.174% | 1.846% | 0.8986% | 3.22% | 1.746% | 2.316% | 1.21% | 1.792% | 0.6514% | 1.598% | 4.481% | 3.395% | 2.508% | 9.218% | 1.968%
| PDM3 |  | 0.4961% | -18.29% | -18.3% | 1.112% | 0.5399% | 12.88% | -6.317% | 3.55% | 0.52% | 2.115% | -3.088% | 0.1907% | -10.22% | 0.3007% | 6.799% | 5.861% | -1.522%
| Z-Score (Absolute) |  | 0 | 4.15 | 4.15 | 0.14 | 0.01 | 2.73 | 1.5 | 0.67 | 0.01 | 0.36 | 0.79 | 0.07 | 2.37 | 0.04 | 1.39 | 1.18 | 0.45
| Analytical results for Re in OREAS 232b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0011 | [12 empty] | 1
| 1 | 2 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0011 | [12 empty] | 38
| 1 | 3 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0006 | [12 empty] | 16
| 1 | 4 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0014 | [12 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0011 | [12 empty] | 103
| 1 | 6 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0009
| Mean | [12 empty] | 0.0010333
| Median | [12 empty] | 0.0011
| Std Dev. | [12 empty] | 0.00026583
| Rel.Std.Dev. | [12 empty] | 25.73%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Rh in OREAS 232b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Rh | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.0009 | [23 empty] | 1
| 1 | 2 | 0.0006 | [23 empty] | 13
| 1 | 3 | 0.0003 | [23 empty] | 16
| 1 | 4 | 0.0009 | [23 empty] | 0.00076667
| 1 | 5 | 0.0008 | [23 empty] | 19
| 1 | 6 | 0.0011
| Mean |  | 0.00076667
| Median |  | 0.00085
| Std Dev. |  | 0.00028048
| Rel.Std.Dev. |  | 36.58%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 232b (Indicative Value < 0.0028 ppm) | [25 empty] | Indicative
Ru | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | < 0.0028 | [23 empty] | 1
| 1 | 2 | < 0.0047 | [23 empty] | 14
| 1 | 3 | 0.0028 | [23 empty] | 16
| 1 | 4 | < 0.0019 | [23 empty] | < 0.0028
| 1 | 5 | < 0.0093 | [23 empty] | 20
| Mean |  | 0.0028
| Median |  | 0.0028
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 232b (Certified Value 0.181 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.17 | 0.19241 | 0.173 | 0.152 | 0.18 | 0.14 | 0.18 | 0.19 | 0.23 | 0.19 | 0.19 | 0.16 | 0.2 | 0.149 | 0.17 | 0.00641 | 0.18 | 0.19 | 0.2049 | 0.189 |  |  |  |  | 1
| 1 | 2 | 0.17 | 0.19281 | 0.174 | 0.14933 | 0.19 | 0.13 | 0.18 | 0.19 | 0.2 | 0.19 | 0.19 | 0.16 | 0.19 | 0.146 | 0.17 | 0.006012 | 0.18 | 0.19 | 0.2057 | 0.1455 |  |  |  |  | 23
| 1 | 3 | 0.17 | 0.19875 | 0.172 | 0.14667 | 0.18 | 0.13 | 0.18 | 0.19 | 0.22 | 0.19 | 0.19 | 0.16 | 0.19 | 0.149 | 0.17 | 0.006002 | 0.18 | 0.19 | 0.2057 | 0.3419 |  |  |  |  | 16
| 1 | 4 | 0.16 | 0.19303 | 0.168 | 0.14933 | 0.18 | 0.14 | 0.18 | 0.19 | 0.2 | 0.2 | 0.18 | 0.17 | 0.19 | 0.146 | 0.17 | 0.005748 | 0.17 | 0.19 | 0.2083 | 2.1559 |  |  |  |  | 0.18064
| 1 | 5 | 0.17 | 0.19311 | 0.172 | 0.14933 | 0.18 | 0.14 | 0.18 | 0.19 | 0.2 | 0.2 | 0.19 | 0.17 | 0.18 | 0.141 | 0.17 | 0.005938 | 0.18 | 0.19 | 0.2097 | 0.2787 |  |  |  |  | 104
| 1 | 6 | 0.18 | 0.19043 | 0.176 | 0.15467 | 0.19 | 0.15 | 0.18 | 0.2 | 0.2 | 0.2 | 0.19 | 0.17 | 0.18 | 0.148 | 0.17 | 0.005753 | 0.17 | 0.19 | 0.207 | 0.2529
| Mean |  | 0.17 | 0.19342 | 0.1725 | 0.15022 | 0.18333 | 0.13833 | 0.18 | 0.19167 | 0.20833 | 0.195 | 0.18833 | 0.165 | 0.18833 | 0.1465 | 0.17 | 0.0059772 | 0.17667 | 0.19 | 0.20688 | 0.56065
| Median |  | 0.17 | 0.19292 | 0.1725 | 0.14933 | 0.18 | 0.14 | 0.18 | 0.19 | 0.2 | 0.195 | 0.19 | 0.165 | 0.19 | 0.147 | 0.17 | 0.00597 | 0.18 | 0.19 | 0.20635 | 0.2658
| Std Dev. |  | 0.0063246 | 0.0027917 | 0.0026646 | 0.0027541 | 0.005164 | 0.0075277 | 3.0405e-17 | 0.0040825 | 0.013292 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0075277 | 0.0030166 | 0 | 0.00024227 | 0.005164 | 3.0405e-17 | 0.0018269 | 0.78452
| Rel.Std.Dev. |  | 3.72% | 1.443% | 1.545% | 1.833% | 2.817% | 5.442% | 1.689e-14% | 2.13% | 6.38% | 2.809% | 2.168% | 3.32% | 3.997% | 2.059% | 0% | 4.053% | 2.923% | 1.6e-14% | 0.8831% | 139.9%
| PDM3 |  | -5.888% | 7.079% | -4.504% | -16.84% | 1.494% | -23.42% | -0.3516% | 6.107% | 15.33% | 7.952% | 4.262% | -8.656% | 4.262% | -18.9% | -5.888% | -96.69% | -2.197% | 5.184% | 14.53% | 210.4%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 0.53 | 1.81 | 0.1 | 2.5 | 0.1 | 0.58 | 1.54 | 0.77 | 0.38 | 0.96 | 0.38 | 2.02 | 0.67 | 10.12 | 0.29 | 0.48 | 1.45 | 21.82
| Analytical results for Sb in OREAS 232b (Certified Value 141 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 0
| 1 | 1 | 159.48 | 121 | 147.0533 | 119.46 | 90.97 | 126 | 132.34 | 153 | 149 | 153 | 129 | 176.26 | 163.4241 | 109 | 185.6 | 116 | 177.32 | 156.5 | 122.84 | 105 | 162.1188 |  |  |  | 1
| 1 | 2 | 157.27 | 119 | 146.7867 | 119.83 | 79.34 | 129 | 144.67 | 151 | 148.5 | 155.5 | 130.5 | 178.18 | 162.9525 | 119 | 174.4 | 117 | 153.24 | 156.1 | 126.44 | 101.4 | 161.7382 |  |  |  | 39
| 1 | 3 | 163.56 | 115 | 148.555 | 120.85 | 84.19 | 130 | 140.23 | 155 | 154 | 159.5 | 134 | 171.03 | 164.8165 | 116 | 169.4 | 120 | 158.25 | 156.1 | 135.1 | 101.5 | 161.7809 |  |  |  | 16
| 1 | 4 | 162.07 | 116 | 148.0967 | 120.47 | 89.05 | 114 | 133.21 | 149.5 | 157.5 | 150 | 138 | 170.47 | 164.692 | 122 | 167.8 | 117 | 159.67 | 156.1 | 134 | 105.5 | 161.8915 |  |  |  | 140.5588
| 1 | 5 | 165.1 | 117 | 146.505 | 119.46 | 85.14 | 129 | 144.72 | 150 | 153.5 | 152.5 | 136.5 | 167.13 | 164.7453 | 119 | 150 | 117 | 170.66 | 155.8 | 136.62 | 109.7 | 161.9285 |  |  |  | 105
| 1 | 6 | 169.47 | 120 | 148.1967 | 121.05 | 82.99 | 126 | 148.5 | 151.5 | 151.5 | 151.5 | 137 | 173.28 | 163.0772 | 115 | 161.5 | 121 | 173.98 | 154.4 | 128.03 | 102.4 | 161.9285
| Mean |  | 162.825 | 118 | 147.5322 | 120.1867 | 85.28 | 125.6667 | 140.6117 | 151.6667 | 152.3333 | 153.6667 | 134.1667 | 172.725 | 163.9513 | 116.6667 | 168.1167 | 118 | 165.52 | 155.8333 | 130.505 | 104.25 | 161.8977
| Median |  | 162.815 | 118 | 147.575 | 120.15 | 84.665 | 127.5 | 142.45 | 151.25 | 152.5 | 152.75 | 135.25 | 172.155 | 164.0581 | 117.5 | 168.6 | 117 | 165.165 | 156.1 | 131.015 | 103.7 | 161.91
| Std Dev. |  | 4.3009 | 2.3664 | 0.85399 | 0.69985 | 4.203 | 5.9554 | 6.617 | 2.0412 | 3.3862 | 3.3862 | 3.6968 | 4.0453 | 0.89076 | 4.5019 | 11.9907 | 2 | 9.7479 | 0.73666 | 5.5159 | 3.1917 | 0.13409
| Rel.Std.Dev. |  | 2.641% | 2.005% | 0.5789% | 0.5823% | 4.928% | 4.739% | 4.706% | 1.346% | 2.223% | 2.204% | 2.755% | 2.342% | 0.5433% | 3.859% | 7.132% | 1.695% | 5.889% | 0.4727% | 4.227% | 3.062% | 0.08283%
| PDM3 |  | 15.84% | -16.05% | 4.961% | -14.49% | -39.33% | -10.59% | 0.03762% | 7.903% | 8.377% | 9.326% | -4.548% | 22.88% | 16.64% | -17% | 19.61% | -16.05% | 17.76% | 10.87% | -7.153% | -25.83% | 15.18%
| Z-Score (Absolute) |  | 0.61 | 1.17 | 0 | 1.08 | 2.47 | 0.87 | 0.27 | 0.16 | 0.19 | 0.24 | 0.53 | 1 | 0.65 | 1.22 | 0.82 | 1.17 | 0.71 | 0.33 | 0.67 | 1.71 | 0.57
| Analytical results for Sc in OREAS 232b (Certified Value 6.73 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 6.86 | 6.1328 | 6.6 | 6.29 | 6 | 7 | 7 | 6.4 | 7.6 | 7.3 | 7 | 6.9 | 6.4 | 7.4086 | 6 | 6.303 | 6.7 | 6.5 | 7 | 6.45 | 10.0334 | 6.6527 |  |  | 1
| 1 | 2 | 6.88 | 6.0527 | 6.4 | 6.3167 | 7 | 7 | 7 | 6.4 | 7.3 | 7.2 | 6.9 | 7 | 6.3 | 7.4311 | 6 | 6.351 | 6.9 | 6.8 | 6.6 | 6.43 | 9.9934 | 6.6973 |  |  | 40
| 1 | 3 | 6.88 | 5.9973 | 6.2 | 6.3067 | 7 | 7 | 7 | 6.5 | 7.7 | 7.5 | 6.9 | 7.3 | 6.2 | 7.6318 | 6 | 6.282 | 6.9 | 6.8 | 6.7 | 6.48 | 10.1586 | 6.7433 |  |  | 16
| 1 | 4 | 6.7 | 5.9567 | 6 | 6.3567 | 6 | 7 | 7 | 6.4 | 7.4 | 7.3 | 6.9 | 7.1 | 6.2 | 7.2987 | 6 | 6.085 | 7.2 | 6.7 | 6.6 | 6.53 | 10.0364 | 6.5479 |  |  | 6.7317
| 1 | 5 | 6.71 | 5.973 | 6.4 | 6.32 | 7 | 7 | 7 | 6.6 | 7.4 | 7.5 | 6.9 | 7.2 | 6.2 | 7.4308 | 6 | 6.701 | 7 | 6.6 | 6.6 | 6.44 | 10.0785 | 6.9288 |  |  | 106
| 1 | 6 | 6.96 | 5.97 | 6.1 | 6.4067 | 6 | 7 | 7 | 6.6 | 7.6 | 7.5 | 7.1 | 7.1 | 6.2 | 7.7018 | 6 | 6.074 | 7.4 | 6.9 | 6.6 | 6.47 | 10.0795 | 6.6744
| Mean |  | 6.8317 | 6.0138 | 6.2833 | 6.3328 | 6.5 | 7 | 7 | 6.4833 | 7.5 | 7.3833 | 6.95 | 7.1 | 6.25 | 7.4838 | 6 | 6.2993 | 7.0167 | 6.7167 | 6.6833 | 6.4667 | 10.0633 | 6.7074
| Median |  | 6.87 | 5.9851 | 6.3 | 6.3183 | 6.5 | 7 | 7 | 6.45 | 7.5 | 7.4 | 6.9 | 7.1 | 6.2 | 7.4309 | 6 | 6.2925 | 6.95 | 6.75 | 6.6 | 6.46 | 10.0574 | 6.6859
| Std Dev. |  | 0.104 | 0.067517 | 0.22286 | 0.042343 | 0.54772 | 0 | 0 | 0.098319 | 0.15492 | 0.13292 | 0.083666 | 0.14142 | 0.083666 | 0.15162 | 0 | 0.22849 | 0.24833 | 0.1472 | 0.16021 | 0.036148 | 0.056708 | 0.12646
| Rel.Std.Dev. |  | 1.522% | 1.123% | 3.547% | 0.6686% | 8.427% | 0% | 0% | 1.516% | 2.066% | 1.8% | 1.204% | 1.992% | 1.339% | 2.026% | 0% | 3.627% | 3.539% | 2.192% | 2.397% | 0.559% | 0.5635% | 1.885%
| PDM3 |  | 1.485% | -10.67% | -6.66% | -5.926% | -3.442% | 3.986% | 3.986% | -3.689% | 11.41% | 9.68% | 3.243% | 5.471% | -7.156% | 11.17% | -10.87% | -6.423% | 4.233% | -0.2232% | -0.7184% | -3.937% | 49.49% | -0.3605%
| Z-Score (Absolute) |  | 0.21 | 1.23 | 0.75 | 0.67 | NA | NA | NA | 0.4 | 1.39 | 1.18 | 0.42 | 0.68 | 0.81 | 1.36 | NA | 0.73 | 0.54 | 0.01 | 0.05 | 0.43 | 5.89 | 0.01
Results from laboratories 1.08, 1.09, 1.10 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 232b (Indicative Value 0.23 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.13 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 1.9 | 0.4 | 0.2 | < 0.2 | < 0.2 | 0.2 | < 1 | < 0.1 | < 1 | < 0.2 | 0.149 | 3.1002 | 0.24313 | [5 empty] | 1
| 1 | 2 | 0.1 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 2.1 | 0.3 | 0.2 | 0.3 | 0.2 | 0.2 | < 1 | 0.2 | < 1 | < 0.2 | 0.1697 | 2.7389 | 0.2647 | [5 empty] | 41
| 1 | 3 | 0.13 | 0.2 | < 5 | < 1 | 0.8 | < 1 | 2 | 0.3 | 0.3 | < 0.2 | 0.3 | 0.1 | < 1 | 0.3 | < 1 | < 0.2 | 0.1642 | 3.583 | 0.27017 | [5 empty] | 16
| 1 | 4 | 0.13 | 0.1 | < 5 | < 1 | 0.7 | < 1 | 2.1 | < 0.2 | 0.3 | 0.2 | 0.3 | < 0.1 | < 1 | 0.4 | < 1 | < 0.2 | 0.149 | 2.1586 | 0.25378 | [5 empty] | 0.23362
| 1 | 5 | 0.13 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 1.9 | 0.3 | 0.2 | 0.3 | 0.2 | 0.3 | < 1 | 0.5 | < 1 | < 0.2 | 0.1711 | 3.0795 | 0.27697 | [5 empty] | 14
| 1 | 6 | 0.14 | 0.3 | < 5 | < 1 | 0.8 | < 1 | 2.1 | 0.2 | 0.2 | 0.3 | 0.2 | 0.1 | < 1 | 0.6 | < 1 | < 0.2 | 0.218 | 2.8387 | 0.28285
| Mean |  | 0.12667 | 0.2 |  |  | 0.73333 |  | 2.0167 | 0.3 | 0.23333 | 0.275 | 0.24 | 0.18 |  | 0.4 |  |  | 0.17017 | 2.9165 | 0.26527
| Median |  | 0.13 | 0.2 |  |  | 0.7 |  | 2.05 | 0.3 | 0.2 | 0.3 | 0.2 | 0.2 |  | 0.4 |  |  | 0.16695 | 2.9591 | 0.26743
| Std Dev. |  | 0.013663 | 0.063246 |  |  | 0.05164 |  | 0.098319 | 0.070711 | 0.05164 | 0.05 | 0.054772 | 0.083666 |  | 0.15811 |  |  | 0.02538 | 0.47256 | 0.014782
| Rel.Std.Dev. |  | 10.79% | 31.62% |  |  | 7.042% |  | 4.875% | 23.57% | 22.13% | 18.18% | 22.82% | 46.48% |  | 39.53% |  |  | 14.91% | 16.2% | 5.572%
| PDM3 |  | -45.78% | -14.39% |  |  | 213.9% |  | 763.2% | 28.41% | -0.1227% | 17.71% | 2.731% | -22.95% |  | 71.22% |  |  | -27.16% | 1148% | 13.55%
| Z-Score (Absolute) |  | 0.67 | 0.32 | 10.76 | 1.12 | 2.25 | 1.12 | 8.43 | 0 | 0.16 | 0.24 | 0.24 | 0.52 | 1.12 | 0.36 | 1.12 | 0.8 | 0.46 | 12.76 | 0.01
| Analytical results for Si in OREAS 232b (Indicative Value 0.405 wt.%) | [25 empty] | Indicative
Si | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | wt.%
|  |  | AR*OES | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.4147 | [23 empty] | 1
| 1 | 2 | 0.4089 | [23 empty] | 9
| 1 | 3 | 0.4065 | [23 empty] | 16
| 1 | 4 | 0.4168 | [23 empty] | 0.40472
| 1 | 5 | 0.3998 | [23 empty] | 15
| 1 | 6 | 0.3816
| Mean |  | 0.40472
| Median |  | 0.4077
| Std Dev. |  | 0.012845
| Rel.Std.Dev. |  | 3.174%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 232b (Indicative Value 4.31 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 5.0521 | 3.4 | 4.156 | 3.6397 | [20 empty] | 1
| 1 | 2 | 5.0572 | 4.3 | 4.358 | 4.3796 | [20 empty] | 10
| 1 | 3 | 5.0341 | 4.1 | 4.458 | 4.3835 | [20 empty] | 16
| 1 | 4 | 4.9581 | 4 | 4.261 | 3.6596 | [20 empty] | 4.3109
| 1 | 5 | 4.9785 | 3.5 | 4.645 | 4.4266 | [20 empty] | 16
| 1 | 6 | 4.9959 | 3.6 | 4.615 | 3.5041
| Mean |  | 5.0127 | 3.8167 | 4.4155 | 3.9989
| Median |  | 5.015 | 3.8 | 4.408 | 4.0196
| Std Dev. |  | 0.041039 | 0.3656 | 0.19432 | 0.43926
| Rel.Std.Dev. |  | 0.8187% | 9.579% | 4.401% | 10.98%
| PDM3 |  | 16.28% | -11.47% | 2.426% | -7.239%
| Z-Score (Absolute) |  | 1.81 | 0.88 | 0.47 | 0.47
| Analytical results for Sn in OREAS 232b (Certified Value 1.62 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.64 | 2.622 | 1.59 | < 10 | 1.5 | 1.7 | 1.8 | 1.5 | 1.6 | 1.6 | 1.6 | 1.7 | 1.7431 | < 1 | 1.5 | < 5 | 1.7 | 1.3288 | 2.9368 | [5 empty] | 1
| 1 | 2 | 1.58 | 2.6265 | 1.71 | < 10 | 1.5 | 1.7 | 1.7 | 1.5 | 1.5 | 1.5 | 1.6 | 1.6 | 1.7085 | < 1 | 1.5 | < 5 | 1.6 | 1.5656 | 2.731 | [5 empty] | 26
| 1 | 3 | 1.63 | 2.5723 | 1.53 | < 10 | 1.5 | 1.6 | 1.8 | 1.6 | 1.6 | 1.6 | 1.6 | 1.7 | 1.727 | < 1 | 1.6 | < 5 | 1.6 | 1.5523 | 3.0546 | [5 empty] | 16
| 1 | 4 | 1.65 | 2.6429 | 1.55 | < 10 | 1.5 | 1.8 | 1.7 | 1.6 | 1.5 | 1.6 | 1.6 | 1.6 | 1.719 | < 1 | 1.6 | < 5 | 1.5 | 1.2688 | 2.7639 | [5 empty] | 1.6152
| 1 | 5 | 1.63 | 2.6283 | 1.59 | < 10 | 1.5 | 1.8 | 1.9 | 1.6 | 1.5 | 1.6 | 1.6 | 1.7 | 1.7235 | < 1 | 1.7 | < 5 | 1.6 | 1.5554 | 2.8264 | [5 empty] | 107
| 1 | 6 | 1.63 | 2.5593 | 1.57 | < 10 | 1.6 | 1.7 | 1.8 | 1.6 | 1.6 | 1.6 | 1.6 | 1.7 | 1.7401 | < 1 | 1.7 | < 5 | 1.6 | 1.2379 | 2.6811
| Mean |  | 1.6267 | 2.6086 | 1.59 |  | 1.5167 | 1.7167 | 1.7833 | 1.5667 | 1.55 | 1.5833 | 1.6 | 1.6667 | 1.7269 |  | 1.6 |  | 1.6 | 1.4181 | 2.8323
| Median |  | 1.63 | 2.6243 | 1.58 |  | 1.5 | 1.7 | 1.8 | 1.6 | 1.55 | 1.6 | 1.6 | 1.7 | 1.7253 |  | 1.6 |  | 1.6 | 1.4405 | 2.7952
| Std Dev. |  | 0.024221 | 0.034088 | 0.063246 |  | 0.040825 | 0.075277 | 0.075277 | 0.05164 | 0.054772 | 0.040825 | 2.4324e-16 | 0.05164 | 0.013038 |  | 0.089443 |  | 0.063246 | 0.15579 | 0.14012
| Rel.Std.Dev. |  | 1.489% | 1.307% | 3.978% |  | 2.692% | 4.385% | 4.221% | 3.296% | 3.534% | 2.578% | 1.52e-14% | 3.098% | 0.755% |  | 5.59% |  | 3.953% | 10.99% | 4.947%
| PDM3 |  | 0.7102% | 61.5% | -1.56% |  | -6.1% | 6.282% | 10.41% | -3.005% | -4.036% | -1.973% | -0.9408% | 3.187% | 6.914% |  | -0.9408% |  | -0.9408% | -12.2% | 75.35%
| Z-Score (Absolute) |  | 0.22 | 8.16 | 0.08 | 27.51 | 0.67 | 0.94 | 1.48 | 0.27 | 0.4 | 0.13 | 0 | 0.54 | 1.03 | 8.9 | 0 | 7.28 | 0 | 1.47 | 9.97
| Analytical results for Sr in OREAS 232b (Certified Value 29 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 28.73 | 29.1128 | 38.3 | 39.3167 | 29.4 | 32 | 39.1 | 29 | 31.4 | 28.1 | 28.4 | 29.4 | 30.2 | 30.4977 | 28 | 28.4 | 27.4 | 25.1 | 32.1 | 22.8 | 125.3507 | 26.4772 |  |  | 1
| 1 | 2 | 28.69 | 29.1027 | 40.8 | 39.3167 | 30.7 | 29 | 39.9 | 28 | 29.8 | 28.3 | 29 | 29.8 | 29.8 | 30.4415 | 29 | 34.24 | 27.5 | 24.9 | 30.5 | 22.8 | 132.9807 | 28.3542 |  |  | 27
| 1 | 3 | 29.37 | 28.9276 | 36.9 | 39.4033 | 30 | 29 | 40.2 | 29 | 31.5 | 28.4 | 29 | 30.1 | 29.7 | 31.1921 | 28 | 26.38 | 27.5 | 26.4 | 30.6 | 22.9 | 126.065 | 28.255 |  |  | 16
| 1 | 4 | 29.24 | 28.5391 | 34.8 | 39.2333 | 30.6 | 30 | 36.4 | 28 | 29.8 | 28.4 | 28.7 | 29.8 | 29.4 | 30.3232 | 29 | 25.22 | 27.8 | 25.5 | 30.3 | 23 | 126.463 | 27.3775 |  |  | 28.9839
| 1 | 5 | 28.25 | 29.1472 | 36.4 | 39.7133 | 30.2 | 31 | 40.8 | 29 | 30.4 | 29.2 | 29.6 | 30.1 | 29.3 | 30.6186 | 28 | 27.27 | 27.4 | 27.7 | 31 | 22.8 | 126.8611 | 29.8757 |  |  | 108
| 1 | 6 | 29.85 | 28.6204 | 38.8 | 39.8733 | 29.2 | 31 | 39.9 | 29 | 30.3 | 28.1 | 29 | 29.8 | 28.9 | 30.871 | 28 | 27.49 | 27.5 | 26.7 | 31.2 | 22.9 | 125.4777 | 30.1709
| Mean |  | 29.0217 | 28.9083 | 37.6667 | 39.4761 | 30.0167 | 30.3333 | 39.3833 | 28.6667 | 30.5333 | 28.4167 | 28.95 | 29.8333 | 29.55 | 30.6573 | 28.3333 | 28.1667 | 27.5167 | 26.05 | 30.95 | 22.8667 | 127.1997 | 28.4184
| Median |  | 28.985 | 29.0151 | 37.6 | 39.36 | 30.1 | 30.5 | 39.9 | 29 | 30.35 | 28.35 | 29 | 29.8 | 29.55 | 30.5582 | 28 | 27.38 | 27.5 | 25.95 | 30.8 | 22.85 | 126.264 | 28.3046
| Std Dev. |  | 0.5737 | 0.26694 | 2.0935 | 0.25657 | 0.61455 | 1.2111 | 1.5613 | 0.5164 | 0.75277 | 0.40702 | 0.39875 | 0.2582 | 0.45056 | 0.32151 | 0.5164 | 3.1647 | 0.1472 | 1.0766 | 0.65345 | 0.08165 | 2.8896 | 1.4197
| Rel.Std.Dev. |  | 1.977% | 0.9234% | 5.558% | 0.6499% | 2.047% | 3.993% | 3.964% | 1.801% | 2.465% | 1.432% | 1.377% | 0.8655% | 1.525% | 1.049% | 1.823% | 11.24% | 0.5349% | 4.133% | 2.111% | 0.3571% | 2.272% | 4.996%
| PDM3 |  | 0.1304% | -0.2608% | 29.96% | 36.2% | 3.563% | 4.656% | 35.88% | -1.094% | 5.346% | -1.957% | -0.1169% | 2.931% | 1.953% | 5.774% | -2.245% | -2.82% | -5.062% | -10.12% | 6.783% | -21.11% | 338.9% | -1.951%
| Z-Score (Absolute) |  | 0.31 | 0.38 | 4.83 | 5.91 | 0.28 | NA | 5.85 | NA | 0.59 | 0.67 | 0.36 | 0.17 | 0 | 0.66 | NA | 0.82 | 1.21 | 2.08 | 0.83 | 3.98 | 58.1 | 0.67
Results from laboratories 1.09, 1.11 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 232b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.05 | 0.0457 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0305 | [10 empty] | 28
| 1 | 3 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0309 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0386 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0736 | [10 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0671
| Mean | [5 empty] | 0.066667 |  |  |  | 0.01 |  |  |  | 0.05 | 0.047733
| Median | [5 empty] | 0.07 |  |  |  | 0.01 |  |  |  | 0.05 | 0.04215
| Std Dev. | [5 empty] | 0.005164 |  |  |  | 0 |  |  |  |  | 0.018504
| Rel.Std.Dev. | [5 empty] | 7.746% |  |  |  | 0% |  |  |  |  | 38.77%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 83.45 | 0 | 1.4 | 0 | 0.84 | 0.67 | 0.51 | 0.67 | 0.67 | 0 | 0.42 | 0.77
| Analytical results for Tb in OREAS 232b (Indicative Value 0.45 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.53935 | 0.4 | 0.439 | 0.41 | 1.44 | [19 empty] | 1
| 1 | 2 | 0.51586 | 0.4 | 0.454 | 0.42 | 1.46 | [19 empty] | 11
| 1 | 3 | 0.54095 | 0.4 | 0.462 | 0.42 | 1.47 | [19 empty] | 16
| 1 | 4 | 0.51867 | 0.4 | 0.442 | 0.42 | 1.43 | [19 empty] | 0.45098
| 1 | 5 | 0.52918 | 0.4 | 0.481 | 0.41 | 1.44 | [19 empty] | 17
| 1 | 6 | 0.53139 | 0.4 | 0.47 | 0.42 | 1.47
| Mean |  | 0.52923 | 0.4 | 0.458 | 0.41667 | 1.4517
| Median |  | 0.53028 | 0.4 | 0.458 | 0.42 | 1.45
| Std Dev. |  | 0.010342 | 6.0809e-17 | 0.01626 | 0.005164 | 0.017224
| Rel.Std.Dev. |  | 1.954% | 1.52e-14% | 3.55% | 1.239% | 1.186%
| PDM3 |  | 17.35% | -11.3% | 1.558% | -7.608% | 221.9%
| Z-Score (Absolute) |  | 0.83 | 0.67 | 0 | 0.48 | 11.55
| Analytical results for Te in OREAS 232b (Indicative Value 0.021 ppm) | [25 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.09 | 5.15 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.04 | 0.0358 | [9 empty] | 1
| 1 | 2 | < 0.02 | 0.05 | 5.013 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.03 | < 0.05 | 0.04 | 0.012 | [9 empty] | 6
| 1 | 3 | 0.02 | 0.05 | 5.6442 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | < 0.05 | 0.03 | 0.0119 | [9 empty] | 16
| 1 | 4 | < 0.02 | 0.06 | 5.1205 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.03 | 0.02 | 0.02 | 0.01 | < 0.02 | < 0.05 | 0.04 | [10 empty] | 0.021167
| 1 | 5 | < 0.02 | 0.02 | 5.1768 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.04 | 0.0239 | [9 empty] | 18
| 1 | 6 | 0.02 | 0.04 | < 5 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | < 0.02 | < 0.05 | 0.03 | 0.0239
| Mean |  | 0.02 | 0.051667 | 5.2209 |  | 0.043333 |  |  | 0.021667 | 0.023333 | 0.021667 | 0.015 | 0.025 |  | 0.036667 | 0.0215
| Median |  | 0.02 | 0.05 | 5.15 |  | 0.04 |  |  | 0.02 | 0.02 | 0.02 | 0.015 | 0.025 |  | 0.04 | 0.0239
| Std Dev. |  | 0 | 0.023166 | 0.24466 |  | 0.005164 |  |  | 0.0040825 | 0.005164 | 0.0040825 | 0.0054772 | 0.0057735 |  | 0.005164 | 0.0099802
| Rel.Std.Dev. |  | 0% | 44.84% | 4.686% |  | 11.92% |  |  | 18.84% | 22.13% | 18.84% | 36.51% | 23.09% |  | 14.08% | 46.42%
| PDM3 |  | -5.512% | 144.1% | 24566% |  | 104.7% |  |  | 2.362% | 10.24% | 2.362% | -29.13% | 18.11% |  | 73.23% | 1.575%
| Z-Score (Absolute) |  | 1.35 | 3.6 | 639.57 | 0 | 2.47 | 0 | 3.37 | 0.45 | 0.22 | 0.45 | 1.35 | 0.67 | 0 | 1.57 | 0.47
| Analytical results for Th in OREAS 232b (Certified Value 11.6 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 10.833 | 12.4766 | 12.2 | 12.41 | 13 | 15 | 11.6 | 12.5 | 11.1 | 11.4 | 11.4 | 11.2 | 11.4291 | 11.32 | 11.4 | 17.63 | 15.7 | 12.2 | 6.4991 | 10.4646 |  |  |  |  | 1
| 1 | 2 | 10.698 | 12.3433 | 12.7 | 12.83 | 11 | 14.3 | 11.9 | 11.9 | 11.2 | 11.4 | 11.3 | 11 | 11.8882 | 11.88 | 11.9 | 16.23 | 15.5 | 11.5 | 8.6819 | 10.4656 |  |  |  |  | 31
| 1 | 3 | 10.941 | 12.2093 | 12.1 | 12.43 | 12 | 14.4 | 12 | 12.3 | 11.8 | 11.8 | 11.4 | 11.4 | 11.856 | 11.95 | 11.1 | 19.01 | 15.4 | 11.6 | 9.0796 | 10.6864 |  |  |  |  | 16
| 1 | 4 | 11.054 | 12.1511 | 12.1 | 12.38 | 11 | 13.6 | 11.1 | 11.8 | 11.4 | 11.1 | 11.4 | 10.8 | 11.3884 | 11.48 | 11.6 | 12.61 | 15 | 11.4 | 7.6428 | 10.9511 |  |  |  |  | 11.6413
| 1 | 5 | 10.806 | 12.1187 | 12.2 | 12.86 | 12 | 14.6 | 12.1 | 11.9 | 11 | 11.8 | 11.4 | 11.5 | 12.1578 | 12.26 | 11.6 | 19.97 | 15.2 | 11.9 | 9.4179 | 10.5127 |  |  |  |  | 110
| 1 | 6 | 11.07 | 12.1695 | 12.1 | 12.31 | 12 | 13.9 | 12.3 | 11.7 | 11.4 | 11.9 | 11.3 | 11.3 | 12.6452 | 11.79 | 11.6 | 19.13 | 15.3 | 11.9 | 7.4154 | 10.5724
| Mean |  | 10.9003 | 12.2447 | 12.2333 | 12.5367 | 11.8333 | 14.3 | 11.8333 | 12.0167 | 11.3167 | 11.5667 | 11.3667 | 11.2 | 11.8941 | 11.78 | 11.5333 | 17.43 | 15.35 | 11.75 | 8.1228 | 10.6088
| Median |  | 10.887 | 12.1894 | 12.15 | 12.42 | 12 | 14.35 | 11.95 | 11.9 | 11.3 | 11.6 | 11.4 | 11.25 | 11.8721 | 11.835 | 11.6 | 18.32 | 15.35 | 11.75 | 8.1623 | 10.5426
| Std Dev. |  | 0.14726 | 0.13786 | 0.23381 | 0.24246 | 0.75277 | 0.498 | 0.4274 | 0.31252 | 0.28577 | 0.31411 | 0.05164 | 0.26077 | 0.47054 | 0.33793 | 0.26583 | 2.7037 | 0.2429 | 0.30166 | 1.12 | 0.18711
| Rel.Std.Dev. |  | 1.351% | 1.126% | 1.911% | 1.934% | 6.361% | 3.482% | 3.612% | 2.601% | 2.525% | 2.716% | 0.4543% | 2.328% | 3.956% | 2.869% | 2.305% | 15.51% | 1.582% | 2.567% | 13.79% | 1.764%
| PDM3 |  | -6.365% | 5.184% | 5.086% | 7.691% | 1.65% | 22.84% | 1.65% | 3.224% | -2.789% | -0.6411% | -2.359% | -3.791% | 2.172% | 1.191% | -0.9275% | 49.73% | 31.86% | 0.9337% | -30.22% | -8.869%
| Z-Score (Absolute) |  | 1.28 | 0.68 | 0.66 | 1.1 | NA | 3.67 | 0.08 | 0.34 | 0.67 | 0.31 | 0.6 | 0.84 | 0.17 | 0 | 0.36 | 8.22 | 5.2 | 0.04 | 5.32 | 1.7
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 232b (Certified Value 0.166 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1482 | 0.14 | 0.14167 | 0.1896 | 0.18 | 0.206 | 0.156 | 0.18 | 0.175 | 0.168 | 0.163 | 0.162 | 0.18126 | 0.156 | 0.15 | 0.15431 | 0.26 | 0.172 | 0.1592 | 0.2964 | 0.15989 |  |  |  | 1
| 1 | 2 | 0.1508 | 0.13 | 0.14167 | 0.187 | 0.17 | 0.21 | 0.154 | 0.178 | 0.179 | 0.168 | 0.166 | 0.161 | 0.18147 | 0.162 | 0.15 | 0.15269 | 0.26 | 0.171 | 0.15804 | 0.2924 | 0.16474 |  |  |  | 32
| 1 | 3 | 0.1518 | 0.13 | 0.14167 | 0.1809 | 0.17 | 0.206 | 0.155 | 0.185 | 0.18 | 0.17 | 0.166 | 0.161 | 0.18263 | 0.162 | 0.15 | 0.15461 | 0.27 | 0.175 | 0.15968 | 0.2997 | 0.16366 |  |  |  | 16
| 1 | 4 | 0.1508 | 0.12 | 0.14167 | 0.1797 | 0.17 | 0.205 | 0.153 | 0.181 | 0.179 | 0.164 | 0.167 | 0.157 | 0.1844 | 0.162 | 0.15 | 0.15608 | 0.27 | 0.173 | 0.16278 | 0.2973 | 0.15992 |  |  |  | 0.16552
| 1 | 5 | 0.1491 | 0.14 | 0.14467 | 0.186 | 0.18 | 0.203 | 0.155 | 0.178 | 0.18 | 0.171 | 0.167 | 0.159 | 0.18737 | 0.156 | 0.15 | 0.1609 | 0.26 | 0.176 | 0.15926 | 0.296 | 0.17164 |  |  |  | 111
| 1 | 6 | 0.1549 | 0.13 | 0.14167 | 0.1802 | 0.18 | 0.206 | 0.155 | 0.182 | 0.181 | 0.167 | 0.167 | 0.157 | 0.18229 | 0.162 | 0.15 | 0.15809 | 0.27 | 0.175 | 0.16054 | 0.2975 | 0.16717
| Mean |  | 0.15093 | 0.13167 | 0.14217 | 0.1839 | 0.175 | 0.206 | 0.15467 | 0.18067 | 0.179 | 0.168 | 0.166 | 0.1595 | 0.18324 | 0.16 | 0.15 | 0.15611 | 0.265 | 0.17367 | 0.15992 | 0.29655 | 0.1645
| Median |  | 0.1508 | 0.13 | 0.14167 | 0.18345 | 0.175 | 0.206 | 0.155 | 0.1805 | 0.1795 | 0.168 | 0.1665 | 0.16 | 0.18246 | 0.162 | 0.15 | 0.15534 | 0.265 | 0.174 | 0.15947 | 0.29685 | 0.1642
| Std Dev. |  | 0.0023389 | 0.0075277 | 0.0012247 | 0.0041675 | 0.0054772 | 0.0022804 | 0.0010328 | 0.0026583 | 0.0020976 | 0.0024495 | 0.0015492 | 0.0021679 | 0.0023121 | 0.0030984 | 0 | 0.0029691 | 0.0054772 | 0.0019664 | 0.0016193 | 0.0024056 | 0.0044947
| Rel.Std.Dev. |  | 1.55% | 5.717% | 0.8615% | 2.266% | 3.13% | 1.107% | 0.6678% | 1.471% | 1.172% | 1.458% | 0.9332% | 1.359% | 1.262% | 1.936% | 0% | 1.902% | 2.067% | 1.132% | 1.013% | 0.8112% | 2.732%
| PDM3 |  | -8.81% | -20.45% | -14.11% | 11.11% | 5.73% | 24.46% | -6.555% | 9.154% | 8.147% | 1.501% | 0.2926% | -3.634% | 10.71% | -3.332% | -9.374% | -5.68% | 60.11% | 4.925% | -3.383% | 79.17% | -0.6118%
| Z-Score (Absolute) |  | 0.78 | 1.78 | 1.24 | 0.93 | 0.47 | 2.07 | 0.59 | 0.76 | 0.67 | 0.1 | 0 | 0.34 | 0.89 | 0.32 | 0.83 | 0.51 | 5.14 | 0.4 | 0.32 | 6.77 | 0.08
| Analytical results for Tl in OREAS 232b (Certified Value 0.49 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.49 | 0.49 | < 10 | 0.52 | 0.44 | 0.53 | 0.53 | 0.48 | 0.48 | 0.47 | 0.52 | 0.5 | 0.54575 | 0.638 | 0.45 | < 5 | 0.51 | 0.4744 | 0.6377 | [5 empty] | 1
| 1 | 2 | 0.49 | 0.51 | < 10 | 0.52 | 0.36 | 0.53 | 0.53 | 0.47 | 0.45 | 0.46 | 0.5 | 0.5 | 0.52593 | 0.697 | 0.46 | < 5 | 0.48 | 0.5006 | 0.6378 | [5 empty] | 33
| 1 | 3 | 0.49 | 0.49 | < 10 | 0.54 | 0.42 | 0.54 | 0.53 | 0.48 | 0.48 | 0.47 | 0.5 | 0.49 | 0.55177 | 0.706 | 0.43 | < 5 | 0.48 | 0.4876 | 0.6227 | [5 empty] | 16
| 1 | 4 | 0.48 | 0.47 | < 10 | 0.53 | 0.49 | 0.51 | 0.51 | 0.45 | 0.47 | 0.46 | 0.51 | 0.49 | 0.55709 | 0.703 | 0.47 | < 5 | 0.49 | 0.4199 | 0.637 | [5 empty] | 0.49302
| 1 | 5 | 0.48 | 0.53 | < 10 | 0.54 | 0.45 | 0.57 | 0.53 | 0.46 | 0.45 | 0.47 | 0.51 | 0.48 | 0.54413 | 0.702 | 0.47 | < 5 | 0.49 | 0.4709 | 0.612 | [5 empty] | 112
| 1 | 6 | 0.49 | 0.52 | < 10 | 0.53 | 0.44 | 0.55 | 0.53 | 0.46 | 0.48 | 0.47 | 0.51 | 0.51 | 0.53544 | 0.718 | 0.46 | < 5 | 0.48 | 0.3693 | 0.6216
| Mean |  | 0.48667 | 0.50167 |  | 0.53 | 0.43333 | 0.53833 | 0.52667 | 0.46667 | 0.46833 | 0.46667 | 0.50833 | 0.495 | 0.54335 | 0.694 | 0.45667 |  | 0.48833 | 0.45378 | 0.62813
| Median |  | 0.49 | 0.5 |  | 0.53 | 0.44 | 0.535 | 0.53 | 0.465 | 0.475 | 0.47 | 0.51 | 0.495 | 0.54494 | 0.7025 | 0.46 |  | 0.485 | 0.47265 | 0.62985
| Std Dev. |  | 0.005164 | 0.022286 |  | 0.0089443 | 0.04274 | 0.020412 | 0.008165 | 0.012111 | 0.01472 | 0.005164 | 0.0075277 | 0.010488 | 0.011241 | 0.02832 | 0.015055 |  | 0.01169 | 0.049677 | 0.010919
| Rel.Std.Dev. |  | 1.061% | 4.442% |  | 1.688% | 9.863% | 3.792% | 1.55% | 2.595% | 3.143% | 1.107% | 1.481% | 2.119% | 2.069% | 4.081% | 3.297% |  | 2.394% | 10.95% | 1.738%
| PDM3 |  | -1.29% | 1.753% |  | 7.5% | -12.11% | 9.19% | 6.824% | -5.346% | -5.008% | -5.346% | 3.105% | 0.4007% | 10.21% | 40.76% | -7.374% |  | -0.9515% | -7.959% | 27.4%
| Z-Score (Absolute) |  | 0.29 | 0 | 86.66 | 0.55 | 1.32 | 0.71 | 0.48 | 0.67 | 0.64 | 0.67 | 0.13 | 0.13 | 0.8 | 3.71 | 0.87 | 38.5 | 0.26 | 0.92 | 2.44
| Analytical results for Tm in OREAS 232b (Indicative Value 0.11 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.13472 | 0.1 | 0.112 | 0.1037 | [20 empty] | 1
| 1 | 2 | 0.13679 | 0.1 | 0.117 | 0.1216 | [20 empty] | 13
| 1 | 3 | 0.13625 | 0.1 | 0.114 | 0.1201 | [20 empty] | 16
| 1 | 4 | 0.12658 | 0.1 | 0.114 | 0.1006 | [20 empty] | 0.11471
| 1 | 5 | 0.13371 | 0.1 | 0.12 | 0.1174 | [20 empty] | 19
| 1 | 6 | 0.12659 | 0.1 | 0.118 | 0.1001
| Mean |  | 0.13244 | 0.1 | 0.11583 | 0.11058
| Median |  | 0.13422 | 0.1 | 0.1155 | 0.11055
| Std Dev. |  | 0.004665 | 1.5202e-17 | 0.0029944 | 0.010152
| Rel.Std.Dev. |  | 3.522% | 1.52e-14% | 2.585% | 9.181%
| PDM3 |  | 15.45% | -12.83% | 0.9756% | -3.601%
| Z-Score (Absolute) |  | 1.64 | 1.13 | 0.22 | 0.22
| Analytical results for U in OREAS 232b (Certified Value 1.4 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.263 | 1.5742 | 1.5 | 1.55 | 1.43 | 1.94 | 1.4 | 1.32 | 1.3 | 1.26 | 1.44 | 1.37 | 1.4059 | 1 | 1.512 | 1.31 | 1.5 | 1.1251 | 3.3188 | [5 empty] | 1
| 1 | 2 | 1.252 | 1.6109 | 1.5 | 1.63 | 1.33 | 1.94 | 1.4 | 1.25 | 1.28 | 1.28 | 1.41 | 1.35 | 1.4034 | 1 | 1.523 | 1.38 | 1.44 | 1.342 | 3.6184 | [5 empty] | 35
| 1 | 3 | 1.291 | 1.588 | 1.4 | 1.56 | 1.27 | 1.9 | 1.4 | 1.29 | 1.34 | 1.29 | 1.41 | 1.43 | 1.4338 | 1 | 1.541 | 1.32 | 1.47 | 1.3084 | 3.6691 | [5 empty] | 16
| 1 | 4 | 1.307 | 1.5704 | 1.4 | 1.57 | 1.35 | 1.82 | 1.4 | 1.24 | 1.32 | 1.22 | 1.43 | 1.35 | 1.4391 | 1 | 1.479 | 1.4 | 1.46 | 1.1075 | 3.3436 | [5 empty] | 1.3951
| 1 | 5 | 1.272 | 1.5515 | 1.5 | 1.65 | 1.41 | 1.92 | 1.4 | 1.28 | 1.3 | 1.28 | 1.46 | 1.37 | 1.4391 | 1 | 1.525 | 1.41 | 1.49 | 1.3057 | 3.3719 | [5 empty] | 113
| 1 | 6 | 1.285 | 1.5206 | 1.5 | 1.52 | 1.42 | 1.86 | 1.5 | 1.26 | 1.3 | 1.31 | 1.42 | 1.34 | 1.3913 | 1 | 1.585 | 1.41 | 1.5 | 1.0726 | 3.3932
| Mean |  | 1.2783 | 1.5693 | 1.4667 | 1.58 | 1.3683 | 1.8967 | 1.4167 | 1.2733 | 1.3067 | 1.2733 | 1.4283 | 1.3683 | 1.4188 | 1 | 1.5275 | 1.3717 | 1.4767 | 1.2102 | 3.4525
| Median |  | 1.2785 | 1.5723 | 1.5 | 1.565 | 1.38 | 1.91 | 1.4 | 1.27 | 1.3 | 1.28 | 1.425 | 1.36 | 1.4198 | 1 | 1.524 | 1.39 | 1.48 | 1.2154 | 3.3826
| Std Dev. |  | 0.019977 | 0.030955 | 0.05164 | 0.0498 | 0.062743 | 0.048028 | 0.040825 | 0.029439 | 0.020656 | 0.030768 | 0.019408 | 0.032506 | 0.021031 | 0 | 0.03495 | 0.04535 | 0.024221 | 0.12071 | 0.15112
| Rel.Std.Dev. |  | 1.563% | 1.973% | 3.521% | 3.152% | 4.585% | 2.532% | 2.882% | 2.312% | 1.581% | 2.416% | 1.359% | 2.376% | 1.482% | 0% | 2.288% | 3.306% | 1.64% | 9.974% | 4.377%
| PDM3 |  | -8.37% | 12.48% | 5.129% | 13.25% | -1.919% | 35.95% | 1.545% | -8.729% | -6.339% | -8.729% | 2.381% | -1.919% | 1.695% | -28.32% | 9.49% | -1.68% | 5.846% | -13.25% | 147.5%
| Z-Score (Absolute) |  | 0.85 | 0.93 | 0.3 | 0.99 | 0.3 | 2.92 | 0.01 | 0.88 | 0.68 | 0.88 | 0.06 | 0.3 | 0.01 | NA | 0.67 | 0.28 | 0.36 | 1.27 | 12.43
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 232b (Certified Value 67 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 69 | 69.5 | 57 | 63.6767 | 63 | 68 | 68 | 69 | 70 | 66 | 67 | 62 | 68 | 70.5034 | 73 | 65 | 68.33 | 67 | 62 | 61.6 | 89.6762 | 66.3607 |  |  | 1
| 1 | 2 | 68.6 | 69.15 | 59 | 63.3833 | 64 | 65 | 69 | 68 | 69 | 66 | 67 | 63 | 67 | 66.813 | 73 | 65 | 65.61 | 66.8 | 62 | 61.4 | 88.7163 | 67.0338 |  |  | 36
| 1 | 3 | 69.1 | 69.7 | 57 | 63.3033 | 63 | 66 | 68 | 70 | 73 | 66 | 68 | 62 | 66 | 69.416 | 72 | 65 | 67.95 | 67 | 63 | 61.7 | 89.1537 | 69.4158 |  |  | 16
| 1 | 4 | 68.5 | 68.91 | 56 | 64.1167 | 62 | 65 | 68 | 67 | 70 | 67 | 66 | 63 | 66 | 69.6378 | 73 | 65 | 66.66 | 66.4 | 62 | 62.3 | 89.5912 | 66.9005 |  |  | 66.6288
| 1 | 5 | 67.3 | 70.97 | 60 | 63.7933 | 63 | 68 | 67 | 70 | 70 | 67 | 67 | 63 | 65 | 67.3977 | 72 | 65 | 66.12 | 67.7 | 62 | 61.6 | 90.6361 | 68.5688 |  |  | 114
| 1 | 6 | 69.4 | 69.84 | 58 | 64.2167 | 64 | 68 | 68 | 70 | 71 | 68 | 66 | 63 | 65 | 66.4308 | 72 | 65 | 66.15 | 67.2 | 63 | 61.9 | 91.681 | 67.0919
| Mean |  | 68.65 | 69.6783 | 57.8333 | 63.7483 | 63.1667 | 66.6667 | 68 | 69 | 70.5 | 66.6667 | 66.8333 | 62.6667 | 66.1667 | 68.3664 | 72.5 | 65 | 66.8033 | 67.0167 | 62.3333 | 61.75 | 89.9091 | 67.5619
| Median |  | 68.8 | 69.6 | 57.5 | 63.735 | 63 | 67 | 68 | 69.5 | 70 | 66.5 | 67 | 63 | 66 | 68.4069 | 72.5 | 65 | 66.405 | 67 | 62 | 61.65 | 89.6337 | 67.0629
| Std Dev. |  | 0.73959 | 0.72059 | 1.472 | 0.3724 | 0.75277 | 1.5055 | 0.63246 | 1.2649 | 1.3784 | 0.8165 | 0.75277 | 0.5164 | 1.169 | 1.6961 | 0.54772 | 0 | 1.094 | 0.43089 | 0.5164 | 0.31464 | 1.0786 | 1.1688
| Rel.Std.Dev. |  | 1.077% | 1.034% | 2.545% | 0.5842% | 1.192% | 2.258% | 0.9301% | 1.833% | 1.955% | 1.225% | 1.126% | 0.824% | 1.767% | 2.481% | 0.7555% | 0% | 1.638% | 0.643% | 0.8284% | 0.5095% | 1.2% | 1.73%
| PDM3 |  | 3.034% | 4.577% | -13.2% | -4.323% | -5.196% | 0.05691% | 2.058% | 3.559% | 5.81% | 0.05691% | 0.307% | -5.947% | -0.6935% | 2.608% | 8.812% | -2.445% | 0.262% | 0.5822% | -6.447% | -7.322% | 34.94% | 1.401%
| Z-Score (Absolute) |  | 0.62 | 0.96 | 3.02 | 1.03 | 1.23 | 0.05 | 0.4 | 0.73 | 1.24 | 0.05 | 0.01 | 1.4 | 0.22 | 0.52 | 1.91 | 0.61 | 0.01 | 0.07 | 1.51 | 1.7 | 7.76 | 0.25
| Analytical results for W in OREAS 232b (Certified Value 0.34 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.35 | 0.3 | 7.2 | 0.5 | 0.33 | 0.4 | 0.4 | 0.35 | 0.38 | 0.32 | 0.37 | 0.31 | < 1 | < 1 | 0.3 | < 5 | 0.33 | 0.6172 | [6 empty] | 1
| 1 | 2 | 0.37 | 0.3 | 8.682 | 0.5 | 0.3 | 0.4 | 0.4 | 0.34 | 0.39 | 0.35 | 0.36 | 0.27 | < 1 | < 1 | 0.3 | < 5 | 0.32 | 0.3886 | [6 empty] | 37
| 1 | 3 | 0.37 | 0.3 | 7.744 | 0.4 | 0.29 | 0.3 | 0.3 | 0.34 | 0.39 | 0.33 | 0.38 | 0.31 | < 1 | < 1 | 0.3 | < 5 | 0.32 | 0.3487 | [6 empty] | 16
| 1 | 4 | 0.33 | 0.2 | 8.956 | 0.4 | 0.4 | 0.3 | 0.3 | 0.33 | 0.41 | 0.34 | 0.36 | 0.27 | < 1 | < 1 | 0.3 | < 5 | 0.32 | 0.3459 | [6 empty] | 0.33663
| 1 | 5 | 0.32 | 0.3 | 7.72 | 0.4 | 0.32 | 0.4 | 0.3 | 0.59 | 0.39 | 0.33 | 0.36 | 0.28 | < 1 | < 1 | 0.3 | < 5 | 0.35 | 0.3458 | [6 empty] | 115
| 1 | 6 | 0.35 | 0.2 | 7.384 | 0.4 | 0.37 | 0.4 | 0.3 | 0.36 | 0.39 | 0.33 | 0.63 | 0.29 | < 1 | < 1 | 0.6 | < 5 | 0.32 | 0.2855
| Mean |  | 0.34833 | 0.26667 | 7.9477 | 0.43333 | 0.335 | 0.36667 | 0.33333 | 0.385 | 0.39167 | 0.33333 | 0.41 | 0.28833 |  |  | 0.35 |  | 0.32667 | 0.38862
| Median |  | 0.35 | 0.3 | 7.732 | 0.4 | 0.325 | 0.4 | 0.3 | 0.345 | 0.39 | 0.33 | 0.365 | 0.285 |  |  | 0.3 |  | 0.32 | 0.3473
| Std Dev. |  | 0.020412 | 0.05164 | 0.71078 | 0.05164 | 0.042308 | 0.05164 | 0.05164 | 0.10095 | 0.0098319 | 0.010328 | 0.10807 | 0.018348 |  |  | 0.12247 |  | 0.012111 | 0.11673
| Rel.Std.Dev. |  | 5.86% | 19.36% | 8.943% | 11.92% | 12.63% | 14.08% | 15.49% | 26.22% | 2.51% | 3.098% | 26.36% | 6.364% |  |  | 34.99% |  | 3.707% | 30.04%
| PDM3 |  | 3.476% | -20.78% | 2261% | 28.73% | -0.4852% | 8.922% | -0.9803% | 14.37% | 16.35% | -0.9803% | 21.79% | -14.35% |  |  | 3.971% |  | -2.961% | 15.44%
| Z-Score (Absolute) |  | 0.44 | 1.73 | 120.14 | 0.91 | 0.65 | 0.15 | 0.67 | 0.15 | 0.25 | 0.67 | 0.54 | 1.39 | 1.97 | 1.97 | 0.41 | 33.7 | 0.78 | 0.2
| Analytical results for Y in OREAS 232b (Certified Value 8.84 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 8.92 | 8.3724 | 8.17 | 8.45 | 9.68 | 8 | 9.4 | 8.5 | 9.36 | 9.12 | 8.61 | 8.83 | 9.6898 | 8 | 7.907 | 8.53 | 9.35 | 7.9135 | 10.5315 | 8.0313 |  |  |  |  | 1
| 1 | 2 | 8.59 | 8.4563 | 8.38 | 8.4533 | 9.97 | 7 | 9.58 | 8.5 | 9.05 | 9.13 | 8.91 | 8.78 | 9.5894 | 9 | 8.165 | 8.44 | 8.7 | 8.0465 | 10.7105 | 8.5489 |  |  |  |  | 38
| 1 | 3 | 8.87 | 8.3912 | 7.78 | 8.42 | 9.85 | 8 | 9.62 | 8.7 | 9.43 | 9.42 | 8.94 | 8.92 | 9.5614 | 8 | 8.166 | 9.08 | 8.86 | 7.9895 | 10.7471 | 8.6799 |  |  |  |  | 16
| 1 | 4 | 8.62 | 8.3738 | 7.97 | 8.54 | 9.77 | 8 | 9.22 | 8.5 | 8.95 | 9.33 | 8.78 | 8.99 | 9.4429 | 9 | 7.933 | 9.05 | 8.81 | 8.018 | 11.1397 | 8.4002 |  |  |  |  | 8.8357
| 1 | 5 | 8.6 | 8.2262 | 8.44 | 8.5067 | 9.87 | 8 | 9.86 | 8.7 | 8.9 | 9.54 | 9.12 | 8.98 | 9.68 | 9 | 8.509 | 9.04 | 9.02 | 7.9325 | 10.4646 | 9.0357 |  |  |  |  | 116
| 1 | 6 | 8.91 | 8.3816 | 8.32 | 8.5467 | 9.66 | 8 | 9.4 | 8.7 | 9.23 | 9.1 | 9.11 | 9.09 | 9.8634 | 9 | 8.245 | 9.63 | 8.9 | 7.98 | 10.9639 | 9.2083
| Mean |  | 8.7517 | 8.3669 | 8.1767 | 8.4861 | 9.8 | 7.8333 | 9.5133 | 8.6 | 9.1533 | 9.2733 | 8.9117 | 8.9317 | 9.6378 | 8.6667 | 8.1542 | 8.9617 | 8.94 | 7.98 | 10.7595 | 8.6507
| Median |  | 8.745 | 8.3777 | 8.245 | 8.48 | 9.81 | 8 | 9.49 | 8.6 | 9.14 | 9.23 | 8.925 | 8.95 | 9.6347 | 9 | 8.1655 | 9.045 | 8.88 | 7.9848 | 10.7288 | 8.6144
| Std Dev. |  | 0.16364 | 0.075724 | 0.25742 | 0.052426 | 0.11933 | 0.40825 | 0.22259 | 0.10954 | 0.2197 | 0.18435 | 0.1959 | 0.11339 | 0.1425 | 0.5164 | 0.2211 | 0.43199 | 0.22672 | 0.050269 | 0.25606 | 0.42821
| Rel.Std.Dev. |  | 1.87% | 0.905% | 3.148% | 0.6178% | 1.218% | 5.212% | 2.34% | 1.274% | 2.4% | 1.988% | 2.198% | 1.269% | 1.479% | 5.958% | 2.711% | 4.82% | 2.536% | 0.6299% | 2.38% | 4.95%
| PDM3 |  | -0.9512% | -5.306% | -7.459% | -3.957% | 10.91% | -11.34% | 7.669% | -2.668% | 3.595% | 4.953% | 0.8596% | 1.086% | 9.078% | -1.913% | -7.714% | 1.426% | 1.18% | -9.685% | 21.77% | -2.094%
| Z-Score (Absolute) |  | 0.29 | 0.95 | 1.28 | 0.75 | 1.5 | NA | 1.01 | 0.55 | 0.4 | 0.6 | 0.02 | 0.02 | 1.23 | NA | 1.31 | 0.07 | 0.03 | 1.61 | 3.15 | 0.46
Results from laboratories 1.09 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 232b (Certified Value 0.83 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | 1.32 | [18 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | [18 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [18 empty] | 2
| 1 | 1 | 0.84967 | 0.8 | 0.703 | 0.7 | 0.95 | 1.0054 | [18 empty] | 1
| 1 | 2 | 0.84747 | 0.8 | 0.735 | 0.8 | 0.98 | 1.1681 | [18 empty] | 39
| 1 | 3 | 0.86265 | 0.8 | 0.729 | 0.8 | 0.96 | 1.1388 | [18 empty] | 16
| 1 | 4 | 0.84154 | 0.8 | 0.723 | 0.8 | 0.96 | 1.1978 | [18 empty] | 0.82623
| 1 | 5 | 0.83643 | 0.8 | 0.763 | 0.8 | 0.97 | 1.1096 | [18 empty] | 117
| 1 | 6 | 0.84221 | 0.8 | 0.754 | 0.8 | 0.98 | 1.1686
| Mean |  | 0.84666 | 0.8 | 0.7345 | 0.78333 | 0.96667 | 1.1314
| Median |  | 0.84484 | 0.8 | 0.732 | 0.8 | 0.965 | 1.1534
| Std Dev. |  | 0.0091224 | 1.2162e-16 | 0.021668 | 0.040825 | 0.012111 | 0.068623
| Rel.Std.Dev. |  | 1.077% | 1.52e-14% | 2.95% | 5.212% | 1.253% | 6.065%
| PDM3 |  | 2.473% | -3.175% | -11.1% | -5.192% | 17% | 36.93%
| Z-Score (Absolute) |  | 0.24 | 0.24 | 0.93 | 0.42 | 1.5 | 3.22
| Analytical results for Zn in OREAS 232b (Certified Value 81 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 1
| 1 | 1 | 76.5 | 81.369 | 80.2 | 75.09 | 82 | 80 | 85 | 79 | 84 | 77 | 82 | 79 | 83.8 | 86.3082 | 82 | 97.32 | 74 | 83.63 | 84.3 | 80 | 93.5424 | 84 | 78.0955 | 80.8692 | 1
| 1 | 2 | 77 | 80.553 | 83.9 | 75.5567 | 84 | 82 | 86 | 78 | 84 | 78 | 83 | 81 | 81.7 | 85.9969 | 82 | 87.37 | 76 | 80.02 | 83.4 | 81 | 93.52 | 85 | 79.3063 | 80.4998 | 40
| 1 | 3 | 75.8 | 81.114 | 81.7 | 73.38 | 81 | 82 | 85 | 80 | 89 | 79 | 83 | 79 | 81.8 | 82.0601 | 81 | 81.42 | 75 | 80.75 | 82.4 | 82 | 93.9568 | 84 | 82.0413 | 82.2977 | 16
| 1 | 4 | 75.6 | 81.653 | 77.4 | 74.72 | 81 | 85 | 85 | 81 | 85 | 79 | 81 | 79 | 80.6 | 82.2309 | 82 | 82.89 | 75 | 81.2 | 85.6 | 81 | 95.3008 | 84 | 86.4079 | 80.5024 | 81.1161
| 1 | 5 | 74.9 | 81.281 | 82.5 | 74.6067 | 84 | 80 | 83 | 80 | 84 | 80 | 83 | 80 | 80.4 | 88.3203 | 81 | 84.35 | 74 | 80.42 | 83.5 | 79 | 94.5168 | 85 | 77.6054 | 83.5305 | 118
| 1 | 6 | 76.4 | 80.52 | 81.9 | 74.3033 | 84 | 85 | 84 | 81 | 85 | 80 | 81 | 79 | 80.1 | 85.3962 | 80 | 88.19 | 75 | 82.8 | 82.4 | 79 | 96.0848 | 87 | 79.8061 | 79.5487
| Mean |  | 76.0333 | 81.0817 | 81.2667 | 74.6094 | 82.6667 | 82.3333 | 84.6667 | 79.8333 | 85.1667 | 78.8333 | 82.1667 | 79.5 | 81.4 | 85.0521 | 81.3333 | 86.9233 | 74.8333 | 81.47 | 83.6 | 80.3333 | 94.4869 | 84.8333 | 80.5438 | 81.2081
| Median |  | 76.1 | 81.1975 | 81.8 | 74.6633 | 83 | 82 | 85 | 80 | 84.5 | 79 | 82.5 | 79 | 81.15 | 85.6966 | 81.5 | 85.86 | 75 | 80.975 | 83.45 | 80.5 | 94.2368 | 84.5 | 79.5562 | 80.6858
| Std Dev. |  | 0.75011 | 0.4571 | 2.242 | 0.74103 | 1.5055 | 2.2509 | 1.0328 | 1.169 | 1.9408 | 1.169 | 0.98319 | 0.83666 | 1.3667 | 2.4569 | 0.8165 | 5.7109 | 0.75277 | 1.4304 | 1.2182 | 1.2111 | 1.0319 | 1.169 | 3.266 | 1.4456
| Rel.Std.Dev. |  | 0.9866% | 0.5638% | 2.759% | 0.9932% | 1.821% | 2.734% | 1.22% | 1.464% | 2.279% | 1.483% | 1.197% | 1.052% | 1.679% | 2.889% | 1.004% | 6.57% | 1.006% | 1.756% | 1.457% | 1.508% | 1.092% | 1.378% | 4.055% | 1.78%
| PDM3 |  | -6.266% | -0.04241% | 0.1857% | -8.021% | 1.912% | 1.501% | 4.377% | -1.581% | 4.994% | -2.814% | 1.295% | -1.992% | 0.35% | 4.852% | 0.2678% | 7.159% | -7.745% | 0.4363% | 3.062% | -0.965% | 16.48% | 4.583% | -0.7056% | 0.1134%
| Z-Score (Absolute) |  | 2.12 | 0.11 | 0.04 | 2.68 | 0.52 | 0.38 | 1.31 | 0.61 | 1.51 | 1 | 0.32 | 0.74 | 0.01 | 1.46 | 0.01 | 2.2 | 2.59 | 0.04 | 0.89 | 0.41 | 5.2 | 1.38 | 0.33 | 0.06
| Analytical results for Zr in OREAS 232b (Certified Value 20 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 18.55 | 28.62 | 10.5 | 21.7786 | 21.3 | 15.7 | 24.9 | 18.2 | 21.2 | 19.8 | 20.4 | 20.6 | 23.768 | 10 | 17.1 | 21.2 | 22.8 | 14.74 | 96.847 | 15.9492 |  |  |  |  | 1
| 1 | 2 | 18.36 | 29.92 | 12.9 | 21.9201 | 22.4 | 13.6 | 23.8 | 18.3 | 20.6 | 19.8 | 20.6 | 20.6 | 23.1473 | 14 | 17.3 | 21.8 | 21.7 | 14.8 | 96.9892 | 18.1367 |  |  |  |  | 41
| 1 | 3 | 19.25 | 30.38 | 11.1 | 21.9273 | 21.6 | 13.4 | 25.1 | 18.7 | 20 | 20.4 | 20.8 | 20.9 | 23.2984 | 13 | 17.7 | 22.5 | 21.7 | 14.63 | 97.1314 | 18.2661 |  |  |  |  | 16
| 1 | 4 | 19.06 | 29.86 | 16.4 | 22.4255 | 21.6 | 13.3 | 22.9 | 18.4 | 19.6 | 27.9 | 19.5 | 21 | 23.8184 | 14 | 17.7 | 21 | 21.1 | 14.56 | 110.471 | 17.4301 |  |  |  |  | 19.9978
| 1 | 5 | 18.46 | 31.78 | 9.8 | 22.3597 | 22.3 | 14.8 | 26 | 18.8 | 20.2 | 20 | 21.2 | 21.2 | 22.6576 | 13 | 17.9 | 22.6 | 23 | 14.77 | 97.3767 | 19.943 |  |  |  |  | 119
| 1 | 6 | 18.46 | 32.05 | 11.6 | 22.4518 | 21.5 | 15.4 | 25.1 | 18.9 | 20.2 | 20.1 | 21 | 20.6 | 23.7446 | 13 | 17.8 | 21.7 | 22.9 | 14.87 | 97.2087 | 20.4345
| Mean |  | 18.69 | 30.435 | 12.05 | 22.1438 | 21.7833 | 14.3667 | 24.6333 | 18.55 | 20.3 | 21.3333 | 20.5833 | 20.8167 | 23.4057 | 12.8333 | 17.5833 | 21.8 | 22.2 | 14.7283 | 99.3373 | 18.3599
| Median |  | 18.505 | 30.15 | 11.35 | 22.1435 | 21.6 | 14.2 | 25 | 18.55 | 20.2 | 20.05 | 20.7 | 20.75 | 23.5215 | 13 | 17.7 | 21.75 | 22.25 | 14.755 | 97.1701 | 18.2014
| Std Dev. |  | 0.37008 | 1.2892 | 2.3755 | 0.30038 | 0.4535 | 1.0671 | 1.1021 | 0.2881 | 0.54772 | 3.2247 | 0.60139 | 0.25626 | 0.45922 | 1.472 | 0.31252 | 0.65422 | 0.8 | 0.11409 | 5.4574 | 1.6459
| Rel.Std.Dev. |  | 1.98% | 4.236% | 19.71% | 1.356% | 2.082% | 7.427% | 4.474% | 1.553% | 2.698% | 15.12% | 2.922% | 1.231% | 1.962% | 11.47% | 1.777% | 3.001% | 3.604% | 0.7746% | 5.494% | 8.965%
| PDM3 |  | -6.54% | 52.19% | -39.74% | 10.73% | 8.929% | -28.16% | 23.18% | -7.24% | 1.511% | 6.679% | 2.928% | 4.095% | 17.04% | -35.83% | -12.07% | 9.012% | 11.01% | -26.35% | 396.7% | -8.19%
| Z-Score (Absolute) |  | 0.63 | 2.86 | 2.61 | 0.39 | 0.29 | 1.92 | 1.14 | 0.67 | 0.15 | 0.15 | 0.07 | 0 | 0.77 | NA | 0.96 | 0.29 | 0.41 | 1.81 | 23.36 | 0.73
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C925: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 232b (Indicative Value 13.85 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.87 | 1
| 1 | 2 | 13.82 | 16
| Mean |  | 13.845 | 16
| Median |  | 13.845 | 13.845
| Std Dev. |  | 0.035355 | 22
| Rel.Std.Dev. |  | 0.2554%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 232b (Indicative Value 1.38 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.38 | 1
| 1 | 2 | 1.38 | 17
| Mean |  | 1.38 | 16
| Median |  | 1.38 | 1.38
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 232b (Indicative Value 5.58 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.6 | 1
| 1 | 2 | 5.55 | 18
| Mean |  | 5.575 | 16
| Median |  | 5.575 | 5.575
| Std Dev. |  | 0.035355 | 24
| Rel.Std.Dev. |  | 0.6342%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 232b (Indicative Value 3.1 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.1 | 1
| 1 | 2 | 3.09 | 19
| Mean |  | 3.095 | 16
| Median |  | 3.095 | 3.095
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.2285%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 232b (Indicative Value 2.75 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.76 | 1
| 1 | 2 | 2.73 | 16
| Mean |  | 2.745 | 16
| Median |  | 2.745 | 2.745
| Std Dev. |  | 0.021213 | 22
| Rel.Std.Dev. |  | 0.7728%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 232b (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 17
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 232b (Indicative Value 1.1 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.11 | 1
| 1 | 2 | 1.09 | 18
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.014142 | 24
| Rel.Std.Dev. |  | 1.286%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 232b (Indicative Value 0.152 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.151 | 1
| 1 | 2 | 0.152 | 19
| Mean |  | 0.1515 | 16
| Median |  | 0.1515 | 0.1515
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 0.4667%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 232b (Indicative Value 69.19 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.26 | 1
| 1 | 2 | 69.11 | 16
| Mean |  | 69.185 | 16
| Median |  | 69.185 | 69.185
| Std Dev. |  | 0.10607 | 22
| Rel.Std.Dev. |  | 0.1533%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 232b (Indicative Value 0.451 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.451 | 17
| Mean |  | 0.4505 | 16
| Median |  | 0.4505 | 0.4505
| Std Dev. |  | 0.00070711 | 23
| Rel.Std.Dev. |  | 0.157%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 232b (Indicative Value 0.81 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.81 | 1
| 1 | 2 | 0.81 | 18
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 232b (Indicative Value 1.8 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.79 | 21
| Mean |  | 1.795 | 16
| Median |  | 1.795 | 1.795
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.3939%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 232b (Indicative Value 0.06 wt.%) |  |  |  | Indicative
C | Round | Replicate | BV Geo |  | 1
| No | No | Perth |  | wt.%
|  |  | IRC |  | 3
3
| 1 | 1 | 0.06 |  | 1
| 1 | 2 | 0.06 |  | 23
| Mean |  | 0.06 |  | 16
| Median |  | 0.06 |  | 0.06
| Std Dev. |  | 0 |  | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 232b (Indicative Value 0.179 wt.%) |  |  |  | Indicative
S | Round | Replicate | Inspectorate | BV Geo | 1
| No | No | Gandhidham | Perth | wt.%
|  |  | IRC | IRC | 3
3
| 1 | 1 | 0.19 | 0.17 | 1
| 1 | 2 | 0.18 | 0.17 | 23
| 1 | 3 | 0.19 |  | 16
| 1 | 4 | 0.19 |  | 0.17917
| 1 | 5 | 0.19 |  | 29
| 1 | 6 | 0.19
| Mean |  | 0.18833 | 0.17
| Median |  | 0.19 | 0.17
| Std Dev. |  | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.168% | 0%
| PDM3 |  | 5.116% | -5.116%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 232b (Indicative Value 0.125 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | < 0.1 | 25
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.125
| Std Dev. |  |  | 31
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 232b (Indicative Value 447 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 447 | 1
| 1 | 2 | 446 | 26
| Mean |  | 446.5 | 16
| Median |  | 446.5 | 446.5
| Std Dev. |  | 0.70711 | 32
| Rel.Std.Dev. |  | 0.1584%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 232b (Indicative Value 700 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 702 | 1
| 1 | 2 | 698 | 27
| Mean |  | 700 | 16
| Median |  | 700 | 700
| Std Dev. |  | 2.8284 | 33
| Rel.Std.Dev. |  | 0.4041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 232b (Indicative Value 2.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.8 | 28
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 232b (Indicative Value 0.34 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.36 | 29
| Mean |  | 0.34 | 16
| Median |  | 0.34 | 0.34
| Std Dev. |  | 0.028284 | 35
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 232b (Indicative Value 0.13 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | 0.2 | 5
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.125
| Std Dev. |  |  | 36
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 232b (Indicative Value 77 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 76.8 | 1
| 1 | 2 | 77.9 | 31
| Mean |  | 77.35 | 16
| Median |  | 77.35 | 77.35
| Std Dev. |  | 0.77782 | 37
| Rel.Std.Dev. |  | 1.006%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 232b (Indicative Value 17.5 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.7 | 1
| 1 | 2 | 17.2 | 32
| Mean |  | 17.45 | 16
| Median |  | 17.45 | 17.45
| Std Dev. |  | 0.35355 | 38
| Rel.Std.Dev. |  | 2.026%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 232b (Indicative Value 150 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 151 | 1
| 1 | 2 | 149 | 33
| Mean |  | 150 | 16
| Median |  | 150 | 150
| Std Dev. |  | 1.4142 | 39
| Rel.Std.Dev. |  | 0.9428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 232b (Indicative Value 8.47 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.38 | 1
| 1 | 2 | 8.56 | 34
| Mean |  | 8.47 | 16
| Median |  | 8.47 | 8.47
| Std Dev. |  | 0.12728 | 40
| Rel.Std.Dev. |  | 1.503%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 232b (Indicative Value 29 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28 | 1
| 1 | 2 | 30 | 35
| Mean |  | 29 | 16
| Median |  | 29 | 29
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 232b (Indicative Value 5.62 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.66 | 1
| 1 | 2 | 5.57 | 9
| Mean |  | 5.615 | 16
| Median |  | 5.615 | 5.615
| Std Dev. |  | 0.06364 | 42
| Rel.Std.Dev. |  | 1.133%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 232b (Indicative Value 3.33 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.21 | 1
| 1 | 2 | 3.44 | 10
| Mean |  | 3.325 | 16
| Median |  | 3.325 | 3.325
| Std Dev. |  | 0.16263 | 43
| Rel.Std.Dev. |  | 4.891%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 232b (Indicative Value 1.34 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.38 | 1
| 1 | 2 | 1.29 | 11
| Mean |  | 1.335 | 16
| Median |  | 1.335 | 1.335
| Std Dev. |  | 0.06364 | 44
| Rel.Std.Dev. |  | 4.767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 232b (Indicative Value 18.2 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.1 | 1
| 1 | 2 | 18.2 | 39
| Mean |  | 18.15 | 16
| Median |  | 18.15 | 18.15
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 232b (Indicative Value 6.08 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.91 | 1
| 1 | 2 | 6.25 | 12
| Mean |  | 6.08 | 16
| Median |  | 6.08 | 6.08
| Std Dev. |  | 0.24042 | 46
| Rel.Std.Dev. |  | 3.954%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 232b (Indicative Value 1.45 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.5 | 1
| 1 | 2 | 1.4 | 6
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0.070711 | 47
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 232b (Indicative Value 7.09 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.01 | 1
| 1 | 2 | 7.16 | 25
| Mean |  | 7.085 | 16
| Median |  | 7.085 | 7.085
| Std Dev. |  | 0.10607 | 31
| Rel.Std.Dev. |  | 1.497%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 232b (Indicative Value 1.16 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.13 | 1
| 1 | 2 | 1.18 | 14
| Mean |  | 1.155 | 16
| Median |  | 1.155 | 1.155
| Std Dev. |  | 0.035355 | 32
| Rel.Std.Dev. |  | 3.061%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 232b (Indicative Value < 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | < 0.05 | 27
| Mean |  |  | 16
| Median |  |  | < 0.05
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 232b (Indicative Value 39.1 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39 | 1
| 1 | 2 | 39.1 | 28
| Mean |  | 39.05 | 16
| Median |  | 39.05 | 39.05
| Std Dev. |  | 0.070711 | 34
| Rel.Std.Dev. |  | 0.1811%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 232b (Indicative Value 0.47 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.48 | 29
| Mean |  | 0.465 | 16
| Median |  | 0.465 | 0.465
| Std Dev. |  | 0.021213 | 35
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 232b (Indicative Value 0.044 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0439 | 1
| 1 | 2 | 0.0444 | 30
| Mean |  | 0.04415 | 16
| Median |  | 0.04415 | 0.04415
| Std Dev. |  | 0.00035355 | 36
| Rel.Std.Dev. |  | 0.8008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 232b (Indicative Value 1.2 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1.4 | 31
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0.28284 | 37
| Rel.Std.Dev. |  | 23.57%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 232b (Indicative Value 15.9 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.9 | 1
| 1 | 2 | 15.9 | 32
| Mean |  | 15.9 | 16
| Median |  | 15.9 | 15.9
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 232b (Indicative Value 35.3 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.1 | 1
| 1 | 2 | 35.5 | 9
| Mean |  | 35.3 | 16
| Median |  | 35.3 | 35.3
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 0.8013%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 232b (Indicative Value 67 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66 | 1
| 1 | 2 | 68 | 34
| Mean |  | 67 | 16
| Median |  | 67 | 67
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 2.111%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 232b (Indicative Value 20 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 35
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 232b (Indicative Value 9.5 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.56 | 1
| 1 | 2 | 9.44 | 11
| Mean |  | 9.5 | 16
| Median |  | 9.5 | 9.5
| Std Dev. |  | 0.084853 | 42
| Rel.Std.Dev. |  | 0.8932%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 232b (Indicative Value 146 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 146 | 1
| 1 | 2 | 146 | 37
| Mean |  | 146 | 16
| Median |  | 146 | 146
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 232b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 232b (Indicative Value 190 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 191 | 1
| 1 | 2 | 189 | 39
| Mean |  | 190 | 16
| Median |  | 190 | 190
| Std Dev. |  | 1.4142 | 45
| Rel.Std.Dev. |  | 0.7443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 232b (Indicative Value 13.8 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.9 | 1
| 1 | 2 | 13.6 | 40
| Mean |  | 13.75 | 16
| Median |  | 13.75 | 13.75
| Std Dev. |  | 0.21213 | 46
| Rel.Std.Dev. |  | 1.543%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 232b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 232b (Indicative Value 7.05 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.24 | 1
| 1 | 2 | 6.86 | 10
| Mean |  | 7.05 | 16
| Median |  | 7.05 | 7.05
| Std Dev. |  | 0.2687 | 31
| Rel.Std.Dev. |  | 3.811%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 232b (Indicative Value 3.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.6 | 26
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 232b (Indicative Value 133 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 133 | 1
| 1 | 2 | 132 | 27
| Mean |  | 132.5 | 16
| Median |  | 132.5 | 132.5
| Std Dev. |  | 0.70711 | 33
| Rel.Std.Dev. |  | 0.5337%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 232b (Indicative Value 1.23 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.19 | 1
| 1 | 2 | 1.27 | 28
| Mean |  | 1.23 | 16
| Median |  | 1.23 | 1.23
| Std Dev. |  | 0.056569 | 34
| Rel.Std.Dev. |  | 4.599%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 232b (Indicative Value 0.94 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.93 | 11
| Mean |  | 0.94 | 16
| Median |  | 0.94 | 0.94
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 1.504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 232b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 6
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 232b (Indicative Value 14.6 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.4 | 1
| 1 | 2 | 14.7 | 31
| Mean |  | 14.55 | 16
| Median |  | 14.55 | 14.55
| Std Dev. |  | 0.21213 | 37
| Rel.Std.Dev. |  | 1.458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 232b (Indicative Value 0.489 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.487 | 1
| 1 | 2 | 0.491 | 32
| Mean |  | 0.489 | 16
| Median |  | 0.489 | 0.489
| Std Dev. |  | 0.0028284 | 38
| Rel.Std.Dev. |  | 0.5784%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 232b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 232b (Indicative Value 0.52 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.51 | 1
| 1 | 2 | 0.52 | 13
| Mean |  | 0.515 | 16
| Median |  | 0.515 | 0.515
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 232b (Indicative Value 3.14 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.16 | 1
| 1 | 2 | 3.12 | 35
| Mean |  | 3.14 | 16
| Median |  | 3.14 | 3.14
| Std Dev. |  | 0.028284 | 41
| Rel.Std.Dev. |  | 0.9008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 232b (Indicative Value 103 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 103 | 1
| 1 | 2 | 103 | 36
| Mean |  | 103 | 16
| Median |  | 103 | 103
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 232b (Indicative Value 2.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.5 | 1
| 1 | 2 | 3 | 37
| Mean |  | 2.75 | 16
| Median |  | 2.75 | 2.75
| Std Dev. |  | 0.35355 | 43
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 232b (Indicative Value 30.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.5 | 1
| 1 | 2 | 30.7 | 38
| Mean |  | 30.6 | 16
| Median |  | 30.6 | 30.6
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 0.4622%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 232b (Indicative Value 3.32 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.33 | 1
| 1 | 2 | 3.31 | 39
| Mean |  | 3.32 | 16
| Median |  | 3.32 | 3.32
| Std Dev. |  | 0.014142 | 45
| Rel.Std.Dev. |  | 0.426%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 232b (Indicative Value 88 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 85 | 40
| Mean |  | 87.5 | 16
| Median |  | 87.5 | 87.5
| Std Dev. |  | 3.5355 | 46
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 232b (Indicative Value 252 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 252 | 1
| 1 | 2 | 252 | 41
| Mean |  | 252 | 16
| Median |  | 252 | 252
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.