## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 237b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 2.2637 | 0.066925 | 2.1298 | 2.3975 | 2.0629 | 2.4644 | 2.956% | 5.913% | 8.869% | 2.1505 | 2.3768
Aqua Regia Digestion (sample mass 10-50g)
Au, ppm | 2.1961 | 0.077678 | 2.0407 | 2.3515 | 1.9631 | 2.4291 | 3.537% | 7.074% | 10.61% | 2.0863 | 2.3059
Cyanide Leach
Au, ppm | 1.979 | 0.10366 | 1.7717 | 2.1863 | 1.668 | 2.29 | 5.238% | 10.48% | 15.71% | 1.8801 | 2.078
PhotonAssay (recommended gross mass* 490±20 g)
Au, ppm | 2.2974 | 0.076283 | 2.1448 | 2.4499 | 2.0685 | 2.5262 | 3.32% | 6.641% | 9.961% | 2.1825 | 2.4122
4-Acid Digestion
Ag, ppm | 0.18317 | 0.035283 | 0.1126 | 0.25373 | 0.077319 | 0.28901 | 19.26% | 38.53% | 57.79% | 0.17401 | 0.19232
Al, wt.% | 7.3159 | 0.23072 | 6.8544 | 7.7773 | 6.6237 | 8.0081 | 3.154% | 6.308% | 9.461% | 6.9501 | 7.6817
As, ppm | 591.043 | 33.5845 | 523.8739 | 658.212 | 490.2894 | 691.7965 | 5.682% | 11.36% | 17.05% | 561.4908 | 620.5951
Ba, ppm | 683.1123 | 18.9199 | 645.2725 | 720.9522 | 626.3526 | 739.8721 | 2.77% | 5.539% | 8.309% | 648.9567 | 717.268
Be, ppm | 2.4654 | 0.18813 | 2.0892 | 2.8417 | 1.9011 | 3.0298 | 7.631% | 15.26% | 22.89% | 2.3422 | 2.5887
Bi, ppm | 0.3537 | 0.02985 | 0.294 | 0.4134 | 0.26415 | 0.44325 | 8.439% | 16.88% | 25.32% | 0.33602 | 0.37139
Ca, wt.% | 0.90308 | 0.028669 | 0.84575 | 0.96042 | 0.81708 | 0.98909 | 3.175% | 6.349% | 9.524% | 0.85793 | 0.94824
Ce, ppm | 80.5252 | 3.3731 | 73.7791 | 87.2714 | 70.406 | 90.6445 | 4.189% | 8.378% | 12.57% | 76.499 | 84.5515
Co, ppm | 17.2918 | 0.5153 | 16.2612 | 18.3224 | 15.7459 | 18.8377 | 2.98% | 5.96% | 8.94% | 16.4272 | 18.1564
Cr, ppm | 122.8375 | 14.6483 | 93.5409 | 152.1341 | 78.8926 | 166.7824 | 11.92% | 23.85% | 35.77% | 116.6956 | 128.9793
Cs, ppm | 9.4201 | 0.33694 | 8.7462 | 10.094 | 8.4093 | 10.4309 | 3.577% | 7.154% | 10.73% | 8.9491 | 9.8911
Cu, ppm | 28.3324 | 1.5256 | 25.2812 | 31.3836 | 23.7556 | 32.9091 | 5.385% | 10.77% | 16.15% | 26.9158 | 29.749
Dy, ppm | 3.6014 | 0.70367 | 2.1941 | 5.0088 | 1.4904 | 5.7124 | 19.54% | 39.08% | 58.62% | 3.4214 | 3.7815
Er, ppm | 1.9965 | 0.48184 | 1.0328 | 2.9601 | 0.55093 | 3.442 | 24.13% | 48.27% | 72.4% | 1.8966 | 2.0963
Eu, ppm | 1.2457 | 0.078541 | 1.0887 | 1.4028 | 1.0101 | 1.4814 | 6.305% | 12.61% | 18.91% | 1.1834 | 1.308
Fe, wt.% | 4.0492 | 0.07209 | 3.9051 | 4.1934 | 3.833 | 4.2655 | 1.78% | 3.561% | 5.341% | 3.8468 | 4.2517
Ga, ppm | 19.3985 | 1.2012 | 16.996 | 21.8009 | 15.7947 | 23.0022 | 6.192% | 12.38% | 18.58% | 18.4285 | 20.3684
Gd, ppm | 5.0623 | 0.32123 | 4.4198 | 5.7047 | 4.0986 | 6.026 | 6.346% | 12.69% | 19.04% | 4.8092 | 5.3154
Hf, ppm | 3.9364 | 0.21486 | 3.5067 | 4.3661 | 3.2918 | 4.581 | 5.458% | 10.92% | 16.37% | 3.7396 | 4.1332
Ho, ppm | 0.68476 | 0.16502 | 0.35471 | 1.0148 | 0.18969 | 1.1798 | 24.1% | 48.2% | 72.3% | 0.65052 | 0.71899
In, ppm | 0.065658 | 0.0066224 | 0.052413 | 0.078903 | 0.045791 | 0.085525 | 10.09% | 20.17% | 30.26% | 0.062375 | 0.068941
K, wt.% | 2.6454 | 0.074971 | 2.4954 | 2.7953 | 2.4205 | 2.8703 | 2.834% | 5.668% | 8.502% | 2.5131 | 2.7776
La, ppm | 38.6398 | 2.0297 | 34.5805 | 42.6991 | 32.5508 | 44.7287 | 5.253% | 10.51% | 15.76% | 36.7078 | 40.5718
Li, ppm | 52.3763 | 3.6661 | 45.0441 | 59.7084 | 41.378 | 63.3745 | 6.999% | 14% | 21% | 49.7575 | 54.9951
Lu, ppm | 0.27762 | 0.074206 | 0.12921 | 0.42603 | 0.055 | 0.50023 | 26.73% | 53.46% | 80.19% | 0.26374 | 0.2915
Mg, wt.% | 1.6983 | 0.051249 | 1.5958 | 1.8008 | 1.5445 | 1.852 | 3.018% | 6.035% | 9.053% | 1.6134 | 1.7832
Mn, wt.% | 0.040752 | 0.0017226 | 0.037307 | 0.044197 | 0.035584 | 0.04592 | 4.227% | 8.454% | 12.68% | 0.038714 | 0.042789
Mo, ppm | 0.97151 | 0.068936 | 0.83364 | 1.1094 | 0.7647 | 1.1783 | 7.096% | 14.19% | 21.29% | 0.92294 | 1.0201
Na, wt.% | 0.7559 | 0.01762 | 0.72066 | 0.79114 | 0.70304 | 0.80876 | 2.331% | 4.662% | 6.993% | 0.7181 | 0.79369
Nb, ppm | 13.606 | 1.3628 | 10.8804 | 16.3317 | 9.5176 | 17.6945 | 10.02% | 20.03% | 30.05% | 12.9257 | 14.2863
Nd, ppm | 34.1903 | 1.2964 | 31.5976 | 36.7831 | 30.3012 | 38.0795 | 3.792% | 7.583% | 11.37% | 32.4808 | 35.8998
Ni, ppm | 64.365 | 3.0777 | 58.2095 | 70.5204 | 55.1318 | 73.5981 | 4.782% | 9.563% | 14.35% | 61.1467 | 67.5832
P, wt.% | 0.066302 | 0.0022645 | 0.061773 | 0.070831 | 0.059508 | 0.073095 | 3.415% | 6.831% | 10.25% | 0.062987 | 0.069617
Pb, ppm | 20.0063 | 0.73574 | 18.5348 | 21.4777 | 17.799 | 22.2135 | 3.678% | 7.355% | 11.03% | 19.0059 | 21.0066
Pr, ppm | 9.2901 | 0.46111 | 8.3679 | 10.2124 | 7.9068 | 10.6735 | 4.963% | 9.927% | 14.89% | 8.8256 | 9.7546
Rb, ppm | 154.9886 | 10.9933 | 133.0021 | 176.9752 | 122.0088 | 187.9684 | 7.093% | 14.19% | 21.28% | 147.2392 | 162.738
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.19308 | 0.010484 | 0.17211 | 0.21405 | 0.16162 | 0.22453 | 5.43% | 10.86% | 16.29% | 0.18342 | 0.20273
Sb, ppm | 460.0851 | 26.0824 | 407.9203 | 512.25 | 381.8379 | 538.3324 | 5.669% | 11.34% | 17.01% | 437.0809 | 483.0894
Sc, ppm | 14.1501 | 0.75858 | 12.6329 | 15.6673 | 11.8744 | 16.4259 | 5.361% | 10.72% | 16.08% | 13.4426 | 14.8576
Sm, ppm | 6.607 | 0.41329 | 5.7804 | 7.4336 | 5.3671 | 7.8469 | 6.255% | 12.51% | 18.77% | 6.2766 | 6.9373
Sn, ppm | 3.6117 | 0.2297 | 3.1523 | 4.0711 | 2.9226 | 4.3008 | 6.36% | 12.72% | 19.08% | 3.4311 | 3.7923
Sr, ppm | 123.2072 | 3.9162 | 115.3749 | 131.0396 | 111.4587 | 134.9557 | 3.179% | 6.357% | 9.536% | 117.0469 | 129.3676
Ta, ppm | 1.0256 | 0.098933 | 0.8277 | 1.2234 | 0.72877 | 1.3224 | 9.647% | 19.29% | 28.94% | 0.97429 | 1.0768
Tb, ppm | 0.66362 | 0.043621 | 0.57637 | 0.75086 | 0.53275 | 0.79448 | 6.573% | 13.15% | 19.72% | 0.63044 | 0.6968
Th, ppm | 14.4502 | 0.65533 | 13.1395 | 15.7608 | 12.4842 | 16.4162 | 4.535% | 9.07% | 13.61% | 13.7277 | 15.1727
Ti, wt.% | 0.45585 | 0.016948 | 0.42195 | 0.48975 | 0.405 | 0.50669 | 3.718% | 7.436% | 11.15% | 0.43306 | 0.47864
Tl, ppm | 0.84289 | 0.056175 | 0.73054 | 0.95524 | 0.67437 | 1.0114 | 6.665% | 13.33% | 19.99% | 0.80075 | 0.88504
Tm, ppm | 0.25243 | 0.042793 | 0.16684 | 0.33801 | 0.12405 | 0.38081 | 16.95% | 33.91% | 50.86% | 0.2398 | 0.26505
U, ppm | 2.7348 | 0.11611 | 2.5026 | 2.967 | 2.3865 | 3.0831 | 4.246% | 8.491% | 12.74% | 2.5981 | 2.8716
V, ppm | 104.1674 | 3.408 | 97.3513 | 110.9835 | 93.9432 | 114.3915 | 3.272% | 6.543% | 9.815% | 98.959 | 109.3758
W, ppm | 2.4385 | 0.28502 | 1.8685 | 3.0086 | 1.5835 | 3.2936 | 11.69% | 23.38% | 35.06% | 2.3166 | 2.5605
Y, ppm | 16.2019 | 2.1949 | 11.8121 | 20.5917 | 9.6172 | 22.7866 | 13.55% | 27.09% | 40.64% | 15.3918 | 17.012
Yb, ppm | 1.7577 | 0.28928 | 1.1792 | 2.3363 | 0.88989 | 2.6256 | 16.46% | 32.91% | 49.37% | 1.6698 | 1.8456
Zn, ppm | 95.0975 | 3.5764 | 87.9448 | 102.2503 | 84.3684 | 105.8267 | 3.761% | 7.521% | 11.28% | 90.3427 | 99.8524
Zr, ppm | 134.6182 | 7.0606 | 120.4971 | 148.7393 | 113.4366 | 155.7999 | 5.245% | 10.49% | 15.73% | 127.8873 | 141.3491
Aqua Regia Digestion
Ag, ppm | 0.17579 | 0.01497 | 0.14584 | 0.20573 | 0.13087 | 0.2207 | 8.516% | 17.03% | 25.55% | 0.167 | 0.18458
Al, wt.% | 2.7461 | 0.15253 | 2.441 | 3.0512 | 2.2885 | 3.2037 | 5.555% | 11.11% | 16.66% | 2.6088 | 2.8834
As, ppm | 583.2871 | 28.5292 | 526.2287 | 640.3456 | 497.6994 | 668.8748 | 4.891% | 9.782% | 14.67% | 554.1228 | 612.4515
Ba, ppm | 107.5927 | 8.1703 | 91.2521 | 123.9334 | 83.0818 | 132.1037 | 7.594% | 15.19% | 22.78% | 102.2131 | 112.9724
Be, ppm | 1.2618 | 0.098175 | 1.0654 | 1.4581 | 0.96725 | 1.5563 | 7.781% | 15.56% | 23.34% | 1.1987 | 1.3249
Bi, ppm | 0.33465 | 0.01905 | 0.29655 | 0.37275 | 0.2775 | 0.3918 | 5.692% | 11.38% | 17.08% | 0.31791 | 0.35138
Ca, wt.% | 0.3122 | 0.019689 | 0.27282 | 0.35158 | 0.25313 | 0.37127 | 6.307% | 12.61% | 18.92% | 0.29659 | 0.32781
Cd, ppm | 0.042794 | 0.013746 | 0.015303 | 0.070285 | 0.0015577 | 0.084031 | 32.12% | 64.24% | 96.36% | 0.040655 | 0.044934
Ce, ppm | 54.4696 | 2.6488 | 49.172 | 59.7671 | 46.5233 | 62.4158 | 4.863% | 9.726% | 14.59% | 51.7461 | 57.193
Co, ppm | 15.7009 | 0.82603 | 14.0488 | 17.3529 | 13.2228 | 18.179 | 5.261% | 10.52% | 15.78% | 14.9158 | 16.4859
Cr, ppm | 107.976 | 5.1306 | 97.7149 | 118.2371 | 92.5843 | 123.3677 | 4.752% | 9.503% | 14.25% | 102.5772 | 113.3748
Cs, ppm | 6.9333 | 0.28853 | 6.3563 | 7.5104 | 6.0678 | 7.7989 | 4.161% | 8.323% | 12.48% | 6.5867 | 7.28
Cu, ppm | 26.9057 | 1.37 | 24.1658 | 29.6456 | 22.7958 | 31.0156 | 5.092% | 10.18% | 15.28% | 25.5604 | 28.251
Fe, wt.% | 3.4333 | 0.1213 | 3.1907 | 3.6759 | 3.0694 | 3.7972 | 3.533% | 7.066% | 10.6% | 3.2617 | 3.605
Ga, ppm | 8.8798 | 0.3929 | 8.094 | 9.6656 | 7.7011 | 10.0585 | 4.425% | 8.849% | 13.27% | 8.4358 | 9.3238
Ge, ppm | 0.10919 | 0.019918 | 0.069354 | 0.14903 | 0.049436 | 0.16894 | 18.24% | 36.48% | 54.72% | 0.10373 | 0.11465
Hf, ppm | 0.4856 | 0.076817 | 0.33197 | 0.63923 | 0.25515 | 0.71605 | 15.82% | 31.64% | 47.46% | 0.46132 | 0.50988
In, ppm | 0.033359 | 0.0045579 | 0.024243 | 0.042475 | 0.019685 | 0.047033 | 13.66% | 27.33% | 40.99% | 0.031691 | 0.035027
K, wt.% | 0.86646 | 0.040699 | 0.78506 | 0.94786 | 0.74436 | 0.98856 | 4.697% | 9.394% | 14.09% | 0.82313 | 0.90978
La, ppm | 27.5305 | 3.4247 | 20.6811 | 34.38 | 17.2564 | 37.8047 | 12.44% | 24.88% | 37.32% | 26.154 | 28.9071
Li, ppm | 42.9679 | 3.0981 | 36.7718 | 49.1641 | 33.6737 | 52.2622 | 7.21% | 14.42% | 21.63% | 40.8195 | 45.1163
Lu, ppm | 0.12039 | 0.02048 | 0.079426 | 0.16135 | 0.058946 | 0.18182 | 17.01% | 34.02% | 51.04% | 0.11437 | 0.1264
Mg, wt.% | 1.3737 | 0.078386 | 1.2169 | 1.5304 | 1.1385 | 1.6088 | 5.706% | 11.41% | 17.12% | 1.305 | 1.4423
Mn, wt.% | 0.02827 | 0.0012978 | 0.025674 | 0.030865 | 0.024376 | 0.032163 | 4.591% | 9.182% | 13.77% | 0.026856 | 0.029683
Mo, ppm | 0.87621 | 0.045437 | 0.78533 | 0.96708 | 0.7399 | 1.0125 | 5.186% | 10.37% | 15.56% | 0.8324 | 0.92002
Na, wt.% | 0.081824 | 0.0080168 | 0.065791 | 0.097858 | 0.057774 | 0.10587 | 9.798% | 19.6% | 29.39% | 0.077733 | 0.085915
Nb, ppm | 0.38659 | 0.11012 | 0.16635 | 0.60683 | 0.056232 | 0.71695 | 28.48% | 56.97% | 85.45% | 0.36726 | 0.40592
Ni, ppm | 59.8278 | 3.2513 | 53.3252 | 66.3305 | 50.0738 | 69.5818 | 5.434% | 10.87% | 16.3% | 56.8364 | 62.8192
P, wt.% | 0.057965 | 0.0024895 | 0.052986 | 0.062944 | 0.050497 | 0.065434 | 4.295% | 8.59% | 12.88% | 0.055067 | 0.060864
Pb, ppm | 9.9661 | 0.92564 | 8.1148 | 11.8174 | 7.1892 | 12.7431 | 9.288% | 18.58% | 27.86% | 9.4678 | 10.4644
Rb, ppm | 84.9409 | 2.9852 | 78.9704 | 90.9114 | 75.9852 | 93.8966 | 3.514% | 7.029% | 10.54% | 80.6938 | 89.1879
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.19179 | 0.018006 | 0.15578 | 0.2278 | 0.13777 | 0.24581 | 9.388% | 18.78% | 28.17% | 0.1822 | 0.20138
Sb, ppm | 342.6233 | 58.3728 | 225.8778 | 459.3689 | 167.505 | 517.7417 | 17.04% | 34.07% | 51.11% | 325.4922 | 359.7545
Sc, ppm | 7.2922 | 0.57213 | 6.148 | 8.4365 | 5.5759 | 9.0086 | 7.846% | 15.69% | 23.54% | 6.9276 | 7.6568
Sn, ppm | 1.6654 | 0.074429 | 1.5165 | 1.8143 | 1.4421 | 1.8887 | 4.469% | 8.938% | 13.41% | 1.5821 | 1.7487
Sr, ppm | 26.1532 | 1.1058 | 23.9417 | 28.3647 | 22.8359 | 29.4705 | 4.228% | 8.456% | 12.68% | 24.8455 | 27.4609
Ta, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 12.2657 | 0.63396 | 10.9978 | 13.5336 | 10.3638 | 14.1676 | 5.169% | 10.34% | 15.51% | 11.6524 | 12.879
Ti, wt.% | 0.14975 | 0.014943 | 0.11986 | 0.17964 | 0.10492 | 0.19458 | 9.979% | 19.96% | 29.94% | 0.14226 | 0.15724
Tl, ppm | 0.50358 | 0.038082 | 0.42741 | 0.57974 | 0.38933 | 0.61782 | 7.562% | 15.12% | 22.69% | 0.4784 | 0.52876
U, ppm | 1.4593 | 0.1167 | 1.2259 | 1.6928 | 1.1092 | 1.8095 | 7.997% | 15.99% | 23.99% | 1.3864 | 1.5323
V, ppm | 69.1929 | 3.0418 | 63.1094 | 75.2764 | 60.0676 | 78.3181 | 4.396% | 8.792% | 13.19% | 65.7332 | 72.6525
W, ppm | 0.49421 | 0.071035 | 0.35214 | 0.63628 | 0.2811 | 0.70732 | 14.37% | 28.75% | 43.12% | 0.4695 | 0.51892
Y, ppm | 9.2314 | 0.63298 | 7.9654 | 10.4973 | 7.3324 | 11.1303 | 6.857% | 13.71% | 20.57% | 8.7698 | 9.6929
Yb, ppm | 0.84496 | 0.094566 | 0.65583 | 1.0341 | 0.56126 | 1.1287 | 11.19% | 22.38% | 33.58% | 0.80271 | 0.88721
Zn, ppm | 84.9363 | 3.176 | 78.5843 | 91.2883 | 75.4083 | 94.4643 | 3.739% | 7.479% | 11.22% | 80.6895 | 89.1831
Zr, ppm | 19.1101 | 3.6945 | 11.721 | 26.4991 | 8.0265 | 30.1936 | 19.33% | 38.67% | 58% | 18.1545 | 20.0656
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 237b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 2.2637 | 2.2404 | 2.2869 | 2.2552 | 2.2721
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 2.1961 | 2.1614 | 2.2308 | 2.1871 | 2.2051
Cyanide Leach
Au, Gold (ppm) | 1.979 | 1.9262 | 2.0318 | 1.9716 | 1.9864
PhotonAssay
Au, Gold (ppm) | 2.2974 | 2.2564 | 2.3383 | 2.2948 | 2.2999
4-Acid Digestion
Ag, Silver (ppm) | 0.18317 | 0.15241 | 0.21392 | 0.17189 | 0.19444
Al, Aluminium (wt.%) | 7.3159 | 7.1165 | 7.5153 | 7.1606 | 7.4712
As, Arsenic (ppm) | 591.043 | 565.994 | 616.092 | 574.8977 | 607.1883
Ba, Barium (ppm) | 683.1123 | 664.0631 | 702.1616 | 671.7568 | 694.4679
Be, Beryllium (ppm) | 2.4654 | 2.3476 | 2.5833 | 2.3882 | 2.5427
Bi, Bismuth (ppm) | 0.3537 | 0.3212 | 0.3862 | 0.32117 | 0.38624
Ca, Calcium (wt.%) | 0.90308 | 0.87653 | 0.92964 | 0.88202 | 0.92414
Ce, Cerium (ppm) | 80.5252 | 76.8645 | 84.1859 | 78.2129 | 82.8375
Co, Cobalt (ppm) | 17.2918 | 16.8069 | 17.7767 | 16.7918 | 17.7918
Cr, Chromium (ppm) | 122.8375 | 115.1673 | 130.5076 | 117.4952 | 128.1797
Cs, Caesium (ppm) | 9.4201 | 9.0456 | 9.7946 | 9.1859 | 9.6543
Cu, Copper (ppm) | 28.3324 | 27.1338 | 29.531 | 27.3421 | 29.3227
Dy, Dysprosium (ppm) | 3.6014 | 3.0743 | 4.1286 | 3.4136 | 3.7893
Er, Erbium (ppm) | 1.9965 | 1.5086 | 2.4844 | 1.8998 | 2.0931
Eu, Europium (ppm) | 1.2457 | 1.155 | 1.3365 | 1.2016 | 1.2899
Fe, Iron (wt.%) | 4.0492 | 3.9446 | 4.1538 | 3.9952 | 4.1033
Ga, Gallium (ppm) | 19.3985 | 18.3789 | 20.418 | 18.8504 | 19.9465
Gd, Gadolinium (ppm) | 5.0623 | 4.7454 | 5.3792 | 4.8764 | 5.2481
Hf, Hafnium (ppm) | 3.9364 | 3.6811 | 4.1917 | 3.7836 | 4.0892
Ho, Holmium (ppm) | 0.68476 | 0.52104 | 0.84847 | 0.65119 | 0.71833
In, Indium (ppm) | 0.065658 | 0.057654 | 0.073662 | 0.061784 | 0.069532
K, Potassium (wt.%) | 2.6454 | 2.5739 | 2.7168 | 2.6021 | 2.6887
La, Lanthanum (ppm) | 38.6398 | 37.008 | 40.2715 | 37.6085 | 39.671
Li, Lithium (ppm) | 52.3763 | 50.1599 | 54.5927 | 50.9766 | 53.776
Lu, Lutetium (ppm) | 0.27762 | 0.22071 | 0.33452 | 0.26255 | 0.29268
Mg, Magnesium (wt.%) | 1.6983 | 1.652 | 1.7446 | 1.667 | 1.7296
Mn, Manganese (wt.%) | 0.040752 | 0.039547 | 0.041957 | 0.040099 | 0.041405
Mo, Molybdenum (ppm) | 0.97151 | 0.9012 | 1.0418 | 0.89558 | 1.0474
Na, Sodium (wt.%) | 0.7559 | 0.7362 | 0.7756 | 0.7413 | 0.7705
Nb, Niobium (ppm) | 13.606 | 12.7717 | 14.4404 | 13.1976 | 14.0145
Nd, Neodymium (ppm) | 34.1903 | 32.9913 | 35.3894 | 33.1673 | 35.2134
Ni, Nickel (ppm) | 64.365 | 62.3704 | 66.3595 | 62.6935 | 66.0364
P, Phosphorus (wt.%) | 0.066302 | 0.064306 | 0.068297 | 0.06514 | 0.067463
Pb, Lead (ppm) | 20.0063 | 19.0116 | 21.0009 | 19.2634 | 20.7491
Pr, Praseodymium (ppm) | 9.2901 | 8.92 | 9.6603 | 8.9718 | 9.6085
Rb, Rubidium (ppm) | 154.9886 | 147.1297 | 162.8476 | 149.3371 | 160.6401
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.19308 | 0.18314 | 0.20302 | 0.18506 | 0.2011
Sb, Antimony (ppm) | 460.0851 | 442.0022 | 478.1681 | 449.2893 | 470.881
Sc, Scandium (ppm) | 14.1501 | 13.4327 | 14.8675 | 13.7157 | 14.5845
Sm, Samarium (ppm) | 6.607 | 6.1938 | 7.0201 | 6.2756 | 6.9384
Sn, Tin (ppm) | 3.6117 | 3.4019 | 3.8216 | 3.4739 | 3.7496
Sr, Strontium (ppm) | 123.2072 | 120.4665 | 125.948 | 120.7842 | 125.6303
Ta, Tantalum (ppm) | 1.0256 | 0.94567 | 1.1055 | 0.9926 | 1.0585
Tb, Terbium (ppm) | 0.66362 | 0.61352 | 0.71371 | 0.62595 | 0.70128
Th, Thorium (ppm) | 14.4502 | 13.8011 | 15.0992 | 13.9656 | 14.9348
Ti, Titanium (wt.%) | 0.45585 | 0.44282 | 0.46888 | 0.4473 | 0.46439
Tl, Thallium (ppm) | 0.84289 | 0.78771 | 0.89807 | 0.81256 | 0.87322
Tm, Thulium (ppm) | 0.25243 | 0.21635 | 0.2885 | 0.24117 | 0.26368
U, Uranium (ppm) | 2.7348 | 2.5939 | 2.8757 | 2.6376 | 2.832
V, Vanadium (ppm) | 104.1674 | 101.6488 | 106.686 | 102.2354 | 106.0994
W, Tungsten (ppm) | 2.4385 | 2.2205 | 2.6566 | 2.2922 | 2.5848
Y, Yttrium (ppm) | 16.2019 | 14.9267 | 17.477 | 15.7014 | 16.7023
Yb, Ytterbium (ppm) | 1.7577 | 1.5468 | 1.9687 | 1.6605 | 1.8549
Zn, Zinc (ppm) | 95.0975 | 92.4459 | 97.7492 | 92.9843 | 97.2108
Zr, Zirconium (ppm) | 134.6182 | 130.1036 | 139.1328 | 130.6496 | 138.5869
Aqua Regia Digestion
Ag, Silver (ppm) | 0.17579 | 0.16281 | 0.18876 | IND | IND
Al, Aluminium (wt.%) | 2.7461 | 2.6425 | 2.8497 | 2.6996 | 2.7926
As, Arsenic (ppm) | 583.2871 | 566.8518 | 599.7225 | 572.7845 | 593.7898
Ba, Barium (ppm) | 107.5927 | 102.6814 | 112.5041 | 105.0772 | 110.1082
Be, Beryllium (ppm) | 1.2618 | 1.1694 | 1.3542 | 1.2075 | 1.316
Bi, Bismuth (ppm) | 0.33465 | 0.31066 | 0.35864 | 0.2884 | 0.3809
Ca, Calcium (wt.%) | 0.3122 | 0.2998 | 0.3246 | 0.30387 | 0.32053
Cd, Cadmium (ppm) | 0.042794 | 0.031826 | 0.053762 | 0.041297 | 0.044292
Ce, Cerium (ppm) | 54.4696 | 51.029 | 57.9102 | 52.4386 | 56.5005
Co, Cobalt (ppm) | 15.7009 | 15.1214 | 16.2803 | 15.2106 | 16.1912
Cr, Chromium (ppm) | 107.976 | 104.7162 | 111.2359 | 105.7237 | 110.2283
Cs, Caesium (ppm) | 6.9333 | 6.717 | 7.1497 | 6.6915 | 7.1752
Cu, Copper (ppm) | 26.9057 | 25.6905 | 28.1209 | 26.1253 | 27.686
Fe, Iron (wt.%) | 3.4333 | 3.3483 | 3.5183 | 3.3902 | 3.4765
Ga, Gallium (ppm) | 8.8798 | 8.487 | 9.2726 | 8.5816 | 9.178
Ge, Germanium (ppm) | 0.10919 | 0.077139 | 0.14124 | IND | IND
Hf, Hafnium (ppm) | 0.4856 | 0.4271 | 0.5441 | 0.45541 | 0.51579
In, Indium (ppm) | 0.033359 | 0.030158 | 0.03656 | 0.030751 | 0.035967
K, Potassium (wt.%) | 0.86646 | 0.83774 | 0.89518 | 0.85375 | 0.87917
La, Lanthanum (ppm) | 27.5305 | 25.5212 | 29.5399 | 26.6759 | 28.3852
Li, Lithium (ppm) | 42.9679 | 40.9565 | 44.9794 | 41.8425 | 44.0934
Lu, Lutetium (ppm) | 0.12039 | 0.092035 | 0.14874 | IND | IND
Mg, Magnesium (wt.%) | 1.3737 | 1.3239 | 1.4234 | 1.3511 | 1.3962
Mn, Manganese (wt.%) | 0.02827 | 0.027382 | 0.029157 | 0.027745 | 0.028795
Mo, Molybdenum (ppm) | 0.87621 | 0.80065 | 0.95176 | 0.81339 | 0.93903
Na, Sodium (wt.%) | 0.081824 | 0.076076 | 0.087573 | 0.078553 | 0.085095
Nb, Niobium (ppm) | 0.38659 | 0.30127 | 0.47191 | 0.34907 | 0.4241
Ni, Nickel (ppm) | 59.8278 | 57.7082 | 61.9474 | 58.6291 | 61.0265
P, Phosphorus (wt.%) | 0.057965 | 0.055943 | 0.059987 | 0.056843 | 0.059088
Pb, Lead (ppm) | 9.9661 | 9.3411 | 10.5912 | 9.5269 | 10.4053
Rb, Rubidium (ppm) | 84.9409 | 82.1262 | 87.7556 | 82.4226 | 87.4592
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.19179 | 0.18114 | 0.20245 | 0.18418 | 0.1994
Sb, Antimony (ppm) | 342.6233 | 311.6656 | 373.5811 | 331.7624 | 353.4843
Sc, Scandium (ppm) | 7.2922 | 6.9184 | 7.6661 | 7.1188 | 7.4656
Sn, Tin (ppm) | 1.6654 | 1.581 | 1.7497 | 1.5749 | 1.7559
Sr, Strontium (ppm) | 26.1532 | 24.9869 | 27.3195 | 25.4027 | 26.9037
Ta, Tantalum (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 12.2657 | 11.7111 | 12.8203 | 11.798 | 12.7335
Ti, Titanium (wt.%) | 0.14975 | 0.1406 | 0.1589 | 0.14583 | 0.15367
Tl, Thallium (ppm) | 0.50358 | 0.48101 | 0.52614 | 0.48649 | 0.52067
U, Uranium (ppm) | 1.4593 | 1.3745 | 1.5441 | 1.4061 | 1.5125
V, Vanadium (ppm) | 69.1929 | 67.1671 | 71.2186 | 67.9817 | 70.404
W, Tungsten (ppm) | 0.49421 | 0.42688 | 0.56154 | 0.45413 | 0.5343
Y, Yttrium (ppm) | 9.2314 | 8.7865 | 9.6763 | 9.0226 | 9.4402
Yb, Ytterbium (ppm) | 0.84496 | 0.70993 | 0.97999 | 0.78255 | 0.90737
Zn, Zinc (ppm) | 84.9363 | 82.3287 | 87.5439 | 83.221 | 86.6517
Zr, Zirconium (ppm) | 19.1101 | 17.0013 | 21.2188 | 18.3803 | 19.8398
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 237b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 0.73816 | Pt | ppb | 0.50004
4-Acid Digestion
Au | ppm | 2.2062 | Hg | ppm | 0.0247 | Rh | ppm | 0.0015
B | ppm | 27.43 | Ir | ppm | 0.00115 | Ru | ppm | 0.0019
Cd | ppm | 0.08679 | Pd | ppb | < 12.8 | Se | ppm | 0.60784
Ge | ppm | 0.21247 | Pt | ppb | 1.1333 | Te | ppm | 0.061952
Aqua Regia Digestion
B | ppm | 15.2614 | Ir | ppm | < 0.0015 | Se | ppm | 0.2439
Dy | ppm | 2.0753 | Nd | ppm | 23.2618 | Si | wt.% | 0.43877
Er | ppm | 0.93801 | Pd | ppb | 26.3167 | Sm | ppm | 4.4191
Eu | ppm | 0.58893 | Pr | ppm | 6.2427 | Tb | ppm | 0.46848
Gd | ppm | 4.0497 | Pt | ppb | 2.1833 | Te | ppm | 0.029248
Hg | ppm | 0.01329 | Rh | ppm | 0.00061667 | Tm | ppm | 0.11783
Ho | ppm | 0.3473 | Ru | ppm | < 0.0037
Borate Fusion XRF
Al2O3 | wt.% | 14.35 | MgO | wt.% | 2.85 | SiO2 | wt.% | 68.18
CaO | wt.% | 1.25 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.482
Fe2O3 | wt.% | 5.795 | Na2O | wt.% | 1.025 | TiO2 | wt.% | 0.815
K2O | wt.% | 3.265 | P2O5 | wt.% | 0.1505
Thermogravimetry
LOI1000 | wt.% | 2.085
Infrared Combustion
C | wt.% | 0.08 | S | wt.% | 0.185
Laser Ablation ICP-MS
Ag | ppm | 0.2 | Hf | ppm | 6.82 | Sm | ppm | 7.25
As | ppm | 598 | Ho | ppm | 1.155 | Sn | ppm | 3.5
Ba | ppm | 699.5 | In | ppm | < 0.05 | Sr | ppm | 122
Be | ppm | 2.9 | La | ppm | 40.55 | Ta | ppm | 1.255
Bi | ppm | 0.39 | Lu | ppm | 0.465 | Tb | ppm | 0.965
Cd | ppm | 0.25 | Mn | wt.% | 0.04355 | Te | ppm | < 0.2
Ce | ppm | 79.55 | Mo | ppm | 1.2 | Th | ppm | 15.15
Co | ppm | 18.2 | Nb | ppm | 16.05 | Ti | wt.% | 0.494
Cr | ppm | 151 | Nd | ppm | 36.95 | Tl | ppm | 0.4
Cs | ppm | 9.46 | Ni | ppm | 69 | Tm | ppm | 0.525
Cu | ppm | 30 | Pb | ppm | 21.5 | U | ppm | 3.285
Dy | ppm | 5.74 | Pr | ppm | 9.65 | V | ppm | 111
Er | ppm | 3.45 | Rb | ppm | 160 | W | ppm | 3.5
Eu | ppm | 1.36 | Re | ppm | < 0.01 | Y | ppm | 31.15
Ga | ppm | 18.8 | Sb | ppm | 479 | Yb | ppm | 3.37
Gd | ppm | 6.015 | Sc | ppm | 14.6 | Zn | ppm | 95
Ge | ppm | 1.55 | Se | ppm | < 5 | Zr | ppm | 240
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 237b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 237b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Carsurin, Cikarang, West Java, Indonesia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 237b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 19 | 19 | 219514 | 0.088601 | 2.4949 | 2.4958
2 | 1 | ANSLu | 1 | 11 | 11 | 219515 | 0.087034 | 2.3806 | 2.4897
2 | 1 | ANSLu | 1 | 3 | 3 | 219516 | 0.087945 | 2.4384 | 2.4928
2 | 1 | ANSLu | 1 | 12 | 12 | 219517 | 0.08284 | 2.452 | 2.4935
2 | 1 | ANSLu | 1 | 17 | 17 | 219518 | 0.086274 | 2.4361 | 2.4926
2 | 1 | ANSLu | 1 | 16 | 16 | 219519 | 0.086789 | 2.5862 | 2.5007
2 | 1 | ANSLu | 1 | 2 | 2 | 219520 | 0.084931 | 2.4576 | 2.4938
2 | 1 | ANSLu | 1 | 15 | 15 | 219521 | 0.086639 | 2.5191 | 2.4971
2 | 1 | ANSLu | 1 | 13 | 13 | 219522 | 0.084196 | 2.5958 | 2.5012
2 | 1 | ANSLu | 1 | 5 | 5 | 219523 | 0.085049 | 2.5194 | 2.4971
2 | 1 | ANSLu | 1 | 7 | 7 | 219524 | 0.084701 | 2.4919 | 2.4956
2 | 1 | ANSLu | 1 | 6 | 6 | 219525 | 0.087375 | 2.474 | 2.4947
2 | 1 | ANSLu | 1 | 14 | 14 | 219526 | 0.086855 | 2.5768 | 2.5002
2 | 1 | ANSLu | 1 | 8 | 8 | 219527 | 0.085186 | 2.5314 | 2.4977
2 | 1 | ANSLu | 1 | 10 | 10 | 219528 | 0.083615 | 2.4716 | 2.4945
2 | 1 | ANSLu | 1 | 18 | 18 | 219529 | 0.082665 | 2.4682 | 2.4943
2 | 1 | ANSLu | 1 | 4 | 4 | 219530 | 0.087185 | 2.5089 | 2.4965
2 | 1 | ANSLu | 1 | 20 | 20 | 219531 | 0.085363 | 2.5566 | 2.4991
2 | 1 | ANSLu | 1 | 1 | 1 | 219532 | 0.085258 | 2.4984 | 2.496
2 | 1 | ANSLu | 1 | 9 | 9 | 219533 | 0.087939 | 2.4586 | 2.4938
Mean | [6 empty] | 0.085822 | 2.4958 | 2.4958
Median | [7 empty] | 2.4934 | 2.4957
SD | [7 empty] | 0.055111 | 0.0029477
RSD | [7 empty] | 2.208% | 0.1181%
RSD Upscale Factor | [7 empty] | 0.053486
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 237b (Certified Value 2.26 ppm) | [32 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*MS | 2
|  |  | 0.085g | 25g | 50g | 50g | 40g | 30g | 40g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 25g | 25g | 3
| 1 | 1 | 2.4984 | 2.307 | 2.39 | 2.224 | 2.24 | 2.11 | 2.23 | 2.31 | 2.22 | 2.19 | 2.119 | 2.29 | 2.22 | 2.31 | 2.35 | 2.312 | 2.1637 | 2.294 | 2.39 | 2.1367 | 2.32 | 2.15 | 2.235 | 2.38 | 2.213 | 2.28 | 2.17 | 2.251 | 2.2176 | 2.32 | 2.1319 | 1
| 1 | 2 | 2.4576 | 2.356 | 2.36 | 2.194 | 2.24 | 2.16 | 2.23 | 2.21 | 2.2 | 2.29 | 2.296 | 2.29 | 2.22 | 2.29 | 2.34 | 2.322 | 2.243 | 2.194 | 2.35 | 2.2787 | 2.32 | 2.11 | 2.285 | 2.32 | 2.193 | 2.19 | 2.23 | 2.345 | 2.256 | 2.27 | 2.1788 | 3
| 1 | 3 | 2.4384 | 2.324 | 2.28 | 2.26 | 2.23 | 2.07 | 2.25 | 2.33 | 2.23 | 2.31 | 2.29 | 2.28 | 2.19 | 2.29 | 2.3 | 2.309 | 2.1774 | 2.26 | 2.48 | 2.328 | 2.35 | 2.19 | 2.268 | 2.37 | 2.223 | 2.26 | 2.257 | 2.339 | 2.2368 | 2.33 | 2.1754 | 16
| 1 | 4 | 2.5089 | 2.379 | 2.32 | 2.272 | 2.24 | 2.18 | 2.23 | 2.28 | 2.21 | 2.31 | 2.276 | 2.31 | 2.19 | 2.29 | 2.33 | 2.337 | 2.1819 | 2.252 | 2.35 | 2.1247 | 2.37 | 2.2 | 2.257 | 2.4 | 2.192 | 2.28 | 2.31 | 2.311 | 2.2272 | 2.35 | 2.1301 | 2.2637
| 1 | 5 | 2.5194 | 2.341 | 2.31 | 2.24 | 2.3 | 2.17 | 2.23 | 2.23 | 2.21 | 2.3 | 2.257 | 2.33 | 2.23 | 2.26 | 2.3 | 2.285 | 2.273 | 2.304 | 2.6 | 2.258 | 2.39 | 2.14 | 2.241 | 2.39 | 2.225 | 2.29 | 2.31 | 2.308 | 2.256 | 2.35 | 2.1759 | 7
| 1 | 6 | 2.474 | 2.353 | 2.35 | 2.206 | 2.22 | 2.2 | 2.25 | 2.29 | 2.22 | 2.26 | 2.305 | 2.33 | 2.21 | 2.26 | 2.23 | 2.317 | 2.2054 | 2.333 | 2.47 | 2.1207 | 2.29 | 2.18 | 2.262 | 2.4 | 2.218 | 2.26 | 2.27 | 2.374 | 2.2848 | 2.33 | 2.1476
|  | 7 | 2.4919
|  | 8 | 2.5314
|  | 9 | 2.4586
|  | 10 | 2.4716
|  | 11 | 2.3806
|  | 12 | 2.452
|  | 13 | 2.5958
|  | 14 | 2.5768
|  | 15 | 2.5191
|  | 16 | 2.5862
|  | 17 | 2.4361
|  | 18 | 2.4682
|  | 19 | 2.4949
|  | 20 | 2.5566
| Mean |  | 2.4958 | 2.3433 | 2.335 | 2.2327 | 2.245 | 2.1483 | 2.2367 | 2.275 | 2.215 | 2.2767 | 2.2572 | 2.305 | 2.21 | 2.2833 | 2.3083 | 2.3137 | 2.2074 | 2.2728 | 2.44 | 2.2078 | 2.34 | 2.1617 | 2.258 | 2.3767 | 2.2107 | 2.26 | 2.2578 | 2.3213 | 2.2464 | 2.325 | 2.1566
| Median |  | 2.4934 | 2.347 | 2.335 | 2.232 | 2.24 | 2.165 | 2.23 | 2.285 | 2.215 | 2.295 | 2.283 | 2.3 | 2.215 | 2.29 | 2.315 | 2.3145 | 2.1937 | 2.277 | 2.43 | 2.1973 | 2.335 | 2.165 | 2.2595 | 2.385 | 2.2155 | 2.27 | 2.2635 | 2.325 | 2.2464 | 2.33 | 2.1615
| Std Dev. |  | 0.055111 | 0.025398 | 0.03937 | 0.030428 | 0.028107 | 0.048751 | 0.010328 | 0.046368 | 0.010488 | 0.046332 | 0.069741 | 0.021679 | 0.016733 | 0.019664 | 0.043551 | 0.017154 | 0.042511 | 0.048688 | 0.096747 | 0.091163 | 0.036878 | 0.034303 | 0.018243 | 0.030111 | 0.014679 | 0.036332 | 0.053068 | 0.042136 | 0.024286 | 0.029496 | 0.022854
| Rel.Std.Dev. |  | 2.208% | 1.084% | 1.686% | 1.363% | 1.252% | 2.269% | 0.4618% | 2.038% | 0.4735% | 2.035% | 3.09% | 0.9405% | 0.7572% | 0.8612% | 1.887% | 0.7414% | 1.926% | 2.142% | 3.965% | 4.129% | 1.576% | 1.587% | 0.8079% | 1.267% | 0.664% | 1.608% | 2.35% | 1.815% | 1.081% | 1.269% | 1.06%
| PDM3 |  | 10.26% | 3.519% | 3.151% | -1.369% | -0.8246% | -5.095% | -1.193% | 0.5007% | -2.15% | 0.5743% | -0.2871% | 1.826% | -2.371% | 0.8689% | 1.973% | 2.209% | -2.486% | 0.405% | 7.79% | -2.469% | 3.372% | -4.506% | -0.2503% | 4.992% | -2.341% | -0.1619% | -0.2576% | 2.548% | -0.7627% | 2.71% | -4.73%
| Z-Score (Absolute) |  | NA | 1.17 | 1.05 | 0.36 | 0.19 | 1.53 | 0.31 | 0.22 | 0.61 | 0.24 | 0.03 | 0.64 | 0.68 | 0.34 | 0.68 | 0.76 | 0.72 | 0.19 | 2.51 | 0.71 | 1.12 | 1.35 | 0.01 | 1.63 | 0.67 | 0.01 | 0.02 | 0.86 | 0.17 | 0.91 | 1.42
| Analytical results for Pd in OREAS 237b (Indicative Value 0.738 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.74054 | [30 empty] | 1
| 1 | 2 | 0.65976 | [30 empty] | 1
| 1 | 3 | 0.83065 | [30 empty] | 16
| 1 | 4 | 0.71467 | [30 empty] | 0.73816
| 1 | 5 | 0.76023 | [30 empty] | 7
| 1 | 6 | 0.72314
| Mean |  | 0.73816
| Median |  | 0.73184
| Std Dev. |  | 0.056517
| Rel.Std.Dev. |  | 7.656%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 237b (Indicative Value 0.5 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.51734 | [30 empty] | 1
| 1 | 2 | 0.49495 | [30 empty] | 1
| 1 | 3 | 0.55935 | [30 empty] | 16
| 1 | 4 | 0.46912 | [30 empty] | 0.50004
| 1 | 5 | 0.4919 | [30 empty] | 7
| 1 | 6 | 0.4676
| Mean |  | 0.50004
| Median |  | 0.49343
| Std Dev. |  | 0.034402
| Rel.Std.Dev. |  | 6.88%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 237b (Certified Value 2.2 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 0.085g | 25g | 50g | 30g | 40g | 30g | 30g | 25g | 30g | 25g | 25g | 25g | 25g | 15g | 50g | 25g | 50g | 25g | 10g | 25g | 20g | 15g | 15g | 50g | 3
| 1 | 1 | 2.4984 | 2.26 | 2.27 | 2.2489 | 2.1 | > 2 | 2.25 | > 0.5 | 2.19 | 2.1 | 2.13 | 2.09 | 2.24 | 2.1682 | 2.23 | 2.25 | 2.15 | 1.74 | 2.17 | 2.256 | 2.3106 | 2.0976 | 2.1833 | 2.26 | 1
| 1 | 2 | 2.4576 | 2.19 | 2.23 | 2.2616 | 2.09 | > 2 | 2.25 | > 0.5 | 2.2 | 2.11 | 2.12 | 2.12 | 2.23 | 2.2518 | 2.21 | 2.28 | 2.17 | 1.78 | 2.18 | 2.253 | 2.3066 | 2.0361 | 2.1654 | 2.25 | 3
| 1 | 3 | 2.4384 | 2.24 | 2.19 | 2.2519 | 2.1 | > 2 | 2.24 | > 0.5 | 2.19 | 2.2 | 2.13 | 2.08 | 2.26 | 2.1984 | 2.22 | 2.35 | 2.05 | 1.64 | 2.22 | 2.244 | 2.3011 | 2.0619 | 2.241 | 2.2517 | 16
| 1 | 4 | 2.5089 | 2.25 | 2.21 | 2.2616 | 2.07 | > 2 | 2.26 | > 0.5 | 2.18 | 2.17 | 2.08 | 2.07 | 2.23 | 2.247 | 2.22 | 2.32 | 2.25 | 1.76 | 2.17 | 2.214 | 2.2638 | 2.0278 | 2.1984 | 2.2842 | 2.1961
| 1 | 5 | 2.5194 | 2.27 | 2.23 | 2.2907 | 2.01 | > 2 | 2.24 | > 0.5 | 2.19 | 2.17 | 2.14 | 1.95 | 2.25 | 2.2915 | 2.19 | 2.35 | 2.2 | 1.76 | 2.26 | 2.204 | 2.2604 | 2.0419 | 2.1698 | 2.28 | 9
| 1 | 6 | 2.474 | 2.24 | 2.21 | 2.3101 | 2.05 | > 2 | 2.24 | > 0.5 | 2.22 | 2.11 | 2.14 | 1.99 | 2.21 | 2.146 | 2.17 | 2.27 | 2.05 | 1.59 | 2.25 | 2.229 | 2.2253 | 2.0506 | 2.1963 | 2.26
|  | 7 | 2.4919
|  | 8 | 2.5314
|  | 9 | 2.4586
|  | 10 | 2.4716
|  | 11 | 2.3806
|  | 12 | 2.452
|  | 13 | 2.5958
|  | 14 | 2.5768
|  | 15 | 2.5191
|  | 16 | 2.5862
|  | 17 | 2.4361
|  | 18 | 2.4682
|  | 19 | 2.4949
|  | 20 | 2.5566
| Mean |  | 2.4958 | 2.2417 | 2.2233 | 2.2708 | 2.07 |  | 2.2467 |  | 2.195 | 2.1433 | 2.1233 | 2.05 | 2.2367 | 2.2172 | 2.2067 | 2.3033 | 2.145 | 1.7117 | 2.2083 | 2.2333 | 2.278 | 2.0526 | 2.1924 | 2.2643
| Median |  | 2.4934 | 2.245 | 2.22 | 2.2616 | 2.08 |  | 2.245 |  | 2.19 | 2.14 | 2.13 | 2.075 | 2.235 | 2.2227 | 2.215 | 2.3 | 2.16 | 1.75 | 2.2 | 2.2365 | 2.2824 | 2.0462 | 2.1898 | 2.26
| Std Dev. |  | 0.055111 | 0.027869 | 0.027325 | 0.024276 | 0.035214 |  | 0.008165 |  | 0.013784 | 0.041793 | 0.022509 | 0.065422 | 0.017512 | 0.055551 | 0.022509 | 0.04274 | 0.080932 | 0.077567 | 0.040702 | 0.021295 | 0.033773 | 0.024969 | 0.027334 | 0.014441
| Rel.Std.Dev. |  | 2.208% | 1.243% | 1.229% | 1.069% | 1.701% |  | 0.3634% |  | 0.628% | 1.95% | 1.06% | 3.191% | 0.7829% | 2.506% | 1.02% | 1.856% | 3.773% | 4.532% | 1.843% | 0.9535% | 1.483% | 1.216% | 1.247% | 0.6378%
| PDM3 |  | 13.65% | 2.075% | 1.24% | 3.402% | -5.742% |  | 2.303% |  | -0.0499% | -2.403% | -3.313% | -6.653% | 1.847% | 0.9587% | 0.4813% | 4.883% | -2.327% | -22.06% | 0.5572% | 1.696% | 3.728% | -6.532% | -0.1698% | 3.106%
| Z-Score (Absolute) |  | NA | 0.4 | 0.18 | 0.75 | 1.67 | NA | 0.46 | NA | 0.16 | 0.78 | 1.02 | 1.91 | 0.34 | 0.11 | 0.02 | 1.14 | 0.76 | 5.98 | 0 | 0.3 | 0.84 | 1.88 | 0.19 | 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 237b (Certified Value 1.98 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*OES | 2
|  |  | 0.085g | 200g | 200g | 200g | 50g | 50g | 20g | 200g | 05g | 30g | 30g | 30g | 30g | 30g | 200g | 50g | 30g | 15g | 3
| 1 | 1 | 2.4984 | 1.97 | 2 | 2.11 | 1.84 | 2.07 | 1.9 | 2.06 | 2.1 | 1.99 | 2.04 | 1.98 | 1.95 | 1.83 | 1.9059 | 2.17 | 1.82 | 1.9665 | 1
| 1 | 2 | 2.4576 | 1.97 | 1.97 | 2.06 | 1.82 | 2.05 | 1.86 | 2.07 | 2.2 | 1.97 | 2.03 | 1.99 | 1.95 | 1.86 | 1.8504 | 2.15 | 1.82 | 2.0045 | 3
| 1 | 3 | 2.4384 | 1.95 | 1.94 | 2.05 | 1.84 | 2.06 | 1.87 | 2.07 | 2.1 | 1.95 | 2.15 | 1.98 | 1.96 | 1.84 | 1.8949 | 2.21 | 1.85 | 2.0045 | 16
| 1 | 4 | 2.5089 | 1.97 | 2.01 | 2.07 | 1.87 | 2.05 | 1.9 | 2.07 | 2.1 | 1.91 | 2.01 | 1.92 | 1.98 | 1.83 | 1.9062 | 2.21 | 1.82 | 2.0045 | 1.979
| 1 | 5 | 2.5194 | 1.97 | 1.9 | 2.05 | 1.88 | 2.06 | 1.82 | 2.07 | 2.1 | 1.93 | 2.01 | 1.96 | 1.97 | 1.86 | 1.8778 | 2.2 | 1.85 | 2.0425 | 11
| 1 | 6 | 2.474 | 1.98 | 2.1 | 2.05 | 1.84 | 2.06 | 1.82 | 2.06 | 2.1 | 1.93 | 2.02 | 1.98 | 1.95 | 1.87 | 1.8985 | 2.162 | 1.8 | 2.0805
|  | 7 | 2.4919
|  | 8 | 2.5314
|  | 9 | 2.4586
|  | 10 | 2.4716
|  | 11 | 2.3806
|  | 12 | 2.452
|  | 13 | 2.5958
|  | 14 | 2.5768
|  | 15 | 2.5191
|  | 16 | 2.5862
|  | 17 | 2.4361
|  | 18 | 2.4682
|  | 19 | 2.4949
|  | 20 | 2.5566
| Mean |  | 2.4958 | 1.9683 | 1.9867 | 2.065 | 1.8483 | 2.0583 | 1.8617 | 2.0667 | 2.1167 | 1.9467 | 2.0433 | 1.9683 | 1.96 | 1.8483 | 1.8889 | 2.1837 | 1.8267 | 2.0172
| Median |  | 2.4934 | 1.97 | 1.985 | 2.055 | 1.84 | 2.06 | 1.865 | 2.07 | 2.1 | 1.94 | 2.025 | 1.98 | 1.955 | 1.85 | 1.8967 | 2.185 | 1.82 | 2.0045
| Std Dev. |  | 0.055111 | 0.0098319 | 0.068605 | 0.023452 | 0.022286 | 0.0075277 | 0.036009 | 0.005164 | 0.040825 | 0.029439 | 0.053541 | 0.025626 | 0.012649 | 0.017224 | 0.021548 | 0.026242 | 0.019664 | 0.039246
| Rel.Std.Dev. |  | 2.208% | 0.4995% | 3.453% | 1.136% | 1.206% | 0.3657% | 1.934% | 0.2499% | 1.929% | 1.512% | 2.62% | 1.302% | 0.6454% | 0.9319% | 1.141% | 1.202% | 1.076% | 1.946%
| PDM3 |  | 26.11% | -0.5393% | 0.3871% | 4.345% | -6.603% | 4.008% | -5.929% | 4.43% | 6.956% | -1.634% | 3.25% | -0.5393% | -0.9604% | -6.603% | -4.551% | 10.34% | -7.698% | 1.928%
| Z-Score (Absolute) |  | NA | 0 | 0.14 | 0.72 | 0.9 | 0.67 | 0.8 | 0.74 | 1.11 | 0.16 | 0.56 | 0 | 0.06 | 0.9 | 0.59 | 1.61 | 1.06 | 0.37
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 237b (Certified Value 2.3 ppm) | [6 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 490±20 g | 490±20 g | 490±20 g | 490±20 g | 3
| 1 | 1 | 2.4984 | 2.06 | 2.38 | 2.33 | 2.27 | 1
| 1 | 2 | 2.4576 | 2.19 | 2.31 | 2.33 | 2.25 | 3
| 1 | 3 | 2.4384 | 2.28 | 2.28 | 2.33 | 2.32 | 16
| 1 | 4 | 2.5089 | 2.35 | 2.3 | 2.33 | 2.23 | 2.2974
| 1 | 5 | 2.5194 | 2.23 | 2.29 | 2.33 | 2.25 | 13
| 1 | 6 | 2.474 | 2.2 | 2.26 | 2.33 | 2.14
| 2 | 7 | 2.4919 | 2.187 | 2.22 | 2.41 | 2.44
| 2 | 8 | 2.5314 | 2.392 | 2.4 | 2.41 | 2.29
| 2 | 9 | 2.4586 | 2.173 | 2.22 | 2.41 | 2.33
| 2 | 10 | 2.4716 | 2.364 | 2.31 | 2.41 | 2.27
| 2 | 11 | 2.3806 | 2.637 | 2.22 | 2.41 | 2.34
| 2 | 12 | 2.452 | 2.462 | 2.32 | 2.41 | 2.41
| 3 | 13 | 2.5958 |  | 2.27 | 2.29
| 3 | 14 | 2.5768 |  | 2.28 | 2.29
| 3 | 15 | 2.5191 |  | 2.39 | 2.29
| 3 | 16 | 2.5862 |  | 2.37 | 2.29
| 3 | 17 | 2.4361 |  | 2.22 | 2.29
| 3 | 18 | 2.4682 |  | 2.3 | 2.29
| 4 | 19 | 2.4949 |  | 2.286 | 2.239
| 4 | 20 | 2.5566 |  | 2.358 | 2.274
| 4 | 21 |  |  | 2.385 | 2.235
| 4 | 22 |  |  | 2.39 | 2.211
| 4 | 23 |  |  | 2.363 | 2.266
| 4 | 24 |  |  | 2.311 | 2.296
| 5 | 25 |  |  | 2.234 | 2.343
| 5 | 26 |  |  | 2.435 | 2.264
| 5 | 27 |  |  | 2.207 | 2.313
| 5 | 28 |  |  | 2.369 | 2.346
| 5 | 29 |  |  | 2.405 | 2.225
| 5 | 30 |  |  |  | 2.443
| 6 | 31 |  |  | 2.186 | 2.367
| 6 | 32 |  |  | 2.4094 | 2.408
| 6 | 33 |  |  | 2.3213 | 2.263
| 6 | 34 |  |  | 2.2175 | 2.386
| 6 | 35 |  |  | 2.3035 | 2.241
| 6 | 36 |  |  | 2.3319 | 2.285
| Mean |  | 2.4958 | 2.2937 | 2.3101 | 2.3218 | 2.295
| Median |  | 2.4934 | 2.255 | 2.31 | 2.3215 | 2.28
| Std Dev. |  | 0.055111 | 0.15618 | 0.068319 | 0.063093 | 0.080735
| Rel.Std.Dev. |  | 2.208% | 6.809% | 2.957% | 2.717% | 3.518%
| PDM3 |  | 8.639% | -0.157% | 0.5536% | 1.064% | -0.1026%
| Z-Score (Absolute) |  | NA | 0.73 | 0.62 | 1.59 | 0.62
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 237b (Certified Value 0.183 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.21 | < 0.2 | 0.18 | 0.704 | 0.2 | < 0.5 | 0.2 | 0.2 | 0.16 | 0.17 | 0.19 | 0.19 | 0.14 | < 1 | 0.35 | 0.27 | < 0.5 | 0.13 | 0.19 | 0.167 | 0.3027 | 1.3409 |  |  |  |  | 1
| 1 | 2 | 0.2 | < 0.2 | 0.19 | 0.68869 | 0.2 | < 0.5 | 0.3 | 0.2 | 0.17 | 0.17 | 0.2 | 0.19 | 0.12 | < 1 | 0.34 | 0.28 | < 0.5 | 0.13 | 0.18 | 0.161 | 0.28291 | 1.2984 |  |  |  |  | 25
| 1 | 3 | 0.19 | < 0.2 | 0.19 | 0.68025 | 0.2 | < 0.5 | 0.3 | 0.2 | 0.16 | 0.18 | 0.18 | 0.19 | 0.14 | < 1 | 0.36 | 0.31 | < 0.5 | 0.14 | 0.18 | 0.17 | 0.30887 | 1.2433 |  |  |  |  | 16
| 1 | 4 | 0.19 | < 0.2 | 0.19 | 0.68501 | 0.2 | < 0.5 | 0.2 | 0.2 | 0.18 | 0.17 | 0.19 | 0.18 | 0.12 | < 1 | 0.34 | 0.27 | < 0.5 | 0.15 | 0.19 | 0.175 | 0.30303 | 1.2581 |  |  |  |  | 0.18317
| 1 | 5 | 0.23 | < 0.2 | 0.19 | 0.6864 | 0.2 | < 0.5 | 0.2 | 0.2 | 0.17 | 0.17 | 0.2 | 0.22 | 0.15 | < 1 | 0.32 | 0.29 | < 0.5 | 0.17 | 0.18 | 0.161 | 0.33025 | 1.2539 |  |  |  |  | 15
| 1 | 6 | 0.19 | < 0.2 | 0.2 | 0.6888 | 0.2 | < 0.5 | 0.2 | 0.2 | 0.18 | 0.16 | 0.18 | 0.19 | 0.11 | < 1 | 0.51 | 0.26 | < 0.5 | 0.17 | 0.2 | 0.163 | 0.39015 | 1.0926
| Mean |  | 0.20167 |  | 0.19 | 0.68886 | 0.2 |  | 0.23333 | 0.2 | 0.17 | 0.17 | 0.19 | 0.19333 | 0.13 |  | 0.37 | 0.28 |  | 0.14833 | 0.18667 | 0.16617 | 0.31965 | 1.2479
| Median |  | 0.195 |  | 0.19 | 0.68754 | 0.2 |  | 0.2 | 0.2 | 0.17 | 0.17 | 0.19 | 0.19 | 0.13 |  | 0.345 | 0.275 |  | 0.145 | 0.185 | 0.165 | 0.30595 | 1.256
| Std Dev. |  | 0.016021 |  | 0.0063246 | 0.0080534 | 3.0405e-17 |  | 0.05164 | 3.0405e-17 | 0.0089443 | 0.0063246 | 0.0089443 | 0.013663 | 0.015492 |  | 0.069857 | 0.017889 |  | 0.018348 | 0.008165 | 0.0056006 | 0.037715 | 0.084236
| Rel.Std.Dev. |  | 7.944% |  | 3.329% | 1.169% | 1.52e-14% |  | 22.13% | 1.52e-14% | 5.261% | 3.72% | 4.708% | 7.067% | 11.92% |  | 18.88% | 6.389% |  | 12.37% | 4.374% | 3.37% | 11.8% | 6.75%
| PDM3 |  | 10.1% |  | 3.731% | 276.1% | 9.19% |  | 27.39% | 9.19% | -7.188% | -7.188% | 3.731% | 5.551% | -29.03% |  | 102% | 52.87% |  | -19.02% | 1.911% | -9.281% | 74.51% | 581.3%
| Z-Score (Absolute) |  | 0.1 | 1.11 | 0.04 | 5.9 | NA | 0.67 | NA | NA | 0.28 | 0.28 | 0.04 | 0 | 0.75 | 3.65 | 2.1 | 1.03 | 0.67 | 0.54 | 0.08 | 0.32 | 1.5 | 12.55
Results from laboratories 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 237b (Certified Value 7.32 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 7.3958 | 7.3168 | 7.16 | 6.51 | 7.33 | 7.612 | 6.24 | 7.4 | 7.07 | 7.19 | 7.45 | 7.33 | 7.31 | 7.268 | 7.0057 | 7.304 | 7.22 | 7.19 | 7.19 | 7.38 | 7.03 | 6.9989 | 7.7087 | 7.4797 |  |  | 1
| 1 | 2 | 7.3389 | 7.3033 | 7.2 | 7.18 | 7.3023 | 7.408 | 6.13 | 7.53 | 7.05 | 7.28 | 7.15 | 7.49 | 7.45 | 7.0448 | 7.0112 | 7.356 | 7.33 | 7.18 | 6.77 | 7.38 | 7.01 | 7.0702 | 7.8129 | 7.4945 |  |  | #N/A
| 1 | 3 | 7.4246 | 7.3164 | 7.28 | 7.33 | 7.28 | 7.477 | 7.07 | 7.54 | 7.06 | 7.24 | 7.32 | 7.6 | 7.47 | 7.3195 | 7.2966 | 7.304 | 7.32 | 7.17 | 7.17 | 7.36 | 6.8 | 7.2499 | 7.6309 | 7.4955 |  |  | 16
| 1 | 4 | 7.5198 | 7.313 | 7.27 | 7.76 | 7.3152 | 7.602 | 8.08 | 7.4 | 7.03 | 7.13 | 7.33 | 7.6 | 7.34 | 7.0185 | 6.9637 | 7.251 | 7.57 | 7.15 | 7.02 | 7.38 | 7.26 | 7.0193 | 7.6136 | 7.504 |  |  | 7.3159
| 1 | 5 | 7.5694 | 7.2877 | 7.3 | 7.48 | 7.3041 | 7.448 | 7.92 | 7.56 | 7.02 | 7.17 | 7.23 | 7.3 | 7.35 | 7.2962 | 7.1675 | 7.304 | 7.43 | 7.19 | 7 | 7.38 | 6.91 | 7.1806 | 7.668 | 7.4929 |  |  | 16
| 1 | 6 | 7.4737 | 7.2827 | 7.39 | 7.82 | 7.33 | 7.321 | 8.21 | 7.61 | 7 | 7.2 | 7.58 | 7.45 | 7.22 | 7.3276 | 6.9955 | 7.515 | 7.58 | 7.19 | 6.91 | 7.45 | 6.91 | 7.2356 | 7.4887 | 7.4907
| Mean |  | 7.4537 | 7.3033 | 7.2667 | 7.3467 | 7.3103 | 7.478 | 7.275 | 7.5067 | 7.0383 | 7.2017 | 7.3433 | 7.4617 | 7.3567 | 7.2124 | 7.0734 | 7.339 | 7.4083 | 7.1783 | 7.01 | 7.3883 | 6.9867 | 7.1258 | 7.6538 | 7.4929
| Median |  | 7.4492 | 7.3082 | 7.275 | 7.405 | 7.3096 | 7.4625 | 7.495 | 7.535 | 7.04 | 7.195 | 7.325 | 7.47 | 7.345 | 7.2821 | 7.0085 | 7.304 | 7.38 | 7.185 | 7.01 | 7.38 | 6.96 | 7.1254 | 7.6494 | 7.4937
| Std Dev. |  | 0.084302 | 0.014935 | 0.080416 | 0.47765 | 0.019083 | 0.11294 | 0.93408 | 0.087101 | 0.026394 | 0.052694 | 0.15384 | 0.12859 | 0.092448 | 0.1418 | 0.13056 | 0.092395 | 0.14525 | 0.016021 | 0.15862 | 0.031252 | 0.15731 | 0.11045 | 0.10764 | 0.0078944
| Rel.Std.Dev. |  | 1.131% | 0.2045% | 1.107% | 6.502% | 0.261% | 1.51% | 12.84% | 1.16% | 0.375% | 0.7317% | 2.095% | 1.723% | 1.257% | 1.966% | 1.846% | 1.259% | 1.961% | 0.2232% | 2.263% | 0.423% | 2.252% | 1.55% | 1.406% | 0.1054%
| PDM3 |  | 1.884% | -0.1714% | -0.6726% | 0.4209% | -0.07668% | 2.216% | -0.5587% | 2.608% | -3.794% | -1.561% | 0.3753% | 1.993% | 0.5576% | -1.414% | -3.315% | 0.3161% | 1.264% | -1.88% | -4.181% | 0.9904% | -4.5% | -2.599% | 4.619% | 2.419%
| Z-Score (Absolute) |  | 0.69 | 0.11 | 0.31 | 0.12 | 0.08 | 0.82 | 0.27 | 0.97 | 1.53 | 0.66 | 0.1 | 0.73 | 0.17 | 0.6 | 1.34 | 0.08 | 0.45 | 0.78 | 1.68 | 0.34 | 1.81 | 1.06 | 1.76 | 0.9
| Analytical results for As in OREAS 237b (Certified Value 591 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 0
0
| 1 | 1 | 592.6 | 548 | 522 | 593.285 | 580 | 580 | 531 | 71.6 | 554 | 602 | 581 | 613 | 601.5 | 540.3361 | 615 | 558.9 | 590 | 624 | 245.2 | 566.4 | 645.8839 | [5 empty] | 1
| 1 | 2 | 589.9 | 565 | 622 | 586.9808 | 569 | 580 | 556 | 70.2 | 569 | 580 | 576 | 610 | 611.2 | 559.8971 | 639 | 560.1 | 566.3 | 634.5 | 260.4 | 562.6 | 683.2759 | [5 empty] | 26
| 1 | 3 | 604.3 | 556 | 605 | 587.7 | 576 | 588 | 569 | 69.7 | 568 | 588 | 596 | 619 | 607.1 | 544.5845 | 683 | 564.4 | 576.6 | 611.7 | 281.8 | 579.2 | 691.3212 | [5 empty] | 16
| 1 | 4 | 588.2 | 588 | 619 | 587.6517 | 578 | 593 | 553 | 65.5 | 544 | 593 | 589 | 606 | 615.9 | 548.2454 | 669 | 563.8 | 564.1 | 616.2 | 223.9 | 554.9 | 604.9365 | [5 empty] | 591.043
| 1 | 5 | 589.1 | 540 | 628 | 594.89 | 583 | 610 | 568 | 67.4 | 575 | 581 | 565 | 642 | 623 | 546.9779 | 639 | 561.8 | 582 | 623.6 | 215.6 | 564 | 690.5898 | [5 empty] | 17
| 1 | 6 | 593.8 | 552 | 621 | 585.38 | 572 | 607 | 565 | 69.4 | 572 | 596 | 585 | 595 | 617.4 | 523.1716 | 619 | 562.6 | 567.8 | 641 | 257.6 | 589.5 | 659.7903
| Mean |  | 592.9833 | 558.1667 | 602.8333 | 589.3146 | 576.3333 | 593 | 557 | 68.9667 | 563.6667 | 590 | 582 | 614.1667 | 612.6833 | 543.8688 | 644 | 561.9333 | 574.4667 | 625.1667 | 247.4167 | 569.4333 | 662.6329
| Median |  | 591.25 | 554 | 620 | 587.6759 | 577 | 590.5 | 560.5 | 69.55 | 568.5 | 590.5 | 583 | 611.5 | 613.55 | 545.7812 | 639 | 562.2 | 572.2 | 623.8 | 251.4 | 565.2 | 671.5331
| Std Dev. |  | 5.9388 | 16.8097 | 40.3258 | 3.8249 | 5.164 | 13.0231 | 14.2969 | 2.175 | 12.0444 | 8.6487 | 10.7703 | 15.8167 | 7.7132 | 12.059 | 27.0629 | 2.1257 | 10.2097 | 10.9852 | 24.5956 | 12.6071 | 33.6102
| Rel.Std.Dev. |  | 1.002% | 3.012% | 6.689% | 0.649% | 0.896% | 2.196% | 2.567% | 3.154% | 2.137% | 1.466% | 1.851% | 2.575% | 1.259% | 2.217% | 4.202% | 0.3783% | 1.777% | 1.757% | 9.941% | 2.214% | 5.072%
| PDM3 |  | 0.3283% | -5.562% | 1.995% | -0.2924% | -2.489% | 0.3311% | -5.76% | -88.33% | -4.632% | -0.1765% | -1.53% | 3.912% | 3.661% | -7.982% | 8.96% | -4.925% | -2.805% | 5.773% | -58.14% | -3.656% | 12.11%
| Z-Score (Absolute) |  | 0.36 | 0.77 | 0.67 | 0.24 | 0.18 | 0.36 | 0.81 | 16.61 | 0.59 | 0.26 | 0 | 1.04 | 0.99 | 1.23 | 2.01 | 0.65 | 0.24 | 1.4 | 10.83 | 0.41 | 2.61
| Analytical results for Au in OREAS 237b (Indicative Value 2.21 ppm) | [27 empty] | Indicative
Au | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 2
2
| 1 | 1 | 2.1748 | [25 empty] | 1
| 1 | 2 | 2.071 | [25 empty] | 3
| 1 | 3 | 2.3348 | [25 empty] | 16
| 1 | 4 | 2.0914 | [25 empty] | 2.2062
| 1 | 5 | 2.2083 | [25 empty] | 9
| 1 | 6 | 2.357
| Mean |  | 2.2062
| Median |  | 2.1916
| Std Dev. |  | 0.11976
| Rel.Std.Dev. |  | 5.428%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for B in OREAS 237b (Indicative Value 27.4 ppm) | [27 empty] | Indicative
B | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.07 | [25 empty] | ppm
|  |  | 4A*OES | [25 empty] | 1
1
| 1 | 1 | 28.45 | [25 empty] | 1
| 1 | 2 | 25.97 | [25 empty] | 4
| 1 | 3 | 28.81 | [25 empty] | 16
| 1 | 4 | 27.09 | [25 empty] | 27.43
| 1 | 5 | 27.6 | [25 empty] | 10
| 1 | 6 | 26.66
| Mean |  | 27.43
| Median |  | 27.345
| Std Dev. |  | 1.0783
| Rel.Std.Dev. |  | 3.931%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 237b (Certified Value 683 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 689.1 | 694.49 | 689 | 709 | 663.76 | 722 | 671 | 681 | 640 | 670 | 700 | 680 | 680 | 684 | 668.2869 | 680 | 615.8 | 668 | 668.1 | 696 | 674 | 672 | 645.7 | 705 | 694.6917 | 606.9658 | 1
| 1 | 2 | 681 | 692.85 | 687 | 737 | 664.224 | 712 | 694 | 677 | 640 | 680 | 670 | 690 | 690 | 673 | 674.6885 | 680 | 619.9 | 671 | 670.1 | 675 | 675 | 658 | 650.6 | 717 | 714.9434 | 624.1406 | 27
| 1 | 3 | 693.9 | 707.98 | 694 | 738 | 665.4276 | 709 | 679 | 712 | 654 | 680 | 680 | 700 | 700 | 685 | 673.7857 | 673 | 643.2 | 658 | 677.4 | 690 | 670 | 645 | 669.5 | 701 | 716.4162 | 600.766 | 16
| 1 | 4 | 687.6 | 699.06 | 694 | 748 | 666.26 | 730 | 709 | 684 | 639 | 680 | 690 | 700 | 680 | 698 | 674.4572 | 678 | 1079 | 701 | 681.2 | 677 | 674 | 671 | 644.2 | 700 | 691.5741 | 624.6855 | 683.1123
| 1 | 5 | 687.2 | 699.09 | 692 | 760 | 665.4592 | 711 | 717 | 664 | 638 | 680 | 670 | 680 | 680 | 698 | 664.5874 | 676 | 639.9 | 701 | 683.7 | 685 | 677 | 667 | 660.7 | 704 | 698.9875 | 629.3918 | 18
| 1 | 6 | 687.5 | 694.25 | 705 | 778 | 665.12 | 729 | 718 | 706 | 646 | 680 | 700 | 690 | 670 | 707 | 667.5539 | 692 | 635.3 | 684 | 672.9 | 673 | 682 | 682 | 664.5 | 688 | 708.6592 | 605.3802
| Mean |  | 687.7167 | 697.9533 | 693.5 | 745 | 665.0418 | 718.8333 | 698 | 687.3333 | 642.8333 | 678.3333 | 685 | 690 | 683.3333 | 690.8333 | 670.56 | 679.8333 | 705.5167 | 680.5 | 675.5667 | 682.6667 | 675.3333 | 665.8333 | 655.8667 | 702.5 | 704.212 | 615.2216
| Median |  | 687.55 | 696.775 | 693 | 743 | 665.2738 | 717 | 701.5 | 682.5 | 640 | 680 | 685 | 690 | 680 | 691.5 | 671.0363 | 679 | 637.6 | 677.5 | 675.15 | 681 | 674.5 | 669 | 655.65 | 702.5 | 703.8234 | 615.5532
| Std Dev. |  | 4.1374 | 5.5642 | 6.2849 | 23.3752 | 0.90829 | 9.411 | 19.9399 | 18.2172 | 6.1455 | 4.0825 | 13.784 | 8.9443 | 10.328 | 12.3518 | 4.3015 | 6.5243 | 183.2967 | 17.9193 | 6.2289 | 9.1797 | 3.9833 | 12.8284 | 10.4974 | 9.3541 | 10.5975 | 12.1974
| Rel.Std.Dev. |  | 0.6016% | 0.7972% | 0.9063% | 3.138% | 0.1366% | 1.309% | 2.857% | 2.65% | 0.956% | 0.6018% | 2.012% | 1.296% | 1.511% | 1.788% | 0.6415% | 0.9597% | 25.98% | 2.633% | 0.922% | 1.345% | 0.5898% | 1.927% | 1.601% | 1.332% | 1.505% | 1.983%
| PDM3 |  | 0.674% | 2.173% | 1.521% | 9.06% | -2.645% | 5.229% | 2.179% | 0.6179% | -5.896% | -0.6996% | 0.2763% | 1.008% | 0.03235% | 1.13% | -1.838% | -0.48% | 3.28% | -0.3824% | -1.105% | -0.06524% | -1.139% | -2.529% | -3.988% | 2.838% | 3.089% | -9.938%
| Z-Score (Absolute) |  | 0.21 | 0.81 | 0.55 | 3.58 | 1.12 | 2.04 | 0.81 | 0.19 | 2.43 | 0.34 | 0.05 | 0.34 | 0.05 | 0.39 | 0.8 | 0.25 | 1.26 | 0.22 | 0.51 | 0.09 | 0.52 | 1.08 | 1.66 | 1.08 | 1.18 | 4.05
| Analytical results for Be in OREAS 237b (Certified Value 2.47 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.45 | 2.5 | 2.4 | < 5 | 2.3 | 2.48 | 2.4 | 2 | 2.32 | 2.56 | 2.69 | 2.33 | 3 | 2.3343 | 3 | 2.084 | 2.7 | 3.2 | 2.59 | 2.41 | 3 | 2.343 | 2.8867 | 2.3714 |  |  | 1
| 1 | 2 | 2.48 | 2.5 | 2.7 | < 5 | 2.3 | 2.55 | 2.5 | 2 | 2.33 | 2.55 | 2.52 | 2.3 | 2 | 2.3339 | 3.1 | 2.123 | 2.8 | 3.19 | 2.42 | 2.27 | 2 | 2.3518 | 2.8311 | 2.4028 |  |  | 28
| 1 | 3 | 2.47 | 2.5 | 2.6 | < 5 | 2.2 | 2.46 | 2.9 | 2 | 2.24 | 2.57 | 2.54 | 2.44 | 3 | 2.4179 | 2.9 | 2.129 | 2.9 | 3.16 | 2.32 | 2.32 | 2 | 2.4189 | 2.8867 | 2.4348 |  |  | 16
| 1 | 4 | 2.56 | 2.5 | 2.5 | < 5 | 2.3 | 2.6 | 2.8 | 2 | 2.32 | 2.56 | 2.53 | 2.33 | 3 | 2.3726 | 3 | 3.547 | 2.8 | 3.13 | 2.39 | 2.45 | 2 | 2.3298 | 2.8311 | 2.3973 |  |  | 2.4654
| 1 | 5 | 2.48 | 2.5 | 2.6 | < 5 | 2.3 | 2.61 | 2.7 | 2 | 2.33 | 2.41 | 2.62 | 2.3 | 2 | 2.2974 | 2.9 | 2.059 | 2.7 | 3.18 | 2.59 | 2.32 | 2 | 2.3881 | 2.8311 | 2.4233 |  |  | 19
| 1 | 6 | 2.51 | 2.5 | 2.5 | < 5 | 2.3 | 2.65 | 2.5 | 2 | 2.34 | 2.61 | 2.48 | 2.26 | 2 | 2.3217 | 3.1 | 2.085 | 2.5 | 3.21 | 2.31 | 2.4 | 2 | 2.4079 | 2.8867 | 2.3973
| Mean |  | 2.4917 | 2.5 | 2.55 |  | 2.2833 | 2.5583 | 2.6333 | 2 | 2.3133 | 2.5433 | 2.5633 | 2.3267 | 2.5 | 2.3463 | 3 | 2.3378 | 2.7333 | 3.1783 | 2.4367 | 2.3617 | 2.1667 | 2.3733 | 2.8589 | 2.4045
| Median |  | 2.48 | 2.5 | 2.55 |  | 2.3 | 2.575 | 2.6 | 2 | 2.325 | 2.56 | 2.535 | 2.315 | 2.5 | 2.3341 | 3 | 2.104 | 2.75 | 3.185 | 2.405 | 2.36 | 2 | 2.37 | 2.8589 | 2.4
| Std Dev. |  | 0.038687 | 0 | 0.10488 |  | 0.040825 | 0.075741 | 0.19664 | 0 | 0.036697 | 0.068605 | 0.077115 | 0.06121 | 0.54772 | 0.042681 | 0.089443 | 0.59295 | 0.13663 | 0.029269 | 0.1258 | 0.068532 | 0.40825 | 0.036792 | 0.030453 | 0.022262
| Rel.Std.Dev. |  | 1.553% | 0% | 4.113% |  | 1.788% | 2.961% | 7.467% | 0% | 1.586% | 2.697% | 3.008% | 2.631% | 21.91% | 1.819% | 2.981% | 25.36% | 4.999% | 0.9209% | 5.163% | 2.902% | 18.84% | 1.55% | 1.065% | 0.9259%
| PDM3 |  | 1.064% | 1.402% | 3.43% |  | -7.386% | 3.768% | 6.81% | -18.88% | -6.17% | 3.159% | 3.971% | -5.629% | 1.402% | -4.833% | 21.68% | -5.176% | 10.87% | 28.92% | -1.167% | -4.209% | -12.12% | -3.739% | 15.96% | -2.473%
| Z-Score (Absolute) |  | 0.04 | 0 | 0.25 | 0 | 1.1 | 0.3 | 0.67 | NA | 0.94 | 0.22 | 0.32 | 0.88 | NA | 0.78 | 2.53 | 0.82 | 1.18 | 3.43 | 0.32 | 0.7 | NA | 0.64 | 1.82 | 0.48
Results from laboratories 1.11, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 237b (Certified Value 0.35 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 0.36 | 0.41824 | 0.4 | 0.32 | < 5 | 0.35 | 0.3 | 0.37 | 0.4 | 0.33 | 0.36 | 0.36 | 0.33 | 0.31 | 0.34228 | 0.5 | 0.331 | 0.35 | < 5 | 0.39 | 0.31 | 0.33 | 0.4139 | < 5 | 1.1734 | 0.34027 | 1
| 1 | 2 | 0.39 | 0.39739 | 0.4 | 0.35 | < 5 | 0.33 | 0.35 | 0.33 | 0.4 | 0.36 | 0.33 | 0.34 | 0.34 | 0.31 | 0.34761 | 0.5 | 0.373 | 0.37 | < 5 | 0.45 | 0.27 | 0.33 | 0.3499 | < 5 | 1.7841 | 0.38488 | 29
| 1 | 3 | 0.34 | 0.40939 | 0.3 | 0.34 | < 5 | 0.37 | 0.38 | 0.36 | 0.4 | 0.32 | 0.33 | 0.35 | 0.32 | 0.33 | 0.35268 | 0.5 | 0.325 | 0.36 | < 5 | 0.36 | 0.32 | 0.32 | 0.3864 | < 5 | 1.4728 | 0.35753 | 16
| 1 | 4 | 0.36 | 0.39007 | 0.3 | 0.59 | < 5 | 0.35 | 0.4 | 0.34 | 0.3 | 0.56 | 0.34 | 0.36 | 0.34 | 0.36 | 0.3275 | 0.5 | 0.555 | 0.38 | < 5 | 0.36 | 0.3 | 0.34 | 0.406 | < 5 | 1.1734 | 0.36062 | 0.3537
| 1 | 5 | 0.38 | 0.39336 | 0.4 | 0.33 | < 5 | 0.36 | 0.37 | 0.35 | 0.4 | 0.33 | 0.32 | 0.35 | 0.34 | 0.37 | 0.35354 | 0.4 | 0.333 | 0.38 | < 5 | 0.35 | 0.31 | 0.34 | 0.3838 | < 5 | 1.0591 | 0.35048 | 20
| 1 | 6 | 0.4 | 0.40369 | 0.4 | 0.35 | < 5 | 0.38 | 0.39 | 0.33 | 0.4 | 0.42 | 0.34 | 0.37 | 0.32 | 0.35 | 0.36268 | 0.4 | 0.494 | 0.34 | < 5 | 0.73 | 0.33 | 0.34 | 0.405 | < 5 | 2.4145 | 0.37996
| Mean |  | 0.37167 | 0.40202 | 0.36667 | 0.38 |  | 0.35667 | 0.365 | 0.34667 | 0.38333 | 0.38667 | 0.33667 | 0.355 | 0.33167 | 0.33833 | 0.34771 | 0.46667 | 0.40183 | 0.36333 |  | 0.44 | 0.30667 | 0.33333 | 0.39083 |  | 1.5129 | 0.36229
| Median |  | 0.37 | 0.40054 | 0.4 | 0.345 |  | 0.355 | 0.375 | 0.345 | 0.4 | 0.345 | 0.335 | 0.355 | 0.335 | 0.34 | 0.35014 | 0.5 | 0.353 | 0.365 |  | 0.375 | 0.31 | 0.335 | 0.3957 |  | 1.3231 | 0.35908
| Std Dev. |  | 0.022286 | 0.010579 | 0.05164 | 0.10354 |  | 0.017512 | 0.036194 | 0.01633 | 0.040825 | 0.092448 | 0.013663 | 0.010488 | 0.0098319 | 0.025626 | 0.012007 | 0.05164 | 0.098433 | 0.01633 |  | 0.1467 | 0.020656 | 0.008165 | 0.02327 |  | 0.51468 | 0.01716
| Rel.Std.Dev. |  | 5.996% | 2.632% | 14.08% | 27.25% |  | 4.91% | 9.916% | 4.711% | 10.65% | 23.91% | 4.058% | 2.954% | 2.964% | 7.574% | 3.453% | 11.07% | 24.5% | 4.494% |  | 33.34% | 6.736% | 2.449% | 5.954% |  | 34.02% | 4.737%
| PDM3 |  | 5.079% | 13.66% | 3.665% | 7.435% |  | 0.8379% | 3.194% | -1.989% | 8.377% | 9.32% | -4.817% | 0.3666% | -6.23% | -4.345% | -1.693% | 31.94% | 13.61% | 2.723% |  | 24.4% | -13.3% | -5.759% | 10.5% |  | 327.7% | 2.428%
| Z-Score (Absolute) |  | 0.17 | 0.94 | NA | 0.38 | 53.99 | 0.21 | 0 | 0.46 | NA | 0.55 | 0.72 | 0.25 | 0.84 | 0.67 | 0.44 | NA | 0.93 | 0.04 | 53.99 | 1.9 | 1.48 | 0.8 | 0.65 | 53.99 | 29.03 | 0.07
Results from laboratories 1.05, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 237b (Certified Value 0.903 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.9063 | 0.89508 | 0.87 | 0.81 | 0.88 | 0.959 | 0.72 | 0.88 | 0.91 | 0.88 | 0.93 | 0.92 | 0.92 | 0.8905 | 0.90124 | 0.822 | 0.88 | 0.91 | 0.9 | 0.89 | 0.86 | 0.86245 | 0.9512 |  |  |  | 1
| 1 | 2 | 0.9066 | 0.89619 | 0.87 | 0.96 | 0.90667 | 0.936 | 0.79 | 0.893 | 0.9 | 0.89 | 0.89 | 0.94 | 0.93 | 0.8588 | 0.87224 | 0.822 | 0.89 | 0.93 | 0.86 | 0.89 | 0.86 | 0.8664 | 0.9751 |  |  |  | #N/A
| 1 | 3 | 0.9123 | 0.89796 | 0.88 | 0.97 | 0.90333 | 0.934 | 0.84 | 0.893 | 0.9 | 0.89 | 0.91 | 0.94 | 0.95 | 0.8823 | 0.89807 | 0.815 | 0.9 | 0.92 | 0.9 | 0.89 | 0.84 | 0.89328 | 0.9481 |  |  |  | 16
| 1 | 4 | 0.9103 | 0.89789 | 0.86 | 0.93 | 0.91 | 0.961 | 0.92 | 0.883 | 0.89 | 0.88 | 0.91 | 0.95 | 0.93 | 0.8527 | 0.89833 | 0.815 | 0.92 | 0.92 | 0.89 | 0.89 | 0.88 | 0.85531 | 0.953 |  |  |  | 0.90308
| 1 | 5 | 0.9026 | 0.89366 | 0.88 | 0.94 | 0.88333 | 0.939 | 0.89 | 0.905 | 0.9 | 0.9 | 0.9 | 0.91 | 0.93 | 0.8803 | 0.88444 | 0.815 | 0.9 | 0.92 | 0.89 | 0.89 | 0.84 | 0.8789 | 0.9605 |  |  |  | 21
| 1 | 6 | 0.9079 | 0.89524 | 0.9 | 0.96 | 0.88 | 0.933 | 0.94 | 0.903 | 0.89 | 0.9 | 0.94 | 0.93 | 0.91 | 0.8843 | 0.87398 | 0.836 | 0.92 | 0.92 | 0.88 | 0.9 | 0.85 | 0.88933 | 0.9348
| Mean |  | 0.90767 | 0.896 | 0.87667 | 0.92833 | 0.89389 | 0.94367 | 0.85 | 0.89283 | 0.89833 | 0.89 | 0.91333 | 0.93167 | 0.92833 | 0.87482 | 0.88805 | 0.82083 | 0.90167 | 0.92 | 0.88667 | 0.89167 | 0.855 | 0.87428 | 0.95378
| Median |  | 0.90725 | 0.89572 | 0.875 | 0.95 | 0.89333 | 0.9375 | 0.865 | 0.893 | 0.9 | 0.89 | 0.91 | 0.935 | 0.93 | 0.8813 | 0.89125 | 0.8185 | 0.9 | 0.92 | 0.89 | 0.89 | 0.855 | 0.87265 | 0.9521
| Std Dev. |  | 0.0033803 | 0.0016941 | 0.013663 | 0.059805 | 0.014207 | 0.012832 | 0.083905 | 0.010128 | 0.0075277 | 0.0089443 | 0.018619 | 0.01472 | 0.013292 | 0.015282 | 0.01297 | 0.0081833 | 0.016021 | 0.0063246 | 0.015055 | 0.0040825 | 0.015166 | 0.015304 | 0.013412
| Rel.Std.Dev. |  | 0.3724% | 0.1891% | 1.558% | 6.442% | 1.589% | 1.36% | 9.871% | 1.134% | 0.838% | 1.005% | 2.039% | 1.58% | 1.432% | 1.747% | 1.46% | 0.997% | 1.777% | 0.6875% | 1.698% | 0.4578% | 1.774% | 1.751% | 1.406%
| PDM3 |  | 0.5076% | -0.7839% | -2.925% | 2.796% | -1.018% | 4.494% | -5.878% | -1.135% | -0.5259% | -1.449% | 1.135% | 3.165% | 2.796% | -3.13% | -1.665% | -9.108% | -0.1568% | 1.873% | -1.818% | -1.264% | -5.324% | -3.19% | 5.614%
| Z-Score (Absolute) |  | 0.49 | 0.07 | 0.61 | 1.22 | 0 | 1.76 | 1.55 | 0.04 | 0.16 | 0.14 | 0.69 | 1.34 | 1.22 | 0.67 | 0.21 | 2.59 | 0.27 | 0.92 | 0.26 | 0.08 | 1.37 | 0.69 | 2.12
| Analytical results for Cd in OREAS 237b (Indicative Value 0.087 ppm) | [27 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.07 | 1.34 | < 0.5 | 0.3 | 0.7995 | 0.07 | 0.11 | 0.12 | 0.2 | 0.05 | 0.04 | 0.03 | 0.04 | 0.1 | 0.10951 | < 0.1 | 0.102 | 0.05 | 7.25 | 0.06 | 0.05 | 0.04 | 0.1339 | < 5 | 0.2349 | 0.05784 | 1
| 1 | 2 | 0.06 | 1.32 | < 0.5 | 0.4 | 0.83 | 0.05 | 0.12 | 0.24 | 0.3 | 0.04 | 0.05 | 0.05 | 0.04 | 0.09 | 0.11277 | < 0.1 | < 0.1 | 0.04 | 7.35 | 0.04 | 0.04 | 0.04 | 0.1111 | < 5 | 0.2367 | 0.06346 | 5
| 1 | 3 | 0.06 | 1.28 | < 0.5 | 0.4 | 0.804 | 0.05 | 0.12 | 0.32 | 0.2 | 0.04 | 0.04 | 0.03 | 0.05 | 0.09 | 0.12388 | < 0.1 | < 0.1 | 0.06 | 7.54 | 0.03 | 0.04 | 0.03 | 0.1285 | < 5 | 0.2403 | 0.05406 | 16
| 1 | 4 | 0.06 | 1.24 | < 0.5 | < 0.3 | 0.80933 | 0.07 | 0.13 | 0.29 | 0.2 | 0.04 | 0.04 | 0.03 | 0.04 | 0.1 | 0.10466 | < 0.1 | 0.213 | 0.06 | 7.69 | 0.08 | 0.04 | 0.04 | 0.1165 | < 5 | 0.2385 | 0.06105 | 0.08679
| 1 | 5 | 0.08 | 1.29 | < 0.5 | 0.4 | 0.82925 | 0.05 | 0.12 | 0.31 | 0.2 | 0.04 | 0.04 | 0.04 | 0.04 | 0.08 | 0.11134 | < 0.1 | 0.105 | 0.04 | 7.79 | 0.04 | 0.04 | 0.05 | 0.1151 | < 5 | 0.2385 | 0.05915 | 11
| 1 | 6 | 0.06 | 1.28 | < 0.5 | 0.3 | 0.82413 | 0.05 | 0.14 | 0.32 | 0.3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.09 | 0.11055 | 0.1 | < 0.1 | 0.05 | 7.26 | 0.03 | 0.05 | 0.04 | 0.1299 | < 5 | 0.242 | 0.06402
| Mean |  | 0.065 | 1.2917 |  | 0.36 | 0.81603 | 0.056667 | 0.12333 | 0.26667 | 0.23333 | 0.041667 | 0.041667 | 0.036667 | 0.041667 | 0.091667 | 0.11212 | 0.1 | 0.14 | 0.05 | 7.48 | 0.046667 | 0.043333 | 0.04 | 0.1225 |  | 0.23848 | 0.05993
| Median |  | 0.06 | 1.285 |  | 0.4 | 0.81673 | 0.05 | 0.12 | 0.3 | 0.2 | 0.04 | 0.04 | 0.035 | 0.04 | 0.09 | 0.11094 | 0.1 | 0.105 | 0.05 | 7.445 | 0.04 | 0.04 | 0.04 | 0.1225 |  | 0.2385 | 0.0601
| Std Dev. |  | 0.0083666 | 0.034881 |  | 0.054772 | 0.013403 | 0.010328 | 0.010328 | 0.077889 | 0.05164 | 0.0040825 | 0.0040825 | 0.008165 | 0.0040825 | 0.0075277 | 0.0063949 |  | 0.063238 | 0.0089443 | 0.22891 | 0.019664 | 0.005164 | 0.0063246 | 0.0093962 |  | 0.0025175 | 0.0037384
| Rel.Std.Dev. |  | 12.87% | 2.7% |  | 15.21% | 1.643% | 18.23% | 8.374% | 29.21% | 22.13% | 9.798% | 9.798% | 22.27% | 9.798% | 8.212% | 5.704% |  | 45.17% | 17.89% | 3.06% | 42.14% | 11.92% | 15.81% | 7.67% |  | 1.056% | 6.238%
| PDM3 |  | -25.11% | 1388% |  | 314.8% | 840.2% | -34.71% | 42.11% | 207.3% | 168.8% | -51.99% | -51.99% | -57.75% | -51.99% | 5.619% | 29.18% | 15.22% | 61.31% | -42.39% | 8519% | -46.23% | -50.07% | -53.91% | 41.15% |  | 174.8% | -30.95%
| Z-Score (Absolute) |  | 0.37 | 15.64 | 2.04 | 3.02 | 9.43 | 0.48 | 0.39 | 2.26 | 1.83 | 0.67 | 0.67 | 0.74 | 0.67 | 0.02 | 0.25 | 0.46 | 0.02 | 0.57 | 96.4 | 0.61 | 0.65 | 0.7 | 0.38 | 31.41 | 1.89 | 0.44
| Analytical results for Ce in OREAS 237b (Certified Value 81 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 83.52 | 90.3764 | 81.7 | 83.4 | 83.35 | 79 | 81.1 | 85.2 | 81.7 | 76 | 78.98 | 76.8674 | 65 | 75.79 | 75.22 | 83.46 | 83.04 | 78.16 | 68.3 | 82.4919 | 74.753 | [5 empty] | 1
| 1 | 2 | 83.22 | 89.1197 | 80.3 | 85 | 85.3 | 78 | 82.9 | 81.5 | 79.2 | 79.3 | 77.18 | 77.7861 | 65 | 76.78 | 76.1 | 81.28 | 79.25 | 78.16 | 68 | 79.8871 | 76.4999 | [5 empty] | 31
| 1 | 3 | 82.81 | 88.1989 | 81.9 | 86 | 86.51 | 80 | 80.8 | 86.1 | 84.4 | 79.7 | 78.71 | 78.315 | 65 | 85.94 | 74.64 | 83.12 | 79.66 | 77.51 | 70.3 | 79.4829 | 75.4736 | [5 empty] | 16
| 1 | 4 | 83.64 | 89.0861 | 82.5 | 86.6 | 93.53 | 77 | 80.1 | 85.2 | 83.4 | 78.5 | 81.77 | 77.5224 | 65 | 77.58 | 79.22 | 82.93 | 80.66 | 79.8 | 66.7 | 77.8213 | 76.7912 | [5 empty] | 80.5252
| 1 | 5 | 80.85 | 89.4392 | 79.1 | 86.7 | 91.03 | 78 | 83.7 | 79.3 | 80.7 | 78.6 | 78.7 | 77.1916 | 64 | 79.64 | 79.13 | 81.16 | 79.68 | 78.36 | 68.7 | 78.5398 | 77.2531 | [5 empty] | 22
| 1 | 6 | 83.16 | 88.9502 | 82.2 | 89.4 | 92.84 | 79 | 84.1 | 87.1 | 82.5 | 75.9 | 81.69 | 76.884 | 67 | 77.9 | 74.69 | 81.99 | 87.78 | 80.08 | 70.6 | 79.6626 | 74.1998
| Mean |  | 82.8667 | 89.1951 | 81.2833 | 86.1833 | 88.76 | 78.5 | 82.1167 | 84.0667 | 81.9833 | 78 | 79.505 | 77.4278 | 65.1667 | 78.9383 | 76.5 | 82.3233 | 81.6783 | 78.6783 | 68.7667 | 79.6476 | 75.8285
| Median |  | 83.19 | 89.1029 | 81.8 | 86.3 | 88.77 | 78.5 | 82 | 85.2 | 82.1 | 78.55 | 78.845 | 77.357 | 65 | 77.74 | 75.66 | 82.46 | 80.17 | 78.26 | 68.5 | 79.5727 | 75.9868
| Std Dev. |  | 1.0301 | 0.71063 | 1.3121 | 1.9984 | 4.2627 | 1.0488 | 1.6666 | 3.0071 | 1.8755 | 1.6492 | 1.8366 | 0.56391 | 0.98319 | 3.6609 | 2.1377 | 0.98466 | 3.2889 | 1.0224 | 1.4692 | 1.5961 | 1.2118
| Rel.Std.Dev. |  | 1.243% | 0.7967% | 1.614% | 2.319% | 4.803% | 1.336% | 2.03% | 3.577% | 2.288% | 2.114% | 2.31% | 0.7283% | 1.509% | 4.638% | 2.794% | 1.196% | 4.027% | 1.299% | 2.137% | 2.004% | 1.598%
| PDM3 |  | 2.908% | 10.77% | 0.9415% | 7.027% | 10.23% | -2.515% | 1.976% | 4.398% | 1.811% | -3.136% | -1.267% | -3.847% | -19.07% | -1.971% | -4.999% | 2.233% | 1.432% | -2.294% | -14.6% | -1.09% | -5.833%
| Z-Score (Absolute) |  | 0.88 | 2.61 | 0.45 | 1.78 | 2.49 | 0.31 | 0.67 | 1.21 | 0.64 | 0.45 | 0.04 | 0.61 | 3.95 | 0.19 | 0.86 | 0.73 | 0.55 | 0.26 | 2.97 | 0 | 1.04
| Analytical results for Co in OREAS 237b (Certified Value 17.3 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 17.1 | 15.3128 | 16 | 17 | 20.0567 | 17.3 | 16 | 17 | 16.6 | 17.8 | 17.6 | 17.3 | 17.7 | 17.5 | 17.6192 | 17.3 | 14.63 | 16 | 23.15 | 17.4 | 17.5 | 18 | 19 | 20 | 20.8798 | 16.9098 | 1
| 1 | 2 | 17.3 | 15.1708 | 16 | 17 | 21.79 | 16.9 | 16 | 16 | 16.1 | 17.9 | 16.8 | 17 | 17.3 | 16.6 | 17.7751 | 17.7 | 14.96 | 17.4 | 23.36 | 16.9 | 16.6 | 17.5 | 19.475 | 20 | 20.2785 | 17.1518 | 32
| 1 | 3 | 17.2 | 15.1611 | 16 | 17 | 22.0733 | 16.9 | 17 | 17 | 16.7 | 17.3 | 17.6 | 17.5 | 18 | 18.3 | 16.3651 | 17.7 | 14.89 | 17.7 | 23.03 | 16.9 | 16.5 | 17.9 | 20.14 | 20 | 20.479 | 16.8699 | 16
| 1 | 4 | 17.5 | 15.084 | 16 | 17 | 22.13 | 17.4 | 17 | 16 | 15.6 | 17.3 | 17.5 | 17.1 | 17.3 | 18.2 | 17.5039 | 17.4 | 25.15 | 17.4 | 23.25 | 17.4 | 16.8 | 18.2 | 19.19 | 20 | 19.9445 | 17.3351 | 17.2918
| 1 | 5 | 17.2 | 15.2874 | 16 | 17 | 22.4133 | 17.2 | 16 | 16 | 16.4 | 17.9 | 16.1 | 17.4 | 17.8 | 17.9 | 16.7508 | 17.2 | 16.56 | 17.1 | 23.38 | 17.3 | 16.4 | 17.9 | 19.665 | 20 | 20.813 | 17.0418 | 23
| 1 | 6 | 17.6 | 15.5975 | 16 | 18 | 21.9133 | 17.4 | 16 | 17 | 16.4 | 17.8 | 17.8 | 17 | 17.5 | 18.4 | 16.7249 | 17.7 | 15.85 | 16.6 | 23.28 | 16.9 | 18.1 | 18 | 19.38 | 19 | 20.8798 | 16.615
| Mean |  | 17.3167 | 15.2689 | 16 | 17.1667 | 21.7294 | 17.1833 | 16.3333 | 16.5 | 16.3 | 17.6667 | 17.2333 | 17.2167 | 17.6 | 17.8167 | 17.1232 | 17.5 | 17.0067 | 17.0333 | 23.2417 | 17.1333 | 16.9833 | 17.9167 | 19.475 | 19.8333 | 20.5458 | 16.9872
| Median |  | 17.25 | 15.2291 | 16 | 17 | 21.9933 | 17.25 | 16 | 16.5 | 16.4 | 17.8 | 17.55 | 17.2 | 17.6 | 18.05 | 17.1274 | 17.55 | 15.405 | 17.25 | 23.265 | 17.1 | 16.7 | 17.95 | 19.4275 | 20 | 20.646 | 16.9758
| Std Dev. |  | 0.19408 | 0.18203 | 0 | 0.40825 | 0.84641 | 0.23166 | 0.5164 | 0.54772 | 0.4 | 0.28752 | 0.6532 | 0.2137 | 0.28284 | 0.67946 | 0.58103 | 0.22804 | 4.0539 | 0.62823 | 0.13258 | 0.2582 | 0.67355 | 0.23166 | 0.39855 | 0.40825 | 0.38262 | 0.24875
| Rel.Std.Dev. |  | 1.121% | 1.192% | 0% | 2.378% | 3.895% | 1.348% | 3.162% | 3.32% | 2.454% | 1.627% | 3.79% | 1.241% | 1.607% | 3.814% | 3.393% | 1.303% | 23.84% | 3.688% | 0.5704% | 1.507% | 3.966% | 1.293% | 2.046% | 2.058% | 1.862% | 1.464%
| PDM3 |  | 0.1438% | -11.7% | -7.471% | -0.7237% | 25.66% | -0.6273% | -5.543% | -4.579% | -5.736% | 2.168% | -0.3381% | -0.4345% | 1.782% | 3.035% | -0.9752% | 1.204% | -1.649% | -1.495% | 34.41% | -0.9165% | -1.784% | 3.614% | 12.63% | 14.7% | 18.82% | -1.761%
| Z-Score (Absolute) |  | 0.21 | 4.97 | NA | NA | 11.37 | 0.13 | NA | NA | 2.36 | 1.1 | 0 | 0.04 | 0.93 | 1.48 | 0.28 | 0.67 | 0.57 | 0.51 | 15.19 | 0.25 | 0.63 | 1.73 | 5.67 | NA | 8.38 | 0.62
Results from laboratories 1.05, 1.06, 1.09, 1.10 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 237b (Certified Value 123 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 129 | 100 | 95 | 126.675 | 114 | 156 | 120 | 100 | 125 | 130 | 126 | 120 | 106 | 128.8886 | 129 | 114.9 | 144.1 | 128 | 119 | 108 | 129.4 | 118 | 96.367 | 144.0021 |  |  | 1
| 1 | 2 | 144 | 110 | 104 | 124.52 | 112 | 137 | 114 | 101 | 125 | 125 | 120 | 121 | 101 | 131.5681 | 134 | 62.63 | 146.1 | 124 | 119 | 111 | 130.1 | 117 | 89.7645 | 148.1445 |  |  | 33
| 1 | 3 | 149 | 110 | 119 | 127.06 | 115 | 168 | 113 | 99 | 126 | 128 | 122 | 124 | 105 | 129.9174 | 132 | 121.7 | 147.9 | 130 | 122 | 116 | 134.3 | 120 | 95.0845 | 145.9173 |  |  | 16
| 1 | 4 | 133 | 110 | 108 | 124.895 | 119 | 157 | 114 | 107 | 126 | 128 | 124 | 122 | 113 | 130.0263 | 133 | 201.6 | 146.6 | 130 | 125 | 116 | 128.6 | 117 | 95.7495 | 148.0362 |  |  | 122.8375
| 1 | 5 | 138 | 110 | 118 | 122.53 | 123 | 147 | 107 | 104 | 126 | 126 | 122 | 121 | 107 | 124.7581 | 133 | 116.7 | 146.6 | 129 | 122 | 110 | 132.6 | 120 | 90.1208 | 147.2717 |  |  | 24
| 1 | 6 | 151 | 100 | 118 | 122.985 | 123 | 132 | 92 | 107 | 127 | 131 | 122 | 120 | 113 | 133.2862 | 136 | 117.6 | 145 | 125 | 116 | 116 | 132.8 | 122 | 91.1183 | 141.8593
| Mean |  | 140.6667 | 106.6667 | 110.3333 | 124.7775 | 117.6667 | 149.5 | 110 | 103 | 125.8333 | 128 | 122.6667 | 121.3333 | 107.5 | 129.7408 | 132.8333 | 122.5217 | 146.05 | 127.6667 | 120.5 | 112.8333 | 131.3 | 119 | 93.0341 | 145.8719
| Median |  | 141 | 110 | 113 | 124.7075 | 117 | 151.5 | 113.5 | 102.5 | 126 | 128 | 122 | 121 | 106.5 | 129.9719 | 133 | 117.15 | 146.35 | 128.5 | 120.5 | 113.5 | 131.35 | 119 | 93.1014 | 146.5945
| Std Dev. |  | 8.8242 | 5.164 | 9.7297 | 1.8523 | 4.7188 | 13.487 | 9.7365 | 3.5214 | 0.75277 | 2.2804 | 2.0656 | 1.5055 | 4.7223 | 2.8836 | 2.3166 | 44.6257 | 1.3368 | 2.582 | 3.1464 | 3.6009 | 2.2486 | 2 | 3.0177 | 2.5063
| Rel.Std.Dev. |  | 6.273% | 4.841% | 8.818% | 1.484% | 4.01% | 9.021% | 8.851% | 3.419% | 0.5982% | 1.782% | 1.684% | 1.241% | 4.393% | 2.223% | 1.744% | 36.42% | 0.9153% | 2.022% | 2.611% | 3.191% | 1.713% | 1.681% | 3.244% | 1.718%
| PDM3 |  | 14.51% | -13.16% | -10.18% | 1.579% | -4.209% | 21.71% | -10.45% | -16.15% | 2.439% | 4.203% | -0.139% | -1.224% | -12.49% | 5.62% | 8.137% | -0.2571% | 18.9% | 3.931% | -1.903% | -8.144% | 6.889% | -3.124% | -24.26% | 18.75%
| Z-Score (Absolute) |  | 1.32 | 1.16 | 0.9 | 0.16 | 0.36 | 1.96 | 0.92 | 1.43 | 0.24 | 0.39 | 0.01 | 0.09 | 1.1 | 0.52 | 0.75 | 0.01 | 1.71 | 0.37 | 0.15 | 0.71 | 0.64 | 0.26 | 2.16 | 1.7
| Analytical results for Cs in OREAS 237b (Certified Value 9.42 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 9.62 | 8.1389 | 9.3 | 9.37 | 9.9 | 8.61 | 9.8 | 9.8 | 9.33 | 9.3 | 9.24 | 9.65 | 9.5 | 10.7 | 8.238 | 9 | 9.73 | 9.14 | 9.2 | 9.5523 | 9.2908 | [5 empty] | 1
| 1 | 2 | 9.57 | 8.2884 | 9.2 | 9.35 | 9.8 | 8.95 | 9.6 | 9.7 | 9.42 | 8.85 | 9.29 | 9.84 | 8.9 | 11.2 | 8.452 | 9 | 9.68 | 8.75 | 9 | 9.1378 | 9.4818 | [5 empty] | 34
| 1 | 3 | 9.81 | 8.239 | 9.2 | 9.68 | 9.8 | 8.69 | 9.5 | 9.4 | 9.21 | 9.29 | 9.51 | 10.1 | 9 | 10.7 | 9.24 | 10 | 9.68 | 8.72 | 9.2 | 9.6672 | 9.3915 | [5 empty] | 16
| 1 | 4 | 9.67 | 8.2739 | 9.2 | 9.85 | 9.7 | 9.58 | 9.5 | 8.7 | 9.15 | 9.19 | 9.15 | 9.57 | 12.3 | 10.6 | 8.549 | 10 | 9.62 | 8.76 | 9.6 | 9.3968 | 9.3028 | [5 empty] | 9.4201
| 1 | 5 | 9.52 | 8.2095 | 9.1 | 9.62 | 9.9 | 9.39 | 9.4 | 9.8 | 9.4 | 8.72 | 9.2 | 10 | 9.7 | 10.5 | 8.656 | 9 | 9.7 | 8.81 | 9.4 | 9.5156 | 9.275 | [5 empty] | 25
| 1 | 6 | 9.84 | 8.2995 | 9.3 | 9.48 | 9.8 | 9.6 | 9.6 | 9.2 | 9.52 | 9.59 | 9.25 | 9.63 | 10.3 | 10.8 | 8.476 | 9 | 9.69 | 9.66 | 9.8 | 9.4794 | 9.3634
| Mean |  | 9.6717 | 8.2415 | 9.2167 | 9.5583 | 9.8167 | 9.1367 | 9.5667 | 9.4333 | 9.3383 | 9.1567 | 9.2733 | 9.7983 | 9.95 | 10.75 | 8.6018 | 9.3333 | 9.6833 | 8.9733 | 9.3667 | 9.4582 | 9.3509
| Median |  | 9.645 | 8.2564 | 9.2 | 9.55 | 9.8 | 9.17 | 9.55 | 9.55 | 9.365 | 9.24 | 9.245 | 9.745 | 9.6 | 10.7 | 8.5125 | 9 | 9.685 | 8.785 | 9.3 | 9.4975 | 9.3331
| Std Dev. |  | 0.12922 | 0.060283 | 0.075277 | 0.19426 | 0.075277 | 0.44433 | 0.13663 | 0.43205 | 0.13819 | 0.31998 | 0.12533 | 0.2172 | 1.2582 | 0.2429 | 0.34166 | 0.5164 | 0.036148 | 0.37028 | 0.29439 | 0.18039 | 0.078264
| Rel.Std.Dev. |  | 1.336% | 0.7315% | 0.8168% | 2.032% | 0.7668% | 4.863% | 1.428% | 4.58% | 1.48% | 3.494% | 1.351% | 2.217% | 12.64% | 2.26% | 3.972% | 5.533% | 0.3733% | 4.126% | 3.143% | 1.907% | 0.837%
| PDM3 |  | 2.671% | -12.51% | -2.16% | 1.467% | 4.21% | -3.009% | 1.556% | 0.1405% | -0.868% | -2.797% | -1.558% | 4.015% | 5.625% | 14.12% | -8.686% | -0.9211% | 2.794% | -4.743% | -0.5673% | 0.4043% | -0.7347%
| Z-Score (Absolute) |  | 0.72 | 3.08 | 0.49 | 0.42 | 1.11 | 0.7 | 0.44 | 0.09 | 0.16 | 0.65 | 0.34 | 1.06 | 1.46 | 3.59 | 2.12 | NA | 0.75 | 1.14 | 0.09 | 0.15 | 0.13
Results from laboratory 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 237b (Certified Value 28.3 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 29.4 | 25.9204 | 28 | 26 | 27.07 | 31 | 26 | 25 | 31.9 | 27.8 | 28.7 | 29.3 | 27.9 | 27.2 | 30.2491 | 26.6 | 26.42 | 29.4 | 29.03 | 30.1 | 28.7 | 27.9 | 29.02 | 29 | 4 | 29.419 | 1
| 1 | 2 | 29.9 | 25.8192 | 28 | 26 | 29.12 | 30 | 26 | 25 | 32.5 | 27.9 | 26.8 | 28.8 | 28.2 | 26.9 | 30.1044 | 27.7 | 26.27 | 29.6 | 30.16 | 28.4 | 28.1 | 27.8 | 29.21 | 30 | 4 | 27.3179 | 35
| 1 | 3 | 29.8 | 25.8836 | 28 | 27 | 28.29 | 29 | 29 | 26 | 30.2 | 27.4 | 28.7 | 30 | 29.5 | 27.4 | 29.4698 | 27 | 27.38 | 29.2 | 30.2 | 29.3 | 27.5 | 27.9 | 30.2 | 29 | 8 | 27.2817 | 16
| 1 | 4 | 29.4 | 25.8468 | 28 | 26 | 27.77 | 30 | 27 | 26 | 28.4 | 27.2 | 27.2 | 29.4 | 27.7 | 27.9 | 29.3006 | 27.7 | 47.17 | 30.6 | 29.79 | 28.9 | 27.2 | 28.9 | 29.15 | 29 | 4 | 27.1992 | 28.3324
| 1 | 5 | 29.9 | 25.5984 | 26 | 26 | 28.1367 | 31 | 31 | 26 | 28.2 | 28.1 | 26.1 | 29.1 | 29 | 28.1 | 28.9186 | 26.2 | 27.74 | 30.1 | 29.79 | 28.9 | 27.2 | 28.2 | 30.16 | 30 | 7 | 28.3434 | 26
| 1 | 6 | 29.7 | 25.81 | 28 | 26 | 28.295 | 30 | 30 | 27 | 29.8 | 28.4 | 29.7 | 28.7 | 27.4 | 28.1 | 28.451 | 28.1 | 27.83 | 30.7 | 30.29 | 28.4 | 29.6 | 28.5 | 30.25 | 31 | 5 | 26.1255
| Mean |  | 29.6833 | 25.8131 | 27.6667 | 26.1667 | 28.1136 | 30.1667 | 28.1667 | 25.8333 | 30.1667 | 27.8 | 27.8667 | 29.2167 | 28.2833 | 27.6 | 29.4156 | 27.2167 | 30.4683 | 29.9333 | 29.8767 | 29 | 28.05 | 28.2 | 29.665 | 29.6667 | 5.3333 | 27.6144
| Median |  | 29.75 | 25.833 | 28 | 26 | 28.2133 | 30 | 28 | 26 | 30 | 27.85 | 27.95 | 29.2 | 28.05 | 27.65 | 29.3852 | 27.35 | 27.56 | 29.85 | 29.975 | 28.9 | 27.8 | 28.05 | 29.685 | 29.5 | 4.5 | 27.2998
| Std Dev. |  | 0.23166 | 0.11293 | 0.8165 | 0.40825 | 0.67595 | 0.75277 | 2.137 | 0.75277 | 1.7648 | 0.44272 | 1.375 | 0.47081 | 0.8085 | 0.50596 | 0.68742 | 0.73598 | 8.2086 | 0.6314 | 0.46638 | 0.63875 | 0.95656 | 0.42895 | 0.59359 | 0.8165 | 1.7512 | 1.1292
| Rel.Std.Dev. |  | 0.7804% | 0.4375% | 2.951% | 1.56% | 2.404% | 2.495% | 7.587% | 2.914% | 5.85% | 1.593% | 4.934% | 1.611% | 2.859% | 1.833% | 2.337% | 2.704% | 26.94% | 2.109% | 1.561% | 2.203% | 3.41% | 1.521% | 2.001% | 2.752% | 32.83% | 4.089%
| PDM3 |  | 4.768% | -8.892% | -2.35% | -7.644% | -0.7722% | 6.474% | -0.5849% | -8.82% | 6.474% | -1.879% | -1.644% | 3.121% | -0.1731% | -2.585% | 3.823% | -3.938% | 7.539% | 5.651% | 5.451% | 2.356% | -0.9967% | -0.4673% | 4.703% | 4.709% | -81.18% | -2.534%
| Z-Score (Absolute) |  | 0.89 | 1.41 | 0.31 | 1.2 | 0.04 | 1.18 | 0.01 | 1.4 | 1.18 | 0.23 | 0.19 | 0.62 | 0.06 | 0.35 | 0.73 | 0.58 | 1.36 | 1.04 | 1.01 | 0.49 | 0.08 | 0.01 | 0.88 | 0.88 | 13.6 | 0.34
| Analytical results for Dy in OREAS 237b (Certified Value 3.6 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 3.1534 | 3.85 | 3.7 | 3.1 | 4.11 | 2.96 | 3.99 | 3 | 2.67 | 4.9644 | [16 empty] | 1
| 1 | 2 | 3.0212 | 4.05 | 4 | 2.6 | 4.25 | 3.073 | 3.94 | 2.9 | 2.67 | 5.0628 | [16 empty] | 9
| 1 | 3 | 3.0229 | 3.75 | 4 | 2.9 | 4.21 | 3.408 | 4.07 | 2.9 | 2.7 | 4.9646 | [16 empty] | 16
| 1 | 4 | 3.0582 | 4 | 4 | 3.7 | 4.1 | 2.934 | 3.84 | 3 | 2.65 | 5.1477 | [16 empty] | 3.6014
| 1 | 5 | 3.1593 | 3.85 | 4 | 3.6 | 4.04 | 3.137 | 3.92 | 3.1 | 2.72 | 5.0054 | [16 empty] | 27
| 1 | 6 | 3.1699 | 3.85 | 4 | 3.2 | 4.24 | 3.049 | 3.84 | 3.2 | 2.72 | 4.943
| Mean |  | 3.0975 | 3.8917 | 3.95 | 3.1833 | 4.1583 | 3.0935 | 3.9333 | 3.0167 | 2.6883 | 5.0146
| Median |  | 3.1058 | 3.85 | 4 | 3.15 | 4.16 | 3.061 | 3.93 | 3 | 2.685 | 4.985
| Std Dev. |  | 0.070889 | 0.11143 | 0.12247 | 0.41673 | 0.086583 | 0.17115 | 0.088919 | 0.1169 | 0.029269 | 0.077807
| Rel.Std.Dev. |  | 2.289% | 2.863% | 3.101% | 13.09% | 2.082% | 5.533% | 2.261% | 3.875% | 1.089% | 1.552%
| PDM3 |  | -13.99% | 8.059% | 9.678% | -11.61% | 15.46% | -14.1% | 9.216% | -16.24% | -25.35% | 39.24%
| Z-Score (Absolute) |  | 0.67 | 0.54 | 0.63 | 0.54 | 0.95 | 0.68 | 0.6 | 0.79 | 1.3 | 2.25
| Analytical results for Er in OREAS 237b (Certified Value 2 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 1.7117 | 2.1 | 2.2 | 1.8 | 2.19 | 1.52 | 2.09 | 1.5 | 1.4316 | 2.914 | [16 empty] | 1
| 1 | 2 | 1.622 | 2 | 2.2 | 1.7 | 2.26 | 1.632 | 2.02 | 1.5 | 1.404 | 2.9899 | [16 empty] | 10
| 1 | 3 | 1.6084 | 2.05 | 2 | 1.5 | 2.25 | 2.092 | 2.16 | 1.5 | 1.4485 | 2.9008 | [16 empty] | 16
| 1 | 4 | 1.5181 | 2.1 | 2.1 | 2 | 2.23 | 1.804 | 2.2 | 1.5 | 1.3958 | 2.9528 | [16 empty] | 1.9965
| 1 | 5 | 1.5908 | 2 | 2.1 | 1.8 | 2.2 | 1.57 | 2 | 1.5 | 1.4068 | 2.8812 | [16 empty] | 28
| 1 | 6 | 1.5731 | 2.05 | 2.1 | 1.7 | 2.28 | 1.62 | 2.23 | 1.6 | 1.436 | 2.8804
| Mean |  | 1.604 | 2.05 | 2.1167 | 1.75 | 2.235 | 1.7063 | 2.1167 | 1.5167 | 1.4204 | 2.9198
| Median |  | 1.5996 | 2.05 | 2.1 | 1.75 | 2.24 | 1.626 | 2.125 | 1.5 | 1.4192 | 2.9074
| Std Dev. |  | 0.063956 | 0.044721 | 0.075277 | 0.16432 | 0.035071 | 0.21193 | 0.095219 | 0.040825 | 0.021059 | 0.043469
| Rel.Std.Dev. |  | 3.987% | 2.182% | 3.556% | 9.39% | 1.569% | 12.42% | 4.499% | 2.692% | 1.483% | 1.489%
| PDM3 |  | -19.66% | 2.682% | 6.021% | -12.34% | 11.95% | -14.53% | 6.021% | -24.03% | -28.85% | 46.25%
| Z-Score (Absolute) |  | 0.88 | 0 | NA | NA | 0.36 | 0.67 | 0.13 | NA | 1.24 | 1.71
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 237b (Certified Value 1.25 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 1.6018 | 1.25 | 1.18 | 1.1 | 1.33 | 1.274 | 1.28 | 1.2 | 1.1016 | 1.1872 | [16 empty] | 1
| 1 | 2 | 1.6644 | 1.25 | 1.26 | 1.1 | 1.37 | 1.3 | 1.22 | 1.1 | 1.042 | 1.2078 | [16 empty] | 11
| 1 | 3 | 1.6833 | 1.2 | 1.3 | 1 | 1.35 | 1.416 | 1.31 | 1.1 | 1.1055 | 1.1922 | [16 empty] | 16
| 1 | 4 | 1.612 | 1.3 | 1.32 | 1.4 | 1.33 | 1.297 | 1.29 | 1.2 | 1.0976 | 1.2151 | [16 empty] | 1.2457
| 1 | 5 | 1.6593 | 1.2 | 1.3 | 1.2 | 1.32 | 1.299 | 1.25 | 1.2 | 1.1002 | 1.2245 | [16 empty] | 29
| 1 | 6 | 1.6344 | 1.25 | 1.36 | 1.2 | 1.36 | 1.283 | 1.22 | 1.2 | 1.0782 | 1.1706
| Mean |  | 1.6425 | 1.2417 | 1.2867 | 1.1667 | 1.3433 | 1.3115 | 1.2617 | 1.1667 | 1.0875 | 1.1996
| Median |  | 1.6469 | 1.25 | 1.3 | 1.15 | 1.34 | 1.298 | 1.265 | 1.2 | 1.0989 | 1.2
| Std Dev. |  | 0.03188 | 0.037639 | 0.061536 | 0.13663 | 0.019664 | 0.052225 | 0.037639 | 0.05164 | 0.02426 | 0.019892
| Rel.Std.Dev. |  | 1.941% | 3.031% | 4.783% | 11.71% | 1.464% | 3.982% | 2.983% | 4.426% | 2.231% | 1.658%
| PDM3 |  | 31.85% | -0.3264% | 3.286% | -6.347% | 7.835% | 5.279% | 1.279% | -6.347% | -12.7% | -3.705%
| Z-Score (Absolute) |  | 4.66 | 0.41 | 0.16 | NA | 0.88 | 0.47 | 0.16 | NA | 2.36 | 0.94
Results from laboratories 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 237b (Certified Value 4.05 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 4.1 | 3.9634 | 3.96 | 3.66 | 4.25 | 4.42 | 4 | 4.03 | 3.88 | 3.91 | 4.15 | 4 | 3.85 | 4.0514 | 4.0213 | 3.938 | 3.9 | 4.05 | 4.14 | 4.12 | 4.06 | 3.9848 | 4.0162 | 4.069 |  |  | 1
| 1 | 2 | 4.11 | 3.9589 | 4.04 | 4.34 | 4.3067 | 4.38 | 4.17 | 4.11 | 3.88 | 3.96 | 3.96 | 4.04 | 3.9 | 3.979 | 4.1508 | 3.952 | 3.96 | 4.11 | 4.02 | 4.11 | 3.96 | 3.9773 | 4.1321 | 4.1136 |  |  | #N/A
| 1 | 3 | 4.11 | 4.0182 | 4.04 | 4.28 | 4.2567 | 4.35 | 4.03 | 4.11 | 3.89 | 3.95 | 4.07 | 4.08 | 3.93 | 4.0739 | 4.068 | 3.91 | 3.96 | 4.15 | 4.09 | 4.09 | 3.96 | 4.1108 | 4.0467 | 4.1151 |  |  | 16
| 1 | 4 | 4.09 | 4.003 | 3.97 | 4.19 | 4.3 | 4.42 | 4.06 | 4.06 | 3.87 | 3.95 | 4.08 | 4.1 | 3.86 | 4.0698 | 4.1382 | 3.91 | 4.1 | 4.11 | 4.1 | 4.1 | 4.11 | 3.9249 | 4.0748 | 4.0906 |  |  | 4.0492
| 1 | 5 | 4.14 | 3.9423 | 4.03 | 4.26 | 4.2967 | 4.32 | 4.1 | 4.18 | 3.86 | 3.95 | 4.01 | 3.94 | 3.85 | 4.082 | 4.0521 | 3.903 | 4.05 | 4.08 | 4.08 | 4.1 | 3.99 | 4.0265 | 4.0734 | 4.0582 |  |  | 30
| 1 | 6 | 4.1 | 3.9634 | 4.06 | 4.4 | 4.24 | 4.29 | 4.06 | 4.18 | 3.86 | 3.96 | 4.19 | 4.01 | 3.79 | 4.1269 | 4.1606 | 4.022 | 4.13 | 4.07 | 4.08 | 4.13 | 4.04 | 4.0702 | 3.9902 | 4.1819
| Mean |  | 4.1083 | 3.9748 | 4.0167 | 4.1883 | 4.275 | 4.3633 | 4.07 | 4.1117 | 3.8733 | 3.9467 | 4.0767 | 4.0283 | 3.8633 | 4.0638 | 4.0985 | 3.9392 | 4.0167 | 4.095 | 4.085 | 4.1083 | 4.02 | 4.0158 | 4.0556 | 4.1047
| Median |  | 4.105 | 3.9634 | 4.035 | 4.27 | 4.2767 | 4.365 | 4.06 | 4.11 | 3.875 | 3.95 | 4.075 | 4.025 | 3.855 | 4.0718 | 4.1031 | 3.924 | 4.005 | 4.095 | 4.085 | 4.105 | 4.015 | 4.0057 | 4.0601 | 4.1021
| Std Dev. |  | 0.017224 | 0.029153 | 0.041312 | 0.26851 | 0.029269 | 0.053166 | 0.05933 | 0.061128 | 0.012111 | 0.018619 | 0.085245 | 0.058109 | 0.048028 | 0.048571 | 0.05869 | 0.044804 | 0.09048 | 0.035637 | 0.038859 | 0.01472 | 0.060332 | 0.067535 | 0.049893 | 0.044233
| Rel.Std.Dev. |  | 0.4192% | 0.7334% | 1.029% | 6.411% | 0.6847% | 1.218% | 1.458% | 1.487% | 0.3127% | 0.4718% | 2.091% | 1.443% | 1.243% | 1.195% | 1.432% | 1.137% | 2.253% | 0.8703% | 0.9513% | 0.3583% | 1.501% | 1.682% | 1.23% | 1.078%
| PDM3 |  | 1.459% | -1.837% | -0.8045% | 3.435% | 5.575% | 7.757% | 0.5126% | 1.542% | -4.344% | -2.533% | 0.6773% | -0.5164% | -4.591% | 0.3603% | 1.216% | -2.718% | -0.8045% | 1.13% | 0.8831% | 1.459% | -0.7222% | -0.8271% | 0.1562% | 1.37%
| Z-Score (Absolute) |  | 0.61 | 1.35 | 0.74 | 1.79 | 3.06 | 4.36 | 0.05 | 0.66 | 2.85 | 1.77 | 0.14 | 0.57 | 3 | 0.05 | 0.46 | 1.88 | 0.74 | 0.41 | 0.27 | 0.61 | 0.69 | 0.75 | 0.17 | 0.56
| Analytical results for Ga in OREAS 237b (Certified Value 19.4 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 19.53 | 24.111 | 19 | 20.2 | 21.57 | 19.6 | 18.62 | 21 | 20 | 18.45 | 20.1 | 19.8 | 18.9 | 18.18 | 18.7702 | 17.2 | 19.1 | 18.63 | 19.24 | 17.98 | 20.8057 | 24 | 18.8064 |  |  |  | 1
| 1 | 2 | 19.42 | 23.5845 | 19.8 | 21 | 21.8133 | 19.3 | 19.92 | 21.1 | 20 | 18.85 | 19.35 | 19.45 | 19.5 | 17.98 | 18.5948 | 17.8 | 19.6 | 18.06 | 17.84 | 17.84 | 18.4198 | 22 | 18.7084 |  |  |  | 39
| 1 | 3 | 19.62 | 24.092 | 20.2 | 21.4 | 21.8133 | 19.6 | 19.6 | 21.1 | 20 | 18.5 | 20.8 | 20.4 | 20.2 | 17.63 | 18.3979 | 17.3 | 20.8 | 18.2 | 18.27 | 17.91 | 20.1358 | 24 | 19.0509 |  |  |  | 16
| 1 | 4 | 19.33 | 23.6075 | 20.2 | 21.7 | 21.54 | 19.7 | 20.78 | 21.4 | 19 | 18.4 | 20 | 19.9 | 19.2 | 18.01 | 18.4372 | 17.5 | 21.3 | 19.2 | 19.54 | 18.65 | 21.0422 | 25 | 18.5286 |  |  |  | 19.3985
| 1 | 5 | 19.68 | 23.522 | 19 | 20.9 | 21.91 | 19.6 | 20.6 | 20.5 | 20 | 18.55 | 19 | 19.8 | 19 | 18.24 | 18.3598 | 17.1 | 19.5 | 17.65 | 17.81 | 18.02 | 20.2446 | 24 | 18.808 |  |  |  | 31
| 1 | 6 | 19.25 | 23.6455 | 19.4 | 21.1 | 21.23 | 19.3 | 20.05 | 21.3 | 19 | 19 | 21.5 | 19.4 | 18.4 | 19.04 | 18.6311 | 17.5 | 18.3 | 18.37 | 20.97 | 18.28 | 19.9557 | 24 | 19.1261
| Mean |  | 19.4717 | 23.7604 | 19.6 | 21.05 | 21.6461 | 19.5167 | 19.9283 | 21.0667 | 19.6667 | 18.625 | 20.125 | 19.7917 | 19.2 | 18.18 | 18.5318 | 17.4 | 19.7667 | 18.3517 | 18.945 | 18.1133 | 20.1006 | 23.8333 | 18.8381
| Median |  | 19.475 | 23.6265 | 19.6 | 21.05 | 21.6917 | 19.6 | 19.985 | 21.1 | 20 | 18.525 | 20.05 | 19.8 | 19.1 | 18.095 | 18.516 | 17.4 | 19.55 | 18.285 | 18.755 | 18 | 20.1902 | 24 | 18.8072
| Std Dev. |  | 0.16774 | 0.2673 | 0.55136 | 0.50892 | 0.2513 | 0.17224 | 0.776 | 0.31411 | 0.5164 | 0.24238 | 0.91965 | 0.36113 | 0.60992 | 0.47214 | 0.15946 | 0.25298 | 1.1057 | 0.52875 | 1.2253 | 0.30277 | 0.92216 | 0.98319 | 0.22035
| Rel.Std.Dev. |  | 0.8615% | 1.125% | 2.813% | 2.418% | 1.161% | 0.8825% | 3.894% | 1.491% | 2.626% | 1.301% | 4.57% | 1.825% | 3.177% | 2.597% | 0.8605% | 1.454% | 5.594% | 2.881% | 6.468% | 1.672% | 4.588% | 4.125% | 1.17%
| PDM3 |  | 0.3774% | 22.49% | 1.039% | 8.514% | 11.59% | 0.6094% | 2.732% | 8.6% | 1.383% | -3.987% | 3.745% | 2.027% | -1.023% | -6.281% | -4.467% | -10.3% | 1.898% | -5.396% | -2.338% | -6.625% | 3.62% | 22.86% | -2.889%
| Z-Score (Absolute) |  | 0.04 | 4.22 | 0.08 | 1.52 | 2.12 | 0 | 0.41 | 1.54 | NA | 0.89 | 0.6 | 0.27 | 0.31 | 1.33 | 0.98 | 2.1 | 0.25 | 1.16 | 0.57 | 1.39 | 0.58 | NA | 0.67
Results from laboratories 1.11 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 237b (Certified Value 5.06 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.7383 | 5.4 | 5.2 | 4.8 | 5 | 5.111 | 5.15 | 4.5 | 4.9646 | 0.82877 | [16 empty] | 1
| 1 | 2 | 6.7423 | 5.4 | 5.4 | 4.3 | 5.2 | 5.248 | 4.95 | 4.4 | 4.6592 | 0.83533 | [16 empty] | 12
| 1 | 3 | 6.7679 | 5.2 | 5.5 | 4.7 | 5.2 | 5.975 | 5.19 | 4.4 | 4.9719 | 0.82111 | [16 empty] | 16
| 1 | 4 | 6.7414 | 5.2 | 5.4 | 5.5 | 5 | 5.165 | 5.29 | 4.6 | 4.8719 | 0.84542 | [16 empty] | 5.0623
| 1 | 5 | 6.6473 | 5.2 | 5.3 | 5.3 | 5 | 5.261 | 5.18 | 4.5 | 4.9346 | 0.8265 | [16 empty] | 32
| 1 | 6 | 6.7215 | 5.4 | 5.6 | 4.9 | 5.2 | 5.143 | 5.21 | 4.7 | 4.8492 | 0.81499
| Mean |  | 6.7265 | 5.3 | 5.4 | 4.9167 | 5.1 | 5.3172 | 5.1617 | 4.5167 | 4.8752 | 0.82869
| Median |  | 6.7398 | 5.3 | 5.4 | 4.85 | 5.1 | 5.2065 | 5.185 | 4.5 | 4.9032 | 0.82764
| Std Dev. |  | 0.041505 | 0.10954 | 0.14142 | 0.43089 | 0.10954 | 0.32762 | 0.11392 | 0.1169 | 0.1168 | 0.010714
| Rel.Std.Dev. |  | 0.617% | 2.067% | 2.619% | 8.764% | 2.148% | 6.162% | 2.207% | 2.588% | 2.396% | 1.293%
| PDM3 |  | 32.87% | 4.696% | 6.671% | -2.876% | 0.7453% | 5.035% | 1.963% | -10.78% | -3.695% | -83.63%
| Z-Score (Absolute) |  | 4.58 | 0.49 | 0.77 | 0.61 | 0.09 | 0.53 | 0.09 | 1.76 | 0.73 | 12.35
| Analytical results for Ge in OREAS 237b (Indicative Value 0.21 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 1.6 | 0.2 | 1.2 | 0.18 | 1.5 | 3.9 | 0.17 | 0.25 | 0.07 | 0.18 | 0.37 | < 0.05 | 0.12 | 0.229 | [12 empty] | 1
| 1 | 2 | 1.7 | 0.2 | 1.2 | 0.18 | 1.4 | 4.2 | 0.17 | 0.24 | 0.09 | 0.17 | 0.47 | < 0.05 | 0.12 | 0.2409 | [12 empty] | 6
| 1 | 3 | 1.8 | 0.3 | 1.2 | 0.18 | 1.7 | 3.9 | 0.17 | 0.25 | 0.07 | 0.19 | 0.39 | < 0.05 | 0.11 | 0.197 | [12 empty] | 16
| 1 | 4 | 1.7 | 0.3 | 1.2 | 0.19 | 2 | 3.9 | 0.15 | 0.24 | 0.09 | 0.18 | 0.69 | < 0.05 | 0.12 | 0.2397 | [12 empty] | 0.21247
| 1 | 5 | 1.6 | 0.3 | 1.3 | 0.19 | 1.7 | 4.4 | 0.17 | 0.2 | 0.08 | 0.22 | 0.39 | < 0.05 | 0.12 | 0.1435 | [12 empty] | 12
| 1 | 6 | 1.8 | 0.3 | 1.3 | 0.2 | 2 | 4.1 | 0.17 | 0.25 | 0.08 | 0.19 | 0.54 | < 0.05 | 0.15 | 0.2327
| Mean |  | 1.7 | 0.26667 | 1.2333 | 0.18667 | 1.7167 | 4.0667 | 0.16667 | 0.23833 | 0.08 | 0.18833 | 0.475 |  | 0.12333 | 0.2138
| Median |  | 1.7 | 0.3 | 1.2 | 0.185 | 1.7 | 4 | 0.17 | 0.245 | 0.08 | 0.185 | 0.43 |  | 0.12 | 0.23085
| Std Dev. |  | 0.089443 | 0.05164 | 0.05164 | 0.008165 | 0.24833 | 0.20656 | 0.008165 | 0.019408 | 0.0089443 | 0.017224 | 0.12325 |  | 0.013663 | 0.037994
| Rel.Std.Dev. |  | 5.261% | 19.36% | 4.187% | 4.374% | 14.47% | 5.079% | 4.899% | 8.143% | 11.18% | 9.145% | 25.95% |  | 11.08% | 17.77%
| PDM3 |  | 700.1% | 25.51% | 480.5% | -12.14% | 708% | 1814% | -21.56% | 12.17% | -62.35% | -11.36% | 123.6% |  | -41.95% | 0.6279%
| Z-Score (Absolute) |  | 7.99 | 0.22 | 5.46 | 0.21 | 8.08 | 20.82 | 0.32 | 0.07 | 0.79 | 0.2 | 1.35 | 1.09 | 0.56 | 0.07
| Analytical results for Hf in OREAS 237b (Certified Value 3.94 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4.08 | 4.0203 | 7.2 | 3.2 | 3.9 | 3.5 | 3.8 | 4 | 3.8 | 4 | 4 | 3.8 | 3.93 | 4.4 | 3.71 | 4.45 | 3.9 | 3.44 | 3.2971 | 2.5966 | 6.2124 | [5 empty] | 1
| 1 | 2 | 3.97 | 4.0991 | 8.2 | 3.3 | 3.8 | 3.7 | 3.7 | 3.9 | 4 | 3.7 | 3.8 | 3.7 | 3.59 | 4.5 | 3.73 | 4.51 | 3.8 | 3.35 | 3.1822 | 2.6216 | 6.229 | [5 empty] | 25
| 1 | 3 | 4.05 | 3.9952 | 7.4 | 4.1 | 3.7 | 3.8 | 4.3 | 4.1 | 4 | 3.9 | 4 | 3.9 | 3.92 | 4.4 | 3.7 | 4.39 | 3.7 | 3.26 | 3.309 | 2.6548 | 6.2422 | [5 empty] | 16
| 1 | 4 | 3.92 | 3.9365 | 7.4 | 4.3 | 3.9 | 3.9 | 4.2 | 3.9 | 3.8 | 3.8 | 3.9 | 3.8 | 4.12 | 4.4 | 3.8 | 4.25 | 3.8 | 3.45 | 3.2564 | 2.6049 | 6.2031 | [5 empty] | 3.9364
| 1 | 5 | 4.04 | 3.9164 | 7.8 | 4.4 | 4 | 3.8 | 3.8 | 4 | 4 | 3.8 | 3.9 | 4.1 | 3.98 | 4.4 | 3.9 | 4.18 | 3.8 | 3.44 | 3.2757 | 2.5966 | 6.2351 | [5 empty] | 33
| 1 | 6 | 4.08 | 3.9607 | 7.2 | 4.6 | 4 | 4 | 4.2 | 3.9 | 4 | 4 | 3.8 | 3.7 | 3.8 | 4.4 | 3.7 | 4.08 | 4.2 | 3.45 | 3.2093 | 2.6964 | 6.3367
| Mean |  | 4.0233 | 3.988 | 7.5333 | 3.9833 | 3.8833 | 3.7833 | 4 | 3.9667 | 3.9333 | 3.8667 | 3.9 | 3.8333 | 3.89 | 4.4167 | 3.7567 | 4.31 | 3.8667 | 3.3983 | 3.2549 | 2.6285 | 6.2431
| Median |  | 4.045 | 3.978 | 7.4 | 4.2 | 3.9 | 3.8 | 4 | 3.95 | 4 | 3.85 | 3.9 | 3.8 | 3.925 | 4.4 | 3.72 | 4.32 | 3.8 | 3.44 | 3.266 | 2.6132 | 6.232
| Std Dev. |  | 0.064704 | 0.066305 | 0.39328 | 0.59133 | 0.1169 | 0.17224 | 0.26077 | 0.08165 | 0.10328 | 0.12111 | 0.089443 | 0.15055 | 0.17978 | 0.040825 | 0.079666 | 0.16697 | 0.17512 | 0.077825 | 0.050014 | 0.039854 | 0.048121
| Rel.Std.Dev. |  | 1.608% | 1.663% | 5.22% | 14.85% | 3.01% | 4.553% | 6.519% | 2.058% | 2.626% | 3.132% | 2.293% | 3.928% | 4.622% | 0.9243% | 2.121% | 3.874% | 4.529% | 2.29% | 1.537% | 1.516% | 0.7708%
| PDM3 |  | 2.208% | 1.312% | 91.38% | 1.192% | -1.348% | -3.889% | 1.616% | 0.7688% | -0.07797% | -1.772% | -0.9248% | -2.618% | -1.179% | 12.2% | -4.566% | 9.491% | -1.772% | -13.67% | -17.31% | -33.23% | 58.6%
| Z-Score (Absolute) |  | 0.83 | 0.59 | 24.5 | 0.56 | 0.11 | 0.79 | 0.67 | 0.45 | 0.22 | 0.22 | 0 | 0.45 | 0.07 | 3.48 | 0.97 | 2.76 | 0.22 | 3.38 | 4.35 | 8.57 | 15.8
| Analytical results for Hg in OREAS 237b (Indicative Value 0.025 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 1.09 | 1.30 | [24 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [24 empty] | 3
3
| 1 | 1 | < 2 | 0.0265 | [24 empty] | 1
| 1 | 2 | < 2 | 0.0227 | [24 empty] | 13
| 1 | 3 | < 2 | 0.0238 | [24 empty] | 16
| 1 | 4 | < 2 | 0.0278 | [24 empty] | 0.0247
| 1 | 5 | < 2 | 0.0218 | [24 empty] | 9
| 1 | 6 | < 2 | 0.0256
| Mean |  |  | 0.0247
| Median |  |  | 0.0247
| Std Dev. |  |  | 0.0023186
| Rel.Std.Dev. |  |  | 9.387%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 237b (Certified Value 0.68 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.55294 | 0.72 | 0.8 | 0.6 | 0.71 | 0.533 | 0.76 | 0.5 | 0.4988 | 0.99859 | [16 empty] | 1
| 1 | 2 | 0.58956 | 0.7 | 0.8 | 0.5 | 0.72 | 0.559 | 0.99 | 0.5 | 0.4642 | 1.0226 | [16 empty] | 14
| 1 | 3 | 0.55602 | 0.72 | 0.8 | 0.6 | 0.74 | 0.73 | 0.76 | 0.5 | 0.5029 | 1.0011 | [16 empty] | 16
| 1 | 4 | 0.57385 | 0.74 | 0.8 | 0.7 | 0.73 | 0.525 | 0.7 | 0.5 | 0.4926 | 1.0106 | [16 empty] | 0.68476
| 1 | 5 | 0.56966 | 0.7 | 0.8 | 0.6 | 0.7 | 0.554 | 0.78 | 0.5 | 0.4932 | 0.99333 | [16 empty] | 34
| 1 | 6 | 0.55733 | 0.72 | 0.8 | 0.7 | 0.78 | 0.538 | 0.72 | 0.6 | 0.5009 | 0.98686
| Mean |  | 0.56656 | 0.71667 | 0.8 | 0.61667 | 0.73 | 0.57317 | 0.785 | 0.51667 | 0.4921 | 1.0022
| Median |  | 0.5635 | 0.72 | 0.8 | 0.6 | 0.725 | 0.546 | 0.76 | 0.5 | 0.496 | 0.99986
| Std Dev. |  | 0.013954 | 0.015055 | 1.2162e-16 | 0.075277 | 0.028284 | 0.077891 | 0.10464 | 0.040825 | 0.014275 | 0.012748
| Rel.Std.Dev. |  | 2.463% | 2.101% | 1.52e-14% | 12.21% | 3.875% | 13.59% | 13.33% | 7.902% | 2.901% | 1.272%
| PDM3 |  | -17.26% | 4.66% | 16.83% | -9.944% | 6.607% | -16.3% | 14.64% | -24.55% | -28.14% | 46.35%
| Z-Score (Absolute) |  | 0.71 | 0 | NA | NA | 0.06 | 0.67 | 0.32 | NA | 1.06 | 1.34
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 237b (Certified Value 0.066 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 0.07 | 0.05 | < 0.1 | < 2 | 0.07 | 0.044 | 0.1 | 0.07 | 0.07 | 0.065 | 0.068 | 0.065 | 0.07 | 0.07 | 0.07 | 0.061 | 0.06 | 0.0825 | 0.06462 | [7 empty] | 1
| 1 | 2 | 0.08 | < 0.05 | < 0.1 | < 2 | 0.07 | 0.046 | 0.08 | 0.07 | 0.073 | 0.063 | 0.069 | 0.065 | 0.06 | 0.07 | 0.062 | 0.057 | 0.06 | 0.0756 | 0.06757 | [7 empty] | 27
| 1 | 3 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.044 | 0.1 | 0.06 | 0.065 | 0.072 | 0.071 | 0.066 | 0.06 | 0.07 | 0.073 | 0.059 | 0.06 | 0.0822 | 0.06488 | [7 empty] | 16
| 1 | 4 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.046 | 0.1 | 0.07 | 0.07 | 0.069 | 0.066 | 0.064 | 0.07 | 0.07 | 0.091 | 0.062 | 0.06 | 0.0753 | 0.06472 | [7 empty] | 0.065658
| 1 | 5 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.038 | 0.11 | 0.05 | 0.073 | 0.056 | 0.067 | 0.065 | 0.07 | 0.07 | 0.079 | 0.057 | 0.06 | 0.0768 | 0.06457 | [7 empty] | 35
| 1 | 6 | 0.06 | 0.05 | < 0.1 | < 2 | 0.06 | 0.046 | 0.12 | 0.07 | 0.069 | 0.066 | 0.062 | 0.062 | 0.07 | 0.06 | 0.072 | 0.066 | 0.06 | 0.0742 | 0.06574
| Mean |  | 0.07 | 0.05 |  |  | 0.063333 | 0.044 | 0.10167 | 0.065 | 0.07 | 0.065167 | 0.067167 | 0.0645 | 0.066667 | 0.068333 | 0.0745 | 0.060333 | 0.06 | 0.077767 | 0.06535
| Median |  | 0.07 | 0.05 |  |  | 0.06 | 0.045 | 0.1 | 0.07 | 0.07 | 0.0655 | 0.0675 | 0.065 | 0.07 | 0.07 | 0.0725 | 0.06 | 0.06 | 0.0762 | 0.0648
| Std Dev. |  | 0.0063246 | 0 |  |  | 0.005164 | 0.0030984 | 0.013292 | 0.0083666 | 0.0029665 | 0.0054924 | 0.0030605 | 0.0013784 | 0.005164 | 0.0040825 | 0.0097724 | 0.0034448 | 0 | 0.0036467 | 0.0011696
| Rel.Std.Dev. |  | 9.035% | 0% |  |  | 8.154% | 7.042% | 13.07% | 12.87% | 4.238% | 8.428% | 4.557% | 2.137% | 7.746% | 5.974% | 13.12% | 5.71% | 0% | 4.689% | 1.79%
| PDM3 |  | 6.613% | -23.85% |  |  | -3.541% | -32.99% | 54.84% | -1.002% | 6.613% | -0.7486% | 2.297% | -1.764% | 1.536% | 4.074% | 13.47% | -8.11% | -8.618% | 18.44% | -0.4694%
| Z-Score (Absolute) |  | 0.67 | 4.64 | 2.23 | 135.54 | 0.29 | 3.1 | 5.27 | 0.05 | 0.67 | 0.03 | 0.26 | 0.12 | 0.19 | 0.43 | 1.33 | 0.73 | 0.78 | 1.8 | 0
| Analytical results for Ir in OREAS 237b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Ir | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0034 | [25 empty] | 1
| 1 | 2 | 0.0003 | [25 empty] | 4
| 1 | 3 | 0.0004 | [25 empty] | 16
| 1 | 4 | 0.0012 | [25 empty] | 0.00115
| 1 | 5 | 0.001 | [25 empty] | 10
| 1 | 6 | 0.0006
| Mean |  | 0.00115
| Median |  | 0.0008
| Std Dev. |  | 0.0011554
| Rel.Std.Dev. |  | 100.5%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 237b (Certified Value 2.65 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 2.7712 | 2.7065 | 2.55 | 1.77 | 2.63 | 2.757 | 2.51 | 2.66 | 2.68 | 2.58 | 2.68 | 2.58 | 2.5 | 2.6969 | 2.6205 | 2.6399 | 2.61 | 2.51 | 2.42 | 2.68 | 2.68 | 2.5844 | 2.9024 | 3.1534 |  |  | 1
| 1 | 2 | 2.8339 | 2.7051 | 2.54 | 3.4 | 2.615 | 2.724 | 2.56 | 2.61 | 2.66 | 2.62 | 2.59 | 2.65 | 2.56 | 2.6673 | 2.6416 | 2.6565 | 2.61 | 2.53 | 2.33 | 2.67 | 2.65 | 2.5977 | 2.9693 | 3.1518 |  |  | #N/A
| 1 | 3 | 2.7716 | 2.7446 | 2.57 | 3.29 | 2.615 | 2.709 | 2.44 | 2.71 | 2.66 | 2.62 | 2.65 | 2.7 | 2.6 | 2.5775 | 2.6486 | 2.6067 | 2.65 | 2.56 | 2.29 | 2.67 | 2.64 | 2.6878 | 2.9069 | 3.1518 |  |  | 16
| 1 | 4 | 2.8039 | 2.6902 | 2.51 | 3.37 | 2.63 | 2.748 | 2.56 | 2.68 | 2.64 | 2.58 | 2.66 | 2.69 | 2.57 | 2.7846 | 2.6243 | 2.6233 | 2.77 | 2.55 | 2.58 | 2.67 | 2.75 | 2.5718 | 2.9056 | 3.1552 |  |  | 2.6454
| 1 | 5 | 2.8521 | 2.7179 | 2.6 | 3.38 | 2.59 | 2.691 | 2.56 | 2.75 | 2.64 | 2.6 | 2.6 | 2.56 | 2.54 | 2.7805 | 2.6053 | 2.5984 | 2.65 | 2.55 | 2.63 | 2.68 | 2.65 | 2.6293 | 2.9209 | 3.1537 |  |  | 36
| 1 | 6 | 2.7667 | 2.6921 | 2.66 | 3.41 | 2.64 | 2.671 | 2.6 | 2.75 | 2.65 | 2.59 | 2.72 | 2.61 | 2.54 | 2.8601 | 2.6062 | 2.698 | 2.72 | 2.54 | 2.4 | 2.7 | 2.67 | 2.6364 | 2.8373 | 3.1526
| Mean |  | 2.7999 | 2.7094 | 2.5717 | 3.1033 | 2.62 | 2.7167 | 2.5383 | 2.6933 | 2.655 | 2.5983 | 2.65 | 2.6317 | 2.5517 | 2.7278 | 2.6244 | 2.6371 | 2.6683 | 2.54 | 2.4417 | 2.6783 | 2.6733 | 2.6179 | 2.9071 | 3.1531
| Median |  | 2.7877 | 2.7058 | 2.56 | 3.375 | 2.6225 | 2.7165 | 2.56 | 2.695 | 2.655 | 2.595 | 2.655 | 2.63 | 2.55 | 2.7387 | 2.6224 | 2.6316 | 2.65 | 2.545 | 2.41 | 2.675 | 2.66 | 2.6135 | 2.9062 | 3.153
| Std Dev. |  | 0.036397 | 0.020026 | 0.052694 | 0.65457 | 0.017607 | 0.033062 | 0.056006 | 0.05465 | 0.015166 | 0.018348 | 0.04899 | 0.057764 | 0.033714 | 0.10071 | 0.017849 | 0.03662 | 0.064005 | 0.017889 | 0.13586 | 0.01169 | 0.040332 | 0.042459 | 0.042322 | 0.0013029
| Rel.Std.Dev. |  | 1.3% | 0.7391% | 2.049% | 21.09% | 0.672% | 1.217% | 2.206% | 2.029% | 0.5712% | 0.7062% | 1.849% | 2.195% | 1.321% | 3.692% | 0.6801% | 1.389% | 2.399% | 0.7043% | 5.564% | 0.4365% | 1.509% | 1.622% | 1.456% | 0.04132%
| PDM3 |  | 5.841% | 2.421% | -2.786% | 17.31% | -0.9593% | 2.695% | -4.046% | 1.813% | 0.3637% | -1.778% | 0.1747% | -0.5183% | -3.542% | 3.116% | -0.7929% | -0.3117% | 0.8678% | -3.983% | -7.701% | 1.246% | 1.057% | -1.039% | 9.892% | 19.19%
| Z-Score (Absolute) |  | 1.79 | 0.69 | 0.98 | 5.47 | 0.39 | 0.78 | 1.39 | 0.5 | 0.03 | 0.66 | 0.03 | 0.25 | 1.22 | 0.91 | 0.34 | 0.19 | 0.19 | 1.37 | 2.56 | 0.31 | 0.25 | 0.42 | 3.09 | 6.08
| Analytical results for La in OREAS 237b (Certified Value 38.6 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 40.58 | 44.3864 | 39.1 | 42.2 | 34.35 | 36.6 | 40 | 40.3 | 41 | 40.3 | 39.2 | 40.2 | 35.7851 | 38 | 37.31 | 36.3 | 35.47 | 38.8 | 38.5 | 39.6 | 33.14 | 35.5857 |  |  |  |  | 1
| 1 | 2 | 39.92 | 44.4203 | 38.4 | 42.2 | 34.21 | 37.9 | 39 | 40.6 | 39.3 | 38.2 | 40.1 | 38.3 | 36.8008 | 39 | 36.85 | 38.2 | 34.48 | 38.1 | 36.9 | 38.1 | 33.08 | 36.5708 |  |  |  |  | 28
| 1 | 3 | 40.64 | 44.4994 | 39 | 42 | 34.15 | 38.5 | 39 | 40.4 | 40.1 | 40.8 | 41.4 | 39.9 | 35.7458 | 39 | 40.57 | 37.7 | 35.11 | 39.2 | 37.3 | 38.3 | 34.44 | 36.1897 |  |  |  |  | 16
| 1 | 4 | 40.19 | 44.1039 | 38.2 | 42.9 | 36.49 | 41.2 | 39 | 39.7 | 41 | 40.3 | 39.2 | 39.7 | 36.9762 | 39 | 37.04 | 39.5 | 35.74 | 38.8 | 37.3 | 39.7 | 33.22 | 36.7648 |  |  |  |  | 38.6398
| 1 | 5 | 39.57 | 44.1717 | 41 | 42.5 | 36.48 | 40.6 | 39 | 40.6 | 38.1 | 39.9 | 40.5 | 41.2 | 36.1502 | 39 | 36.82 | 38.4 | 35.56 | 38 | 37.1 | 39 | 33.84 | 37.3181 |  |  |  |  | 37
| 1 | 6 | 40.78 | 44.0587 | 40.5 | 44.6 | 34.93 | 41.5 | 39 | 41.8 | 42.3 | 39.3 | 38.4 | 40.3 | 35.8315 | 40 | 36.48 | 36.3 | 35.88 | 38.4 | 40.9 | 40.3 | 33.96 | 35.8441
| Mean |  | 40.28 | 44.2734 | 39.3667 | 42.7333 | 35.1017 | 39.3833 | 39.1667 | 40.5667 | 40.3 | 39.8 | 39.8 | 39.9333 | 36.2149 | 39 | 37.5117 | 37.7333 | 35.3733 | 38.55 | 38 | 39.1667 | 33.6133 | 36.3789
| Median |  | 40.385 | 44.279 | 39.05 | 42.35 | 34.64 | 39.55 | 39 | 40.5 | 40.55 | 40.1 | 39.65 | 40.05 | 35.9908 | 39 | 36.945 | 37.95 | 35.515 | 38.6 | 37.3 | 39.3 | 33.53 | 36.3803
| Std Dev. |  | 0.47147 | 0.18471 | 1.1361 | 0.96678 | 1.1066 | 1.9994 | 0.40825 | 0.68896 | 1.4738 | 0.92952 | 1.0789 | 0.95219 | 0.54383 | 0.63246 | 1.523 | 1.2565 | 0.51044 | 0.46368 | 1.5271 | 0.85713 | 0.55102 | 0.63543
| Rel.Std.Dev. |  | 1.17% | 0.4172% | 2.886% | 2.262% | 3.152% | 5.077% | 1.042% | 1.698% | 3.657% | 2.335% | 2.711% | 2.384% | 1.502% | 1.622% | 4.06% | 3.33% | 1.443% | 1.203% | 4.019% | 2.188% | 1.639% | 1.747%
| PDM3 |  | 4.245% | 14.58% | 1.881% | 10.59% | -9.157% | 1.924% | 1.364% | 4.987% | 4.297% | 3.003% | 3.003% | 3.348% | -6.276% | 0.9323% | -2.92% | -2.346% | -8.454% | -0.2323% | -1.656% | 1.364% | -13.01% | -5.851%
| Z-Score (Absolute) |  | 0.65 | 2.99 | 0.12 | 2.09 | 2.38 | 0.13 | 0 | 0.82 | 0.66 | 0.37 | 0.37 | 0.45 | 1.73 | 0.1 | 0.97 | 0.84 | 2.22 | 0.36 | 0.68 | 0 | 3.26 | 1.63
| Analytical results for Li in OREAS 237b (Certified Value 52 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 52.5 | 56.392 | 50 | 56 | 49.56 | 49.1 | 54 | 48.1 | 75.2 | 55 | 49 | 56.7 | 46.8 | 54.7 | 53.5976 | 53 | 43.48 | 49.3 | 55.16 | 53 | 53.1 | 53.6 | 56.1 | 45.2015 |  |  | 1
| 1 | 2 | 53.5 | 56.6049 | 50 | 56 | 52.2833 | 51.3 | 55 | 49 | 71.1 | 54 | 48.8 | 55.9 | 46.9 | 53.8 | 53.6521 | 54 | 43.83 | 53 | 56.04 | 51.6 | 49.7 | 53 | 56.6 | 46.4932 |  |  | #N/A
| 1 | 3 | 53.6 | 56.3849 | 50 | 56 | 51.58 | 49.3 | 52 | 51.6 | 69.6 | 53.1 | 48.7 | 58.9 | 49.7 | 55.1 | 52.3577 | 53 | 45.15 | 54.1 | 55.04 | 51 | 50.1 | 53.2 | 58.6 | 45.3483 |  |  | 16
| 1 | 4 | 53.5 | 56.4566 | 60 | 57 | 52.1033 | 49.8 | 55 | 49.5 | 66.6 | 52.8 | 47.7 | 56.8 | 47 | 55.3 | 54.4372 | 53 | 75.59 | 56.4 | 56.28 | 51.6 | 50.5 | 54.4 | 55.9 | 44.1289 |  |  | 52.3763
| 1 | 5 | 54 | 56.9826 | 50 | 57 | 50.3533 | 51.4 | 54 | 49.3 | 84.6 | 53.5 | 45 | 56.6 | 49.7 | 56 | 53.3314 | 53 | 44.28 | 54.8 | 55.66 | 49.4 | 49.7 | 54.1 | 57.4 | 45.7066 |  |  | 38
| 1 | 6 | 54.2 | 56.0978 | 60 | 56 | 49.8767 | 50 | 55 | 51 | 69.8 | 55.2 | 48.9 | 55 | 46.9 | 56.5 | 53.7924 | 55 | 44.5 | 53.1 | 53.64 | 50.9 | 50.8 | 53.9 | 58 | 45.6148
| Mean |  | 53.55 | 56.4865 | 53.3333 | 56.3333 | 50.9594 | 50.15 | 54.1667 | 49.75 | 72.8167 | 53.9333 | 48.0167 | 56.65 | 47.8333 | 55.2333 | 53.5281 | 53.5 | 49.4717 | 53.45 | 55.3033 | 51.25 | 50.65 | 53.7 | 57.1 | 45.4155
| Median |  | 53.55 | 56.4243 | 50 | 56 | 50.9667 | 49.9 | 54.5 | 49.4 | 70.45 | 53.75 | 48.75 | 56.65 | 46.95 | 55.2 | 53.6249 | 53 | 44.39 | 53.6 | 55.41 | 51.3 | 50.3 | 53.75 | 57 | 45.4815
| Std Dev. |  | 0.58907 | 0.29372 | 5.164 | 0.5164 | 1.1785 | 0.98539 | 1.169 | 1.3065 | 6.412 | 0.9913 | 1.551 | 1.2942 | 1.4473 | 0.95429 | 0.68192 | 0.83666 | 12.8081 | 2.3873 | 0.94661 | 1.176 | 1.2771 | 0.53666 | 1.0807 | 0.77352
| Rel.Std.Dev. |  | 1.1% | 0.52% | 9.682% | 0.9167% | 2.313% | 1.965% | 2.158% | 2.626% | 8.806% | 1.838% | 3.23% | 2.285% | 3.026% | 1.728% | 1.274% | 1.564% | 25.89% | 4.466% | 1.712% | 2.295% | 2.521% | 0.9994% | 1.893% | 1.703%
| PDM3 |  | 2.241% | 7.847% | 1.827% | 7.555% | -2.705% | -4.251% | 3.418% | -5.014% | 39.03% | 2.973% | -8.324% | 8.16% | -8.674% | 5.455% | 2.199% | 2.145% | -5.546% | 2.05% | 5.589% | -2.15% | -3.296% | 2.527% | 9.019% | -13.29%
| Z-Score (Absolute) |  | 0.01 | 0.75 | 0.05 | 0.71 | 0.64 | 0.84 | 0.16 | 0.94 | 4.84 | 0.11 | 1.38 | 0.79 | 1.43 | 0.43 | 0 | 0 | 1.01 | 0.02 | 0.45 | 0.57 | 0.72 | 0.05 | 0.9 | 2.03
| Analytical results for Lu in OREAS 237b (Certified Value 0.28 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.26651 | 0.32 | 0.3 | 0.3 | 0.234 | 0.29 | 0.34 | 0.2 | 0.1601 | 0.40866 | [16 empty] | 1
| 1 | 2 | 0.26784 | 0.3 | 0.3 | 0.2 | 0.243 | 0.27 | 0.79 | 0.2 | 0.1387 | 0.4165 | [16 empty] | 29
| 1 | 3 | 0.27091 | 0.32 | 0.3 | 0.3 | 0.269 | 0.27 | 0.34 | 0.2 | 0.1511 | 0.40577 | [16 empty] | 16
| 1 | 4 | 0.2808 | 0.34 | 0.3 | 0.3 | 0.238 | 0.28 | 0.33 | 0.2 | 0.1237 | 0.40829 | [16 empty] | 0.27762
| 1 | 5 | 0.27416 | 0.3 | 0.3 | 0.3 | 0.246 | 0.27 | 0.32 | 0.2 | 0.1152 | 0.40132 | [16 empty] | 39
| 1 | 6 | 0.27368 | 0.32 | 0.4 | 0.3 | 0.235 | 0.27 | 0.31 | 0.2 | 0.1374 | 0.40321
| Mean |  | 0.27232 | 0.31667 | 0.31667 | 0.28333 | 0.24417 | 0.275 | 0.405 | 0.2 | 0.1377 | 0.40729
| Median |  | 0.27229 | 0.32 | 0.3 | 0.3 | 0.2405 | 0.27 | 0.335 | 0.2 | 0.13805 | 0.40703
| Std Dev. |  | 0.0051562 | 0.015055 | 0.040825 | 0.040825 | 0.013014 | 0.0083666 | 0.18897 | 3.0405e-17 | 0.016641 | 0.0053348
| Rel.Std.Dev. |  | 1.893% | 4.754% | 12.89% | 14.41% | 5.33% | 3.042% | 46.66% | 1.52e-14% | 12.09% | 1.31%
| PDM3 |  | -1.91% | 14.07% | 14.07% | 2.059% | -12.05% | -0.9428% | 45.88% | -27.96% | -50.4% | 46.71%
| Z-Score (Absolute) |  | 0.12 | 0.67 | 0.67 | 0.07 | 0.63 | 0.07 | 2.26 | 1.42 | 2.54 | 2.3
| Analytical results for Mg in OREAS 237b (Certified Value 1.7 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 1.6578 | 1.7475 | 1.64 | 1.49 | 1.69 | 1.777 | 1.62 | 1.72 | 1.67 | 1.61 | 1.68 | 1.64 | 1.61 | 1.7023 | 1.724 | 1.7609 | 1.65 | 1.73 | 1.78 | 1.67 | 1.7 | 1.6939 | 1.9529 | 1.6485 |  |  | 1
| 1 | 2 | 1.6528 | 1.7485 | 1.65 | 1.75 | 1.66 | 1.761 | 1.7 | 1.75 | 1.66 | 1.63 | 1.62 | 1.67 | 1.63 | 1.6709 | 1.6828 | 1.767 | 1.68 | 1.73 | 1.74 | 1.68 | 1.66 | 1.6994 | 2.033 | 1.6389 |  |  | #N/A
| 1 | 3 | 1.6522 | 1.7763 | 1.66 | 1.81 | 1.66 | 1.75 | 1.66 | 1.75 | 1.66 | 1.64 | 1.65 | 1.69 | 1.65 | 1.7091 | 1.7258 | 1.7368 | 1.67 | 1.72 | 1.77 | 1.66 | 1.67 | 1.7535 | 1.9877 | 1.6371 |  |  | 16
| 1 | 4 | 1.6723 | 1.746 | 1.73 | 1.8 | 1.72 | 1.775 | 1.73 | 1.73 | 1.65 | 1.61 | 1.66 | 1.69 | 1.62 | 1.7346 | 1.7199 | 1.7489 | 1.74 | 1.74 | 1.78 | 1.67 | 1.74 | 1.6848 | 1.9919 | 1.6251 |  |  | 1.6983
| 1 | 5 | 1.6885 | 1.7298 | 1.66 | 1.74 | 1.72 | 1.736 | 1.75 | 1.77 | 1.67 | 1.63 | 1.63 | 1.63 | 1.61 | 1.7238 | 1.7031 | 1.7489 | 1.69 | 1.72 | 1.76 | 1.67 | 1.68 | 1.7283 | 1.9841 | 1.6117 |  |  | 40
| 1 | 6 | 1.6747 | 1.7395 | 1.69 | 1.84 | 1.69 | 1.726 | 1.76 | 1.78 | 1.65 | 1.64 | 1.7 | 1.66 | 1.59 | 1.7424 | 1.687 | 1.7971 | 1.71 | 1.72 | 1.76 | 1.69 | 1.7 | 1.7368 | 1.9597 | 1.6988
| Mean |  | 1.6664 | 1.7479 | 1.6717 | 1.7383 | 1.69 | 1.7542 | 1.7033 | 1.75 | 1.66 | 1.6267 | 1.6567 | 1.6633 | 1.6183 | 1.7139 | 1.7071 | 1.7599 | 1.69 | 1.7267 | 1.765 | 1.6733 | 1.6917 | 1.7161 | 1.9849 | 1.6434
| Median |  | 1.665 | 1.7467 | 1.66 | 1.775 | 1.69 | 1.7555 | 1.715 | 1.75 | 1.66 | 1.63 | 1.655 | 1.665 | 1.615 | 1.7165 | 1.7115 | 1.7549 | 1.685 | 1.725 | 1.765 | 1.67 | 1.69 | 1.7138 | 1.9859 | 1.638
| Std Dev. |  | 0.01451 | 0.015538 | 0.033116 | 0.12734 | 0.026833 | 0.020702 | 0.05465 | 0.022804 | 0.0089443 | 0.013663 | 0.030111 | 0.025033 | 0.020412 | 0.025866 | 0.019037 | 0.021028 | 0.031623 | 0.008165 | 0.015166 | 0.010328 | 0.028577 | 0.027304 | 0.028391 | 0.029983
| Rel.Std.Dev. |  | 0.8707% | 0.8889% | 1.981% | 7.326% | 1.588% | 1.18% | 3.208% | 1.303% | 0.5388% | 0.8399% | 1.818% | 1.505% | 1.261% | 1.509% | 1.115% | 1.195% | 1.871% | 0.4729% | 0.8592% | 0.6172% | 1.689% | 1.591% | 1.43% | 1.824%
| PDM3 |  | -1.879% | 2.923% | -1.568% | 2.358% | -0.4883% | 3.29% | 0.2968% | 3.045% | -2.255% | -4.218% | -2.451% | -2.059% | -4.708% | 0.916% | 0.5183% | 3.63% | -0.4883% | 1.671% | 3.928% | -1.47% | -0.3902% | 1.049% | 16.88% | -3.235%
| Z-Score (Absolute) |  | 0.59 | 0.95 | 0.49 | 0.77 | 0.14 | 1.07 | 0.11 | 0.99 | 0.71 | 1.33 | 0.77 | 0.64 | 1.49 | 0.31 | 0.18 | 1.18 | 0.14 | 0.55 | 1.27 | 0.45 | 0.11 | 0.35 | 5.41 | 1.02
| Analytical results for Mn in OREAS 237b (Certified Value 0.041 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.043 | 0.041162 | 0.044 | 0.0363 | 0.042381 | 0.0413 | 0.0398 | 0.0379 | 0.0398 | 0.0387 | 0.0404 | 0.0398 | 0.0394 | 0.0411 | 0.040903 | 0.0387 | 0.0372 | 0.03827 | 0.0416 | 0.0393 | 0.0416 | 0.041679 | 0.0434 | 0.0404 |  |  | 1
| 1 | 2 | 0.0426 | 0.04113 | 0.0436 | 0.0433 | 0.042651 | 0.04 | 0.0396 | 0.0385 | 0.0398 | 0.0393 | 0.039 | 0.04 | 0.0396 | 0.0406 | 0.04136 | 0.0387 | 0.0377 | 0.03837 | 0.0404 | 0.0394 | 0.0407 | 0.042031 | 0.0438 | 0.0406 |  |  | 30
| 1 | 3 | 0.0435 | 0.04137 | 0.0438 | 0.0425 | 0.042793 | 0.0403 | 0.0389 | 0.0385 | 0.0394 | 0.0396 | 0.0397 | 0.0408 | 0.0401 | 0.0415 | 0.041023 | 0.0387 | 0.0378 | 0.03848 | 0.0413 | 0.0391 | 0.0407 | 0.043263 | 0.0427 | 0.0406 |  |  | 16
| 1 | 4 | 0.043 | 0.041255 | 0.0444 | 0.0437 | 0.042762 | 0.0414 | 0.0414 | 0.0379 | 0.0401 | 0.0387 | 0.0401 | 0.041 | 0.0399 | 0.0417 | 0.041345 | 0.0379 | 0.0388 | 0.03858 | 0.0413 | 0.0393 | 0.0423 | 0.041943 | 0.0427 | 0.0402 |  |  | 0.040752
| 1 | 5 | 0.0434 | 0.041007 | 0.042 | 0.0432 | 0.042519 | 0.0411 | 0.0408 | 0.0388 | 0.0393 | 0.0391 | 0.0395 | 0.0393 | 0.0402 | 0.0416 | 0.041465 | 0.0379 | 0.0382 | 0.03782 | 0.041 | 0.0394 | 0.0407 | 0.042581 | 0.0431 | 0.04 |  |  | 41
| 1 | 6 | 0.0439 | 0.041198 | 0.0434 | 0.0453 | 0.042405 | 0.0406 | 0.0414 | 0.0393 | 0.0405 | 0.0398 | 0.0413 | 0.0399 | 0.0393 | 0.0421 | 0.040549 | 0.0395 | 0.039 | 0.03802 | 0.0409 | 0.0397 | 0.0413 | 0.043043 | 0.0425 | 0.0411
| Mean |  | 0.043233 | 0.041187 | 0.043533 | 0.042383 | 0.042585 | 0.040783 | 0.040317 | 0.038483 | 0.039817 | 0.0392 | 0.04 | 0.040133 | 0.03975 | 0.041433 | 0.041108 | 0.038567 | 0.038117 | 0.038257 | 0.041083 | 0.039367 | 0.041217 | 0.042423 | 0.043033 | 0.040483
| Median |  | 0.0432 | 0.04118 | 0.0437 | 0.04325 | 0.042585 | 0.04085 | 0.0403 | 0.0385 | 0.0398 | 0.0392 | 0.0399 | 0.03995 | 0.03975 | 0.04155 | 0.041184 | 0.0387 | 0.038 | 0.03832 | 0.04115 | 0.03935 | 0.041 | 0.042306 | 0.0429 | 0.0405
| Std Dev. |  | 0.00045898 | 0.00012207 | 0.00082624 | 0.0031231 | 0.0001773 | 0.00057067 | 0.0010362 | 0.00053821 | 0.0004446 | 0.00045607 | 0.0008 | 0.00064395 | 0.00037283 | 0.00052026 | 0.00034882 | 0.00060222 | 0.00068823 | 0.0002878 | 0.00041673 | 0.00019664 | 0.00065243 | 0.00064071 | 0.00049666 | 0.00038166
| Rel.Std.Dev. |  | 1.062% | 0.2964% | 1.898% | 7.369% | 0.4163% | 1.399% | 2.57% | 1.399% | 1.117% | 1.163% | 2% | 1.605% | 0.9379% | 1.256% | 0.8486% | 1.561% | 1.806% | 0.7523% | 1.014% | 0.4995% | 1.583% | 1.51% | 1.154% | 0.9428%
| PDM3 |  | 6.089% | 1.068% | 6.825% | 4.004% | 4.499% | 0.07733% | -1.068% | -5.567% | -2.295% | -3.808% | -1.845% | -1.518% | -2.458% | 1.672% | 0.873% | -5.362% | -6.466% | -6.123% | 0.8135% | -3.399% | 1.141% | 4.102% | 5.599% | -0.6588%
| Z-Score (Absolute) |  | 1.63 | 0.35 | 1.82 | 1.1 | 1.22 | 0.09 | 0.2 | 1.35 | 0.51 | 0.9 | 0.4 | 0.31 | 0.55 | 0.5 | 0.3 | 1.28 | 1.58 | 1.49 | 0.28 | 0.79 | 0.37 | 1.12 | 1.51 | 0.09
| Analytical results for Mo in OREAS 237b (Certified Value 0.97 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 1 | 1 | < 1 | < 2 | 1 | 1.13 | 0.7 | 2.8 | 0.92 | 1.07 | 0.96 | 0.94 | 0.95 | 0.95737 | 1.08 | 1.309 | 0.92 | < 1 | 0.97 | 1.02 | 0.88 | 0.8209 | < 5 | 2.3465 | 0.93617 |  | 1
| 1 | 2 | 1 | 1 | < 1 | < 2 | 0.9 | 1.03 | 1 | 2.5 | 0.89 | 0.88 | 0.94 | 0.92 | 0.92 | 0.99472 | 1.02 | 1.124 | 1.02 | < 1 | 1.03 | 0.97 | 0.89 | 0.6859 | < 5 | 2.124 | 0.95895 |  | 31
| 1 | 3 | 1 | 1 | < 1 | < 2 | 1 | 1.1 | 0.9 | 2.7 | 0.92 | 0.92 | 1 | 0.94 | 0.91 | 1.0121 | 1.09 | 1.308 | 1.02 | < 1 | 0.88 | 0.99 | 0.9 | 0.8198 | < 5 | 2.9583 | 0.89732 |  | 16
| 1 | 4 | 1 | 1 | < 1 | < 2 | 0.9 | 1.11 | 0.8 | 2.6 | 0.95 | 0.89 | 0.98 | 0.92 | 0.93 | 1.0362 | 1.15 | 2.366 | 1.04 | < 1 | 1.15 | 1.06 | 0.9 | 0.7744 | < 5 | 2.7914 | 0.94082 |  | 0.97151
| 1 | 5 | 1 | 1 | < 1 | < 2 | 1 | 1.07 | 0.9 | 2.9 | 0.99 | 0.9 | 0.98 | 1.02 | 0.95 | 0.97501 | 1.09 | 1.374 | 0.97 | < 1 | 0.9 | 1.05 | 0.89 | 0.8036 | < 5 | 2.5133 | 0.99918 |  | 42
| 1 | 6 | 1 | 1 | < 1 | < 2 | 1 | 0.97 | 0.9 | 3.2 | 0.93 | 0.95 | 0.92 | 0.89 | 0.93 | 0.94991 | 1.14 | 1.307 | 0.94 | < 1 | 1 | 1.11 | 0.96 | 0.7237 | < 5 | 2.6246 | 0.91445
| Mean |  | 1 | 1 |  |  | 0.96667 | 1.0683 | 0.86667 | 2.7833 | 0.93333 | 0.935 | 0.96333 | 0.93833 | 0.93167 | 0.98756 | 1.095 | 1.4647 | 0.985 |  | 0.98833 | 1.0333 | 0.90333 | 0.77138 |  | 2.5597 | 0.94115
| Median |  | 1 | 1 |  |  | 1 | 1.085 | 0.9 | 2.75 | 0.925 | 0.91 | 0.97 | 0.93 | 0.93 | 0.98486 | 1.09 | 1.3085 | 0.995 |  | 0.985 | 1.035 | 0.895 | 0.789 |  | 2.569 | 0.93849
| Std Dev. |  | 0 | 0 |  |  | 0.05164 | 0.05947 | 0.10328 | 0.24833 | 0.033862 | 0.07064 | 0.029439 | 0.044008 | 0.016021 | 0.033196 | 0.046797 | 0.44951 | 0.048888 |  | 0.09786 | 0.050859 | 0.028752 | 0.055541 |  | 0.3014 | 0.035618
| Rel.Std.Dev. |  | 0% | 0% |  |  | 5.342% | 5.567% | 11.92% | 8.922% | 3.628% | 7.555% | 3.056% | 4.69% | 1.72% | 3.361% | 4.274% | 30.69% | 4.963% |  | 9.902% | 4.922% | 3.183% | 7.2% |  | 11.78% | 3.785%
| PDM3 |  | 2.932% | 2.932% |  |  | -0.4988% | 9.966% | -10.79% | 186.5% | -3.93% | -3.758% | -0.8419% | -3.415% | -4.101% | 1.652% | 12.71% | 50.76% | 1.388% |  | 1.731% | 6.363% | -7.018% | -20.6% |  | 163.5% | -3.125%
| Z-Score (Absolute) |  | 0.2 | 0.2 | 6.33 | 0.2 | 0.24 | 1.09 | 1.54 | 23.47 | 0.67 | 0.65 | 0.28 | 0.61 | 0.7 | 0.03 | 1.44 | 6.26 | 0 | 6.33 | 0.04 | 0.63 | 1.07 | 2.79 | 19.77 | 20.55 | 0.57
| Analytical results for Na in OREAS 237b (Certified Value 0.756 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.764 | 0.78895 | 0.75 | 0.71 | 0.78 | 0.75 | 0.72 | 0.772 | 0.76 | 0.73 | 0.76 | 0.73 | 0.76 | 0.757 | 0.72324 | 0.742 | 0.74 | 0.76 | 0.76 | 0.74 | 0.758 | 0.74431 | 0.8091 |  |  |  | 1
| 1 | 2 | 0.7715 | 0.78791 | 0.74 | 0.8 | 0.801 | 0.746 | 0.73 | 0.787 | 0.75 | 0.74 | 0.74 | 0.75 | 0.77 | 0.742 | 0.72681 | 0.749 | 0.74 | 0.75 | 0.74 | 0.74 | 0.746 | 0.74872 | 0.808 |  |  |  | #N/A
| 1 | 3 | 0.7756 | 0.78807 | 0.74 | 0.79 | 0.81433 | 0.77 | 0.72 | 0.786 | 0.75 | 0.74 | 0.75 | 0.76 | 0.77 | 0.755 | 0.71187 | 0.742 | 0.76 | 0.76 | 0.76 | 0.74 | 0.743 | 0.77342 | 0.7976 |  |  |  | 16
| 1 | 4 | 0.7837 | 0.78186 | 0.75 | 0.81 | 0.82 | 0.748 | 0.75 | 0.777 | 0.76 | 0.72 | 0.75 | 0.76 | 0.76 | 0.757 | 0.71507 | 0.742 | 0.79 | 0.76 | 0.76 | 0.74 | 0.775 | 0.74274 | 0.776 |  |  |  | 0.7559
| 1 | 5 | 0.779 | 0.78262 | 0.76 | 0.79 | 0.7915 | 0.771 | 0.74 | 0.797 | 0.73 | 0.73 | 0.74 | 0.74 | 0.76 | 0.763 | 0.72046 | 0.734 | 0.75 | 0.75 | 0.75 | 0.74 | 0.76 | 0.76293 | 0.7958 |  |  |  | 43
| 1 | 6 | 0.7745 | 0.7883 | 0.76 | 0.84 | 0.80733 | 0.774 | 0.78 | 0.797 | 0.73 | 0.73 | 0.77 | 0.75 | 0.75 | 0.771 | 0.72314 | 0.757 | 0.79 | 0.76 | 0.74 | 0.75 | 0.751 | 0.76528 | 0.811
| Mean |  | 0.77472 | 0.78628 | 0.75 | 0.79 | 0.80236 | 0.75983 | 0.74 | 0.786 | 0.74667 | 0.73167 | 0.75167 | 0.74833 | 0.76167 | 0.7575 | 0.7201 | 0.74433 | 0.76167 | 0.75667 | 0.75167 | 0.74167 | 0.7555 | 0.75623 | 0.79958
| Median |  | 0.77505 | 0.78799 | 0.75 | 0.795 | 0.80417 | 0.76 | 0.735 | 0.7865 | 0.75 | 0.73 | 0.75 | 0.75 | 0.76 | 0.757 | 0.7218 | 0.742 | 0.755 | 0.76 | 0.755 | 0.74 | 0.7545 | 0.75582 | 0.8028
| Std Dev. |  | 0.0067092 | 0.0031623 | 0.0089443 | 0.043359 | 0.014826 | 0.013091 | 0.022804 | 0.010198 | 0.013663 | 0.0075277 | 0.01169 | 0.01169 | 0.0075277 | 0.0095864 | 0.0056111 | 0.0078145 | 0.023166 | 0.005164 | 0.0098319 | 0.0040825 | 0.011606 | 0.01267 | 0.013162
| Rel.Std.Dev. |  | 0.866% | 0.4022% | 1.193% | 5.488% | 1.848% | 1.723% | 3.082% | 1.297% | 1.83% | 1.029% | 1.555% | 1.562% | 0.9883% | 1.266% | 0.7792% | 1.05% | 3.041% | 0.6825% | 1.308% | 0.5504% | 1.536% | 1.675% | 1.646%
| PDM3 |  | 2.49% | 4.02% | -0.7802% | 4.512% | 6.147% | 0.5207% | -2.103% | 3.982% | -1.221% | -3.206% | -0.5597% | -1.001% | 0.7633% | 0.212% | -4.736% | -1.53% | 0.7633% | 0.1018% | -0.5597% | -1.883% | -0.05255% | 0.04446% | 5.779%
| Z-Score (Absolute) |  | 1.3 | 2.12 | 0.44 | 2.38 | 3.25 | 0.25 | 1.14 | 2.1 | 0.67 | 1.73 | 0.32 | 0.56 | 0.38 | 0.09 | 2.55 | 0.84 | 0.38 | 0.03 | 0.32 | 1.03 | 0.05 | 0 | 3.06
| Analytical results for Nb in OREAS 237b (Certified Value 13.6 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
1
| 1 | 1 | 14.39 | 12.2681 | 0.3 | 15.81 | 13.6 | 12 | 12.5 | 12.8 | 14.3 | 14.3 | 14 | 13.2 | 14.24 | 14.42 | 15.5 | 11.6 | 15.51 | 13.9 | 12.15 | 10.3082 | 15.0573 | [5 empty] | 1
| 1 | 2 | 14.65 | 12.2259 | 0.3 | 15.21 | 12.9 | 11.3 | 11.6 | 12.9 | 14.2 | 14 | 14.1 | 13.4 | 13.44 | 14.1464 | 15.6 | 12.3 | 14.76 | 13.6 | 12.09 | 9.9328 | 14.8932 | [5 empty] | 32
| 1 | 3 | 14.88 | 12.2134 | 0.3 | 15.26 | 13.1 | 12.1 | 12.8 | 13.1 | 14.5 | 14.2 | 14.4 | 13.9 | 13.95 | 14.5736 | 15.1 | 13.3 | 15.31 | 13.5 | 11.87 | 10.2974 | 15.0498 | [5 empty] | 16
| 1 | 4 | 14.27 | 12.3015 | 0.3 | 14.73 | 13.2 | 12.8 | 12.3 | 12.6 | 14.3 | 13.6 | 13.9 | 13.6 | 14.42 | 14.4026 | 15.3 | 12.9 | 15.37 | 13.7 | 11.94 | 10.174 | 15.0943 | [5 empty] | 13.606
| 1 | 5 | 14.2 | 12.232 | 1.2 | 15.41 | 13.6 | 12.9 | 11.9 | 13.1 | 14.4 | 13.1 | 14 | 13.7 | 14.65 | 14.1915 | 14.8 | 12.6 | 15.24 | 13.5 | 12.32 | 10.1947 | 15.1557 | [5 empty] | 44
| 1 | 6 | 14.97 | 12.308 | 1.4 | 16.05 | 13.1 | 13.9 | 14.1 | 13.4 | 14.7 | 15 | 14.2 | 13.5 | 14.92 | 14.1962 | 15.4 | 12.7 | 14.91 | 15 | 12 | 10.0617 | 15.066
| Mean |  | 14.56 | 12.2582 | 0.63333 | 15.4117 | 13.25 | 12.5 | 12.5333 | 12.9833 | 14.4 | 14.0333 | 14.1 | 13.55 | 14.27 | 14.3217 | 15.2833 | 12.5667 | 15.1833 | 13.8667 | 12.0617 | 10.1615 | 15.0527
| Median |  | 14.52 | 12.25 | 0.3 | 15.335 | 13.15 | 12.45 | 12.4 | 13 | 14.35 | 14.1 | 14.05 | 13.55 | 14.33 | 14.2994 | 15.35 | 12.65 | 15.275 | 13.65 | 12.045 | 10.1843 | 15.0616
| Std Dev. |  | 0.32286 | 0.040471 | 0.52026 | 0.46786 | 0.2881 | 0.90111 | 0.87788 | 0.27869 | 0.17889 | 0.64704 | 0.17889 | 0.2429 | 0.52566 | 0.16917 | 0.29269 | 0.5785 | 0.288 | 0.57504 | 0.16167 | 0.14389 | 0.08717
| Rel.Std.Dev. |  | 2.217% | 0.3302% | 82.15% | 3.036% | 2.174% | 7.209% | 7.004% | 2.147% | 1.242% | 4.611% | 1.269% | 1.793% | 3.684% | 1.181% | 1.915% | 4.603% | 1.897% | 4.147% | 1.34% | 1.416% | 0.5791%
| PDM3 |  | 7.011% | -9.906% | -95.35% | 13.27% | -2.617% | -8.129% | -7.884% | -4.577% | 5.835% | 3.141% | 3.63% | -0.4118% | 4.88% | 5.26% | 12.33% | -7.639% | 11.59% | 1.916% | -11.35% | -25.32% | 10.63%
| Z-Score (Absolute) |  | 0.39 | 0.91 | 7.52 | 0.88 | 0.35 | 0.78 | 0.76 | 0.5 | 0.3 | 0.09 | 0.13 | 0.18 | 0.23 | 0.26 | 0.81 | 0.74 | 0.75 | 0 | 1.03 | 2.11 | 0.67
| Analytical results for Nd in OREAS 237b (Certified Value 34.2 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 1
2
| 1 | 1 | 33.3518 | 34.5 | 33.8 | 33.3 | 35.3 | 33.1 | 34.39 | 34.4 | 29.24 | 32.7877 | [16 empty] | 1
| 1 | 2 | 33.0411 | 34.7 | 36 | 30.1 | 36 | 33.59 | 33.17 | 33.6 | 29.25 | 33.471 | [16 empty] | 9
| 1 | 3 | 33.0149 | 34.9 | 36.3 | 31.4 | 35.4 | 36.56 | 34.52 | 33.4 | 30.27 | 32.7895 | [16 empty] | 16
| 1 | 4 | 33.2705 | 35.5 | 36.2 | 39.3 | 35.3 | 33.57 | 34.42 | 34.4 | 29.26 | 33.2944 | [16 empty] | 34.1903
| 1 | 5 | 33.2589 | 34.6 | 36 | 35.4 | 34.5 | 34.28 | 33.74 | 33.9 | 29.66 | 33.9346 | [16 empty] | 45
| 1 | 6 | 32.8504 | 36 | 36 | 36.3 | 35.4 | 33.31 | 33.54 | 34.8 | 29.97 | 33.0131
| Mean |  | 33.1313 | 35.0333 | 35.7167 | 34.3 | 35.3167 | 34.0683 | 33.9633 | 34.0833 | 29.6083 | 33.2151
| Median |  | 33.15 | 34.8 | 36 | 34.35 | 35.35 | 33.58 | 34.065 | 34.15 | 29.46 | 33.1537
| Std Dev. |  | 0.19228 | 0.59217 | 0.94745 | 3.3841 | 0.47924 | 1.284 | 0.55838 | 0.53821 | 0.43742 | 0.44589
| Rel.Std.Dev. |  | 0.5804% | 1.69% | 2.653% | 9.866% | 1.357% | 3.769% | 1.644% | 1.579% | 1.477% | 1.342%
| PDM3 |  | -3.098% | 2.466% | 4.464% | 0.3208% | 3.294% | -0.3568% | -0.6639% | -0.313% | -13.4% | -2.852%
| Z-Score (Absolute) |  | 0.71 | 0.72 | 1.23 | 0.17 | 0.93 | 0.01 | 0.08 | 0.01 | 3.34 | 0.64
| Analytical results for Ni in OREAS 237b (Certified Value 64 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 64 | 60.09 | 65 | 67 | 59.6439 | 65 | 70 | 66 | 66.9 | 63.1 | 70.2 | 66.8 | 62 | 62.1 | 67.0386 | 66.4 | 57.08 | 64 | 58.85 | 64.5 | 67 | 64.4 | 62.81 | 66 | 59.8737 | 64.5146 | 1
| 1 | 2 | 64.5 | 59.12 | 65 | 66 | 60.0267 | 64 | 78 | 62 | 68.6 | 64.1 | 64 | 65.9 | 63.8 | 61.5 | 66.971 | 67.6 | 54.52 | 67 | 58.19 | 62.5 | 63.7 | 66.8 | 63.47 | 67 | 60.5591 | 64.2893 | 34
| 1 | 3 | 64.8 | 59.83 | 65 | 69 | 60.6933 | 63 | 76 | 63 | 66.6 | 63.4 | 69.1 | 67.5 | 64.3 | 61.7 | 66.6 | 68.5 | 61.16 | 64 | 60.01 | 62.6 | 62.9 | 66.6 | 64.46 | 68 | 59.6595 | 63.1918 | 16
| 1 | 4 | 65 | 59.14 | 70 | 66 | 59.5467 | 65 | 71 | 63 | 66.8 | 62.8 | 65.7 | 67 | 62.8 | 63.3 | 66.8399 | 66.8 | 96.95 | 66 | 58.52 | 64.1 | 64.2 | 67.4 | 62.59 | 67 | 60.1307 | 65.0284 | 64.365
| 1 | 5 | 64.4 | 59.9 | 70 | 66 | 61.03 | 65 | 74 | 64 | 66.8 | 64.1 | 64.3 | 66.5 | 63.2 | 63.2 | 68.144 | 67 | 59.22 | 67 | 59.71 | 63.1 | 62.5 | 66.8 | 63.69 | 67 | 62.5298 | 64.9478 | 46
| 1 | 6 | 64.9 | 60.29 | 70 | 65 | 59.9056 | 64 | 63 | 65 | 67.2 | 64.8 | 67.4 | 66.7 | 62.1 | 64.6 | 66.2086 | 65.8 | 61.39 | 65 | 60.45 | 62.3 | 68.6 | 68.4 | 63.8 | 67 | 61.2017 | 60.6058
| Mean |  | 64.6 | 59.7283 | 67.5 | 66.5 | 60.141 | 64.3333 | 72 | 63.8333 | 67.15 | 63.7167 | 66.7833 | 66.7333 | 63.0333 | 62.7333 | 66.967 | 67.0167 | 65.0533 | 65.5 | 59.2883 | 63.1833 | 64.8167 | 66.7333 | 63.47 | 67 | 60.6591 | 63.7629
| Median |  | 64.65 | 59.865 | 67.5 | 66 | 59.9662 | 64.5 | 72.5 | 63.5 | 66.85 | 63.75 | 66.55 | 66.75 | 63 | 62.65 | 66.9054 | 66.9 | 60.19 | 65.5 | 59.28 | 62.85 | 63.95 | 66.8 | 63.58 | 67 | 60.3449 | 64.4019
| Std Dev. |  | 0.37417 | 0.49036 | 2.7386 | 1.3784 | 0.594 | 0.8165 | 5.3292 | 1.472 | 0.73689 | 0.74677 | 2.5498 | 0.53166 | 0.91797 | 1.1843 | 0.65061 | 0.94322 | 15.84 | 1.3784 | 0.89854 | 0.91305 | 2.4392 | 1.3186 | 0.68519 | 0.63246 | 1.0676 | 1.6817
| Rel.Std.Dev. |  | 0.5792% | 0.821% | 4.057% | 2.073% | 0.9877% | 1.269% | 7.402% | 2.306% | 1.097% | 1.172% | 3.818% | 0.7967% | 1.456% | 1.888% | 0.9715% | 1.407% | 24.35% | 2.104% | 1.516% | 1.445% | 3.763% | 1.976% | 1.08% | 0.944% | 1.76% | 2.637%
| PDM3 |  | 0.3652% | -7.204% | 4.871% | 3.317% | -6.562% | -0.04913% | 11.86% | -0.826% | 4.327% | -1.007% | 3.757% | 3.68% | -2.069% | -2.535% | 4.043% | 4.12% | 1.069% | 1.763% | -7.887% | -1.836% | 0.7018% | 3.68% | -1.39% | 4.094% | -5.758% | -0.9353%
| Z-Score (Absolute) |  | 0.04 | 1.68 | 0.94 | 0.6 | 1.54 | 0.13 | 2.46 | 0.3 | 0.82 | 0.33 | 0.7 | 0.68 | 0.56 | 0.67 | 0.76 | 0.78 | 0.12 | 0.27 | 1.83 | 0.51 | 0.04 | 0.68 | 0.42 | 0.77 | 1.37 | 0.32
| Analytical results for P in OREAS 237b (Certified Value 0.066 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0667 | 0.063166 | 0.06 | 0.062 | 0.065002 | 0.068 | 0.0636 | 0.064 | 0.066 | 0.065 | 0.068 | 0.066 | 0.065 | 0.0666 | 0.063611 | 0.0689 | 0.07 | 0.07 | 0.0652 | 0.0677 | 0.066 | 0.064 | 0.0648 | 0.0725 |  |  | 1
| 1 | 2 | 0.0661 | 0.062405 | 0.06 | 0.063 | 0.064796 | 0.07 | 0.0661 | 0.066 | 0.066 | 0.066 | 0.066 | 0.069 | 0.066 | 0.066 | 0.064093 | 0.0689 | 0.07 | 0.07 | 0.0631 | 0.0675 | 0.065 | 0.0652 | 0.0669 | 0.0728 |  |  | #N/A
| 1 | 3 | 0.067 | 0.0634 | 0.06 | 0.063 | 0.064982 | 0.07 | 0.0627 | 0.065 | 0.067 | 0.066 | 0.067 | 0.07 | 0.068 | 0.0669 | 0.065704 | 0.0672 | 0.06 | 0.07 | 0.066 | 0.0672 | 0.067 | 0.067 | 0.0648 | 0.0727 |  |  | 16
| 1 | 4 | 0.0658 | 0.062326 | 0.065 | 0.062 | 0.065217 | 0.07 | 0.0678 | 0.065 | 0.066 | 0.065 | 0.067 | 0.068 | 0.065 | 0.0676 | 0.066531 | 0.0685 | 0.07 | 0.07 | 0.0636 | 0.0677 | 0.068 | 0.0631 | 0.065 | 0.072 |  |  | 0.066302
| 1 | 5 | 0.0647 | 0.062434 | 0.06 | 0.064 | 0.064715 | 0.069 | 0.0674 | 0.066 | 0.066 | 0.065 | 0.066 | 0.067 | 0.067 | 0.0674 | 0.065379 | 0.0685 | 0.07 | 0.07 | 0.0635 | 0.0682 | 0.067 | 0.0652 | 0.0652 | 0.0727 |  |  | 47
| 1 | 6 | 0.0652 | 0.062516 | 0.06 | 0.062 | 0.065122 | 0.068 | 0.0685 | 0.066 | 0.066 | 0.066 | 0.069 | 0.068 | 0.065 | 0.0691 | 0.068014 | 0.0707 | 0.07 | 0.07 | 0.0643 | 0.0686 | 0.068 | 0.0661 | 0.0635 | 0.0734
| Mean |  | 0.065917 | 0.062708 | 0.060833 | 0.062667 | 0.064972 | 0.069167 | 0.066017 | 0.065333 | 0.066167 | 0.0655 | 0.067167 | 0.068 | 0.066 | 0.067267 | 0.065555 | 0.068783 | 0.068333 | 0.07 | 0.064283 | 0.067817 | 0.066833 | 0.0651 | 0.065033 | 0.072683
| Median |  | 0.06595 | 0.062475 | 0.06 | 0.0625 | 0.064992 | 0.0695 | 0.06675 | 0.0655 | 0.066 | 0.0655 | 0.067 | 0.068 | 0.0655 | 0.06715 | 0.065542 | 0.0687 | 0.07 | 0.07 | 0.06395 | 0.0677 | 0.067 | 0.0652 | 0.0649 | 0.0727
| Std Dev. |  | 0.00087502 | 0.00045565 | 0.0020412 | 0.0008165 | 0.00019027 | 0.00098319 | 0.002371 | 0.0008165 | 0.00040825 | 0.00054772 | 0.001169 | 0.0014142 | 0.0012649 | 0.0010652 | 0.0016104 | 0.0011286 | 0.0040825 | 0 | 0.0011197 | 0.00050365 | 0.001169 | 0.0014029 | 0.001093 | 0.0004535
| Rel.Std.Dev. |  | 1.327% | 0.7266% | 3.355% | 1.303% | 0.2928% | 1.421% | 3.592% | 1.25% | 0.617% | 0.8362% | 1.741% | 2.08% | 1.917% | 1.584% | 2.457% | 1.641% | 5.974% | 0% | 1.742% | 0.7427% | 1.749% | 2.155% | 1.681% | 0.6239%
| PDM3 |  | -0.5807% | -5.421% | -8.248% | -5.483% | -2.005% | 4.321% | -0.4299% | -1.461% | -0.2037% | -1.209% | 1.305% | 2.561% | -0.455% | 1.455% | -1.125% | 3.743% | 3.064% | 5.578% | -3.044% | 2.285% | 0.8018% | -1.812% | -1.913% | 9.625%
| Z-Score (Absolute) |  | 0.05 | 1.84 | 2.89 | 1.86 | 0.58 | 1.76 | 0 | 0.38 | 0.09 | 0.28 | 0.65 | 1.11 | 0 | 0.7 | 0.25 | 1.56 | 1.3 | 2.23 | 0.96 | 1.01 | 0.46 | 0.51 | 0.54 | 3.72
| Analytical results for Pb in OREAS 237b (Certified Value 20 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 21.1 | 20 | 18 | 18.971 | 20 | 13.8 | 20 | 16.1 | 19.5 | 20.1 | 20.2 | 19.8 | 19.82 | 21.8505 | 19.7 | 19 | 20.2 | 20.14 | 21.4 | 19.4 | 20.63 | 18.1 | 15 | 22.3317 | 20.288 |  | 1
| 1 | 2 | 20.9 | 20 | 16 | 18.564 | 20 | 16 | 19 | 16.1 | 19.9 | 19.2 | 20 | 19.8 | 18.56 | 22.0921 | 20 | 20.28 | 20.4 | 19.55 | 20.9 | 18.3 | 19.67 | 16.6 | 16 | 20.1429 | 20.0575 |  | 35
| 1 | 3 | 21 | 20 | 19 | 18.7893 | 20 | 13.3 | 20 | 18 | 20 | 19.8 | 20.8 | 21.1 | 19.24 | 22.6229 | 19.6 | 21.75 | 19.8 | 20.15 | 21.6 | 18.6 | 19.55 | 17.7 | 17 | 21.884 | 20.1681 |  | 16
| 1 | 4 | 20.7 | 21 | 19 | 19.5087 | 20 | 17.1 | 20 | 16.7 | 20 | 19.4 | 20.5 | 20.1 | 21.05 | 22.7013 | 19.4 | 45.12 | 21.2 | 19.3 | 21.3 | 19.5 | 20.34 | 18.1 | 14 | 23.5256 | 20.0825 |  | 20.0063
| 1 | 5 | 20.6 | 20 | 18 | 19.1493 | 20 | 13 | 20 | 16.4 | 19.8 | 18.8 | 19.8 | 20.2 | 19.66 | 21.8905 | 19 | 20.48 | 20.4 | 19.92 | 21.3 | 17.8 | 20.3 | 19.1 | 16 | 22.1327 | 20.4354 |  | 48
| 1 | 6 | 20.7 | 20 | 20 | 18.32 | 20 | 17.7 | 21 | 16.2 | 19.9 | 19.7 | 20.4 | 19.7 | 20.22 | 21.8307 | 19.5 | 20.75 | 19.6 | 20.52 | 20.6 | 21.5 | 20.46 | 19.3 | 15 | 23.2769 | 20.2301
| Mean |  | 20.8333 | 20.1667 | 18.3333 | 18.8837 | 20 | 15.15 | 20 | 16.5833 | 19.85 | 19.5 | 20.2833 | 20.1167 | 19.7583 | 22.1647 | 19.5333 | 24.5633 | 20.2667 | 19.93 | 21.1833 | 19.1833 | 20.1583 | 18.15 | 15.5 | 22.2156 | 20.2103
| Median |  | 20.8 | 20 | 18.5 | 18.8802 | 20 | 14.9 | 20 | 16.3 | 19.9 | 19.55 | 20.3 | 19.95 | 19.74 | 21.9913 | 19.55 | 20.615 | 20.3 | 20.03 | 21.3 | 19 | 20.32 | 18.1 | 15.5 | 22.2322 | 20.1991
| Std Dev. |  | 0.19664 | 0.40825 | 1.3663 | 0.42371 | 0 | 2.0443 | 0.63246 | 0.73052 | 0.18708 | 0.46476 | 0.36009 | 0.51929 | 0.84849 | 0.39713 | 0.33267 | 10.1094 | 0.56095 | 0.44317 | 0.3656 | 1.3075 | 0.44161 | 0.98336 | 1.0488 | 1.2057 | 0.14041
| Rel.Std.Dev. |  | 0.9439% | 2.024% | 7.452% | 2.244% | 0% | 13.49% | 3.162% | 4.405% | 0.9425% | 2.383% | 1.775% | 2.581% | 4.294% | 1.792% | 1.703% | 41.16% | 2.768% | 2.224% | 1.726% | 6.816% | 2.191% | 5.418% | 6.767% | 5.427% | 0.6948%
| PDM3 |  | 4.134% | 0.8018% | -8.362% | -5.611% | -0.03128% | -24.27% | -0.03128% | -17.11% | -0.781% | -2.531% | 1.385% | 0.5519% | -1.239% | 10.79% | -2.364% | 22.78% | 1.302% | -0.3812% | 5.884% | -4.113% | 0.7601% | -9.278% | -22.52% | 11.04% | 1.02%
| Z-Score (Absolute) |  | 1.12 | 0.22 | 2.25 | 1.51 | 0 | 6.54 | 0 | 4.61 | 0.2 | 0.67 | 0.38 | 0.16 | 0.33 | 2.92 | 0.63 | 6.15 | 0.36 | 0.09 | 1.6 | 1.1 | 0.21 | 2.49 | 6.07 | 2.99 | 0.28
| Analytical results for Pd in OREAS 237b (Indicative Value < 12.8 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 1
1
| 1 | 1 | < 9.6 | [25 empty] | 1
| 1 | 2 | < 18.2 | [25 empty] | 1
| 1 | 3 | 1.1 | [25 empty] | 16
| 1 | 4 | < 12.8 | [25 empty] | < 12.8
| 1 | 5 | < 6.4 | [25 empty] | 11
| 1 | 6 | < 15
| Mean |  | 1.1
| Median |  | 1.1
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 237b (Certified Value 9.29 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 8.8293 | 9.25 | 8.9 | 9.4 | 10 | 8.813 | 9.42 | 9.4 | 7.45 | 8.7334 | [16 empty] | 1
| 1 | 2 | 8.782 | 9.15 | 9.7 | 8.1 | 10.2 | 9.043 | 9.3 | 9.1 | 7.36 | 8.9304 | [16 empty] | 11
| 1 | 3 | 8.6666 | 9.3 | 9.7 | 8.5 | 10.1 | 9.89 | 9.61 | 9.2 | 7.64 | 8.8212 | [16 empty] | 16
| 1 | 4 | 8.8112 | 9.15 | 9.9 | 11.3 | 9.9 | 9.141 | 9.4 | 9.5 | 7.29 | 8.8516 | [16 empty] | 9.2901
| 1 | 5 | 8.8004 | 9.3 | 9.8 | 9.7 | 9.8 | 9.109 | 9.4 | 9.5 | 7.38 | 8.9941 | [16 empty] | 49
| 1 | 6 | 8.7395 | 9.5 | 9.6 | 10 | 10 | 9.082 | 9.19 | 9.7 | 7.55 | 8.8317
| Mean |  | 8.7715 | 9.275 | 9.6 | 9.5 | 10 | 9.1797 | 9.3867 | 9.4 | 7.445 | 8.8604
| Median |  | 8.7912 | 9.275 | 9.7 | 9.55 | 10 | 9.0955 | 9.4 | 9.45 | 7.415 | 8.8416
| Std Dev. |  | 0.059809 | 0.12942 | 0.35777 | 1.1402 | 0.14142 | 0.36708 | 0.13967 | 0.21909 | 0.13004 | 0.090953
| Rel.Std.Dev. |  | 0.6818% | 1.395% | 3.727% | 12% | 1.414% | 3.999% | 1.488% | 2.331% | 1.747% | 1.027%
| PDM3 |  | -5.582% | -0.1629% | 3.335% | 2.259% | 7.641% | -1.189% | 1.039% | 1.183% | -19.86% | -4.626%
| Z-Score (Absolute) |  | 1.72 | 0.17 | 0.83 | 0.52 | 2.06 | 0.47 | 0.17 | 0.21 | 5.8 | 1.45
| Analytical results for Pt in OREAS 237b (Indicative Value 1.13 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 2
2
| 1 | 1 | 0.4 | [25 empty] | 1
| 1 | 2 | < 1.4 | [25 empty] | 1
| 1 | 3 | 0.8 | [25 empty] | 16
| 1 | 4 | 2.3 | [25 empty] | 1.1333
| 1 | 5 | 2.5 | [25 empty] | 12
| 1 | 6 | 0.1
| Mean |  | 1.22
| Median |  | 0.8
| Std Dev. |  | 1.1077
| Rel.Std.Dev. |  | 90.8%
| PDM3 |  | 7.647%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 237b (Certified Value 155 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 163.38 | 139.6557 | 154 | 134 | 163.1 | 168.2 | 148 | 158.8 | 159 | 157.5 | 154.5 | 160.5 | 201.3 | 149.0847 | 158 | 139.5 | 161 | 144.64 | 151.5 | 154.7 | 165.7912 | 118.2765 |  |  |  |  | 1
| 1 | 2 | 163.04 | 139.1068 | 152 | 171 | 160.1 | 168.8 | 139 | 156.1 | 160.5 | 152 | 153.5 | 160 | 202.4 | 149.2397 | 160 | 140.6 | 166 | 139.82 | 144.5 | 153.2 | 157.2484 | 136.5333 |  |  |  |  | 37
| 1 | 3 | 164.52 | 136.6005 | 152 | 172 | 159.4 | 166.5 | 149 | 159.5 | 158.5 | 159 | 159 | 172 | 192.2 | 148.7028 | 157 | 143.8 | 174 | 135.79 | 142.4 | 159.7 | 165.6029 | 125.7496 |  |  |  |  | 16
| 1 | 4 | 164.18 | 139.0914 | 159 | 178 | 163.3 | 175.1 | 149 | 151.7 | 159 | 154.5 | 150 | 159.5 | 254.8 | 145.109 | 157 | 244.5 | 174 | 153.54 | 144.7 | 159.4 | 160.8605 | 133.5489 |  |  |  |  | 154.9886
| 1 | 5 | 165.63 | 138.0229 | 155 | 178 | 162.7 | 173.3 | 142 | 159.5 | 161.5 | 148 | 153.5 | 166.5 | 214.1 | 147.8299 | 155 | 141.7 | 164 | 155.93 | 144.2 | 154.3 | 163.1185 | 134.0724 |  |  |  |  | 50
| 1 | 6 | 162.54 | 138.6813 | 155 | 179 | 161.7 | 176.3 | 150 | 155.8 | 166.5 | 167 | 153 | 158 | 238.2 | 144.9477 | 157 | 143.4 | 157 | 141.04 | 160 | 157.6 | 162.2109 | 129.8875
| Mean |  | 163.8817 | 138.5264 | 154.5 | 168.6667 | 161.7167 | 171.3667 | 146.1667 | 156.9 | 160.8333 | 156.3333 | 153.9167 | 162.75 | 217.1667 | 147.4856 | 157.3333 | 158.9167 | 166 | 145.1267 | 147.8833 | 156.4833 | 162.4721 | 129.678
| Median |  | 163.78 | 138.8863 | 154.5 | 175 | 162.2 | 171.05 | 148.5 | 157.45 | 159.75 | 156 | 153.5 | 160.25 | 208.25 | 148.2663 | 157 | 142.55 | 165 | 142.84 | 144.6 | 156.15 | 162.6647 | 131.7182
| Std Dev. |  | 1.1237 | 1.088 | 2.5884 | 17.3167 | 1.6351 | 4.0574 | 4.535 | 3.0338 | 2.9944 | 6.5396 | 2.9226 | 5.3921 | 24.336 | 1.9659 | 1.633 | 41.9588 | 6.8993 | 7.9971 | 6.7098 | 2.7867 | 3.1996 | 6.7344
| Rel.Std.Dev. |  | 0.6857% | 0.7854% | 1.675% | 10.27% | 1.011% | 2.368% | 3.103% | 1.934% | 1.862% | 4.183% | 1.899% | 3.313% | 11.21% | 1.333% | 1.038% | 26.4% | 4.156% | 5.51% | 4.537% | 1.781% | 1.969% | 5.193%
| PDM3 |  | 5.738% | -10.62% | -0.3153% | 8.825% | 4.341% | 10.57% | -5.692% | 1.233% | 3.771% | 0.8676% | -0.6916% | 5.008% | 40.12% | -4.841% | 1.513% | 2.534% | 7.105% | -6.363% | -4.584% | 0.9644% | 4.828% | -16.33%
| Z-Score (Absolute) |  | 0.74 | 2.02 | 0.28 | 1.26 | 0.5 | 1.55 | 1.19 | 0.02 | 0.4 | 0.09 | 0.35 | 0.61 | 6.53 | 1.05 | 0.02 | 0.2 | 0.97 | 1.3 | 1 | 0.07 | 0.58 | 2.98
| Analytical results for Re in OREAS 237b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | 0.003 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | 0.001 | < 0.002 | < 0.002 | < 0.0001 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.003 | < 0.001 | < 0.002 | < 0.002 | < 0.0001 | [10 empty] | 38
| 1 | 3 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.005 | 0.003 | < 0.002 | < 0.002 | 0.0002 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.005 | 0.003 | < 0.002 | < 0.002 | 0.0014 | [10 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.007 | 0.001 | < 0.002 | < 0.002 | [11 empty] | 51
| 1 | 6 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.003 | 0.001 | < 0.002 | < 0.002 | 0.0003
| Mean |  | 0.003 |  |  |  | 0.0036667 |  |  |  |  | 0.002 |  | 0.0041667 | 0.0018 |  |  | 0.00063333
| Median |  | 0.003 |  |  |  | 0.004 |  |  |  |  | 0.002 |  | 0.004 | 0.001 |  |  | 0.0003
| Std Dev. | [5 empty] | 0.0005164 |  |  |  |  | 0 |  | 0.0018348 | 0.0010954 |  |  | 0.00066583
| Rel.Std.Dev. | [5 empty] | 14.08% |  |  |  |  | 0% |  | 44.04% | 60.86% |  |  | 105.1%
PDM3
| Z-Score (Absolute) |  | 0.08 | 49.14 | 0.93 | 23.85 | 2.28 | 23.85 | 1.1 | 0.42 | 0.42 | 0.08 | 0.42 | 2.78 | 0.17 | 0.42 | 0.42 | 1.03
| Analytical results for Rh in OREAS 237b (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Rh | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0022 | [25 empty] | 1
| 1 | 2 | 0.0006 | [25 empty] | 13
| 1 | 3 | 0.0013 | [25 empty] | 16
| 1 | 4 | 0.0011 | [25 empty] | 0.0015
| 1 | 5 | 0.0021 | [25 empty] | 9
| 1 | 6 | 0.0017
| Mean |  | 0.0015
| Median |  | 0.0015
| Std Dev. |  | 0.00061644
| Rel.Std.Dev. |  | 41.1%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 237b (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Ru | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0021 | [25 empty] | 1
| 1 | 2 | < 0.0004 | [25 empty] | 14
| 1 | 3 | 0.0004 | [25 empty] | 16
| 1 | 4 | 0.0061 | [25 empty] | 0.0019
| 1 | 5 | < 0.0044 | [25 empty] | 10
| 1 | 6 | 0.0004
| Mean |  | 0.00225
| Median |  | 0.00125
| Std Dev. |  | 0.0026889
| Rel.Std.Dev. |  | 119.5%
| PDM3 |  | 18.42%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 237b (Certified Value 0.193 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.21 | 0.185 | 0.19 | 0.2 | 0.206 | 0.14 | 0.195 | 0.16 | 0.19 | 0.2 | 0.18 | 0.19 | 0.2029 | 0.198 | 0.19 | 0.15 | 0.2 | 0.2 | 0.18 | 0.17679 | 0.2203 | [5 empty] | 1
| 1 | 2 | 0.19 | 0.185 | 0.19 | 0.2 | 0.204 | 0.15 | 0.196 | 0.17 | 0.19 | 0.2 | 0.18 | 0.19 | 0.1967 | 0.193 | 0.19 | 0.15 | 0.19 | 0.2 | 0.18 | 0.17679 | 0.2131 | [5 empty] | 23
| 1 | 3 | 0.2 | 0.18 | 0.19 | 0.19 | 0.201 | 0.16 | 0.197 | 0.17 | 0.19 | 0.2 | 0.19 | 0.2 | 0.205 | 0.192 | 0.2 | 0.14 | 0.2 | 0.19 | 0.18 | 0.18383 | 0.196 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.185 | 0.19 | 0.2 | 0.204 | 0.17 | 0.193 | 0.17 | 0.18 | 0.2 | 0.19 | 0.19 | 0.2112 | 0.194 | 0.21 | 0.15 | 0.19 | 0.19 | 0.19 | 0.17451 | 0.1942 | [5 empty] | 0.19308
| 1 | 5 | 0.2 | 0.175 | 0.19 | 0.19 | 0.204 | 0.17 | 0.205 | 0.17 | 0.19 | 0.2 | 0.18 | 0.2 | 0.2009 | 0.189 | 0.21 | 0.14 | 0.19 | 0.2 | 0.18 | 0.18088 | 0.2177 | [5 empty] | 52
| 1 | 6 | 0.21 | 0.185 | 0.19 | 0.19 | 0.2 | 0.18 | 0.2 | 0.18 | 0.19 | 0.21 | 0.19 | 0.19 | 0.2163 | 0.195 | 0.19 | 0.15 | 0.19 | 0.2 | 0.18 | 0.18401 | 0.2077
| Mean |  | 0.20167 | 0.1825 | 0.19 | 0.195 | 0.20317 | 0.16167 | 0.19767 | 0.17 | 0.18833 | 0.20167 | 0.185 | 0.19333 | 0.2055 | 0.1935 | 0.19833 | 0.14667 | 0.19333 | 0.19667 | 0.18167 | 0.17947 | 0.20817
| Median |  | 0.2 | 0.185 | 0.19 | 0.195 | 0.204 | 0.165 | 0.1965 | 0.17 | 0.19 | 0.2 | 0.185 | 0.19 | 0.20395 | 0.1935 | 0.195 | 0.15 | 0.19 | 0.2 | 0.18 | 0.17884 | 0.2104
| Std Dev. |  | 0.0075277 | 0.0041833 | 3.0405e-17 | 0.0054772 | 0.0022286 | 0.01472 | 0.004274 | 0.0063246 | 0.0040825 | 0.0040825 | 0.0054772 | 0.005164 | 0.0071378 | 0.0030166 | 0.0098319 | 0.005164 | 0.005164 | 0.005164 | 0.0040825 | 0.0040118 | 0.011007
| Rel.Std.Dev. |  | 3.733% | 2.292% | 1.6e-14% | 2.809% | 1.097% | 9.105% | 2.162% | 3.72% | 2.168% | 2.024% | 2.961% | 2.671% | 3.473% | 1.559% | 4.957% | 3.521% | 2.671% | 2.626% | 2.247% | 2.235% | 5.288%
| PDM3 |  | 4.449% | -5.478% | -1.594% | 0.9958% | 5.225% | -16.27% | 2.377% | -11.95% | -2.457% | 4.449% | -4.184% | 0.1325% | 6.434% | 0.2189% | 2.722% | -24.04% | 0.1325% | 1.859% | -5.91% | -7.047% | 7.815%
| Z-Score (Absolute) |  | 0.67 | 0.88 | 0.27 | 0.13 | 0.8 | 2.56 | 0.35 | 1.89 | 0.4 | 0.67 | 0.67 | 0 | 0.98 | 0.01 | 0.4 | 3.78 | 0 | 0.27 | 0.94 | 1.12 | 1.2
| Analytical results for Sb in OREAS 237b (Certified Value 460 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 504.6 | 24.8 | 440.1025 | 486.8 | 459.7 | 406 | 429 | 451 | 521 | 442 | 456 | 464.73 | 480.7717 | 377.8 | 421 | 435.5 | 449.95 | 454.7 | 471.33 | 464.487 | 489 | 343.5159 | 461.096 |  |  |  | 1
| 1 | 2 | 490.69 | 19.8 | 436.8 | 479.3 | 463.13 | 401 | 421.7 | 455 | 503 | 464 | 467 | 466.66 | 473.2591 | 367.7 | 444 | 429.5 | 437.69 | 450.3 | 445.53 | 464.487 | 509 | 372.3246 | 467.1401 |  |  |  | 39
| 1 | 3 | 511.69 | 27.7 | 439.64 | 477.7 | 456.67 | 422 | 429.9 | 463 | 516 | 465 | 460 | 461.83 | 461.8144 | 379.8 | 452 | 435.5 | 442.49 | 437.4 | 451.27 | 475.508 | 497 | 369.5614 | 461.6028 |  |  |  | 16
| 1 | 4 | 497.19 | 22.6 | 445.39 | 491.5 | 472.68 | 406 | 410.7 | 450 | 506 | 473 | 453 | 466.34 | 450.4425 | 634.6 | 456 | 433.7 | 430.4 | 442.9 | 468.85 | 459.03 | 498 | 368.5764 | 458.9742 |  |  |  | 460.0851
| 1 | 5 | 499.47 | 14.7 | 444.48 | 491.2 | 467.35 | 422 | 430.7 | 471 | 509 | 452 | 482 | 472.17 | 451.8346 | 369 | 434 | 329.1 | 443.38 | 440.2 | 449.09 | 466.413 | 502 | 373.3992 | 463.9269 |  |  |  | 53
| 1 | 6 | 510.94 | 14.2 | 444.555 | 486.8 | 475.82 | 429 | 434.7 | 459 | 522 | 461 | 453 | 476.27 | 483.9035 | 378.2 | 438 | 426.7 | 434.4 | 451.6 | 460.44 | 492.949 | 489 | 374.2563 | 460.8442
| Mean |  | 502.43 | 20.6333 | 441.8279 | 485.55 | 465.8917 | 414.3333 | 426.1167 | 458.1667 | 512.8333 | 459.5 | 461.8333 | 468 | 467.0043 | 417.85 | 440.8333 | 415 | 439.7183 | 446.1833 | 457.7517 | 470.479 | 497.3333 | 366.939 | 462.264
| Median |  | 502.035 | 21.2 | 442.2912 | 486.8 | 465.24 | 414 | 429.45 | 457 | 512.5 | 462.5 | 458 | 466.5 | 467.5368 | 378 | 441 | 431.6 | 440.09 | 446.6 | 455.855 | 465.45 | 497.5 | 370.943 | 461.3494
| Std Dev. |  | 8.2088 | 5.4489 | 3.47 | 5.8504 | 7.4536 | 11.3959 | 8.6543 | 7.9603 | 7.9854 | 10.9316 | 11.1967 | 5.2744 | 14.461 | 106.3069 | 12.7502 | 42.2283 | 7.0001 | 6.9649 | 10.7813 | 12.2443 | 7.7115 | 11.6832 | 2.8699
| Rel.Std.Dev. |  | 1.634% | 26.41% | 0.7854% | 1.205% | 1.6% | 2.75% | 2.031% | 1.737% | 1.557% | 2.379% | 2.424% | 1.127% | 3.097% | 25.44% | 2.892% | 10.18% | 1.592% | 1.561% | 2.355% | 2.603% | 1.551% | 3.184% | 0.6208%
| PDM3 |  | 9.204% | -95.52% | -3.968% | 5.535% | 1.262% | -9.944% | -7.383% | -0.417% | 11.46% | -0.1272% | 0.38% | 1.72% | 1.504% | -9.18% | -4.184% | -9.799% | -4.427% | -3.022% | -0.5072% | 2.259% | 8.096% | -20.25% | 0.4736%
| Z-Score (Absolute) |  | 1.72 | 17.02 | 0.64 | 1.07 | 0.3 | 1.71 | 1.25 | 0 | 2.13 | 0.05 | 0.14 | 0.38 | 0.34 | 1.57 | 0.67 | 1.68 | 0.72 | 0.47 | 0.02 | 0.48 | 1.52 | 3.55 | 0.16
| Analytical results for Sc in OREAS 237b (Certified Value 14.2 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.3 | 13.2229 | 15 | 14 | 10.2253 | 15 | 13.4 | 14 | 14.3 | 14.7 | 15.2 | 14.6 | 14.3 | 12.9 | 13.9772 | 13.7 | 11.95 | 13.4 | 14.4 | 13.4 | 14.2 | 14.049 | 14.8845 | 10.6655 |  |  | 1
| 1 | 2 | 14.3 | 13.1406 | 15 | 14 | 11.1007 | 15 | 11.5 | 14 | 14.3 | 15.2 | 14.3 | 14.7 | 15.4 | 12.7 | 13.8643 | 14.4 | 11.96 | 14 | 14.2 | 12.4 | 14 | 14.1015 | 14.8273 | 11.5129 |  |  | 40
| 1 | 3 | 14.2 | 13.1473 | 15 | 14 | 11.21 | 14 | 12.9 | 14 | 14.4 | 14.5 | 15.5 | 15.1 | 15.2 | 13.1 | 13.9464 | 13.7 | 14.61 | 15 | 13.8 | 12.6 | 14.4 | 14.553 | 14.8273 | 10.3128 |  |  | 16
| 1 | 4 | 14.3 | 12.9844 | 16 | 14 | 11.9633 | 14 | 14.6 | 14 | 14.3 | 14.7 | 15.2 | 14.5 | 14.7 | 13 | 14.0476 | 14.1 | 12.9 | 14.9 | 14.3 | 13.4 | 14.5 | 13.9545 | 14.6559 | 11.2051 |  |  | 14.1501
| 1 | 5 | 14.5 | 12.9862 | 15 | 14 | 9.6153 | 14 | 14.6 | 14 | 14.4 | 15.1 | 14.1 | 14.7 | 14.9 | 13.1 | 13.7712 | 13.6 | 14.95 | 13.9 | 13.6 | 12.3 | 14.4 | 14.322 | 14.4845 | 11.114 |  |  | 54
| 1 | 6 | 14.4 | 12.9972 | 15 | 14 | 11.676 | 14 | 14.9 | 14 | 14.3 | 15.4 | 15.6 | 14.5 | 14.5 | 13.3 | 13.8439 | 13.8 | 12.85 | 13.3 | 13.9 | 14.5 | 14.4 | 14.385 | 14.9416 | 10.7892
| Mean |  | 14.3333 | 13.0798 | 15.1667 | 14 | 10.9651 | 14.3333 | 13.65 | 14 | 14.3333 | 14.9333 | 14.9833 | 14.6833 | 14.8333 | 13.0167 | 13.9084 | 13.8833 | 13.2033 | 14.0833 | 14.0333 | 13.1 | 14.3167 | 14.2275 | 14.7702 | 10.9333
| Median |  | 14.3 | 13.0689 | 15 | 14 | 11.1553 | 14 | 14 | 14 | 14.3 | 14.9 | 15.2 | 14.65 | 14.8 | 13.05 | 13.9054 | 13.75 | 12.875 | 13.95 | 14.05 | 13 | 14.4 | 14.2118 | 14.8273 | 10.9516
| Std Dev. |  | 0.10328 | 0.10336 | 0.40825 | 0 | 0.88874 | 0.5164 | 1.3126 | 0 | 0.05164 | 0.35024 | 0.63061 | 0.22286 | 0.41793 | 0.20412 | 0.10047 | 0.30605 | 1.2933 | 0.72503 | 0.31411 | 0.83905 | 0.18348 | 0.22889 | 0.1695 | 0.42895
| Rel.Std.Dev. |  | 0.7206% | 0.7902% | 2.692% | 0% | 8.105% | 3.603% | 9.616% | 0% | 0.3603% | 2.345% | 4.209% | 1.518% | 2.818% | 1.568% | 0.7224% | 2.204% | 9.795% | 5.148% | 2.238% | 6.405% | 1.282% | 1.609% | 1.148% | 3.923%
| PDM3 |  | 1.295% | -7.564% | 7.184% | -1.061% | -22.51% | 1.295% | -3.534% | -1.061% | 1.295% | 5.535% | 5.888% | 3.768% | 4.828% | -8.01% | -1.708% | -1.885% | -6.691% | -0.472% | -0.8253% | -7.421% | 1.177% | 0.5469% | 4.382% | -22.73%
| Z-Score (Absolute) |  | 0.36 | 1.28 | 1.45 | 0.08 | 4.04 | 0.36 | 0.53 | 0.08 | 0.36 | 1.14 | 1.21 | 0.82 | 1.01 | 1.36 | 0.2 | 0.23 | 1.12 | 0.03 | 0.03 | 1.25 | 0.34 | 0.22 | 0.93 | 4.08
| Analytical results for Se in OREAS 237b (Indicative Value 0.61 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 3 | < 1 | 1 | < 1 | < 1 | 0.8 | < 0.1 | < 2 | 0.6 | < 1 | 0.3 | 0.239 | 3.1861 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 5 | 0.5 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 1.3 | < 0.1 | < 2 | < 0.5 | < 1 | < 0.3 | 0.2247 | 3.4516 | [7 empty] | 41
| 1 | 3 | < 0.5 | < 5 | 0.6 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | < 2 | 1.1 | < 1 | < 0.3 | 0.2081 | 3.524 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 5 | 0.7 | < 5 | < 2 | < 1 | 1 | < 1 | 2 | < 1 | < 1 | 1.2 | < 0.1 | < 2 | < 0.5 | < 1 | 0.4 | 0.2378 | 3.0895 | [7 empty] | 0.60784
| 1 | 5 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | < 2 | < 0.5 | < 1 | 0.4 | 0.2485 | 3.352 | [7 empty] | 11
| 1 | 6 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.7 | < 0.1 | < 2 | 0.5 | < 1 | 0.4 | 0.1712 | 3.355
| Mean |  |  |  | 0.8 |  |  |  | 2 |  | 1.1667 |  |  | 0.96667 |  |  | 0.73333 |  | 0.375 | 0.22155 | 3.3264
| Median |  |  |  | 0.7 |  |  |  | 2 |  | 1 |  |  | 0.9 |  |  | 0.6 |  | 0.4 | 0.23125 | 3.3535
| Std Dev. |  |  |  | 0.28284 |  |  |  | 1.4142 |  | 0.40825 |  |  | 0.23381 |  |  | 0.32146 |  | 0.05 | 0.028358 | 0.16248
| Rel.Std.Dev. |  |  |  | 35.36% |  |  |  | 70.71% |  | 34.99% |  |  | 24.19% |  |  | 43.83% |  | 13.33% | 12.8% | 4.885%
| PDM3 |  |  |  | 31.61% |  |  |  | 229% |  | 91.94% |  |  | 59.03% |  |  | 20.65% |  | -38.31% | -63.55% | 447.2%
| Z-Score (Absolute) |  | 0.61 | 4.84 | 0.42 | 4.84 | 1.21 | 0 | 1.21 | 0 | 1.61 | 0 | 0 | 1.13 | 1.09 | 1.21 | 0.02 | 0 | 0.48 | 0.67 | 6.84
| Analytical results for Sm in OREAS 237b (Certified Value 6.61 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.4557 | 6.6 | 6.9 | 5.8 | 6.9 | 5.928 | 6.77 | 6.4 | 7.1816 | 6.3354 | [16 empty] | 1
| 1 | 2 | 6.4244 | 6.55 | 6.7 | 5.2 | 7 | 6.162 | 6.56 | 6.2 | 6.9184 | 6.5509 | [16 empty] | 10
| 1 | 3 | 6.2559 | 6.7 | 6.9 | 5.6 | 7 | 6.698 | 6.52 | 6.2 | 7.5012 | 6.1869 | [16 empty] | 16
| 1 | 4 | 6.4166 | 6.95 | 7.1 | 7.3 | 6.9 | 6.136 | 6.84 | 6.4 | 7.2756 | 6.499 | [16 empty] | 6.607
| 1 | 5 | 6.4429 | 6.7 | 7.2 | 6.2 | 6.8 | 6.375 | 6.4 | 6.4 | 7.1722 | 6.3896 | [16 empty] | 55
| 1 | 6 | 6.5051 | 6.95 | 7.1 | 6.2 | 7 | 6.146 | 6.93 | 6.4 | 7.473 | 6.2477
| Mean |  | 6.4168 | 6.7417 | 6.9833 | 6.05 | 6.9333 | 6.2408 | 6.67 | 6.3333 | 7.2537 | 6.3682
| Median |  | 6.4336 | 6.7 | 7 | 6 | 6.95 | 6.154 | 6.665 | 6.4 | 7.2286 | 6.3625
| Std Dev. |  | 0.084783 | 0.17151 | 0.18348 | 0.72042 | 0.08165 | 0.26498 | 0.20688 | 0.10328 | 0.21638 | 0.14099
| Rel.Std.Dev. |  | 1.321% | 2.544% | 2.627% | 11.91% | 1.178% | 4.246% | 3.102% | 1.631% | 2.983% | 2.214%
| PDM3 |  | -2.879% | 2.039% | 5.696% | -8.43% | 4.94% | -5.542% | 0.9539% | -4.142% | 9.788% | -3.613%
| Z-Score (Absolute) |  | 0.33 | 0.52 | 1.16 | 1.3 | 1.03 | 0.8 | 0.33 | 0.55 | 1.87 | 0.46
| Analytical results for Sn in OREAS 237b (Certified Value 3.61 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.8 | 3.7 | 2 | < 10 | 3.7 | 3.5 | 3.5 | 4.1 | 3.4 | 3.7 | 3.8 | 3.4 | 3.6 | 3.7471 | 3.63 | 3.33 | 3.5 | < 5 | 3.7 | 3.6 | 3.3 | 2.8853 | < 5 | 3.2861 | 3.7682 |  | 1
| 1 | 2 | 3.8 | 3 | 2 | < 10 | 3.8 | 3.6 | 2.8 | 3.7 | 3.4 | 3.6 | 3.7 | 3.5 | 3.3 | 3.7979 | 3.67 | 3.357 | 3.8 | < 5 | 3.7 | 3.5 | 3.3 | 2.735 | < 5 | 3.4759 | 3.7557 |  | 26
| 1 | 3 | 3.8 | 3 | 2 | < 10 | 3.8 | 4 | 3.5 | 4 | 3.4 | 3.8 | 3.8 | 3.6 | 3.6 | 3.714 | 3.62 | 3.521 | 3.9 | < 5 | 3.8 | 3.4 | 3.3 | 2.9139 | < 5 | 3.2972 | 3.7573 |  | 16
| 1 | 4 | 3.8 | 3.1 | 2 | < 10 | 3.8 | 4.4 | 3.5 | 3.7 | 3.4 | 3.8 | 3.8 | 3.5 | 3.8 | 3.6811 | 3.86 | 5.98 | 4 | < 5 | 3.7 | 3.5 | 3.4 | 2.7686 | < 5 | 3.2628 | 3.6726 |  | 3.6117
| 1 | 5 | 3.8 | 3 | 2 | < 10 | 3.7 | 4.1 | 4.9 | 4 | 3.5 | 3.6 | 3.8 | 3.6 | 3.8 | 3.8431 | 3.84 | 3.518 | 3.7 | < 5 | 3.7 | 3.4 | 3.3 | 2.8516 | < 5 | 3.3571 | 3.7194 |  | 56
| 1 | 6 | 3.6 | 3.4 | 2 | < 10 | 3.7 | 3.7 | 3.8 | 4.1 | 3.4 | 4 | 3.8 | 3.4 | 3.5 | 3.5641 | 3.73 | 3.427 | 3.6 | < 5 | 3.7 | 3.8 | 3.3 | 2.7941 | < 5 | 3.3893 | 3.7555
| Mean |  | 3.7667 | 3.2 | 2 |  | 3.75 | 3.8833 | 3.6667 | 3.9333 | 3.4167 | 3.75 | 3.7833 | 3.5 | 3.6 | 3.7246 | 3.725 | 3.8555 | 3.75 |  | 3.7167 | 3.5333 | 3.3167 | 2.8247 |  | 3.3447 | 3.7381
| Median |  | 3.8 | 3.05 | 2 |  | 3.75 | 3.85 | 3.5 | 4 | 3.4 | 3.75 | 3.8 | 3.5 | 3.6 | 3.7305 | 3.7 | 3.4725 | 3.75 |  | 3.7 | 3.5 | 3.3 | 2.8228 |  | 3.3272 | 3.7556
| Std Dev. |  | 0.08165 | 0.28983 | 0 |  | 0.054772 | 0.34303 | 0.68896 | 0.18619 | 0.040825 | 0.15166 | 0.040825 | 0.089443 | 0.18974 | 0.097642 | 0.10445 | 1.0438 | 0.18708 |  | 0.040825 | 0.15055 | 0.040825 | 0.069975 |  | 0.079692 | 0.036123
| Rel.Std.Dev. |  | 2.168% | 9.057% | 0% |  | 1.461% | 8.833% | 18.79% | 4.734% | 1.195% | 4.044% | 1.079% | 2.556% | 5.27% | 2.622% | 2.804% | 27.07% | 4.989% |  | 1.098% | 4.261% | 1.231% | 2.477% |  | 2.383% | 0.9663%
| PDM3 |  | 4.29% | -11.4% | -44.62% |  | 3.828% | 7.52% | 1.521% | 8.904% | -5.401% | 3.828% | 4.751% | -3.094% | -0.3249% | 3.124% | 3.136% | 6.749% | 3.828% |  | 2.905% | -2.171% | -8.17% | -21.79% |  | -7.393% | 3.499%
| Z-Score (Absolute) |  | 0.21 | 2.36 | NA | 5.8 | 0.13 | 0.74 | 0.24 | 0.96 | 1.38 | 0.13 | 0.28 | 1 | 0.55 | 0.02 | 0.02 | 0.61 | 0.13 | 5.53 | 0.02 | 0.85 | 1.83 | 4.06 | 5.53 | 1.7 | 0.08
| Analytical results for Sr in OREAS 237b (Certified Value 123 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 125.54 | 127.351 | 123 | 111 | 118.6328 | 126.7 | 121 | 122 | 121 | 122.5 | 127.5 | 122.5 | 121.5 | 126 | 121.4116 | 123 | 111.2 | 119 | 117.6 | 122.3 | 126.6 | 118 | 139.3 | 132 | 131.9919 | 116.7203 | 1
| 1 | 2 | 124.96 | 126.8806 | 124 | 126 | 118.2316 | 124.9 | 115 | 117 | 120 | 124 | 121.5 | 124 | 123 | 124 | 122.1181 | 124 | 111.5 | 117 | 119 | 118.8 | 123.8 | 118 | 141.5 | 134 | 124.7242 | 118.2081 | 27
| 1 | 3 | 127.67 | 128.0076 | 122 | 126 | 120.7393 | 124.7 | 120 | 125 | 120 | 122 | 125.5 | 129 | 126 | 127 | 119.0571 | 123 | 115.4 | 120 | 118.6 | 123.2 | 123.6 | 120 | 144.6 | 131 | 124.7747 | 116.7202 | 16
| 1 | 4 | 127.39 | 128.0664 | 133 | 131 | 120.9787 | 126.2 | 129 | 119 | 120 | 121.5 | 125 | 124.5 | 122.5 | 124 | 119.025 | 122 | 194.3 | 126 | 118.5 | 124.6 | 123.6 | 122 | 138.7 | 131 | 121.797 | 117.7123 | 123.2072
| 1 | 5 | 127.39 | 126.91 | 123 | 129 | 118.9605 | 126 | 128 | 115 | 120 | 123 | 120 | 124.5 | 123.5 | 125 | 118.9155 | 122 | 114.1 | 120 | 119.8 | 120.9 | 123.3 | 120 | 142.5 | 132 | 122.9073 | 115.7766 | 57
| 1 | 6 | 125.23 | 127.057 | 127 | 129 | 117.7927 | 126.6 | 130 | 125 | 120 | 124.5 | 129 | 125 | 120.5 | 127 | 119.1993 | 126 | 114.4 | 126 | 117.4 | 119.2 | 125.7 | 123 | 144.4 | 129 | 125.2289 | 116.672
| Mean |  | 126.3633 | 127.3788 | 125.3333 | 125.3333 | 119.2226 | 125.85 | 123.8333 | 120.5 | 120.1667 | 122.9167 | 124.75 | 124.9167 | 122.8333 | 125.5 | 119.9544 | 123.3333 | 126.8167 | 121.3333 | 118.4833 | 121.5 | 124.4333 | 120.1667 | 141.8333 | 131.5 | 125.2373 | 116.9683
| Median |  | 126.465 | 127.204 | 123.5 | 127.5 | 118.7967 | 126.1 | 124.5 | 120.5 | 120 | 122.75 | 125.25 | 124.5 | 122.75 | 125.5 | 119.1282 | 123 | 114.25 | 120 | 118.55 | 121.6 | 123.7 | 120 | 142 | 131.5 | 124.7494 | 116.7202
| Std Dev. |  | 1.2448 | 0.53675 | 4.1312 | 7.2847 | 1.3287 | 0.85499 | 6.047 | 4.1833 | 0.40825 | 1.1583 | 3.446 | 2.1775 | 1.8886 | 1.3784 | 1.4229 | 1.5055 | 33.102 | 3.7771 | 0.89088 | 2.2839 | 1.3692 | 2.0412 | 2.4913 | 1.6432 | 3.56 | 0.86274
| Rel.Std.Dev. |  | 0.9851% | 0.4214% | 3.296% | 5.812% | 1.114% | 0.6794% | 4.883% | 3.472% | 0.3397% | 0.9423% | 2.762% | 1.743% | 1.537% | 1.098% | 1.186% | 1.221% | 26.1% | 3.113% | 0.7519% | 1.88% | 1.1% | 1.699% | 1.757% | 1.25% | 2.843% | 0.7376%
| PDM3 |  | 2.562% | 3.386% | 1.726% | 1.726% | -3.234% | 2.145% | 0.5082% | -2.197% | -2.468% | -0.2358% | 1.252% | 1.387% | -0.3035% | 1.861% | -2.64% | 0.1023% | 2.93% | -1.521% | -3.834% | -1.386% | 0.9951% | -2.468% | 15.12% | 6.731% | 1.648% | -5.064%
| Z-Score (Absolute) |  | 0.62 | 0.9 | 0.33 | 0.33 | 1.36 | 0.48 | 0.08 | 1.01 | 1.1 | 0.34 | 0.17 | 0.22 | 0.36 | 0.38 | 1.16 | 0.22 | 0.74 | 0.78 | 1.57 | 0.73 | 0.08 | 1.1 | 4.91 | 2.04 | 0.31 | 1.99
| Analytical results for Ta in OREAS 237b (Certified Value 1.03 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.07 | 0.9446 | < 0.1 | < 5 | 1.03 | 1.06 | 0.84 | 0.8 | 1 | 1.08 | 1.05 | 1.01 | 1.1 | 1.1439 | 1.01 | 0.96 | 1.17 | 1.06 | 0.9 | 0.6681 | 1.1751 | [5 empty] | 1
| 1 | 2 | 1.09 | 0.90428 | < 0.1 | < 5 | 1.03 | 0.97 | 0.88 | 0.8 | 1.02 | 0.99 | 1.01 | 1 | 0.9 | 1.1734 | 1.02 | 0.94 | 1.16 | 1.05 | 0.8 | 0.624 | 1.1964 | [5 empty] | 28
| 1 | 3 | 1.1 | 0.86143 | < 0.1 | < 5 | 1 | 1.04 | 0.96 | 0.7 | 1.04 | 1.08 | 1.06 | 1.06 | 1 | 1.13 | 1 | 0.98 | 1.2 | 1.04 | 0.8 | 0.6885 | 1.1924 | [5 empty] | 16
| 1 | 4 | 1.06 | 0.89853 | < 0.1 | 5.4133 | 1.01 | 1.01 | 0.89 | 0.7 | 1.01 | 1.07 | 1.03 | 1.01 | 1.1 | 1.1577 | 1.01 | 0.99 | 1.18 | 1.04 | 0.8 | 0.6581 | 1.181 | [5 empty] | 1.0256
| 1 | 5 | 1.05 | 0.93342 | < 0.1 | < 5 | 1 | 1 | 0.83 | 0.7 | 1.04 | 1 | 1.02 | 1.02 | 1 | 1.1571 | 0.99 | 0.99 | 1.18 | 1.22 | 0.8 | 0.6346 | 1.1816 | [5 empty] | 58
| 1 | 6 | 1.11 | 0.88894 | < 0.1 | < 5 | 1.01 | 1.1 | 1.11 | 0.7 | 1.04 | 1.12 | 1.01 | 0.99 | 1 | 1.152 | 1.01 | 0.96 | 1.16 | 1.15 | 0.9 | 0.6259 | 1.1918
| Mean |  | 1.08 | 0.9052 |  | 5.4133 | 1.0133 | 1.03 | 0.91833 | 0.73333 | 1.025 | 1.0567 | 1.03 | 1.015 | 1.0167 | 1.1523 | 1.0067 | 0.97 | 1.175 | 1.0933 | 0.83333 | 0.64987 | 1.1864
| Median |  | 1.08 | 0.90141 |  | 5.4133 | 1.01 | 1.025 | 0.885 | 0.7 | 1.03 | 1.075 | 1.025 | 1.01 | 1 | 1.1546 | 1.01 | 0.97 | 1.175 | 1.055 | 0.8 | 0.64635 | 1.1867
| Std Dev. |  | 0.023664 | 0.030247 |  |  | 0.013663 | 0.046476 | 0.10458 | 0.05164 | 0.017607 | 0.050859 | 0.020976 | 0.02429 | 0.075277 | 0.014595 | 0.010328 | 0.02 | 0.015166 | 0.074744 | 0.05164 | 0.025959 | 0.0083081
| Rel.Std.Dev. |  | 2.191% | 3.341% |  |  | 1.348% | 4.512% | 11.39% | 7.042% | 1.718% | 4.813% | 2.037% | 2.393% | 7.404% | 1.267% | 1.026% | 2.062% | 1.291% | 6.836% | 6.197% | 3.995% | 0.7003%
| PDM3 |  | 5.308% | -11.74% |  | 427.8% | -1.193% | 0.4324% | -10.46% | -28.49% | -0.05516% | 3.033% | 0.4324% | -1.03% | -0.8677% | 12.36% | -1.843% | -5.418% | 14.57% | 6.608% | -18.74% | -36.63% | 15.68%
| Z-Score (Absolute) |  | 0.56 | 0.98 | 8.5 | 17.32 | 0.03 | 0.12 | 0.86 | 2.49 | 0.07 | 0.35 | 0.12 | 0.01 | 0 | 1.19 | 0.09 | 0.41 | 1.39 | 0.67 | 1.61 | 3.23 | 1.49
| Analytical results for Tb in OREAS 237b (Certified Value 0.66 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.72956 | 0.7 | 0.7 | 0.6 | 0.66 | 0.639 | 0.61 | 0.72 | 0.6 | 1.38 | 5.3494 | [15 empty] | 1
| 1 | 2 | 0.69575 | 0.68 | 0.7 | 0.6 | 0.72 | 0.67 | 0.61 | 0.68 | 0.6 | 2.28 | 5.6277 | [15 empty] | 11
| 1 | 3 | 0.68441 | 0.68 | 0.7 | 0.6 | 0.7 | 0.769 | 0.62 | 0.74 | 0.5 | 2.22 | 5.3702 | [15 empty] | 16
| 1 | 4 | 0.69708 | 0.72 | 0.7 | 0.7 | 0.68 | 0.648 | 0.64 | 0.71 | 0.6 | 2.21 | 5.4806 | [15 empty] | 0.66362
| 1 | 5 | 0.65794 | 0.68 | 0.7 | 0.6 | 0.66 | 0.66 | 0.63 | 0.71 | 0.6 | 2.27 | 5.4989 | [15 empty] | 59
| 1 | 6 | 0.70497 | 0.7 | 0.7 | 0.6 | 0.68 | 0.671 | 0.61 | 0.68 | 0.6 | 2.28 | 5.3037
| Mean |  | 0.69495 | 0.69333 | 0.7 | 0.61667 | 0.68333 | 0.67617 | 0.62 | 0.70667 | 0.58333 | 2.1067 | 5.4384
| Median |  | 0.69642 | 0.69 | 0.7 | 0.6 | 0.68 | 0.665 | 0.615 | 0.71 | 0.6 | 2.245 | 5.4254
| Std Dev. |  | 0.023595 | 0.01633 | 1.2162e-16 | 0.040825 | 0.023381 | 0.047157 | 0.012649 | 0.023381 | 0.040825 | 0.3573 | 0.12001
| Rel.Std.Dev. |  | 3.395% | 2.355% | 1.737e-14% | 6.62% | 3.422% | 6.974% | 2.04% | 3.309% | 6.999% | 16.96% | 2.207%
| PDM3 |  | 4.722% | 4.478% | 5.483% | -7.075% | 2.971% | 1.891% | -6.573% | 6.487% | -12.1% | 217.5% | 719.5%
| Z-Score (Absolute) |  | 0.06 | 0 | 0.26 | 3.01 | 0.39 | 0.67 | 2.88 | 0.52 | 4.32 | 55.52 | 186.39
| Analytical results for Te in OREAS 237b (Indicative Value 0.062 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | < 0.1 | 8.346 | < 0.1 | 0.12 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.03102 | [8 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 8.0707 | < 0.1 | 0.11 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.1 | 0.07 | < 0.05 | 0.01496 | [8 empty] | 6
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 8.922 | < 0.1 | 0.08 | < 0.1 | 0.2 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.01 | [8 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 9.086 | < 0.1 | 0.09 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.02995 | [8 empty] | 0.061952
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 9.038 | < 0.1 | 0.07 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.01507 | [8 empty] | 12
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 9.0507 | < 0.1 | 0.08 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.06 | 0.07 | < 0.05 | 0.021
| Mean |  |  |  |  | 8.7522 |  | 0.091667 | 0.2 | 0.2 |  |  |  | 0.056 |  |  | 0.08 | 0.066667 |  | 0.020333
| Median |  |  |  |  | 8.98 |  | 0.085 | 0.2 | 0.2 |  |  |  | 0.06 |  |  | 0.08 | 0.07 |  | 0.018035
| Std Dev. |  |  |  |  | 0.43371 |  | 0.019408 | 0 |  |  |  |  | 0.0054772 |  |  | 0.028284 | 0.005164 |  | 0.0086079
| Rel.Std.Dev. |  |  |  |  | 4.955% |  | 21.17% | 0% |  |  |  |  | 9.781% |  |  | 35.36% | 7.746% |  | 42.33%
| PDM3 |  |  |  |  | 14027% |  | 47.96% | 222.8% | 222.8% |  |  |  | -9.608% |  |  | 29.13% | 7.61% |  | -67.18%
| Z-Score (Absolute) |  | 1.35 | 1.35 | 0 | 234.72 | 0 | 1.12 | 1.35 | 0.67 | 0.67 | 0.67 | 0.67 | 0.02 | 0.67 | 0.67 | 0.36 | 0.45 | 0.67 | 0.8
| Analytical results for Th in OREAS 237b (Certified Value 14.5 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.99 | 15.3681 | 17.2 | 16 | 14.84 | 12.59 | 15.1 | 14.8 | 14.6 | 14.95 | 14.25 | 15.2 | 16.1 | 13.8097 | 13.5 | 13.94 | 14.1 | 14.2 | 15.5 | 14.3 | 13.6607 | 15.2048 | 13.4572 |  |  |  | 1
| 1 | 2 | 14.53 | 15.0728 | 17 | 15.6 | 14.61 | 12.64 | 14.6 | 14.4 | 14.85 | 13.55 | 13.85 | 15.5 | 15 | 14.7684 | 13.9 | 13.95 | 13.6 | 13.8 | 14.9 | 13.5 | 13.0335 | 14.9092 | 14.0628 |  |  |  | 31
| 1 | 3 | 15.01 | 15.0981 | 17.7 | 15.5 | 14.6 | 13.55 | 14.8 | 15 | 14.8 | 14.3 | 14.75 | 15.7 | 15.3 | 14.1177 | 13.6 | 15.32 | 13.8 | 14.3 | 14.9 | 13.8 | 13.7328 | 14.791 | 13.9062 |  |  |  | 16
| 1 | 4 | 14.84 | 14.9256 | 17.7 | 15.4 | 14.86 | 14.78 | 14.6 | 14.3 | 14.75 | 15 | 14.35 | 15.15 | 21 | 14.207 | 13.6 | 14.03 | 14.8 | 14.3 | 15.1 | 14 | 13.3889 | 14.7319 | 13.9284 |  |  |  | 14.4502
| 1 | 5 | 14.61 | 15.0231 | 17.6 | 15.2 | 14.79 | 14.55 | 14.4 | 14.5 | 14.8 | 13.6 | 14.05 | 15.05 | 17.4 | 13.9575 | 13.5 | 14.17 | 14.5 | 13.6 | 15.2 | 13.8 | 13.474 | 14.7713 | 14.1904 |  |  |  | 60
| 1 | 6 | 15.04 | 15.0942 | 17.2 | 15.6 | 14.79 | 14.93 | 15.4 | 14.6 | 14.95 | 14.25 | 14.2 | 14.85 | 18.2 | 14.5345 | 13.6 | 13.91 | 13.7 | 13.9 | 16.4 | 14.5 | 13.4395 | 14.8698 | 14.0532
| Mean |  | 14.8367 | 15.097 | 17.4 | 15.55 | 14.7483 | 13.84 | 14.8167 | 14.6 | 14.7917 | 14.275 | 14.2417 | 15.2417 | 17.1667 | 14.2325 | 13.6167 | 14.22 | 14.0833 | 14.0167 | 15.3333 | 13.9833 | 13.4549 | 14.8797 | 13.933
| Median |  | 14.915 | 15.0835 | 17.4 | 15.55 | 14.79 | 14.05 | 14.7 | 14.55 | 14.8 | 14.275 | 14.225 | 15.175 | 16.75 | 14.1624 | 13.6 | 13.99 | 13.95 | 14.05 | 15.15 | 13.9 | 13.4567 | 14.8304 | 13.9908
| Std Dev. |  | 0.21924 | 0.1476 | 0.30332 | 0.26646 | 0.11444 | 1.0642 | 0.37103 | 0.26077 | 0.11583 | 0.6267 | 0.304 | 0.3089 | 2.242 | 0.35961 | 0.1472 | 0.54699 | 0.47924 | 0.29269 | 0.56804 | 0.3656 | 0.24583 | 0.17212 | 0.25484
| Rel.Std.Dev. |  | 1.478% | 0.9777% | 1.743% | 1.714% | 0.776% | 7.689% | 2.504% | 1.786% | 0.7831% | 4.39% | 2.135% | 2.027% | 13.06% | 2.527% | 1.081% | 3.847% | 3.403% | 2.088% | 3.705% | 2.615% | 1.827% | 1.157% | 1.829%
| PDM3 |  | 2.675% | 4.476% | 20.41% | 7.611% | 2.063% | -4.223% | 2.536% | 1.037% | 2.363% | -1.212% | -1.443% | 5.477% | 18.8% | -1.507% | -5.768% | -1.593% | -2.539% | -3% | 6.112% | -3.231% | -6.888% | 2.972% | -3.579%
| Z-Score (Absolute) |  | 0.31 | 0.65 | 3.65 | 1.24 | 0.19 | 0.99 | 0.28 | 0 | 0.25 | 0.42 | 0.47 | 0.84 | 3.35 | 0.48 | 1.28 | 0.5 | 0.67 | 0.76 | 0.96 | 0.8 | 1.49 | 0.36 | 0.87
| Analytical results for Ti in OREAS 237b (Certified Value 0.456 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.456 | 0.4411 | 0.435 | 0.371 | 0.34 | 0.502 | 0.44 | 0.428 | 0.462 | 0.447 | 0.465 | 0.443 | 0.448 | 0.4651 | 0.46189 | 0.5144 | 0.45 | 0.45 | 0.45 | 0.48 | 0.449 | 0.37436 | 0.4828 | 0.4294 | 0.47497 |  | 1
| 1 | 2 | 0.4594 | 0.4395 | 0.44 | 0.401 | 0.34667 | 0.492 | 0.45 | 0.431 | 0.466 | 0.45 | 0.448 | 0.451 | 0.451 | 0.4559 | 0.46339 | 0.5144 | 0.44 | 0.46 | 0.44 | 0.49 | 0.44 | 0.3764 | 0.4892 | 0.4403 | 0.4627 |  | 32
| 1 | 3 | 0.4646 | 0.44075 | 0.44 | 0.378 | 0.35 | 0.494 | 0.43 | 0.434 | 0.47 | 0.451 | 0.456 | 0.455 | 0.455 | 0.4641 | 0.46287 | 0.5132 | 0.45 | 0.46 | 0.46 | 0.49 | 0.438 | 0.38621 | 0.4801 | 0.4332 | 0.46735 |  | 16
| 1 | 4 | 0.4466 | 0.44025 | 0.43 | 0.399 | 0.35667 | 0.502 | 0.46 | 0.433 | 0.462 | 0.451 | 0.45 | 0.459 | 0.446 | 0.4641 | 0.45722 | 0.512 | 0.46 | 0.45 | 0.46 | 0.49 | 0.451 | 0.37348 | 0.4806 | 0.4347 | 0.47144 |  | 0.45585
| 1 | 5 | 0.4542 | 0.43756 | 0.435 | 0.443 | 0.35333 | 0.492 | 0.44 | 0.445 | 0.463 | 0.451 | 0.447 | 0.444 | 0.449 | 0.4661 | 0.46365 | 0.509 | 0.45 | 0.46 | 0.45 | 0.48 | 0.448 | 0.38154 | 0.4811 | 0.4259 | 0.46434 |  | 61
| 1 | 6 | 0.4595 | 0.43823 | 0.44 | 0.408 | 0.34667 | 0.488 | 0.45 | 0.463 | 0.462 | 0.452 | 0.47 | 0.45 | 0.445 | 0.4692 | 0.46195 | 0.5257 | 0.46 | 0.45 | 0.45 | 0.49 | 0.439 | 0.38741 | 0.4758 | 0.4378 | 0.45293
| Mean |  | 0.45672 | 0.43956 | 0.43667 | 0.4 | 0.34889 | 0.495 | 0.445 | 0.439 | 0.46417 | 0.45033 | 0.456 | 0.45033 | 0.449 | 0.46408 | 0.46183 | 0.51478 | 0.45167 | 0.455 | 0.45167 | 0.48667 | 0.44417 | 0.3799 | 0.4816 | 0.43355 | 0.46562
| Median |  | 0.4577 | 0.43988 | 0.4375 | 0.4 | 0.34833 | 0.493 | 0.445 | 0.4335 | 0.4625 | 0.451 | 0.453 | 0.4505 | 0.4485 | 0.4646 | 0.46241 | 0.5138 | 0.45 | 0.455 | 0.45 | 0.49 | 0.444 | 0.37897 | 0.48085 | 0.43395 | 0.46584
| Std Dev. |  | 0.0061039 | 0.0014168 | 0.0040825 | 0.025456 | 0.0058373 | 0.0057619 | 0.010488 | 0.0131 | 0.0032506 | 0.0017512 | 0.0095708 | 0.006186 | 0.0036332 | 0.0044328 | 0.0023688 | 0.0057126 | 0.0075277 | 0.0054772 | 0.0075277 | 0.005164 | 0.0057764 | 0.0060494 | 0.0043895 | 0.0053069 | 0.0076872
| Rel.Std.Dev. |  | 1.336% | 0.3223% | 0.9349% | 6.364% | 1.673% | 1.164% | 2.357% | 2.984% | 0.7003% | 0.3889% | 2.099% | 1.374% | 0.8092% | 0.9552% | 0.5129% | 1.11% | 1.667% | 1.204% | 1.667% | 1.061% | 1.301% | 1.592% | 0.9114% | 1.224% | 1.651%
| PDM3 |  | 0.1903% | -3.572% | -4.208% | -12.25% | -23.46% | 8.589% | -2.38% | -3.696% | 1.825% | -1.21% | 0.0331% | -1.21% | -1.502% | 1.806% | 1.312% | 12.93% | -0.9175% | -0.1863% | -0.9175% | 6.76% | -2.563% | -16.66% | 5.649% | -4.892% | 2.144%
| Z-Score (Absolute) |  | 0.27 | 0.66 | 0.81 | 2.81 | 5.58 | 2.35 | 0.36 | 0.69 | 0.68 | 0.07 | 0.24 | 0.07 | 0.14 | 0.67 | 0.55 | 3.43 | 0 | 0.18 | 0 | 1.9 | 0.41 | 3.9 | 1.63 | 0.98 | 0.76
| Analytical results for Tl in OREAS 237b (Certified Value 0.84 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.88 | 0.79543 | 0.8 | 0.84 | 0.85 | 1.01 | 0.81 | 0.9 | 0.87 | 0.82 | 0.83 | 0.79 | 0.83 | 0.88325 | 0.773 | 0.8 | 0.88 | 0.81 | 0.81 | 0.94105 | 0.8687 | 0.86358 |  |  |  |  | 1
| 1 | 2 | 0.9 | 0.77884 | 0.8 | 0.9 | 0.83 | 0.98 | 0.71 | 0.9 | 0.85 | 0.77 | 0.81 | 0.8 | 0.73 | 0.8446 | 0.771 | 0.81 | 0.82 | 0.76 | 0.78 | 0.91454 | 0.9108 | 0.87214 |  |  |  |  | 33
| 1 | 3 | 0.86 | 0.77949 | 0.8 | 0.89 | 0.84 | 0.91 | 0.84 | 0.9 | 0.88 | 0.83 | 0.78 | 0.82 | 0.81 | 0.85457 | 0.799 | 0.8 | 0.88 | 0.77 | 0.79 | 0.95909 | 0.8779 | 0.8783 |  |  |  |  | 16
| 1 | 4 | 0.87 | 0.78647 | 0.8 | 0.9 | 0.86 | 1.01 | 0.71 | 0.8 | 0.82 | 0.81 | 0.79 | 0.79 | 0.86 | 0.87264 | 1.342 | 0.83 | 0.86 | 0.82 | 0.82 | 0.92411 | 0.887 | 0.86665 |  |  |  |  | 0.84289
| 1 | 5 | 0.87 | 0.83916 | 0.8 | 0.88 | 0.84 | 0.96 | 0.76 | 0.8 | 0.88 | 0.79 | 0.83 | 0.81 | 0.82 | 0.81915 | 0.78 | 0.81 | 0.85 | 0.74 | 0.8 | 0.93335 | 0.9047 | 0.87022 |  |  |  |  | 62
| 1 | 6 | 0.88 | 0.84346 | 0.8 | 0.91 | 0.84 | 0.92 | 0.76 | 0.8 | 0.85 | 0.85 | 0.82 | 0.79 | 0.81 | 0.90513 | 0.782 | 0.79 | 0.84 | 0.88 | 0.82 | 0.9438 | 0.837 | 0.88113
| Mean |  | 0.87667 | 0.80381 | 0.8 | 0.88667 | 0.84333 | 0.965 | 0.765 | 0.85 | 0.85833 | 0.81167 | 0.81 | 0.8 | 0.81 | 0.86322 | 0.8745 | 0.80667 | 0.855 | 0.79667 | 0.80333 | 0.93599 | 0.88102 | 0.872
| Median |  | 0.875 | 0.79095 | 0.8 | 0.895 | 0.84 | 0.97 | 0.76 | 0.85 | 0.86 | 0.815 | 0.815 | 0.795 | 0.815 | 0.8636 | 0.781 | 0.805 | 0.855 | 0.79 | 0.805 | 0.9372 | 0.88245 | 0.87118
| Std Dev. |  | 0.013663 | 0.029693 | 1.2162e-16 | 0.025033 | 0.010328 | 0.043243 | 0.05244 | 0.054772 | 0.023166 | 0.028577 | 0.020976 | 0.012649 | 0.043359 | 0.030365 | 0.22924 | 0.013663 | 0.023452 | 0.050859 | 0.01633 | 0.015674 | 0.026768 | 0.0067194
| Rel.Std.Dev. |  | 1.558% | 3.694% | 1.52e-14% | 2.823% | 1.225% | 4.481% | 6.855% | 6.444% | 2.699% | 3.521% | 2.59% | 1.581% | 5.353% | 3.518% | 26.21% | 1.694% | 2.743% | 6.384% | 2.033% | 1.675% | 3.038% | 0.7706%
| PDM3 |  | 4.007% | -4.637% | -5.089% | 5.194% | 0.0525% | 14.49% | -9.241% | 0.8434% | 1.832% | -3.704% | -3.902% | -5.089% | -3.902% | 2.412% | 3.75% | -4.298% | 1.437% | -5.484% | -4.693% | 11.05% | 4.523% | 3.454%
| Z-Score (Absolute) |  | 0.47 | 0.81 | NA | 0.64 | 0.12 | 2.02 | 1.5 | 0 | 0.15 | 0.67 | 0.7 | 0.88 | 0.7 | 0.23 | 0.43 | 0.76 | 0.09 | 0.94 | 0.82 | 1.51 | 0.55 | 0.39
Results from laboratory 1.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 237b (Certified Value 0.25 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.23298 | 0.3 | 0.3 | 0.3 | 0.217 | 0.28 | 0.2 | 0.2093 | 0.424 | [17 empty] | 1
| 1 | 2 | 0.23396 | 0.3 | 0.3 | 0.2 | 0.228 | 0.29 | 0.2 | 0.1937 | 0.4371 | [17 empty] | 13
| 1 | 3 | 0.23932 | 0.3 | 0.3 | 0.2 | 0.326 | 0.3 | 0.2 | 0.2075 | 0.41491 | [17 empty] | 16
| 1 | 4 | 0.25167 | 0.3 | 0.3 | 0.3 | 0.214 | 0.28 | 0.2 | 0.208 | 0.42896 | [17 empty] | 0.25243
| 1 | 5 | 0.24242 | 0.3 | 0.3 | 0.3 | 0.223 | 0.28 | 0.2 | 0.2086 | 0.42394 | [17 empty] | 63
| 1 | 6 | 0.23323 | 0.3 | 0.3 | 0.3 | 0.216 | 0.29 | 0.2 | 0.2043 | 0.41353
| Mean |  | 0.23893 | 0.3 | 0.3 | 0.26667 | 0.23733 | 0.28667 | 0.2 | 0.20523 | 0.42374
| Median |  | 0.23664 | 0.3 | 0.3 | 0.3 | 0.22 | 0.285 | 0.2 | 0.20775 | 0.42397
| Std Dev. |  | 0.0073112 | 0 | 0 | 0.05164 | 0.043743 | 0.008165 | 3.0405e-17 | 0.0059085 | 0.0088098
| Rel.Std.Dev. |  | 3.06% | 0% | 0% | 19.36% | 18.43% | 2.848% | 1.52e-14% | 2.879% | 2.079%
| PDM3 |  | -5.346% | 18.85% | 18.85% | 5.642% | -5.979% | 13.56% | -20.77% | -18.7% | 67.87%
| Z-Score (Absolute) |  | 0.56 | 0.67 | 0.67 | 0 | 0.59 | 0.4 | 1.35 | 1.24 | 3.18
| Analytical results for U in OREAS 237b (Certified Value 2.73 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.82 | 2.8816 | 2.8 | 2.9 | 2.83 | 2.6 | 2.82 | 2.8 | 2.5 | 2.7 | 2.8 | 2.6 | 2.8 | 2.6847 | 2.67 | 2.505 | 2.61 | 2.91 | 2.8 | 2.6 | 2.681 | 4.3069 | 3.0008 |  |  |  | 1
| 1 | 2 | 2.72 | 2.7966 | 2.8 | 2.9 | 2.79 | 2.7 | 2.98 | 2.8 | 2.6 | 2.7 | 2.7 | 2.7 | 2.7 | 2.7765 | 2.72 | 2.512 | 2.72 | 2.79 | 2.6 | 2.5 | 2.5926 | 4.1683 | 3.0942 |  |  |  | 35
| 1 | 3 | 2.82 | 2.7574 | 2.8 | 3 | 2.77 | 2.8 | 2.97 | 2.9 | 2.6 | 2.8 | 2.8 | 2.8 | 2.7 | 2.5667 | 2.82 | 2.666 | 2.72 | 2.84 | 2.7 | 2.5 | 2.7054 | 3.9305 | 3.0444 |  |  |  | 16
| 1 | 4 | 2.78 | 2.7674 | 2.8 | 2.9 | 2.87 | 2.8 | 2.71 | 2.7 | 2.6 | 2.8 | 2.8 | 2.6 | 3.2 | 2.835 | 2.68 | 4.428 | 2.64 | 2.87 | 2.8 | 2.6 | 2.6397 | 3.9305 | 3.0623 |  |  |  | 2.7348
| 1 | 5 | 2.84 | 2.7577 | 2.8 | 2.9 | 2.84 | 2.7 | 2.8 | 2.8 | 2.5 | 2.6 | 2.8 | 2.6 | 2.7 | 2.717 | 2.62 | 2.667 | 2.62 | 2.81 | 2.6 | 2.6 | 2.6578 | 4.0077 | 3.1003 |  |  |  | 64
| 1 | 6 | 2.81 | 2.8463 | 2.7 | 2.9 | 2.81 | 2.8 | 2.97 | 2.9 | 2.6 | 2.8 | 2.8 | 2.6 | 2.8 | 2.5344 | 2.85 | 2.559 | 2.61 | 2.77 | 3.1 | 2.6 | 2.6442 | 3.9714 | 3.0764
| Mean |  | 2.7983 | 2.8012 | 2.7833 | 2.9167 | 2.8183 | 2.7333 | 2.875 | 2.8167 | 2.5667 | 2.7333 | 2.7833 | 2.65 | 2.8167 | 2.6857 | 2.7267 | 2.8895 | 2.6533 | 2.8317 | 2.7667 | 2.5667 | 2.6534 | 4.0526 | 3.0631
| Median |  | 2.815 | 2.782 | 2.8 | 2.9 | 2.82 | 2.75 | 2.895 | 2.8 | 2.6 | 2.75 | 2.8 | 2.6 | 2.75 | 2.7009 | 2.7 | 2.6125 | 2.63 | 2.825 | 2.75 | 2.6 | 2.651 | 3.9895 | 3.0693
| Std Dev. |  | 0.043089 | 0.051909 | 0.040825 | 0.040825 | 0.036009 | 0.08165 | 0.11397 | 0.075277 | 0.05164 | 0.08165 | 0.040825 | 0.083666 | 0.19408 | 0.11712 | 0.090259 | 0.7571 | 0.052789 | 0.052313 | 0.18619 | 0.05164 | 0.038611 | 0.15263 | 0.036777
| Rel.Std.Dev. |  | 1.54% | 1.853% | 1.467% | 1.4% | 1.278% | 2.987% | 3.964% | 2.673% | 2.012% | 2.987% | 1.467% | 3.157% | 6.89% | 4.361% | 3.31% | 26.2% | 1.99% | 1.847% | 6.73% | 2.012% | 1.455% | 3.766% | 1.201%
| PDM3 |  | 2.322% | 2.426% | 1.774% | 6.649% | 3.054% | -0.05427% | 5.126% | 2.993% | -6.149% | -0.05427% | 1.774% | -3.101% | 2.993% | -1.795% | -0.298% | 5.656% | -2.98% | 3.541% | 1.165% | -6.149% | -2.975% | 48.18% | 12%
| Z-Score (Absolute) |  | 0.18 | 0.21 | 0 | 1.59 | 0.42 | 0.59 | 1.09 | 0.4 | 2.58 | 0.59 | 0 | 1.59 | 0.4 | 1.16 | 0.67 | 1.26 | 1.55 | 0.58 | 0.2 | 2.58 | 1.55 | 15.1 | 3.33
| Analytical results for V in OREAS 237b (Certified Value 104 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 107 | 89.8 | 100 | 89 | 114.3764 | 107 | 102 | 101 | 99 | 101 | 107 | 102 | 101 | 105 | 106.1452 | 102 | 104 | 99.57 | 108 | 107 | 105 | 95.92 | 106 | 106.5167 | 104.5539 |  | 1
| 1 | 2 | 108 | 90.95 | 105 | 109 | 114.6056 | 107 | 104 | 100 | 99 | 103 | 104 | 103 | 101 | 104 | 103.8531 | 103 | 106 | 99.74 | 105 | 108 | 103 | 96.03 | 108 | 105.7123 | 104.0519 |  | 36
| 1 | 3 | 112 | 88 | 100 | 110 | 116.4943 | 105 | 100 | 102 | 99 | 103 | 107 | 103 | 102 | 106 | 106.0401 | 102 | 105 | 99.59 | 105 | 107 | 103 | 98.78 | 105 | 104.4254 | 105.4756 |  | 16
| 1 | 4 | 108 | 88.63 | 105 | 110 | 115.9299 | 105 | 104 | 101 | 101 | 101 | 107 | 104 | 102 | 108 | 104.7792 | 103 | 107 | 99.51 | 105 | 108 | 107 | 95.15 | 105 | 108.0449 | 106.0897 |  | 104.1674
| 1 | 5 | 108 | 87.97 | 100 | 115 | 115.6087 | 107 | 105 | 99 | 99 | 102 | 106 | 100 | 102 | 109 | 103.3589 | 102 | 107 | 100.3 | 105 | 107 | 105 | 97.79 | 106 | 107.0797 | 104.3824 |  | 65
| 1 | 6 | 112 | 88.48 | 105 | 112 | 113.59 | 107 | 105 | 95 | 101 | 103 | 111 | 103 | 99 | 109 | 102.5197 | 104 | 106 | 100.1 | 105 | 108 | 105 | 98.34 | 103 | 107.2406 | 100.0967
| Mean |  | 109.1667 | 88.9717 | 102.5 | 107.5 | 115.1008 | 106.3333 | 103.3333 | 99.6667 | 99.6667 | 102.1667 | 107 | 102.5 | 101.1667 | 106.8333 | 104.4494 | 102.6667 | 105.8333 | 99.8017 | 105.5 | 107.5 | 104.6667 | 97.0017 | 105.5 | 106.5033 | 104.1084
| Median |  | 108 | 88.555 | 102.5 | 110 | 115.1071 | 107 | 104 | 100.5 | 99 | 102.5 | 107 | 103 | 101.5 | 107 | 104.3162 | 102.5 | 106 | 99.665 | 105 | 107.5 | 105 | 96.91 | 105.5 | 106.7982 | 104.4682
| Std Dev. |  | 2.2286 | 1.1752 | 2.7386 | 9.3113 | 1.0899 | 1.0328 | 1.9664 | 2.5033 | 1.0328 | 0.98319 | 2.2804 | 1.3784 | 1.169 | 2.137 | 1.4686 | 0.8165 | 1.169 | 0.32394 | 1.2247 | 0.54772 | 1.5055 | 1.4912 | 1.6432 | 1.2803 | 2.1055
| Rel.Std.Dev. |  | 2.041% | 1.321% | 2.672% | 8.662% | 0.9469% | 0.9713% | 1.903% | 2.512% | 1.036% | 0.9623% | 2.131% | 1.345% | 1.156% | 2% | 1.406% | 0.7953% | 1.105% | 0.3246% | 1.161% | 0.5095% | 1.438% | 1.537% | 1.558% | 1.202% | 2.022%
| PDM3 |  | 4.799% | -14.59% | -1.601% | 3.199% | 10.5% | 2.079% | -0.8007% | -4.321% | -4.321% | -1.921% | 2.719% | -1.601% | -2.881% | 2.559% | 0.2707% | -1.441% | 1.599% | -4.191% | 1.279% | 3.199% | 0.4793% | -6.879% | 1.279% | 2.242% | -0.05666%
| Z-Score (Absolute) |  | 1.39 | 4.57 | 0.58 | 0.9 | 3.15 | 0.56 | 0.33 | 1.41 | 1.41 | 0.67 | 0.75 | 0.58 | 0.97 | 0.7 | 0 | 0.53 | 0.41 | 1.37 | 0.31 | 0.9 | 0.06 | 2.2 | 0.31 | 0.61 | 0.1
| Analytical results for W in OREAS 237b (Certified Value 2.44 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 2.5 | 2.5064 | 1 | < 0.1 | 17.1871 | 2.5 | 2.2 | 1.9 | 2.3 | 2.5 | 2.3 | 2.5 | 2.2 | 2.4 | 2.4752 | 3.5 | 2.3 | < 5 | 3.06 | 2.9 | 2.1 | 1.5536 | < 5 | 2.9482 |  |  | 1
| 1 | 2 | 2.6 | 2.5243 | 1 | < 0.1 | 18.8506 | 2.3 | 2.1 | 2.3 | 2.3 | 2.5 | 2.3 | 2.3 | 2.3 | 2.2 | 2.4127 | 3.6 | 2.3 | < 5 | 2.91 | 2.9 | 1.9 | 1.4623 | < 5 | 2.966 |  |  | 37
| 1 | 3 | 2.6 | 2.3717 | 1 | < 0.1 | 17.6891 | 2.5 | 2.2 | 2.1 | 2.1 | 2.5 | 2.5 | 2.6 | 2.3 | 2.4 | 2.4667 | 3.6 | 2.3 | < 5 | 3.07 | 2.8 | 1.9 | 1.6181 | < 5 | 2.943 |  |  | 16
| 1 | 4 | 2.4 | 2.4783 | 1 | < 0.1 | 18.4591 | 2.4 | 2.3 | 1.4 | 2 | 2.5 | 2.4 | 2.6 | 2.3 | 3 | 2.4994 | 3.5 | 2.3 | < 5 | 2.94 | 2.9 | 1.9 | 1.5515 | < 5 | 2.8708 |  |  | 2.4385
| 1 | 5 | 2.4 | 2.405 | 1 | < 0.1 | 16.6814 | 2.4 | 2.3 | 2.1 | 2 | 2.5 | 2.2 | 2.4 | 2.3 | 2.5 | 2.505 | 3.4 | 2.2 | < 5 | 2.92 | 2.8 | 1.9 | 1.5813 | < 5 | 2.9741 |  |  | 66
| 1 | 6 | 2.7 | 2.4931 | 1 | 0.2 | 19.1617 | 2.4 | 2.4 | 2.2 | 2.2 | 2.6 | 2.5 | 2.4 | 2.3 | 2.5 | 2.4503 | 3.6 | 2.3 | < 5 | 3 | 3.2 | 2 | 1.5292 | < 5 | 2.9598
| Mean |  | 2.5333 | 2.4631 | 1 | 0.2 | 18.0048 | 2.4167 | 2.25 | 2 | 2.15 | 2.5167 | 2.3667 | 2.4667 | 2.2833 | 2.5 | 2.4682 | 3.5333 | 2.2833 |  | 2.9833 | 2.9167 | 1.95 | 1.5493 |  | 2.9436
| Median |  | 2.55 | 2.4857 | 1 | 0.2 | 18.0741 | 2.4 | 2.25 | 2.1 | 2.15 | 2.5 | 2.35 | 2.45 | 2.3 | 2.45 | 2.4709 | 3.55 | 2.3 |  | 2.97 | 2.9 | 1.9 | 1.5526 |  | 2.954
| Std Dev. |  | 0.12111 | 0.060799 | 0 |  | 0.97774 | 0.075277 | 0.10488 | 0.32249 | 0.13784 | 0.040825 | 0.12111 | 0.12111 | 0.040825 | 0.26833 | 0.034006 | 0.08165 | 0.040825 |  | 0.070616 | 0.1472 | 0.083666 | 0.052448 |  | 0.037458
| Rel.Std.Dev. |  | 4.781% | 2.468% | 0% |  | 5.43% | 3.115% | 4.661% | 16.12% | 6.411% | 1.622% | 5.117% | 4.91% | 1.788% | 10.73% | 1.378% | 2.311% | 1.788% |  | 2.367% | 5.047% | 4.291% | 3.385% |  | 1.272%
| PDM3 |  | 3.888% | 1.009% | -58.99% | -91.8% | 638.3% | -0.8965% | -7.731% | -17.98% | -11.83% | 3.204% | -2.947% | 1.154% | -6.364% | 2.521% | 1.216% | 44.9% | -6.364% |  | 22.34% | 19.61% | -20.03% | -36.46% |  | 20.71%
| Z-Score (Absolute) |  | 0.25 | 0.01 | NA | 8.8 | 57.15 | 0.18 | 0.8 | 1.72 | 1.16 | 0.18 | 0.37 | 0 | 0.67 | 0.12 | 0.01 | 3.92 | 0.67 | 0.12 | 1.9 | 1.66 | 1.9 | 3.37 | 0.12 | 1.75
Results from laboratory 1.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 237b (Certified Value 16.2 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 16.1 | 13.1301 | 18.5 | 19.2 | 9.64 | 18 | 12 | 15.9 | 13.9 | 17.7 | 16.3 | 16.4 | 16.9 | 15 | 15.8188 | 19.6 | 12.47 | 14 | 19.08 | 16.8 | 14.6 | 12.61 | 14.5755 | 25.918 |  |  | 1
| 1 | 2 | 16.36 | 12.9419 | 17.7 | 20.6 | 10.49 | 17.2 | 12 | 16.5 | 14.6 | 17.8 | 15.7 | 16.4 | 17 | 14.6 | 16.3099 | 20.7 | 13.17 | 14.7 | 18.8 | 16.5 | 14.3 | 12.31 | 15.3452 | 26.39 |  |  | 38
| 1 | 3 | 16.42 | 13.0262 | 18 | 20.5 | 10.425 | 16.6 | 13 | 17.7 | 14.4 | 17.5 | 16.6 | 17 | 17.4 | 15 | 16.4758 | 20.8 | 14.59 | 15.2 | 18.88 | 15.9 | 14.4 | 12.73 | 14.9347 | 25.0413 |  |  | 16
| 1 | 4 | 16.25 | 12.948 | 18.8 | 20.9 | 11.02 | 18.2 | 14 | 16.7 | 13.9 | 17.3 | 16.3 | 16.3 | 16.3 | 14.6 | 15.9383 | 20.4 | 12.63 | 15.4 | 18.91 | 16 | 14.9 | 12.24 | 14.8834 | 26.3185 |  |  | 16.2019
| 1 | 5 | 16.22 | 12.8057 | 18.3 | 20.8 | 10.185 | 17 | 14 | 16.4 | 14.8 | 17.8 | 15.3 | 16.4 | 16.5 | 15.3 | 15.6138 | 19.9 | 12.92 | 14.8 | 18.2 | 15.9 | 15.1 | 12.55 | 15.2426 | 25.9481 |  |  | 67
| 1 | 6 | 16.22 | 12.9214 | 18.5 | 21 | 10.88 | 16.9 | 14 | 18.5 | 14 | 18.2 | 17.1 | 16.6 | 16.1 | 15 | 16.4896 | 20.8 | 12.56 | 14.6 | 18.01 | 17.9 | 14.9 | 12.72 | 15.4479 | 25.4671
| Mean |  | 16.2617 | 12.9622 | 18.3 | 20.5 | 10.44 | 17.3167 | 13.1667 | 16.95 | 14.2667 | 17.7167 | 16.2167 | 16.5167 | 16.7 | 14.9167 | 16.1077 | 20.3667 | 13.0567 | 14.7833 | 18.6467 | 16.5 | 14.7 | 12.5267 | 15.0716 | 25.8472
| Median |  | 16.235 | 12.945 | 18.4 | 20.7 | 10.4575 | 17.1 | 13.5 | 16.6 | 14.2 | 17.75 | 16.3 | 16.4 | 16.7 | 15 | 16.1241 | 20.55 | 12.775 | 14.75 | 18.84 | 16.25 | 14.75 | 12.58 | 15.0886 | 25.933
| Std Dev. |  | 0.11356 | 0.10868 | 0.39497 | 0.66332 | 0.49754 | 0.64005 | 0.98319 | 0.96281 | 0.38816 | 0.30605 | 0.64005 | 0.25626 | 0.4858 | 0.27142 | 0.36831 | 0.50859 | 0.79442 | 0.4916 | 0.43357 | 0.77717 | 0.31623 | 0.20753 | 0.33045 | 0.51496
| Rel.Std.Dev. |  | 0.6984% | 0.8384% | 2.158% | 3.236% | 4.766% | 3.696% | 7.467% | 5.68% | 2.721% | 1.727% | 3.947% | 1.551% | 2.909% | 1.82% | 2.287% | 2.497% | 6.084% | 3.325% | 2.325% | 4.71% | 2.151% | 1.657% | 2.193% | 1.992%
| PDM3 |  | 0.3689% | -20% | 12.95% | 26.53% | -35.56% | 6.881% | -18.73% | 4.617% | -11.94% | 9.349% | 0.09118% | 1.943% | 3.074% | -7.933% | -0.5815% | 25.71% | -19.41% | -8.756% | 15.09% | 1.84% | -9.27% | -22.68% | -6.977% | 59.53%
| Z-Score (Absolute) |  | 0 | 1.5 | 0.93 | 1.93 | 2.66 | 0.48 | NA | 0.31 | 0.91 | 0.66 | 0.02 | 0.12 | 0.2 | 0.61 | 0.07 | 1.87 | 1.46 | 0.67 | 1.09 | 0.11 | 0.71 | 1.7 | 0.54 | 4.37
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 237b (Certified Value 1.76 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.32 | 1.34 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 1.4758 | 2.05 | 2.1 | 1.5 | 2.05 | 1.501 | 1.5 | 2.13 | 1.6 | 1.42 | 1.6423 | 2.8307 | [14 empty] | 1
| 1 | 2 | 1.5302 | 1.95 | 2.2 | 1.7 | 2.11 | 1.612 | 1.6 | 2.05 | 1.5 | 1.33 | 1.7273 | 2.8856 | [14 empty] | 39
| 1 | 3 | 1.5867 | 2 | 2.2 | 1.5 | 2.19 | 2.208 | 1.6 | 2.12 | 1.5 | 1.4 | 1.6848 | 2.7526 | [14 empty] | 16
| 1 | 4 | 1.6266 | 2.1 | 2.3 | 2.9 | 2.03 | 1.526 | 1.6 | 2.02 | 1.5 | 1.31 | 1.6848 | 2.8078 | [14 empty] | 1.7577
| 1 | 5 | 1.5033 | 2 | 2.3 | 1.9 | 2 | 1.809 | 1.6 | 2.07 | 1.6 | 1.36 | 1.6848 | 2.7856 | [14 empty] | 68
| 1 | 6 | 1.4731 | 2.05 | 2.3 | 1.6 | 2.08 | 1.546 | 1.5 | 2.04 | 1.5 | 1.37 | 1.7273 | 2.7933
| Mean |  | 1.5326 | 2.025 | 2.2333 | 1.85 | 2.0767 | 1.7003 | 1.5667 | 2.0717 | 1.5333 | 1.365 | 1.6919 | 2.8093
| Median |  | 1.5167 | 2.025 | 2.25 | 1.65 | 2.065 | 1.579 | 1.6 | 2.06 | 1.5 | 1.365 | 1.6848 | 2.8005
| Std Dev. |  | 0.062276 | 0.05244 | 0.08165 | 0.53572 | 0.067429 | 0.2725 | 0.05164 | 0.04446 | 0.05164 | 0.041352 | 0.031993 | 0.045435
| Rel.Std.Dev. |  | 4.063% | 2.59% | 3.656% | 28.96% | 3.247% | 16.03% | 3.296% | 2.146% | 3.368% | 3.029% | 1.891% | 1.617%
| PDM3 |  | -12.81% | 15.21% | 27.06% | 5.25% | 18.15% | -3.265% | -10.87% | 17.86% | -12.77% | -22.34% | -3.746% | 59.82%
| Z-Score (Absolute) |  | 0.66 | 0.68 | 1.25 | 0.2 | 0.83 | 0.2 | 0.57 | 0.81 | 0.66 | 1.12 | 0.23 | 2.83
| Analytical results for Zn in OREAS 237b (Certified Value 95 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 100 | 89.356 | 90 | 101 | 96.4481 | 97 | 98 | 93 | 96 | 94 | 101 | 92 | 94 | 96.9 | 93.7727 | 99 | 95.32 | 93 | 105.8 | 96 | 94 | 91 | 96.294 | 110 | 87.3624 | 91.9898 | 1
| 1 | 2 | 99 | 89.408 | 92 | 102 | 94.0086 | 96 | 96 | 94 | 98 | 94 | 95 | 92 | 95 | 95.1 | 94.3612 | 103 | 94.64 | 95 | 108.9 | 92 | 92 | 90.6 | 96.382 | 100 | 86.6091 | 90.6942 | 40
| 1 | 3 | 101 | 89.9488 | 90 | 102 | 97.19 | 96 | 95 | 95 | 97 | 94 | 97 | 93 | 96 | 96.8 | 93.9097 | 114 | 97.54 | 93 | 110.2 | 91 | 94 | 92 | 99.055 | 98 | 85.6097 | 92.2381 | 16
| 1 | 4 | 98 | 89.9592 | 94 | 102 | 93.31 | 95 | 99 | 94 | 97 | 93 | 97 | 94 | 94 | 97.4 | 94.3619 | 100 | 167.4 | 94 | 104.5 | 93 | 94 | 97 | 96.019 | 100 | 86.5802 | 91.5182 | 95.0975
| 1 | 5 | 99 | 89.1064 | 90 | 101 | 96.3814 | 95 | 95 | 96 | 97 | 94 | 96 | 90 | 95 | 98.3 | 94.7021 | 101 | 119.3 | 94 | 104 | 95 | 94 | 91.3 | 98.406 | 100 | 91.2443 | 91.3183 | 69
| 1 | 6 | 98 | 89.8448 | 94 | 100 | 94.3085 | 97 | 105 | 96 | 98 | 94 | 100 | 92 | 93 | 99.2 | 93.1777 | 101 | 97.37 | 93 | 109.8 | 92 | 95 | 92.7 | 98.978 | 96 | 91.056 | 90.9971
| Mean |  | 99.1667 | 89.6039 | 91.6667 | 101.3333 | 95.2744 | 96 | 98 | 94.6667 | 97.1667 | 93.8333 | 97.6667 | 92.1667 | 94.5 | 97.2833 | 94.0476 | 103 | 111.9283 | 93.6667 | 107.2 | 93.1667 | 93.8333 | 92.4333 | 97.5223 | 100.6667 | 88.0769 | 91.4593
| Median |  | 99 | 89.6264 | 91 | 101.5 | 95.345 | 96 | 97 | 94.5 | 97 | 94 | 97 | 92 | 94.5 | 97.15 | 94.1354 | 101 | 97.455 | 93.5 | 107.35 | 92.5 | 94 | 91.65 | 97.394 | 100 | 86.9857 | 91.4183
| Std Dev. |  | 1.169 | 0.36071 | 1.9664 | 0.8165 | 1.5916 | 0.89443 | 3.7947 | 1.2111 | 0.75277 | 0.40825 | 2.3381 | 1.3292 | 1.0488 | 1.4049 | 0.54346 | 5.5498 | 28.7233 | 0.8165 | 2.7619 | 1.9408 | 0.98319 | 2.3585 | 1.4365 | 4.8442 | 2.4454 | 0.58486
| Rel.Std.Dev. |  | 1.179% | 0.4026% | 2.145% | 0.8058% | 1.671% | 0.9317% | 3.872% | 1.279% | 0.7747% | 0.4351% | 2.394% | 1.442% | 1.11% | 1.444% | 0.5779% | 5.388% | 25.66% | 0.8717% | 2.576% | 2.083% | 1.048% | 2.552% | 1.473% | 4.812% | 2.776% | 0.6395%
| PDM3 |  | 4.279% | -5.777% | -3.608% | 6.557% | 0.186% | 0.949% | 3.052% | -0.4531% | 2.176% | -1.329% | 2.702% | -3.082% | -0.6283% | 2.298% | -1.104% | 8.31% | 17.7% | -1.505% | 12.73% | -2.03% | -1.329% | -2.802% | 2.55% | 5.856% | -7.383% | -3.826%
| Z-Score (Absolute) |  | 1.08 | 1.38 | 0.85 | 1.64 | 0.08 | 0.26 | 0.78 | 0.08 | 0.56 | 0.29 | 0.69 | 0.72 | 0.12 | 0.59 | 0.24 | 2.06 | 4.36 | 0.34 | 3.14 | 0.46 | 0.29 | 0.65 | 0.66 | 1.46 | 1.77 | 0.9
| Analytical results for Zr in OREAS 237b (Certified Value 135 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 142.2 | 137.3372 | 138 | 112 | 117.05 | 132.4 | 137.7 | 133 | 134.1 | 134 | 137.5 | 136.5 | 135.5 | 130.7 | 131.785 | 143 | 128 | 169.5 | 142.3 | 128.1 | 123.9 | 118.03 | 143 | 134.3555 | 229.5273 |  | 1
| 1 | 2 | 140.4 | 136.6218 | 141 | 126 | 113.8394 | 129.1 | 129.3 | 135 | 137.4 | 137 | 133 | 137 | 139.5 | 130.2 | 134.4849 | 147 | 129 | 168 | 136.2 | 127.8 | 122.9 | 117.799 | 148 | 135.7654 | 225.2677 |  | 41
| 1 | 3 | 144.1 | 137.4548 | 140 | 143 | 113.3867 | 127.8 | 130.7 | 137 | 139.9 | 136 | 139.5 | 139 | 140.5 | 130.7 | 128.5043 | 144 | 128 | 166.5 | 143 | 127.3 | 122.5 | 122.243 | 145 | 134.5065 | 226.3663 |  | 16
| 1 | 4 | 143.3 | 136.0534 | 144 | 152 | 116.15 | 133.4 | 137.8 | 134 | 139.9 | 133 | 135 | 138.5 | 139 | 134.1 | 131.2331 | 143 | 132 | 168.5 | 138.1 | 127.2 | 127.3 | 117.667 | 143 | 133.9526 | 228.305 |  | 134.6182
| 1 | 5 | 144.1 | 134.2404 | 136 | 159 | 117.7 | 129.8 | 125.2 | 131 | 136.6 | 136 | 129 | 136.5 | 140.5 | 130 | 132.2895 | 143 | 129 | 167.4 | 136.5 | 127.9 | 126.2 | 120.725 | 146 | 133.3987 | 226.6979 |  | 70
| 1 | 6 | 145 | 134.652 | 135 | 164 | 114.71 | 131.5 | 134.8 | 145 | 138 | 141.5 | 144.5 | 136.5 | 134.5 | 131.8 | 132.246 | 147 | 129 | 165.3 | 135 | 129.8 | 124.4 | 120.835 | 144 | 137.0243 | 228.7306
| Mean |  | 143.1833 | 136.0599 | 139 | 142.6667 | 115.4727 | 130.6667 | 132.5833 | 135.8333 | 137.65 | 136.25 | 136.4167 | 137.3333 | 138.25 | 131.25 | 131.7571 | 144.5 | 129.1667 | 167.5333 | 138.5167 | 128.0167 | 124.5333 | 119.5498 | 144.8333 | 134.8338 | 227.4825
| Median |  | 143.7 | 136.3376 | 139 | 147.5 | 115.43 | 130.65 | 132.75 | 134.5 | 137.7 | 136 | 136.25 | 136.75 | 139.25 | 130.7 | 132.0155 | 143.5 | 129 | 167.7 | 137.3 | 127.85 | 124.15 | 119.3775 | 144.5 | 134.431 | 227.5015
| Std Dev. |  | 1.6558 | 1.3554 | 3.3466 | 20.1362 | 1.7611 | 2.1257 | 5.0428 | 4.916 | 2.1916 | 2.9623 | 5.3797 | 1.1255 | 2.6029 | 1.5294 | 1.9401 | 1.9748 | 1.472 | 1.4895 | 3.3582 | 0.9411 | 1.879 | 1.9599 | 1.9408 | 1.3295 | 1.6237
| Rel.Std.Dev. |  | 1.156% | 0.9962% | 2.408% | 14.11% | 1.525% | 1.627% | 3.803% | 3.619% | 1.592% | 2.174% | 3.944% | 0.8195% | 1.883% | 1.165% | 1.472% | 1.367% | 1.14% | 0.8891% | 2.424% | 0.7351% | 1.509% | 1.639% | 1.34% | 0.986% | 0.7138%
| PDM3 |  | 6.363% | 1.071% | 3.255% | 5.979% | -14.22% | -2.935% | -1.512% | 0.9026% | 2.252% | 1.212% | 1.336% | 2.017% | 2.698% | -2.502% | -2.125% | 7.341% | -4.05% | 24.45% | 2.896% | -4.904% | -7.491% | -11.19% | 7.588% | 0.1602% | 68.98%
| Z-Score (Absolute) |  | 0.94 | 0.03 | 0.37 | 0.87 | 2.8 | 0.75 | 0.49 | 0.06 | 0.19 | 0 | 0.02 | 0.15 | 0.27 | 0.67 | 0.61 | 1.11 | 0.96 | 4.22 | 0.31 | 1.11 | 1.58 | 2.25 | 1.16 | 0.19 | 12.3
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 237b (Certified Value 0.176 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.16 | 0.202 | < 0.5 | 0.18 | 0.14 | 0.19 | 0.2 | 0.18 | 0.18 | 0.17 | 0.18 | 0.159 | < 1 | 0.3 | 0.17 | 0.17 | < 0.5 | 0.19 | 0.2 | < 1 | 1.2605 |  |  |  | 1
| 1 | 2 | 0.17 | 0.196 | < 0.5 | 0.17 | 0.17 | 0.19 | 0.1 | 0.19 | 0.19 | 0.17 | 0.18 | 0.132 | < 1 | 0.2 | 0.17 | 0.11 | < 0.5 | 0.18 | 0.19 | < 1 | 1.3269 |  |  |  | 25
| 1 | 3 | 0.18 | 0.186 | < 0.5 | 0.18 | 0.16 | 0.2 | 0.2 | 0.19 | 0.19 | 0.18 | 0.17 | 0.161 | < 1 | 0.2 | 0.17 | 0.16 | < 0.5 | 0.18 | 0.16 | < 1 | 1.3933 |  |  |  | 16
| 1 | 4 | 0.17 | 0.183 | < 0.5 | 0.17 | 0.15 | 0.21 | 0.2 | 0.18 | 0.19 | 0.18 | 0.17 | 0.151 | < 1 | 0.2 | 0.18 | 0.15 | < 0.5 | 0.19 | 0.2 | < 1 | 1.2709 |  |  |  | 0.17579
| 1 | 5 | 0.17 | 0.183 | < 0.5 | 0.17 | 0.14 | 0.19 | 0.2 | 0.18 | 0.19 | 0.18 | 0.18 | 0.165 | < 1 | 0.2 | 0.17 | 0.16 | < 0.5 | 0.18 | 0.16 | < 1 | 1.3321 |  |  |  | 72
| 1 | 6 | 0.17 | 0.182 | < 0.5 | 0.18 | 0.15 | 0.19 | 0.1 | 0.18 | 0.18 | 0.17 | 0.18 | 0.15 | < 1 | 0.3 | 0.19 | 0.19 | < 0.5 | 0.19 | 0.17 | < 1 | 1.2901
| Mean |  | 0.17 | 0.18867 |  | 0.175 | 0.15167 | 0.195 | 0.16667 | 0.18333 | 0.18667 | 0.175 | 0.17667 | 0.153 |  | 0.23333 | 0.175 | 0.15667 |  | 0.185 | 0.18 |  | 1.3123
| Median |  | 0.17 | 0.1845 |  | 0.175 | 0.15 | 0.19 | 0.2 | 0.18 | 0.19 | 0.175 | 0.18 | 0.155 |  | 0.2 | 0.17 | 0.16 |  | 0.185 | 0.18 |  | 1.3085
| Std Dev. |  | 0.0063246 | 0.0083347 |  | 0.0054772 | 0.01169 | 0.0083666 | 0.05164 | 0.005164 | 0.005164 | 0.0054772 | 0.005164 | 0.011815 |  | 0.05164 | 0.0083666 | 0.026583 |  | 0.0054772 | 0.018974 |  | 0.049104
| Rel.Std.Dev. |  | 3.72% | 4.418% |  | 3.13% | 7.708% | 4.291% | 30.98% | 2.817% | 2.766% | 3.13% | 2.923% | 7.722% |  | 22.13% | 4.781% | 16.97% |  | 2.961% | 10.54% |  | 3.742%
| PDM3 |  | -3.291% | 7.328% |  | -0.447% | -13.72% | 10.93% | -5.188% | 4.294% | 6.19% | -0.447% | 0.5012% | -12.96% |  | 32.74% | -0.447% | -10.88% |  | 5.242% | 2.397% |  | 646.5%
| Z-Score (Absolute) |  | 0.77 | 0.31 | 3.85 | 0.48 | 1.83 | 0.67 | NA | 0 | 0.19 | 0.48 | 0.39 | 1.75 | 18.3 | NA | 0.48 | 1.54 | 3.85 | 0.1 | 0.19 | 18.3 | 65.25
Results from laboratories 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 237b (Certified Value 2.75 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 2.564 | 3.3544 | 2.88 | 3.0433 | 2.81 | 2.9 | 3.36 | 2.67 | 2.67 | 2.86 | 2.77 | 2.77 | 2.98 | 2.8453 | 2.54 | 2.62 | 2.71 | 2.65 | 2.4841 | 4.9795 |  |  |  |  | 1
| 1 | 2 | 2.573 | 3.3422 | 2.79 | 3.0333 | 2.809 | 2.77 | 3.27 | 2.69 | 2.66 | 2.95 | 2.74 | 2.82 | 2.72 | 2.8124 | 2.535 | 2.65 | 2.69 | 2.63 | 2.4037 | 4.9837 |  |  |  |  | #N/A
| 1 | 3 | 2.581 | 3.3127 | 2.93 | 3.02 | 2.919 | 2.79 | 3.42 | 2.71 | 2.64 | 2.92 | 2.75 | 2.78 | 2.84 | 2.7303 | 2.572 | 2.68 | 2.71 | 2.64 | 2.5024 | 4.9841 |  |  |  |  | 16
| 1 | 4 | 2.607 | 3.2933 | 3.12 | 3.02 | 2.866 | 2.91 | 3.35 | 2.67 | 2.7 | 2.92 | 2.77 | 2.65 | 2.71 | 2.6993 | 2.62 | 2.65 | 2.76 | 2.48 | 2.5304 | 4.9866 |  |  |  |  | 2.7461
| 1 | 5 | 2.597 | 3.2931 | 2.97 | 3.0467 | 2.85 | 2.9 | 3.4 | 2.71 | 2.63 | 2.88 | 2.8 | 2.77 | 3.05 | 2.6937 | 2.577 | 2.68 | 2.69 | 2.59 | 2.4524 | 4.9838 |  |  |  |  | 73
| 1 | 6 | 2.612 | 3.3163 | 2.8 | 3.06 | 2.92 | 2.82 | 3.33 | 2.73 | 2.67 | 2.8 | 2.76 | 2.77 | 2.8 | 2.8236 | 2.445 | 2.64 | 2.72 | 2.61 | 2.5173 | 4.9835
| Mean |  | 2.589 | 3.3187 | 2.915 | 3.0372 | 2.8623 | 2.8483 | 3.355 | 2.6967 | 2.6617 | 2.8883 | 2.765 | 2.76 | 2.85 | 2.7674 | 2.5482 | 2.6533 | 2.7133 | 2.6 | 2.4817 | 4.9835
| Median |  | 2.589 | 3.3145 | 2.905 | 3.0383 | 2.858 | 2.86 | 3.355 | 2.7 | 2.665 | 2.9 | 2.765 | 2.77 | 2.82 | 2.7714 | 2.556 | 2.65 | 2.71 | 2.62 | 2.4932 | 4.9837
| Std Dev. |  | 0.019298 | 0.025175 | 0.12276 | 0.015835 | 0.049569 | 0.062423 | 0.053198 | 0.024221 | 0.024833 | 0.053821 | 0.020736 | 0.057271 | 0.13856 | 0.067366 | 0.059064 | 0.023381 | 0.02582 | 0.06261 | 0.046938 | 0.0022844
| Rel.Std.Dev. |  | 0.7454% | 0.7586% | 4.211% | 0.5214% | 1.732% | 2.192% | 1.586% | 0.8982% | 0.933% | 1.863% | 0.75% | 2.075% | 4.862% | 2.434% | 2.318% | 0.8812% | 0.9516% | 2.408% | 1.891% | 0.04584%
| PDM3 |  | -5.721% | 20.85% | 6.151% | 10.6% | 4.233% | 3.723% | 22.17% | -1.8% | -3.075% | 5.18% | 0.6883% | 0.5062% | 3.784% | 0.7771% | -7.208% | -3.378% | -1.193% | -5.32% | -9.628% | 81.48%
| Z-Score (Absolute) |  | 1.02 | 3.17 | 0.85 | 1.56 | 0.55 | 0.47 | 3.38 | 0.4 | 0.6 | 0.7 | 0.01 | 0.04 | 0.48 | 0.01 | 1.25 | 0.65 | 0.3 | 0.95 | 1.63 | 12.72
| Analytical results for As in OREAS 237b (Certified Value 583 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 588.48 | 605.58 | 532 | 569.5073 | 562 | 594.5 | 581 | 608 | 603 | 607 | 598 | 602 | 625 | 590.6253 | 553 | 529 | 665.08 | 587.2 | 555 | 540 | 563 | 708.185 | 565.498 |  | 1
| 1 | 2 | 584.85 | 604.83 | 547 | 570.4779 | 560 | 590.3 | 574 | 608.6 | 594 | 634 | 601 | 621 | 582.3 | 579.5561 | 546 | 527 | 657.19 | 590.6 | 543.5 | 540.6 | 575 | 675.6314 | 564.6757 |  | 26
| 1 | 3 | 584.01 | 600.63 | 560 | 565.0849 | 572 | 607.9 | 595 | 618 | 590 | 614 | 609 | 599 | 609.1 | 574.482 | 566 | 534 | 682.03 | 586.7 | 554.1 | 547.8 | 568 | 683.6551 | 569.0745 |  | 16
| 1 | 4 | 589.05 | 599.18 | 553 | 564.76 | 569 | 590.2 | 579 | 615.1 | 598 | 625 | 608 | 591 | 598.9 | 576.3105 | 569 | 549 | 650.46 | 590.6 | 532.5 | 537.4 | 575 | 598.2882 | 570.0684 |  | 583.2871
| 1 | 5 | 599.88 | 607.77 | 551 | 572.0774 | 565 | 595.3 | 581 | 610.8 | 586 | 610 | 604 | 601 | 649.3 | 581.416 | 573 | 501 | 653.39 | 594.5 | 557.1 | 525.8 | 566 | 683.0686 | 571.1473 |  | 74
| 1 | 6 | 591.21 | 604.79 | 546 | 570.0494 | 571 | 601.9 | 580 | 617.3 | 597 | 590 | 591 | 597 | 590 | 589.9795 | 540 | 529 | 650.31 | 589.9 | 563.4 | 547.9 | 568 | 652.6699 | 571.4365
| Mean |  | 589.58 | 603.7967 | 548.1667 | 568.6595 | 566.5 | 596.6833 | 581.6667 | 612.9667 | 594.6667 | 613.3333 | 601.8333 | 601.8333 | 609.1 | 582.0616 | 557.8333 | 528.1667 | 659.7433 | 589.9167 | 550.9333 | 539.9167 | 569.1667 | 666.9164 | 568.6501
| Median |  | 588.765 | 604.81 | 549 | 569.7784 | 567 | 594.9 | 580.5 | 612.95 | 595.5 | 612 | 602.5 | 600 | 604 | 580.486 | 559.5 | 529 | 655.29 | 590.25 | 554.55 | 540.3 | 568 | 679.35 | 569.5714
| Std Dev. |  | 5.7208 | 3.2362 | 9.3684 | 3.0207 | 4.9295 | 6.964 | 7.0333 | 4.4067 | 6.0553 | 15.2272 | 6.7355 | 10.1669 | 24.7222 | 6.8301 | 13.4375 | 15.5488 | 12.2248 | 2.8167 | 11.091 | 8.1482 | 4.8751 | 38.034 | 2.8956
| Rel.Std.Dev. |  | 0.9703% | 0.536% | 1.709% | 0.5312% | 0.8702% | 1.167% | 1.209% | 0.7189% | 1.018% | 2.483% | 1.119% | 1.689% | 4.059% | 1.173% | 2.409% | 2.944% | 1.853% | 0.4775% | 2.013% | 1.509% | 0.8565% | 5.703% | 0.5092%
| PDM3 |  | 1.079% | 3.516% | -6.021% | -2.508% | -2.878% | 2.297% | -0.2778% | 5.088% | 1.951% | 5.151% | 3.18% | 3.18% | 4.425% | -0.2101% | -4.364% | -9.45% | 13.11% | 1.137% | -5.547% | -7.436% | -2.421% | 14.34% | -2.509%
| Z-Score (Absolute) |  | 0 | 0.46 | 1.33 | 0.67 | 0.74 | 0.23 | 0.26 | 0.75 | 0.16 | 0.77 | 0.39 | 0.39 | 0.63 | 0.24 | 1.02 | 1.98 | 2.26 | 0.01 | 1.25 | 1.6 | 0.66 | 2.49 | 0.67
| Analytical results for B in OREAS 237b (Indicative Value 15.3 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [13 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [13 empty] | 1
| 1 | 1 | < 5 | 5 | 26.48 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21 | [13 empty] | 1
| 1 | 2 | < 5 | 5 | 24.87 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 19 | [13 empty] | 4
| 1 | 3 | < 5 | 5 | 25.885 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 20 | [13 empty] | 16
| 1 | 4 | < 5 | 6 | 26.055 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 20 | [13 empty] | 15.2614
| 1 | 5 | < 5 | 5 | 25.99 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 20 | [13 empty] | 14
| 1 | 6 | < 5 | 5 | 24.9938 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21
| Mean |  |  | 5.1667 | 25.7123 |  |  | 10 |  |  |  |  | 20.1667
| Median |  |  | 5 | 25.9375 |  |  | 10 |  |  |  |  | 20
| Std Dev. |  |  | 0.40825 | 0.63868 |  |  | 0 |  |  |  |  | 0.75277
| Rel.Std.Dev. |  |  | 7.902% | 2.484% |  |  | 0% |  |  |  |  | 3.733%
| PDM3 |  |  | -66.15% | 68.48% |  |  | -34.48% |  |  |  |  | 32.14%
| Z-Score (Absolute) |  | 10.11 | 0.67 | 83.8 | 0 | 0 | 20.23 | 0 | 0 | 20.23 | 0 | 61.36
| Analytical results for Ba in OREAS 237b (Certified Value 108 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 95.91 | 120.6579 | 113 | 112.895 | 102 | 104 | 122 | 101 | 100 | 100 | 110 | 100 | 111.6 | 103.9194 | 98 | 103.9 | 81 | 120.43 | 118.9 | 107 | 102.37 | 697.8949 |  |  | 1
| 1 | 2 | 96.1 | 118.0967 | 116 | 113.93 | 107 | 105 | 123 | 101 | 100 | 100 | 110 | 100 | 103.2 | 100.6548 | 98 | 102.3 | 82 | 113.1 | 121.5 | 107 | 99.65 | 603.4863 |  |  | 27
| 1 | 3 | 98.88 | 120.433 | 116 | 112.535 | 108 | 105 | 124 | 103 | 100 | 100 | 110 | 100 | 106.3 | 107.4231 | 99 | 102.5 | 83 | 117.69 | 123.9 | 107 | 103.36 | 604.142 |  |  | 16
| 1 | 4 | 96.62 | 120.2607 | 114 | 112.525 | 103 | 110 | 125 | 103 | 100 | 100 | 110 | 100 | 102.3 | 101.1355 | 101 | 105.1 | 82 | 115.28 | 124.9 | 100 | 104.12 | 593.6251 |  |  | 107.5927
| 1 | 5 | 96.61 | 126.3264 | 120 | 112.69 | 105 | 110 | 124 | 104 | 100 | 100 | 110 | 100 | 113.2 | 100.8632 | 99 | 104 | 84 | 106.47 | 118.5 | 105 | 100.19 | 596.8023 |  |  | 75
| 1 | 6 | 98.32 | 126.5165 | 116 | 113.445 | 107 | 108 | 129 | 106 | 100 | 100 | 110 | 100 | 104.6 | 102.9016 | 96 | 104.2 | 82 | 107.55 | 119.9 | 106 | 102.47 | 600.9392
| Mean |  | 97.0733 | 122.0485 | 115.8333 | 113.0033 | 105.3333 | 107 | 124.5 | 103 | 100 | 100 | 110 | 100 | 106.8667 | 102.8163 | 98.5 | 103.6667 | 82.3333 | 113.42 | 121.2667 | 105.3333 | 102.0267 | 616.1483
| Median |  | 96.615 | 120.5454 | 116 | 112.7925 | 106 | 106.5 | 124 | 103 | 100 | 100 | 110 | 100 | 105.45 | 102.0186 | 98.5 | 103.95 | 82 | 114.19 | 120.7 | 106.5 | 102.42 | 602.2128
| Std Dev. |  | 1.2279 | 3.5106 | 2.4014 | 0.56777 | 2.4221 | 2.6833 | 2.429 | 1.8974 | 0 | 0 | 0 | 0 | 4.5227 | 2.5984 | 1.6432 | 1.0708 | 1.0328 | 5.5443 | 2.6576 | 2.7325 | 1.7605 | 40.2489
| Rel.Std.Dev. |  | 1.265% | 2.876% | 2.073% | 0.5024% | 2.299% | 2.508% | 1.951% | 1.842% | 0% | 0% | 0% | 0% | 4.232% | 2.527% | 1.668% | 1.033% | 1.254% | 4.888% | 2.192% | 2.594% | 1.725% | 6.532%
| PDM3 |  | -9.777% | 13.44% | 7.659% | 5.029% | -2.1% | -0.5509% | 15.71% | -4.269% | -7.057% | -7.057% | 2.237% | -7.057% | -0.6748% | -4.439% | -8.451% | -3.649% | -23.48% | 5.416% | 12.71% | -2.1% | -5.173% | 472.7%
| Z-Score (Absolute) |  | 1.04 | 2.11 | 1.33 | 0.97 | 0 | 0.21 | 2.42 | 0.3 | 0.67 | 0.67 | 0.59 | 0.67 | 0.19 | 0.32 | 0.86 | 0.21 | 2.91 | 1.02 | 2.01 | 0 | 0.42 | 64.58
| Analytical results for Be in OREAS 237b (Certified Value 1.26 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.22 | 1.8223 | 1.3 | < 5 | 1.2 | 1.37 | 1.49 | 1 | 1.37 | 1.28 | 1.34 | 1.29 | 1.3555 | 1.3 | 1.204 | 1.3 | 1.22 | < 2 | 1.19 | 0.966 | 1.1244 | 1.2119 |  |  | 1
| 1 | 2 | 1.22 | 1.7335 | 1.3 | < 5 | 1.2 | 1.26 | 1.41 | 2 | 1.38 | 1.31 | 1.37 | 1.29 | 1.3912 | 1.3 | 1.166 | 1.2 | 1.18 | < 2 | 1.11 | 0.934 | 1.1028 | 1.1717 |  |  | 28
| 1 | 3 | 1.19 | 1.6171 | 1.3 | < 5 | 1.1 | 1.26 | 1.48 | 1 | 1.3 | 1.23 | 1.32 | 1.31 | 1.298 | 1.3 | 1.177 | 1.4 | 1.19 | < 2 | 1.11 | 0.97 | 1.1246 | 1.2549 |  |  | 16
| 1 | 4 | 1.18 | 1.6983 | 1.4 | < 5 | 1.1 | 1.35 | 1.59 | 1 | 1.44 | 1.28 | 1.32 | 1.26 | 1.4278 | 1.3 | 1.209 | 1.3 | 1.18 | < 2 | 1.07 | 0.991 | 1.103 | 1.2342 |  |  | 1.2618
| 1 | 5 | 1.21 | 1.803 | 1.2 | < 5 | 1.1 | 1.31 | 1.44 | 1 | 1.31 | 1.2 | 1.37 | 1.35 | 1.2591 | 1.3 | 1.172 | 1.3 | 1.17 | < 2 | 1.17 | 0.952 | 1.1031 | 1.1942 |  |  | 76
| 1 | 6 | 1.2 | 1.7564 | 1.3 | < 5 | 1.2 | 1.3 | 1.53 | 1 | 1.34 | 1.24 | 1.3 | 1.33 | 1.3657 | 1.2 | 1.181 | 1.3 | 1.15 | < 2 | 1.13 | 0.984 | 1.1249 | 1.2552
| Mean |  | 1.2033 | 1.7384 | 1.3 |  | 1.15 | 1.3083 | 1.49 | 1.1667 | 1.3567 | 1.2567 | 1.3367 | 1.305 | 1.3495 | 1.2833 | 1.1848 | 1.3 | 1.1817 |  | 1.13 | 0.96617 | 1.1138 | 1.2203
| Median |  | 1.205 | 1.745 | 1.3 |  | 1.15 | 1.305 | 1.485 | 1 | 1.355 | 1.26 | 1.33 | 1.3 | 1.3606 | 1.3 | 1.179 | 1.3 | 1.18 |  | 1.12 | 0.968 | 1.1138 | 1.2231
| Std Dev. |  | 0.01633 | 0.074645 | 0.063246 |  | 0.054772 | 0.04535 | 0.064187 | 0.40825 | 0.05164 | 0.040332 | 0.028752 | 0.032094 | 0.061618 | 0.040825 | 0.017589 | 0.063246 | 0.023166 |  | 0.043818 | 0.020904 | 0.011869 | 0.033837
| Rel.Std.Dev. |  | 1.357% | 4.294% | 4.865% |  | 4.763% | 3.466% | 4.308% | 34.99% | 3.806% | 3.209% | 2.151% | 2.459% | 4.566% | 3.181% | 1.484% | 4.865% | 1.96% |  | 3.878% | 2.164% | 1.066% | 2.773%
| PDM3 |  | -4.632% | 37.78% | 3.03% |  | -8.859% | 3.69% | 18.09% | -7.538% | 7.521% | -0.4048% | 5.935% | 3.426% | 6.955% | 1.709% | -6.098% | 3.03% | -6.349% |  | -10.44% | -23.43% | -11.73% | -3.285%
| Z-Score (Absolute) |  | 0.67 | 3.84 | 0.14 | 10.26 | 1.12 | 0.21 | 1.74 | NA | 0.62 | 0.22 | 0.45 | 0.18 | 0.56 | 0 | 0.83 | 0.14 | 0.86 | 2.39 | 1.29 | 2.67 | 1.43 | 0.53
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 237b (Certified Value 0.33 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.31 | 0.32569 | 0.34 | < 2 | 0.43 | 0.32 | 0.34 | 0.3 | 0.36 | 0.33 | 0.33 | 0.37 | 0.3 | 0.35984 | 5 | 0.404 | 0.33 | 36.58 | < 5 | 0.34 | 0.5496 | 1.1953 |  |  | 1
| 1 | 2 | 0.31 | 0.331 | 0.36 | < 2 | 0.42 | 0.34 | 0.33 | 0.3 | 0.32 | 0.33 | 0.32 | 0.36 | 0.29 | 0.34516 | 2 | 0.441 | 0.35 | 31.46 | < 5 | 0.31 | 0.3737 | 1.8176 |  |  | 29
| 1 | 3 | 0.32 | 0.32175 | 0.36 | < 2 | 0.41 | 0.35 | 0.32 | 0.3 | 0.3 | 0.35 | 0.31 | 0.36 | 0.33 | 0.36621 | 3 | 0.358 | 0.32 | 20.73 | < 5 | 0.34 | 0.4959 | 1.5006 |  |  | 16
| 1 | 4 | 0.33 | 0.33501 | 0.34 | < 2 | 0.41 | 0.33 | 0.31 | 0.3 | 0.38 | 0.32 | 0.33 | 0.37 | 0.3 | 0.35279 | 3 | 0.384 | 0.33 | 29.75 | < 5 | 0.35 | 0.4812 | 1.1957 |  |  | 0.33465
| 1 | 5 | 0.33 | 0.34742 | 0.34 | < 2 | 0.42 | 0.32 | 0.33 | 0.3 | 0.36 | 0.33 | 0.32 | 0.35 | 0.33 | 0.33767 | 4 | 0.435 | 0.34 | 31.78 | < 5 | 0.33 | 0.4153 | 1.0794 |  |  | 77
| 1 | 6 | 0.32 | 0.36006 | 0.36 | < 2 | 0.43 | 0.34 | 0.35 | 0.3 | 0.31 | 0.33 | 0.31 | 0.36 | 0.32 | 0.33983 | 4 | 0.368 | 0.35 | 26.51 | < 5 | 0.31 | 0.3783 | 1.1999
| Mean |  | 0.32 | 0.33682 | 0.35 |  | 0.42 | 0.33333 | 0.33 | 0.3 | 0.33833 | 0.33167 | 0.32 | 0.36167 | 0.31167 | 0.35025 | 3.5 | 0.39833 | 0.33667 | 29.4683 |  | 0.33 | 0.449 | 1.3314
| Median |  | 0.32 | 0.33301 | 0.35 |  | 0.42 | 0.335 | 0.33 | 0.3 | 0.34 | 0.33 | 0.32 | 0.36 | 0.31 | 0.34898 | 3.5 | 0.394 | 0.335 | 30.605 |  | 0.335 | 0.44825 | 1.1978
| Std Dev. |  | 0.0089443 | 0.014427 | 0.010954 |  | 0.0089443 | 0.012111 | 0.014142 | 0 | 0.032506 | 0.0098319 | 0.0089443 | 0.0075277 | 0.017224 | 0.011364 | 1.0488 | 0.034494 | 0.012111 | 5.3866 |  | 0.016733 | 0.070938 | 0.27668
| Rel.Std.Dev. |  | 2.795% | 4.283% | 3.13% |  | 2.13% | 3.633% | 4.285% | 0% | 9.608% | 2.964% | 2.795% | 2.081% | 5.526% | 3.245% | 29.97% | 8.66% | 3.597% | 18.28% |  | 5.071% | 15.8% | 20.78%
| PDM3 |  | -4.377% | 0.6507% | 4.588% |  | 25.51% | -0.3925% | -1.389% | -10.35% | 1.102% | -0.8905% | -4.377% | 8.074% | -6.867% | 4.663% | 945.9% | 19.03% | 0.6036% | 8706% |  | -1.389% | 34.17% | 297.9%
| Z-Score (Absolute) |  | 0.78 | 0.24 | 0.19 | 21.23 | 2.45 | 0.35 | 0.46 | NA | 0.19 | 0.4 | 0.78 | 0.57 | 1.05 | 0.2 | NA | 1.75 | 0.24 | 942.66 | 69.78 | 0.46 | 3.39 | 31.95
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 237b (Certified Value 0.312 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.299 | 0.3209 | 0.37 | 0.355 | 0.31 | 0.32 | 0.32 | 0.33 | 0.3 | 0.32 | 0.32 | 0.31 | 0.33 | 0.33329 | 0.293 | 0.27 | 0.33786 | 0.3 | 0.31 | 0.2697 |  |  |  |  | 1
| 1 | 2 | 0.299 | 0.31987 | 0.37 | 0.35467 | 0.31 | 0.31 | 0.32 | 0.33 | 0.3 | 0.33 | 0.32 | 0.32 | 0.3 | 0.33327 | 0.293 | 0.28 | 0.32511 | 0.3 | 0.31 | 0.2662 |  |  |  |  | #N/A
| 1 | 3 | 0.306 | 0.32415 | 0.37 | 0.35467 | 0.32 | 0.31 | 0.33 | 0.33 | 0.29 | 0.32 | 0.32 | 0.31 | 0.31 | 0.33155 | 0.3 | 0.28 | 0.32473 | 0.31 | 0.3 | 0.2737 |  |  |  |  | 16
| 1 | 4 | 0.299 | 0.32009 | 0.37 | 0.34967 | 0.31 | 0.32 | 0.32 | 0.33 | 0.3 | 0.33 | 0.32 | 0.3 | 0.3 | 0.33207 | 0.3 | 0.28 | 0.31765 | 0.31 | 0.29 | 0.2734 |  |  |  |  | 0.3122
| 1 | 5 | 0.303 | 0.3278 | 0.38 | 0.35067 | 0.32 | 0.32 | 0.33 | 0.33 | 0.29 | 0.32 | 0.33 | 0.31 | 0.34 | 0.32976 | 0.3 | 0.28 | 0.32513 | 0.3 | 0.3 | 0.2646 |  |  |  |  | 78
| 1 | 6 | 0.302 | 0.32085 | 0.37 | 0.357 | 0.32 | 0.31 | 0.32 | 0.33 | 0.3 | 0.31 | 0.32 | 0.31 | 0.31 | 0.32594 | 0.293 | 0.28 | 0.31949 | 0.29 | 0.31 | 0.2744
| Mean |  | 0.30133 | 0.32228 | 0.37167 | 0.35361 | 0.315 | 0.315 | 0.32333 | 0.33 | 0.29667 | 0.32167 | 0.32167 | 0.31 | 0.315 | 0.33098 | 0.2965 | 0.27833 | 0.325 | 0.30167 | 0.30333 | 0.27033
| Median |  | 0.3005 | 0.32087 | 0.37 | 0.35467 | 0.315 | 0.315 | 0.32 | 0.33 | 0.3 | 0.32 | 0.32 | 0.31 | 0.31 | 0.33181 | 0.2965 | 0.28 | 0.32492 | 0.3 | 0.305 | 0.27155
| Std Dev. |  | 0.0028752 | 0.0031155 | 0.0040825 | 0.0028238 | 0.0054772 | 0.0054772 | 0.005164 | 0 | 0.005164 | 0.0075277 | 0.0040825 | 0.0063246 | 0.016432 | 0.0027916 | 0.0038341 | 0.0040825 | 0.0070703 | 0.0075277 | 0.008165 | 0.0041865
| Rel.Std.Dev. |  | 0.9542% | 0.9667% | 1.098% | 0.7986% | 1.739% | 1.739% | 1.597% | 0% | 1.741% | 2.34% | 1.269% | 2.04% | 5.216% | 0.8434% | 1.293% | 1.467% | 2.176% | 2.495% | 2.692% | 1.549%
| PDM3 |  | -3.481% | 3.227% | 19.05% | 13.26% | 0.897% | 0.897% | 3.566% | 5.702% | -4.975% | 3.032% | 3.032% | -0.7046% | 0.897% | 6.015% | -5.029% | -10.85% | 4.099% | -3.374% | -2.84% | -13.41%
| Z-Score (Absolute) |  | 0.74 | 0.39 | 3.06 | 2.08 | 0 | 0 | 0.45 | 0.81 | 0.99 | 0.36 | 0.36 | 0.27 | 0 | 0.86 | 0.99 | 1.98 | 0.54 | 0.72 | 0.63 | 2.41
| Analytical results for Cd in OREAS 237b (Certified Value 0.043 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.035 | 1.17 | 0.03 | 0.6121 | 0.05 | 0.07 | 0.06 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.048529 | < 1 | < 0.1 | 0.02 | 20.99 | 11 | 0.03 | 0.87 | 0.3355 |  |  | 1
| 1 | 2 | 0.035 | 1.2 | 0.05 | 0.62343 | 0.05 | 0.07 | 0.06 | < 0.1 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.048721 | < 1 | < 0.1 | 0.03 | 20.44 | 11 | 0.04 | 0.8 | 0.3381 |  |  | 5
| 1 | 3 | 0.033 | 1.15 | 0.02 | 0.64 | 0.05 | 0.07 | 0.07 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.047964 | < 1 | < 0.1 | 0.04 | 21.33 | 11.3 | 0.04 | 0.88 | 0.3432 |  |  | 16
| 1 | 4 | 0.036 | 1.19 | < 0.01 | 0.61819 | 0.05 | 0.08 | 0.06 | < 0.1 | 0.03 | 0.04 | 0.03 | 0.03 | 0.05 | 0.045984 | < 1 | < 0.1 | 0.03 | 21.36 | 11.2 | 0.04 | 0.78 | 0.3407 |  |  | 0.042794
| 1 | 5 | 0.037 | 1.22 | 0.04 | 0.617 | 0.05 | 0.08 | 0.06 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.03 | 0.04 | 0.047806 | < 1 | < 0.1 | 0.03 | 21.28 | 11.1 | 0.04 | 0.79 | 0.3407 |  |  | 79
| 1 | 6 | 0.031 | 1.1 | 0.07 | 0.61 | 0.05 | 0.07 | 0.06 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.03 | 0.04 | 0.047777 | < 1 | < 0.1 | 0.04 | 19.72 | 11 | 0.04 | 0.77 | 0.3458
| Mean |  | 0.0345 | 1.1717 | 0.042 | 0.62012 | 0.05 | 0.073333 | 0.061667 |  | 0.033333 | 0.036667 | 0.03 | 0.035 | 0.041667 | 0.047797 |  |  | 0.031667 | 20.8533 | 11.1 | 0.038333 | 0.815 | 0.34067
| Median |  | 0.035 | 1.18 | 0.04 | 0.61759 | 0.05 | 0.07 | 0.06 |  | 0.03 | 0.04 | 0.03 | 0.035 | 0.04 | 0.047885 |  |  | 0.03 | 21.135 | 11.05 | 0.04 | 0.795 | 0.3407
| Std Dev. |  | 0.0021679 | 0.042622 | 0.019235 | 0.010829 | 7.6012e-18 | 0.005164 | 0.0040825 |  | 0.005164 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.00096998 |  |  | 0.0075277 | 0.65421 | 0.12649 | 0.0040825 | 0.047645 | 0.0036346
| Rel.Std.Dev. |  | 6.284% | 3.638% | 45.8% | 1.746% | 1.52e-14% | 7.042% | 6.62% |  | 15.49% | 14.08% | 0% | 15.65% | 9.798% | 2.029% |  |  | 23.77% | 3.137% | 1.14% | 10.65% | 5.846% | 1.067%
| PDM3 |  | -19.38% | 2638% | -1.856% | 1349% | 16.84% | 71.36% | 44.1% |  | -22.11% | -14.32% | -29.9% | -18.21% | -2.635% | 11.69% |  |  | -26% | 48629% | 25838% | -10.42% | 1804% | 696.1%
| Z-Score (Absolute) |  | 0.67 | 48.8 | 0.62 | 24.8 | 0 | 1.02 | 0.51 | 0 | 0.73 | 0.58 | 0.87 | 0.65 | 0.36 | 0.1 | 19.58 | 0 | 0.8 | 905.02 | NA | 0.51 | 33.28 | 12.65
| Analytical results for Ce in OREAS 237b (Certified Value 54 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 55.302 | 83.4007 | 48.2 | 65.65 | 50.24 | 69.9 | 55 | 55.1 | 54.6 | 56.6 | 56.1 | 53.3167 | 54.87 | 54.39 | 55.19 | 43.427 | 0.2472 | [7 empty] | 1
| 1 | 2 | 54.666 | 83.3243 | 54.7 | 67.65 | 48.73 | 67.7 | 55 | 56.3 | 56 | 56.2 | 58.4 | 55.7065 | 53.74 | 53.03 | 54.06 | 36.7162 | 0.2394 | [7 empty] | 31
| 1 | 3 | 55.106 | 83.3314 | 55.5 | 67.15 | 40.41 | 73.9 | 56 | 55.1 | 51.2 | 55.3 | 58.7 | 54.4455 | 57.11 | 55.42 | 53.5 | 42.6772 | 0.2382 | [7 empty] | 16
| 1 | 4 | 57.13 | 83.2072 | 52 | 66.5 | 50.02 | 71.5 | 54 | 55.5 | 52.1 | 55.8 | 55.2 | 53.1064 | 59.23 | 53.01 | 49.81 | 44.9484 | 0.2332 | [7 empty] | 54.4696
| 1 | 5 | 55.993 | 83.1565 | 51.7 | 66.42 | 50.75 | 71.3 | 57 | 52.8 | 53.4 | 57.4 | 59.8 | 55.7258 | 55.16 | 58.46 | 51.5 | 37.1151 | 0.2354 | [7 empty] | 80
| 1 | 6 | 56.475 | 83.4856 | 49.4 | 68.13 | 48.45 | 71.2 | 56 | 53.8 | 53.7 | 54.7 | 58.2 | 53.3173 | 60.35 | 56.6 | 51.76 | 33.7544 | 0.2388
| Mean |  | 55.7787 | 83.3176 | 51.9167 | 66.9167 | 48.1 | 70.9167 | 55.5 | 54.7667 | 53.5 | 56 | 57.7333 | 54.2697 | 56.7433 | 55.1517 | 52.6367 | 39.773 | 0.2387
| Median |  | 55.6475 | 83.3278 | 51.85 | 66.825 | 49.375 | 71.25 | 55.5 | 55.1 | 53.55 | 56 | 58.3 | 53.8814 | 56.135 | 54.905 | 52.63 | 39.8961 | 0.2385
| Std Dev. |  | 0.92462 | 0.12124 | 2.8562 | 0.90434 | 3.8715 | 2.042 | 1.0488 | 1.258 | 1.7181 | 0.95708 | 1.7294 | 1.2152 | 2.6215 | 2.1346 | 1.9668 | 4.499 | 0.0047779
| Rel.Std.Dev. |  | 1.658% | 0.1455% | 5.501% | 1.351% | 8.049% | 2.879% | 1.89% | 2.297% | 3.211% | 1.709% | 2.995% | 2.239% | 4.62% | 3.87% | 3.737% | 11.31% | 2.002%
| PDM3 |  | 2.403% | 52.96% | -4.687% | 22.85% | -11.69% | 30.2% | 1.892% | 0.5455% | -1.78% | 2.81% | 5.992% | -0.3669% | 4.174% | 1.252% | -3.365% | -26.98% | -99.56%
| Z-Score (Absolute) |  | 0.17 | 7.55 | 0.87 | 3.15 | 1.89 | 4.23 | 0.09 | 0.1 | 0.44 | 0.23 | 0.69 | 0.24 | 0.43 | 0 | 0.67 | 4.12 | 14.72
| Analytical results for Co in OREAS 237b (Certified Value 15.7 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 15.2 | 16.713 | 16.6 | 17.0968 | 14.8 | 15.5 | 16.1 | 14 | 16 | 15.8 | 15.6 | 15.3 | 16.1 | 16.4899 | 14 | 14.1 | 15.5 | 24.98 | 19.4 | 15 | 16.245 | 14.3884 | 15.4803 |  | 1
| 1 | 2 | 15.32 | 16.641 | 16.9 | 17.1376 | 15.6 | 15.3 | 15.7 | 13.7 | 16.3 | 15.9 | 15.8 | 15.6 | 15.5 | 16.7979 | 13 | 12.8 | 15.3 | 18.87 | 20 | 15.3 | 16.15 | 16.1395 | 15.4908 |  | 32
| 1 | 3 | 14.95 | 16.965 | 16.8 | 17.0053 | 15.1 | 15.7 | 16.6 | 14 | 15.9 | 15.2 | 15.2 | 15.2 | 16.6 | 16.0932 | 13 | 12.6 | 15.9 | 20.66 | 19.7 | 14.8 | 16.625 | 14.1149 | 15.4847 |  | 16
| 1 | 4 | 15.27 | 16.635 | 17 | 17.2547 | 15.1 | 15.5 | 16.2 | 14 | 16.5 | 15.7 | 15.2 | 15 | 15.7 | 16.271 | 14 | 13.26 | 15.4 | 21.11 | 20.2 | 14.9 | 16.53 | 13.7479 | 15.0968 |  | 15.7009
| 1 | 5 | 15.12 | 16.785 | 16.9 | 17.1012 | 15.1 | 15.5 | 15.9 | 14 | 15.8 | 15.6 | 15.8 | 15.5 | 16.1 | 15.6587 | 13 | 14.04 | 14.8 | 22.71 | 19.7 | 15 | 16.34 | 16.0522 | 15.2851 |  | 81
| 1 | 6 | 15.33 | 16.659 | 16.9 | 17.1937 | 15.4 | 15.5 | 15.8 | 13.9 | 16 | 15.4 | 15.1 | 15.2 | 15.9 | 16.1528 | 13 | 13 | 14.9 | 19.34 | 19.5 | 15.6 | 16.815 | 14.3956 | 15.6779
| Mean |  | 15.1983 | 16.733 | 16.85 | 17.1316 | 15.1833 | 15.5 | 16.05 | 13.9333 | 16.0833 | 15.6 | 15.45 | 15.3 | 15.9833 | 16.2439 | 13.3333 | 13.3 | 15.3 | 21.2783 | 19.75 | 15.1 | 16.4508 | 14.8064 | 15.4193
| Median |  | 15.235 | 16.686 | 16.9 | 17.1194 | 15.1 | 15.5 | 16 | 14 | 16 | 15.65 | 15.4 | 15.25 | 16 | 16.2119 | 13 | 13.13 | 15.35 | 20.885 | 19.7 | 15 | 16.435 | 14.392 | 15.4825
| Std Dev. |  | 0.14497 | 0.12686 | 0.13784 | 0.086163 | 0.27869 | 0.12649 | 0.32711 | 0.12111 | 0.26394 | 0.26077 | 0.32094 | 0.21909 | 0.38166 | 0.38514 | 0.5164 | 0.63548 | 0.40497 | 2.268 | 0.30166 | 0.29665 | 0.25075 | 1.0268 | 0.201
| Rel.Std.Dev. |  | 0.9539% | 0.7582% | 0.818% | 0.5029% | 1.835% | 0.8161% | 2.038% | 0.8692% | 1.641% | 1.672% | 2.077% | 1.432% | 2.388% | 2.371% | 3.873% | 4.778% | 2.647% | 10.66% | 1.527% | 1.965% | 1.524% | 6.935% | 1.304%
| PDM3 |  | -3.201% | 6.574% | 7.319% | 9.112% | -3.296% | -1.279% | 2.224% | -11.26% | 2.436% | -0.6425% | -1.598% | -2.553% | 1.799% | 3.459% | -15.08% | -15.29% | -2.553% | 35.52% | 25.79% | -3.827% | 4.777% | -5.697% | -1.794%
| Z-Score (Absolute) |  | 0.46 | 1.54 | 1.7 | 2.06 | 0.48 | 0.07 | 0.65 | 2.11 | 0.7 | 0.07 | 0.13 | 0.33 | 0.57 | 0.91 | NA | 2.94 | 0.33 | 7.48 | 5.48 | 0.59 | 1.18 | 0.97 | 0.17
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 237b (Certified Value 108 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 112 | 106.5425 | 98 | 114.6633 | 98 | 114 | 113 | 108 | 108 | 112 | 109 | 105 | 112.6 | 109.779 | 106 | 116.6 | 101 | 114.56 | 110 | 103.1 | 92.3921 | 108.9536 |  |  | 1
| 1 | 2 | 112.1 | 107.0365 | 99 | 113.5967 | 101 | 110 | 111 | 108 | 106 | 114 | 108 | 107 | 103.7 | 108.3164 | 107 | 115.3 | 99 | 106.98 | 109 | 99.7 | 86.0705 | 109.4605 |  |  | 33
| 1 | 3 | 112.7 | 108.262 | 100 | 113.86 | 98 | 111 | 113 | 111 | 105 | 115 | 108 | 105 | 108.9 | 104.163 | 106 | 111 | 100 | 111.3 | 108 | 103.5 | 91.1807 | 111.0719 |  |  | 16
| 1 | 4 | 112 | 103.759 | 99 | 113.7467 | 100 | 113 | 111 | 110 | 108 | 114 | 109 | 101 | 106.3 | 113.7477 | 109 | 119.2 | 99 | 109.18 | 105 | 104.9 | 91.8276 | 110.0837 |  |  | 107.976
| 1 | 5 | 111.4 | 106.571 | 100 | 114.41 | 99 | 112 | 113 | 111 | 106 | 114 | 110 | 104 | 116.6 | 103.0292 | 107 | 121.2 | 99 | 102.32 | 109 | 100.8 | 86.4381 | 109.0607 |  |  | 82
| 1 | 6 | 111.1 | 106.0105 | 97 | 114.6433 | 101 | 111 | 111 | 113 | 107 | 111 | 108 | 105 | 107.7 | 107.6388 | 103 | 113.6 | 100 | 104.49 | 109 | 104.1 | 87.4035 | 110.4416
| Mean |  | 111.8833 | 106.3636 | 98.8333 | 114.1533 | 99.5 | 111.8333 | 112 | 110.1667 | 106.6667 | 113.3333 | 108.6667 | 104.5 | 109.3 | 107.779 | 106.3333 | 116.15 | 99.6667 | 108.1383 | 108.3333 | 102.6833 | 89.2188 | 109.8453
| Median |  | 112 | 106.5567 | 99 | 114.135 | 99.5 | 111.5 | 112 | 110.5 | 106.5 | 114 | 108.5 | 105 | 108.3 | 107.9776 | 106.5 | 115.95 | 99.5 | 108.08 | 109 | 103.3 | 89.2921 | 109.7721
| Std Dev. |  | 0.56362 | 1.486 | 1.169 | 0.47484 | 1.3784 | 1.472 | 1.0954 | 1.9408 | 1.2111 | 1.5055 | 0.8165 | 1.9748 | 4.6316 | 3.8877 | 1.9664 | 3.7082 | 0.8165 | 4.4905 | 1.7512 | 2.0104 | 2.8867 | 0.83376
| Rel.Std.Dev. |  | 0.5038% | 1.397% | 1.183% | 0.416% | 1.385% | 1.316% | 0.9781% | 1.762% | 1.135% | 1.328% | 0.7514% | 1.89% | 4.238% | 3.607% | 1.849% | 3.193% | 0.8192% | 4.153% | 1.616% | 1.958% | 3.235% | 0.759%
| PDM3 |  | 3.619% | -1.493% | -8.467% | 5.721% | -7.85% | 3.572% | 3.727% | 2.029% | -1.213% | 4.962% | 0.6396% | -3.219% | 1.226% | -0.1825% | -1.521% | 7.57% | -7.696% | 0.1503% | 0.3309% | -4.902% | -17.37% | 1.731%
| Z-Score (Absolute) |  | 0.68 | 0.35 | 1.75 | 1.1 | 1.63 | 0.67 | 0.7 | 0.36 | 0.29 | 0.95 | 0.08 | 0.7 | 0.2 | 0.09 | 0.35 | 1.47 | 1.6 | 0.02 | 0.02 | 1.03 | 3.54 | 0.3
| Analytical results for Cs in OREAS 237b (Certified Value 6.93 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.20 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 6.79 | 6.0193 | 6.65 | 6.75 | 6.73 | 8 | 6.8 | 6.93 | 7.37 | 6.99 | 7.13 | 7.21 | 6.59 | 7.31 | 7.7975 | [9 empty] | 1
| 1 | 2 | 6.81 | 5.8416 | 6.53 | 7.13 | 6.85 | 7.72 | 6.6 | 7.33 | 7.31 | 7.14 | 7.21 | 6.805 | 6.56 | 7.22 | 6.6172 | [9 empty] | 34
| 1 | 3 | 6.81 | 6.1605 | 7.04 | 6.92 | 6.67 | 8.41 | 6.8 | 7 | 6.97 | 6.98 | 7.28 | 6.622 | 6.86 | 7.2 | 7.7544 | [9 empty] | 16
| 1 | 4 | 6.8 | 6.0922 | 6.52 | 6.89 | 6.96 | 8.21 | 6.7 | 7.1 | 7.18 | 6.91 | 7.14 | 7.264 | 6.91 | 7 | 7.8682 | [9 empty] | 6.9333
| 1 | 5 | 6.81 | 6.3389 | 6.75 | 6.89 | 6.86 | 7.96 | 6.8 | 6.96 | 7.13 | 6.93 | 7.41 | 7.176 | 6.71 | 7.36 | 6.529 | [9 empty] | 83
| 1 | 6 | 6.94 | 6.3795 | 6.58 | 7.07 | 6.77 | 7.88 | 6.8 | 7.03 | 7.11 | 6.97 | 7.15 | 6.528 | 6.59 | 7.22 | 6.0895
| Mean |  | 6.8267 | 6.1387 | 6.6783 | 6.9417 | 6.8067 | 8.03 | 6.75 | 7.0583 | 7.1783 | 6.9867 | 7.22 | 6.9342 | 6.7033 | 7.2183 | 7.1093
| Median |  | 6.81 | 6.1264 | 6.615 | 6.905 | 6.81 | 7.98 | 6.8 | 7.015 | 7.155 | 6.975 | 7.18 | 6.9905 | 6.65 | 7.22 | 7.1858
| Std Dev. |  | 0.056095 | 0.20163 | 0.19671 | 0.13747 | 0.10405 | 0.24544 | 0.083666 | 0.14552 | 0.14456 | 0.081158 | 0.10881 | 0.32325 | 0.15069 | 0.12368 | 0.78544
| Rel.Std.Dev. |  | 0.8217% | 3.285% | 2.946% | 1.98% | 1.529% | 3.057% | 1.239% | 2.062% | 2.014% | 1.162% | 1.507% | 4.662% | 2.248% | 1.713% | 11.05%
| PDM3 |  | -1.538% | -11.46% | -3.678% | 0.1202% | -1.827% | 15.82% | -2.644% | 1.803% | 3.534% | 0.7692% | 4.135% | 0.012% | -3.317% | 4.111% | 2.538%
| Z-Score (Absolute) |  | 0.4 | 2.83 | 0.93 | 0 | 0.47 | 3.83 | 0.67 | 0.41 | 0.83 | 0.16 | 0.98 | 0.03 | 0.84 | 0.97 | 0.59
| Analytical results for Cu in OREAS 237b (Certified Value 26.9 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 26.3 | 24.395 | 26.2 | 25.9733 | 28 | 28 | 27 | 25 | 27.6 | 28.2 | 27.2 | 28.6 | 27.81 | 27.2828 | 28 | 25.91 | 24.6 | 27.21 | 31.3 | 28 | 28.0915 | 26 | 1 | 27.6722 | 1
| 1 | 2 | 27.19 | 24.278 | 27.5 | 26.12 | 28 | 28 | 25 | 26.5 | 28 | 27.7 | 27.5 | 29.1 | 26.15 | 29.5326 | 28 | 24.66 | 24.8 | 25.42 | 31.3 | 28.1 | 27.4455 | 27 | 1 | 27.3748 | 35
| 1 | 3 | 27.17 | 24.737 | 28.4 | 26.1 | 28 | 27 | 28 | 25.1 | 27.5 | 26.7 | 28.2 | 28.7 | 27.41 | 28.6477 | 28 | 25.56 | 24.8 | 25.79 | 31.7 | 27 | 28.1105 | 27 | 1 | 27.7239 | 16
| 1 | 4 | 26.87 | 24.197 | 27.1 | 25.82 | 28 | 25 | 27 | 25.3 | 28.2 | 27.3 | 27 | 28.6 | 26.86 | 28.8348 | 29 | 29.48 | 24.9 | 25.99 | 32.9 | 27.8 | 28.5095 | 27 | 1 | 27.6036 | 26.9057
| 1 | 5 | 27.08 | 24.755 | 29.3 | 25.4367 | 28 | 24 | 26 | 25.6 | 27.1 | 26.3 | 27.4 | 29.3 | 29.11 | 27.6993 | 29 | 26.28 | 24.7 | 24.65 | 31.1 | 27.6 | 27.417 | 27 | 1 | 27.3612 | 84
| 1 | 6 | 26.42 | 24.728 | 28.7 | 25.9767 | 28 | 24 | 25 | 25.7 | 27.3 | 26.5 | 26.7 | 28.8 | 26.7 | 28.9776 | 27 | 24.74 | 25.8 | 25.77 | 30.3 | 28.6 | 28.538 | 26 | 1 | 27.1707
| Mean |  | 26.8383 | 24.515 | 27.8667 | 25.9044 | 28 | 26 | 26.3333 | 25.5333 | 27.6167 | 27.1167 | 27.3333 | 28.85 | 27.34 | 28.4958 | 28.1667 | 26.105 | 24.9333 | 25.805 | 31.4333 | 27.85 | 28.0187 | 26.6667 | 1 | 27.4844
| Median |  | 26.975 | 24.5615 | 27.95 | 25.975 | 28 | 26 | 26.5 | 25.45 | 27.55 | 27 | 27.3 | 28.75 | 27.135 | 28.7413 | 28 | 25.735 | 24.8 | 25.78 | 31.3 | 27.9 | 28.101 | 27 | 1 | 27.4892
| Std Dev. |  | 0.38933 | 0.25454 | 1.1431 | 0.25323 | 0 | 1.8974 | 1.2111 | 0.5465 | 0.41673 | 0.74409 | 0.51251 | 0.2881 | 1.0407 | 0.84273 | 0.75277 | 1.7724 | 0.43665 | 0.8357 | 0.85479 | 0.53572 | 0.49293 | 0.5164 | 0 | 0.2156
| Rel.Std.Dev. |  | 1.451% | 1.038% | 4.102% | 0.9775% | 0% | 7.298% | 4.599% | 2.14% | 1.509% | 2.744% | 1.875% | 0.9986% | 3.806% | 2.957% | 2.673% | 6.79% | 1.751% | 3.239% | 2.719% | 1.924% | 1.759% | 1.936% | 0% | 0.7844%
| PDM3 |  | -0.2503% | -8.885% | 3.572% | -3.721% | 4.067% | -3.366% | -2.127% | -5.101% | 2.643% | 0.7842% | 1.589% | 7.226% | 1.614% | 5.91% | 4.687% | -2.976% | -7.331% | -4.091% | 16.83% | 3.51% | 4.137% | -0.8883% | -96.28% | 2.151%
| Z-Score (Absolute) |  | 0.28 | 1.99 | 0.47 | 0.97 | 0.57 | 0.9 | 0.66 | 1.24 | 0.29 | 0.08 | 0.08 | 1.2 | 0.08 | 0.93 | 0.69 | 0.82 | 1.69 | 1.04 | 3.1 | 0.46 | 0.58 | 0.41 | 19.29 | 0.19
| Analytical results for Dy in OREAS 237b (Indicative Value 2.08 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 2.2666 | 1.8 | 2.03 | 2.1104 | [20 empty] | 1
| 1 | 2 | 2.2216 | 2.1 | 2.008 | 1.7797 | [20 empty] | 9
| 1 | 3 | 2.3019 | 2.2 | 2.088 | 2.1211 | [20 empty] | 16
| 1 | 4 | 2.2737 | 2.2 | 2.124 | 2.1275 | [20 empty] | 2.0753
| 1 | 5 | 2.3552 | 1.9 | 2.037 | 1.7888 | [20 empty] | 15
| 1 | 6 | 2.2743 | 1.9 | 2.165 | 1.6347
| Mean |  | 2.2822 | 2.0167 | 2.0753 | 1.927
| Median |  | 2.274 | 2 | 2.0625 | 1.9496
| Std Dev. |  | 0.044185 | 0.17224 | 0.061037 | 0.21806
| Rel.Std.Dev. |  | 1.936% | 8.541% | 2.941% | 11.32%
| PDM3 |  | 9.969% | -2.826% | 0.001153% | -7.145%
| Z-Score (Absolute) |  | 2.15 | 0.27 | 0.27 | 1.08
| Analytical results for Er in OREAS 237b (Indicative Value 0.94 ppm) | [25 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.9992 | 0.9 | 0.926 | 0.9195 | [20 empty] | 1
| 1 | 2 | 0.99187 | 0.9 | 0.907 | 0.7907 | [20 empty] | 10
| 1 | 3 | 0.99316 | 1 | 0.942 | 0.9185 | [20 empty] | 16
| 1 | 4 | 1.0142 | 0.9 | 0.952 | 0.9467 | [20 empty] | 0.93801
| 1 | 5 | 1.0818 | 1 | 0.927 | 0.7883 | [20 empty] | 16
| 1 | 6 | 1.1017 | 0.9 | 0.985 | 0.7276
| Mean |  | 1.0303 | 0.93333 | 0.93983 | 0.84855
| Median |  | 1.0067 | 0.9 | 0.9345 | 0.8546
| Std Dev. |  | 0.048657 | 0.05164 | 0.026933 | 0.090737
| Rel.Std.Dev. |  | 4.723% | 5.533% | 2.866% | 10.69%
| PDM3 |  | 9.841% | -0.4986% | 0.1944% | -9.537%
| Z-Score (Absolute) |  | 1.38 | 0.05 | 0.05 | 1.3
| Analytical results for Eu in OREAS 237b (Indicative Value 0.59 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.7217 | 0.5 | 0.584 | 0.5442 | [20 empty] | 1
| 1 | 2 | 0.69578 | 0.5 | 0.574 | 0.4736 | [20 empty] | 11
| 1 | 3 | 0.69385 | 0.6 | 0.594 | 0.5459 | [20 empty] | 16
| 1 | 4 | 0.73806 | 0.5 | 0.615 | 0.5686 | [20 empty] | 0.58893
| 1 | 5 | 0.71925 | 0.6 | 0.579 | 0.4679 | [20 empty] | 17
| 1 | 6 | 0.77093 | 0.5 | 0.626 | 0.4225
| Mean |  | 0.72326 | 0.53333 | 0.59533 | 0.50378
| Median |  | 0.72047 | 0.5 | 0.589 | 0.5089
| Std Dev. |  | 0.028742 | 0.05164 | 0.020877 | 0.057296
| Rel.Std.Dev. |  | 3.974% | 9.682% | 3.507% | 11.37%
| PDM3 |  | 22.81% | -9.44% | 1.088% | -14.46%
| Z-Score (Absolute) |  | 2.34 | 0.46 | 0.46 | 0.89
| Analytical results for Fe in OREAS 237b (Certified Value 3.43 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.304 | 3.5597 | 3.44 | 3.54 | 3.35 | 3.49 | 3.85 | 3.46 | 3.37 | 3.47 | 3.41 | 3.28 | 3.69 | 3.5679 | 3.357 | 3.2 | 3.5696 | 3.59 | 3.51 | 3.2204 | 4.0637 |  |  |  | 1
| 1 | 2 | 3.371 | 3.5252 | 3.4 | 3.5633 | 3.36 | 3.44 | 3.79 | 3.48 | 3.34 | 3.55 | 3.4 | 3.33 | 3.4 | 3.5861 | 3.364 | 3.2 | 3.4211 | 3.58 | 3.5 | 3.152 | 4.1087 |  |  |  | #N/A
| 1 | 3 | 3.425 | 3.557 | 3.47 | 3.5333 | 3.42 | 3.42 | 3.86 | 3.5 | 3.3 | 3.54 | 3.41 | 3.29 | 3.54 | 3.5655 | 3.385 | 3.2 | 3.427 | 3.61 | 3.48 | 3.2582 | 4.1106 |  |  |  | 16
| 1 | 4 | 3.339 | 3.5621 | 3.5 | 3.56 | 3.46 | 3.49 | 3.76 | 3.46 | 3.4 | 3.53 | 3.4 | 3.16 | 3.46 | 3.575 | 3.441 | 3.18 | 3.3866 | 3.6 | 3.37 | 3.2664 | 4.0865 |  |  |  | 3.4333
| 1 | 5 | 3.35 | 3.5521 | 3.49 | 3.58 | 3.36 | 3.48 | 3.88 | 3.5 | 3.32 | 3.51 | 3.42 | 3.26 | 3.78 | 3.5494 | 3.399 | 3.19 | 3.4575 | 3.58 | 3.49 | 3.1639 | 4.0546 |  |  |  | 85
| 1 | 6 | 3.383 | 3.5163 | 3.43 | 3.59 | 3.45 | 3.4 | 3.79 | 3.49 | 3.37 | 3.4 | 3.39 | 3.27 | 3.5 | 3.5241 | 3.322 | 3.2 | 3.408 | 3.62 | 3.51 | 3.2751 | 4.1785
| Mean |  | 3.362 | 3.5454 | 3.455 | 3.5611 | 3.4 | 3.4533 | 3.8217 | 3.4817 | 3.35 | 3.5 | 3.405 | 3.265 | 3.5617 | 3.5613 | 3.378 | 3.195 | 3.445 | 3.5967 | 3.4767 | 3.2227 | 4.1004
| Median |  | 3.3605 | 3.5546 | 3.455 | 3.5617 | 3.39 | 3.46 | 3.82 | 3.485 | 3.355 | 3.52 | 3.405 | 3.275 | 3.52 | 3.5667 | 3.3745 | 3.2 | 3.424 | 3.595 | 3.495 | 3.2393 | 4.0976
| Std Dev. |  | 0.041299 | 0.019559 | 0.038341 | 0.021976 | 0.049396 | 0.038816 | 0.047924 | 0.018348 | 0.036878 | 0.056569 | 0.010488 | 0.056833 | 0.14483 | 0.021873 | 0.040576 | 0.0083666 | 0.065359 | 0.01633 | 0.053541 | 0.053623 | 0.04452
| Rel.Std.Dev. |  | 1.228% | 0.5517% | 1.11% | 0.6171% | 1.453% | 1.124% | 1.254% | 0.527% | 1.101% | 1.616% | 0.308% | 1.741% | 4.066% | 0.6142% | 1.201% | 0.2619% | 1.897% | 0.454% | 1.54% | 1.664% | 1.086%
| PDM3 |  | -2.077% | 3.264% | 0.6313% | 3.722% | -0.9707% | 0.5827% | 11.31% | 1.408% | -2.427% | 1.942% | -0.8251% | -4.903% | 3.738% | 3.728% | -1.611% | -6.942% | 0.3388% | 4.757% | 1.262% | -6.136% | 19.43%
| Z-Score (Absolute) |  | 0.67 | 0.66 | 0 | 0.77 | 0.4 | 0.01 | 2.66 | 0.19 | 0.76 | 0.33 | 0.36 | 1.38 | 0.77 | 0.77 | 0.56 | 1.89 | 0.07 | 1.03 | 0.16 | 1.68 | 4.68
| Analytical results for Ga in OREAS 237b (Certified Value 8.88 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.27 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 8.9 | 7.8747 | 8.24 | 9.544 | 8.8 | 11 | 8.6 | 10 | 8.43 | 8.99 | 8.88 | 8.57 | 8.9 | 8.9692 | 9.3 | 11.3 | 9.93 | 8.7487 | 8.2114 | [5 empty] | 1
| 1 | 2 | 8.9 | 7.6838 | 8.73 | 9.224 | 9.1 | 12 | 8.2 | 9 | 8.58 | 9.09 | 8.97 | 8.81 | 9.3 | 8.9316 | 8.4 | 11.2 | 9.56 | 7.117 | 8.1308 | [5 empty] | 39
| 1 | 3 | 9.2 | 7.9477 | 8.66 | 9.2613 | 9 | 10 | 8.9 | 10 | 8.47 | 8.59 | 8.78 | 8.62 | 9 | 9.0878 | 9.4 | 11.8 | 9.64 | 8.557 | 8.4696 | [5 empty] | 16
| 1 | 4 | 9 | 7.9024 | 8.7 | 9.392 | 8.8 | 9 | 8.8 | 10 | 8.7 | 8.84 | 8.64 | 8.51 | 9 | 8.6636 | 9.5 | 11.3 | 9.31 | 8.8183 | 8.425 | [5 empty] | 8.8798
| 1 | 5 | 9 | 7.9365 | 8.85 | 9.4053 | 8.9 | 11 | 8.5 | 10 | 8.35 | 8.39 | 9.03 | 8.95 | 9 | 8.8512 | 8.7 | 11.3 | 9.94 | 7.1226 | 7.955 | [5 empty] | 86
| 1 | 6 | 9.1 | 7.8798 | 8.81 | 9.3493 | 9 | 10 | 8.4 | 10 | 8.43 | 8.47 | 8.75 | 8.64 | 9.2 | 9.0158 | 8.7 | 11.1 | 9.6 | 6.6949 | 8.5543
| Mean |  | 9.0167 | 7.8708 | 8.665 | 9.3627 | 8.9333 | 10.5 | 8.5667 | 9.8333 | 8.4933 | 8.7283 | 8.8417 | 8.6833 | 9.0667 | 8.9199 | 9 | 11.3333 | 9.6633 | 7.8431 | 8.291
| Median |  | 9 | 7.8911 | 8.715 | 9.3707 | 8.95 | 10.5 | 8.55 | 10 | 8.45 | 8.715 | 8.83 | 8.63 | 9 | 8.9504 | 9 | 11.3 | 9.62 | 7.8398 | 8.3182
| Std Dev. |  | 0.1169 | 0.096216 | 0.2197 | 0.11423 | 0.12111 | 1.0488 | 0.2582 | 0.40825 | 0.12596 | 0.28708 | 0.1458 | 0.16488 | 0.15055 | 0.14858 | 0.45607 | 0.24221 | 0.23989 | 0.9639 | 0.22983
| Rel.Std.Dev. |  | 1.297% | 1.222% | 2.536% | 1.22% | 1.356% | 9.989% | 3.014% | 4.152% | 1.483% | 3.289% | 1.649% | 1.899% | 1.661% | 1.666% | 5.067% | 2.137% | 2.482% | 12.29% | 2.772%
| PDM3 |  | 1.542% | -11.36% | -2.419% | 5.438% | 0.6031% | 18.25% | -3.526% | 10.74% | -4.352% | -1.705% | -0.4292% | -2.212% | 2.105% | 0.4513% | 1.354% | 27.63% | 8.824% | -11.67% | -6.63%
| Z-Score (Absolute) |  | 0.35 | 2.57 | 0.55 | 1.23 | 0.13 | 4.12 | 0.8 | NA | 0.99 | 0.39 | 0.1 | 0.5 | 0.47 | 0.1 | 0.3 | 6.24 | 1.99 | 2.64 | 1.5
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 237b (Indicative Value 4.05 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 5.739 | 2.9 | 3.671 | 3.8533 | [20 empty] | 1
| 1 | 2 | 5.8322 | 3.3 | 3.614 | 3.2886 | [20 empty] | 12
| 1 | 3 | 5.7156 | 3.1 | 3.761 | 3.7688 | [20 empty] | 16
| 1 | 4 | 5.6642 | 3.3 | 3.855 | 3.9763 | [20 empty] | 4.0497
| 1 | 5 | 5.867 | 3.1 | 3.671 | 3.2995 | [20 empty] | 18
| 1 | 6 | 5.8867 | 3.1 | 3.97 | 2.9588
| Mean |  | 5.7841 | 3.1333 | 3.757 | 3.5242
| Median |  | 5.7856 | 3.1 | 3.716 | 3.5342
| Std Dev. |  | 0.090367 | 0.15055 | 0.13434 | 0.39955
| Rel.Std.Dev. |  | 1.562% | 4.805% | 3.576% | 11.34%
| PDM3 |  | 42.83% | -22.63% | -7.227% | -12.98%
| Z-Score (Absolute) |  | 4.64 | 1.1 | 0.25 | 0.25
| Analytical results for Ge in OREAS 237b (Certified Value 0.11 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.05 | < 0.1 | 0.12 | 1.7 | 0.2 | 0.12 | 0.07 | 0.12 | 0.2 | 0.12 | 0.0855 | [13 empty] | 1
| 1 | 2 | < 0.05 | < 0.1 | 0.11 | 1.7 | 0.21 | 0.13 | 0.07 | 0.11 | 0.2 | 0.12 | 0.0646 | [13 empty] | 6
| 1 | 3 | < 0.05 | < 0.1 | 0.12 | 1.7 | 0.2 | 0.11 | 0.07 | 0.12 | 0.2 | 0.12 | 0.0842 | [13 empty] | 16
| 1 | 4 | < 0.05 | < 0.1 | 0.12 | 1.6 | 0.21 | 0.12 | 0.07 | 0.1 | 0.2 | 0.13 | 0.0894 | [13 empty] | 0.10919
| 1 | 5 | < 0.05 | < 0.1 | 0.12 | 1.7 | 0.22 | 0.11 | 0.08 | 0.12 | 0.2 | 0.16 | 0.0611 | [13 empty] | 87
| 1 | 6 | < 0.05 | < 0.1 | 0.12 | 1.7 | 0.2 | 0.12 | 0.08 | 0.11 | 0.2 | 0.12 | 0.0589
| Mean |  |  |  | 0.11833 | 1.6833 | 0.20667 | 0.11833 | 0.073333 | 0.11333 | 0.2 | 0.12833 | 0.07395
| Median |  |  |  | 0.12 | 1.7 | 0.205 | 0.12 | 0.07 | 0.115 | 0.2 | 0.12 | 0.0744
| Std Dev. |  |  |  | 0.0040825 | 0.040825 | 0.008165 | 0.0075277 | 0.005164 | 0.008165 | 3.0405e-17 | 0.016021 | 0.013829
| Rel.Std.Dev. |  |  |  | 3.45% | 2.425% | 3.951% | 6.361% | 7.042% | 7.204% | 1.52e-14% | 12.48% | 18.7%
| PDM3 |  |  |  | 8.374% | 1442% | 89.27% | 8.374% | -32.84% | 3.795% | 83.17% | 17.53% | -32.27%
| Z-Score (Absolute) |  | 1.51 | 1.08 | 0.09 | NA | 1.6 | 0.09 | 0.68 | 0 | NA | 0.26 | 0.67
Results from laboratories 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 237b (Certified Value 0.49 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.48 | 0.77411 | 0.2 | 0.63 | 0.55 | 0.7 | 0.5 | 0.5 | 0.46 | 0.48 | 0.52 | 0.51 | 0.32 | 0.2893 | 0.4279 | [9 empty] | 1
| 1 | 2 | 0.47 | 0.76783 | 0.4 | 0.67 | 0.49 | 0.73 | 0.5 | 0.52 | 0.47 | 0.48 | 0.53 | 0.49 | 0.34 | 0.2721 | 0.432 | [9 empty] | 25
| 1 | 3 | 0.51 | 0.80703 | 0.4 | 0.67 | 0.52 | 0.7 | 0.5 | 0.51 | 0.46 | 0.46 | 0.54 | 0.52 | 0.31 | 0.3012 | 0.4376 | [9 empty] | 16
| 1 | 4 | 0.5 | 0.79313 | 0.2 | 0.64 | 0.45 | 0.74 | 0.5 | 0.51 | 0.44 | 0.46 | 0.53 | 0.48 | 0.32 | 0.3017 | 0.4294 | [9 empty] | 0.4856
| 1 | 5 | 0.47 | 0.83715 | 0.2 | 0.62 | 0.5 | 0.75 | 0.5 | 0.49 | 0.47 | 0.48 | 0.53 | 0.54 | 0.31 | 0.2601 | 0.4281 | [9 empty] | 88
| 1 | 6 | 0.49 | 0.86084 | 0.2 | 0.65 | 0.43 | 0.72 | 0.5 | 0.48 | 0.44 | 0.46 | 0.51 | 0.53 | 0.31 | 0.2224 | 0.4446
| Mean |  | 0.48667 | 0.80668 | 0.26667 | 0.64667 | 0.49 | 0.72333 | 0.5 | 0.50167 | 0.45667 | 0.47 | 0.52667 | 0.51167 | 0.31833 | 0.27447 | 0.43327
| Median |  | 0.485 | 0.80008 | 0.2 | 0.645 | 0.495 | 0.725 | 0.5 | 0.505 | 0.46 | 0.47 | 0.53 | 0.515 | 0.315 | 0.2807 | 0.4307
| Std Dev. |  | 0.01633 | 0.036374 | 0.10328 | 0.020656 | 0.044272 | 0.020656 | 0 | 0.01472 | 0.013663 | 0.010954 | 0.010328 | 0.023166 | 0.01169 | 0.030329 | 0.0066223
| Rel.Std.Dev. |  | 3.355% | 4.509% | 38.73% | 3.194% | 9.035% | 2.856% | 0% | 2.934% | 2.992% | 2.331% | 1.961% | 4.528% | 3.672% | 11.05% | 1.528%
| PDM3 |  | 0.2197% | 66.12% | -45.09% | 33.17% | 0.9061% | 48.96% | 2.965% | 3.309% | -5.958% | -3.213% | 8.457% | 5.368% | -34.45% | -43.48% | -10.78%
| Z-Score (Absolute) |  | 0.06 | 5.82 | 4.11 | 2.88 | 0 | 4.29 | 0.18 | 0.21 | 0.61 | 0.37 | 0.67 | 0.4 | 3.16 | 3.96 | 1.04
| Analytical results for Hg in OREAS 237b (Indicative Value 0.013 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.101 | 0.01 | 0.02 | 0.01 | 0.01 | 0.022 | < 1 | < 0.01 | < 5 | 0.009 | 0.0172 | [9 empty] | 1
| 1 | 2 | < 0.01 | 0.06 | < 0.05 | < 2 | 0.102 | 0.01 | 0.02 | 0.01 | 0.01 | 0.018 | < 1 | < 0.01 | < 5 | 0.013 | 0.0129 | [9 empty] | 13
| 1 | 3 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.099 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.021 | < 1 | < 0.01 | < 5 | 0.01 | 0.0193 | [9 empty] | 16
| 1 | 4 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.09 | 0.01 | 0.02 | 0.01 | 0.01 | 0.02 | < 1 | < 0.01 | < 5 | 0.006 | 0.0154 | [9 empty] | 0.01329
| 1 | 5 | < 0.01 | 0.06 | < 0.05 | < 2 | 0.099 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.025 | < 1 | < 0.01 | < 5 | 0.009 | 0.0127 | [9 empty] | 19
| 1 | 6 | < 0.01 | 0.03 | < 0.05 | < 2 | 0.098 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.022 | < 1 | < 0.01 | < 5 | 0.007 | 0.0087
| Mean |  |  | 0.05 |  |  | 0.098167 | 0.01 | 0.018333 | 0.01 | 0.01 | 0.021333 |  |  |  | 0.009 | 0.014367
| Median |  |  | 0.05 |  |  | 0.099 | 0.01 | 0.02 | 0.01 | 0.01 | 0.0215 |  |  |  | 0.009 | 0.01415
| Std Dev. |  |  | 0.010954 |  |  | 0.0042622 | 0 | 0.0040825 | 0 | 0 | 0.0023381 |  |  |  | 0.0024495 | 0.0037543
| Rel.Std.Dev. |  |  | 21.91% |  |  | 4.342% | 0% | 22.27% | 0% | 0% | 10.96% |  |  |  | 27.22% | 26.13%
| PDM3 |  |  | 276.2% |  |  | 638.6% | -24.76% | 37.94% | -24.76% | -24.76% | 60.52% |  |  |  | -32.28% | 8.097%
| Z-Score (Absolute) |  | 0.83 | 1.97 | 0.41 | 61.1 | 4.97 | 0.52 | 0 | 0.52 | 0.67 | 0.19 | 29.98 | 0.83 | 154.47 | 0.58 | 0.25
| Analytical results for Ho in OREAS 237b (Indicative Value 0.35 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.39623 | 0.3 | 0.346 | 0.356 | [20 empty] | 1
| 1 | 2 | 0.38731 | 0.3 | 0.34 | 0.3051 | [20 empty] | 14
| 1 | 3 | 0.38582 | 0.3 | 0.353 | 0.3497 | [20 empty] | 16
| 1 | 4 | 0.38946 | 0.4 | 0.359 | 0.3608 | [20 empty] | 0.3473
| 1 | 5 | 0.41173 | 0.3 | 0.346 | 0.2999 | [20 empty] | 20
| 1 | 6 | 0.40125 | 0.3 | 0.369 | 0.2788
| Mean |  | 0.3953 | 0.31667 | 0.35217 | 0.32505
| Median |  | 0.39284 | 0.3 | 0.3495 | 0.3274
| Std Dev. |  | 0.0099412 | 0.040825 | 0.010534 | 0.034679
| Rel.Std.Dev. |  | 2.515% | 12.89% | 2.991% | 10.67%
| PDM3 |  | 13.82% | -8.819% | 1.403% | -6.405%
| Z-Score (Absolute) |  | 2.15 | 0.83 | 0.52 | 0.52
| Analytical results for In in OREAS 237b (Certified Value 0.033 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.03 | 0.03 | < 5 | < 0.05 | 0.027 | < 0.05 | 0.04 | 0.037 | 0.035 | 0.034 | 0.035 | 0.03 | 0.032 | 0.0455 | [10 empty] | 1
| 1 | 2 | 0.03 | 0.03 | < 5 | < 0.05 | 0.025 | < 0.05 | 0.04 | 0.037 | 0.036 | 0.034 | 0.035 | 0.03 | 0.033 | 0.0379 | [10 empty] | 27
| 1 | 3 | 0.03 | 0.04 | < 5 | < 0.05 | 0.03 | < 0.05 | 0.04 | 0.035 | 0.034 | 0.033 | 0.034 | 0.03 | 0.034 | 0.0392 | [10 empty] | 16
| 1 | 4 | 0.03 | 0.03 | < 5 | < 0.05 | 0.022 | < 0.05 | 0.04 | 0.037 | 0.033 | 0.032 | 0.034 | 0.03 | 0.033 | 0.042 | [10 empty] | 0.033359
| 1 | 5 | 0.03 | 0.03 | < 5 | < 0.05 | 0.024 | < 0.05 | 0.04 | 0.037 | 0.033 | 0.033 | 0.035 | 0.03 | 0.036 | 0.0302 | [10 empty] | 89
| 1 | 6 | 0.03 | 0.05 | < 5 | < 0.05 | 0.023 | < 0.05 | 0.04 | 0.036 | 0.033 | 0.033 | 0.036 | 0.03 | 0.034 | 0.0309
| Mean |  | 0.03 | 0.035 |  |  | 0.025167 |  | 0.04 | 0.0365 | 0.034 | 0.033167 | 0.034833 | 0.03 | 0.033667 | 0.037617
| Median |  | 0.03 | 0.03 |  |  | 0.0245 |  | 0.04 | 0.037 | 0.0335 | 0.033 | 0.035 | 0.03 | 0.0335 | 0.03855
| Std Dev. |  | 0 | 0.0083666 |  |  | 0.0029269 |  | 0 | 0.00083666 | 0.0012649 | 0.00075277 | 0.00075277 | 0 | 0.0013663 | 0.0060674
| Rel.Std.Dev. |  | 0% | 23.9% |  |  | 11.63% |  | 0% | 2.292% | 3.72% | 2.27% | 2.161% | 0% | 4.058% | 16.13%
| PDM3 |  | -10.07% | 4.919% |  |  | -24.56% |  | 19.91% | 9.415% | 1.921% | -0.5768% | 4.419% | -10.07% | 0.922% | 12.76%
| Z-Score (Absolute) |  | 0.68 | 0.21 | 436.66 | 1.56 | 1.53 | 1.56 | 1.09 | 0.47 | 0.03 | 0.12 | 0.18 | 0.68 | 0.03 | 0.67
| Analytical results for Ir in OREAS 237b (Indicative Value < 0.0015 ppm) | [25 empty] | Indicative
Ir | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | < 0.0015 | [23 empty] | 1
| 1 | 2 | < 0.0013 | [23 empty] | 4
| 1 | 3 | < 0.0019 | [23 empty] | 16
| 1 | 4 | < 0.0015 | [23 empty] | < 0.0015
| 1 | 5 | < 0.0019 | [23 empty] | 14
| 1 | 6 | < 0.0018
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 237b (Certified Value 0.866 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.823 | 1.0389 | 0.78 | 0.8915 | 0.858 | 0.91 | 0.958 | 0.83 | 0.84 | 0.92 | 0.9 | 0.84 | 0.92 | 0.86927 | 0.855 | 0.8 | 0.89432 | 0.87 | 0.89 | 0.8803 | 1.0078 |  |  |  | 1
| 1 | 2 | 0.825 | 1.0405 | 0.78 | 0.8915 | 0.857 | 0.86 | 0.937 | 0.83 | 0.84 | 0.95 | 0.89 | 0.86 | 0.85 | 0.868 | 0.855 | 0.81 | 0.86027 | 0.87 | 0.89 | 0.8532 | 1.0074 |  |  |  | #N/A
| 1 | 3 | 0.837 | 1.0308 | 0.78 | 0.8765 | 0.883 | 0.87 | 0.955 | 0.84 | 0.83 | 0.95 | 0.9 | 0.85 | 0.88 | 0.85442 | 0.863 | 0.81 | 0.86152 | 0.87 | 0.89 | 0.888 | 1.0075 |  |  |  | 16
| 1 | 4 | 0.832 | 1.0197 | 0.79 | 0.891 | 0.863 | 0.91 | 0.937 | 0.83 | 0.86 | 0.93 | 0.9 | 0.81 | 0.86 | 0.87117 | 0.88 | 0.8 | 0.85415 | 0.87 | 0.85 | 0.9031 | 1.0087 |  |  |  | 0.86646
| 1 | 5 | 0.837 | 1.0421 | 0.82 | 0.8915 | 0.868 | 0.9 | 0.961 | 0.84 | 0.83 | 0.92 | 0.9 | 0.83 | 0.95 | 0.84572 | 0.872 | 0.81 | 0.85752 | 0.87 | 0.88 | 0.8704 | 1.0083 |  |  |  | 90
| 1 | 6 | 0.837 | 1.0026 | 0.78 | 0.8965 | 0.888 | 0.89 | 0.94 | 0.84 | 0.85 | 0.9 | 0.89 | 0.84 | 0.88 | 0.85927 | 0.822 | 0.81 | 0.84598 | 0.87 | 0.89 | 0.8965 | 1.008
| Mean |  | 0.83183 | 1.0291 | 0.78833 | 0.88975 | 0.8695 | 0.89 | 0.948 | 0.835 | 0.84167 | 0.92833 | 0.89667 | 0.83833 | 0.89 | 0.86131 | 0.85783 | 0.80667 | 0.86229 | 0.87 | 0.88167 | 0.88192 | 1.008
| Median |  | 0.8345 | 1.0349 | 0.78 | 0.8915 | 0.8655 | 0.895 | 0.9475 | 0.835 | 0.84 | 0.925 | 0.9 | 0.84 | 0.88 | 0.86364 | 0.859 | 0.81 | 0.8589 | 0.87 | 0.89 | 0.88415 | 1.0079
| Std Dev. |  | 0.0064005 | 0.015426 | 0.016021 | 0.0068099 | 0.013096 | 0.020976 | 0.011171 | 0.0054772 | 0.01169 | 0.019408 | 0.005164 | 0.017224 | 0.037947 | 0.01 | 0.020094 | 0.005164 | 0.016645 | 1.2162e-16 | 0.016021 | 0.018212 | 0.00049295
| Rel.Std.Dev. |  | 0.7694% | 1.499% | 2.032% | 0.7654% | 1.506% | 2.357% | 1.178% | 0.656% | 1.389% | 2.091% | 0.5759% | 2.055% | 4.264% | 1.161% | 2.342% | 0.6402% | 1.93% | 1.398e-14% | 1.817% | 2.065% | 0.04891%
| PDM3 |  | -3.996% | 18.77% | -9.016% | 2.688% | 0.3512% | 2.717% | 9.411% | -3.631% | -2.861% | 7.141% | 3.487% | -3.246% | 2.717% | -0.5941% | -0.9953% | -6.901% | -0.4805% | 0.4089% | 1.755% | 1.784% | 16.33%
| Z-Score (Absolute) |  | 0.97 | 4.02 | 2.07 | 0.5 | 0.01 | 0.51 | 1.97 | 0.89 | 0.72 | 1.48 | 0.67 | 0.8 | 0.51 | 0.22 | 0.31 | 1.6 | 0.19 | 0 | 0.3 | 0.3 | 3.49
| Analytical results for La in OREAS 237b (Certified Value 27.5 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 27.959 | 36.5204 | 23.5 | 25.57 | 33.51 | 26 | 34.2 | 28 | 27.5 | 26.7 | 27.4 | 28.2 | 29.9 | 26.3313 | 17 | 27 | 26.7 | 32.9 | 24.3 | 24.5167 |  |  |  |  | 1
| 1 | 2 | 27.552 | 35.4412 | 25.9 | 25.5833 | 34.15 | 21 | 32.6 | 28 | 28.1 | 26.6 | 28.3 | 28.8 | 27.6 | 26.5925 | 18 | 26.47 | 26.8 | 32.8 | 23.2 | 20.8913 |  |  |  |  | 28
| 1 | 3 | 27.411 | 37.2574 | 27.1 | 25.5333 | 34.04 | 21 | 35.8 | 27 | 27.5 | 25 | 27.6 | 29.1 | 28.4 | 26.3583 | 17 | 28.04 | 26.4 | 32.3 | 23.1 | 24.1878 |  |  |  |  | 16
| 1 | 4 | 28.344 | 36.9135 | 25.3 | 25.2367 | 33.43 | 26 | 34.5 | 28 | 27.6 | 24.7 | 27.8 | 27.8 | 26.4 | 26.1185 | 18 | 29 | 26.4 | 33.1 | 21.6 | 25.566 |  |  |  |  | 27.5305
| 1 | 5 | 27.62 | 36.6663 | 25.4 | 25.49 | 34.03 | 26 | 34.6 | 28 | 26.1 | 25.5 | 28.3 | 29.3 | 29.8 | 27.1941 | 18 | 27.01 | 29.8 | 32.9 | 22.1 | 20.9705 |  |  |  |  | 91
| 1 | 6 | 27.925 | 37.3453 | 24.3 | 25.8567 | 33.92 | 25 | 34.1 | 29 | 26.8 | 25.3 | 27.7 | 29.1 | 27.6 | 26.2475 | 17 | 29.5 | 26.8 | 31.8 | 22.2 | 19.1878
| Mean |  | 27.8018 | 36.6907 | 25.25 | 25.545 | 33.8467 | 24.1667 | 34.3 | 28 | 27.2667 | 25.6333 | 27.85 | 28.7167 | 28.2833 | 26.4737 | 17.5 | 27.8367 | 27.15 | 32.6333 | 22.75 | 22.5533
| Median |  | 27.7725 | 36.7899 | 25.35 | 25.5517 | 33.975 | 25.5 | 34.35 | 28 | 27.5 | 25.4 | 27.75 | 28.95 | 28 | 26.3448 | 17.5 | 27.525 | 26.75 | 32.85 | 22.65 | 22.5791
| Std Dev. |  | 0.34119 | 0.69139 | 1.2518 | 0.19873 | 0.30177 | 2.4833 | 1.0315 | 0.63246 | 0.70616 | 0.83347 | 0.37283 | 0.59133 | 1.3717 | 0.38574 | 0.54772 | 1.2176 | 1.3111 | 0.48854 | 0.97724 | 2.5376
| Rel.Std.Dev. |  | 1.227% | 1.884% | 4.958% | 0.7779% | 0.8916% | 10.28% | 3.007% | 2.259% | 2.59% | 3.251% | 1.339% | 2.059% | 4.85% | 1.457% | 3.13% | 4.374% | 4.829% | 1.497% | 4.296% | 11.25%
| PDM3 |  | 0.9854% | 33.27% | -8.284% | -7.212% | 22.94% | -12.22% | 24.59% | 1.705% | -0.9585% | -6.891% | 1.16% | 4.308% | 2.734% | -3.839% | -36.43% | 1.112% | -1.382% | 18.54% | -17.36% | -18.08%
| Z-Score (Absolute) |  | 0.09 | 3.17 | 0.79 | 0.69 | 2.19 | 1.17 | 2.35 | 0.16 | 0.09 | 0.66 | 0.11 | 0.41 | 0.26 | 0.37 | 3.48 | 0.1 | 0.13 | 1.77 | 1.66 | 1.73
| Analytical results for Li in OREAS 237b (Certified Value 43 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [6 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 41.74 | 48.73 | 46.5 | 46.17 | 39.8 | 46 | 49.3 | 42.9 | 44.2 | 43.2 | 42.3 | 43.8 | 43.6539 | 39.29 | 40 | 45.2 | 37.4 | 39.5937 | [6 empty] | 1
| 1 | 2 | 41.26 | 48.89 | 45.1 | 45.37 | 41.4 | 43 | 47 | 42.1 | 45.1 | 43.4 | 42 | 43.1 | 41.3878 | 38.76 | 40 | 42.5 | 35.9 | 39.6223 | [6 empty] | #N/A
| 1 | 3 | 40.48 | 49.44 | 46.6 | 45.75 | 39.7 | 44 | 51.2 | 46.2 | 44.4 | 42.1 | 40.9 | 44.6 | 40.0347 | 38.69 | 40 | 42.7 | 37.6 | 41.5074 | [6 empty] | 16
| 1 | 4 | 41.98 | 47.98 | 47.4 | 45.71 | 41.1 | 46 | 47.8 | 43.1 | 45.5 | 41.9 | 40.6 | 43.5 | 41.796 | 40.09 | 40 | 41.8 | 38.2 | 40.3768 | [6 empty] | 42.9679
| 1 | 5 | 41.78 | 49.19 | 44.9 | 45.535 | 40.4 | 46 | 47.5 | 45.4 | 43.8 | 40.7 | 42.2 | 44.5 | 44.3377 | 39.68 | 40 | 42.8 | 36.7 | 39.1815 | [6 empty] | 92
| 1 | 6 | 41.83 | 48.51 | 45.3 | 45.9 | 41.6 | 46 | 47.4 | 44.7 | 44.1 | 41 | 41.8 | 44.2 | 39.8804 | 39.18 | 40 | 41.6 | 38.1 | 40.9124
| Mean |  | 41.5117 | 48.79 | 45.9667 | 45.7392 | 40.6667 | 45.1667 | 48.3667 | 44.0667 | 44.5167 | 42.05 | 41.6333 | 43.95 | 41.8484 | 39.2817 | 40 | 42.7667 | 37.3167 | 40.199
| Median |  | 41.76 | 48.81 | 45.9 | 45.73 | 40.75 | 46 | 47.65 | 43.9 | 44.3 | 42 | 41.9 | 44 | 41.5919 | 39.235 | 40 | 42.6 | 37.5 | 39.9995
| Std Dev. |  | 0.5608 | 0.51587 | 1.0073 | 0.27944 | 0.81894 | 1.3292 | 1.5983 | 1.6058 | 0.64936 | 1.1041 | 0.71181 | 0.58907 | 1.8349 | 0.53753 | 0 | 1.2879 | 0.87958 | 0.89297
| Rel.Std.Dev. |  | 1.351% | 1.057% | 2.191% | 0.6109% | 2.014% | 2.943% | 3.305% | 3.644% | 1.459% | 2.626% | 1.71% | 1.34% | 4.385% | 1.368% | 0% | 3.011% | 2.357% | 2.221%
| PDM3 |  | -3.389% | 13.55% | 6.979% | 6.45% | -5.356% | 5.117% | 12.56% | 2.557% | 3.604% | -2.136% | -3.106% | 2.286% | -2.605% | -8.579% | -6.907% | -0.4684% | -13.15% | -6.444%
| Z-Score (Absolute) |  | 0.28 | 1.99 | 1.11 | 1.04 | 0.54 | 0.86 | 1.86 | 0.52 | 0.66 | 0.11 | 0.24 | 0.48 | 0.17 | 0.98 | 0.75 | 0.11 | 1.59 | 0.69
| Analytical results for Lu in OREAS 237b (Certified Value 0.12 ppm) | [25 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.12562 | 0.1 | 0.101 | 0.11 | 0.1778 | [19 empty] | 1
| 1 | 2 | 0.129 | 0.1 | 0.101 | 0.11 | 0.1506 | [19 empty] | 29
| 1 | 3 | 0.12552 | 0.1 | 0.106 | 0.12 | 0.1527 | [19 empty] | 16
| 1 | 4 | 0.13145 | 0.1 | 0.108 | 0.11 | 0.1368 | [19 empty] | 0.12039
| 1 | 5 | 0.13246 | 0.1 | 0.102 | 0.12 | 0.1543 | [19 empty] | 93
| 1 | 6 | 0.13362 | 0.1 | 0.11 | 0.12 | 0.1437
| Mean |  | 0.12961 | 0.1 | 0.10467 | 0.115 | 0.15265
| Median |  | 0.13022 | 0.1 | 0.104 | 0.115 | 0.15165
| Std Dev. |  | 0.0034821 | 1.5202e-17 | 0.0038816 | 0.0054772 | 0.013934
| Rel.Std.Dev. |  | 2.687% | 1.52e-14% | 3.709% | 4.763% | 9.128%
| PDM3 |  | 7.663% | -16.93% | -13.06% | -4.474% | 26.8%
| Z-Score (Absolute) |  | 0.67 | 0.69 | 0.48 | 0 | 1.74
| Analytical results for Mg in OREAS 237b (Certified Value 1.37 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.188 | 1.3318 | 1.4 | 1.3967 | 1.31 | 1.44 | 1.48 | 1.29 | 1.33 | 1.38 | 1.38 | 1.32 | 1.45 | 1.4442 | 1.399 | 1.22 | 1.3749 | 1.41 | 1.4 | 1.4725 | 1.4899 |  |  |  | 1
| 1 | 2 | 1.207 | 1.3231 | 1.35 | 1.39 | 1.31 | 1.4 | 1.49 | 1.3 | 1.33 | 1.4 | 1.37 | 1.34 | 1.37 | 1.4568 | 1.393 | 1.22 | 1.3268 | 1.4 | 1.4 | 1.4365 | 1.4814 |  |  |  | #N/A
| 1 | 3 | 1.223 | 1.3291 | 1.4 | 1.3867 | 1.34 | 1.45 | 1.53 | 1.3 | 1.31 | 1.4 | 1.37 | 1.31 | 1.42 | 1.4578 | 1.423 | 1.22 | 1.3165 | 1.41 | 1.39 | 1.4865 | 1.4799 |  |  |  | 16
| 1 | 4 | 1.23 | 1.2955 | 1.47 | 1.38 | 1.32 | 1.47 | 1.48 | 1.3 | 1.34 | 1.4 | 1.38 | 1.27 | 1.39 | 1.5026 | 1.453 | 1.22 | 1.2975 | 1.42 | 1.34 | 1.5002 | 1.4693 |  |  |  | 1.3737
| 1 | 5 | 1.227 | 1.3291 | 1.37 | 1.3933 | 1.32 | 1.45 | 1.46 | 1.3 | 1.32 | 1.38 | 1.38 | 1.3 | 1.51 | 1.4582 | 1.429 | 1.22 | 1.3229 | 1.4 | 1.39 | 1.4485 | 1.4573 |  |  |  | 94
| 1 | 6 | 1.252 | 1.3023 | 1.36 | 1.4067 | 1.34 | 1.44 | 1.49 | 1.3 | 1.33 | 1.34 | 1.36 | 1.32 | 1.41 | 1.4673 | 1.399 | 1.22 | 1.3044 | 1.38 | 1.39 | 1.4986 | 1.5362
| Mean |  | 1.2212 | 1.3185 | 1.3917 | 1.3922 | 1.3233 | 1.4417 | 1.4883 | 1.2983 | 1.3267 | 1.3833 | 1.3733 | 1.31 | 1.425 | 1.4645 | 1.416 | 1.22 | 1.3238 | 1.4033 | 1.385 | 1.4738 | 1.4857
| Median |  | 1.225 | 1.3261 | 1.385 | 1.3917 | 1.32 | 1.445 | 1.485 | 1.3 | 1.33 | 1.39 | 1.375 | 1.315 | 1.415 | 1.458 | 1.411 | 1.22 | 1.3197 | 1.405 | 1.39 | 1.4795 | 1.4807
| Std Dev. |  | 0.021757 | 0.015587 | 0.043551 | 0.0091084 | 0.013663 | 0.023166 | 0.023166 | 0.0040825 | 0.010328 | 0.023381 | 0.008165 | 0.023664 | 0.049699 | 0.020098 | 0.023212 | 0 | 0.027366 | 0.013663 | 0.022583 | 0.026482 | 0.027187
| Rel.Std.Dev. |  | 1.782% | 1.182% | 3.129% | 0.6542% | 1.032% | 1.607% | 1.557% | 0.3144% | 0.7785% | 1.69% | 0.5945% | 1.806% | 3.488% | 1.372% | 1.639% | 0% | 2.067% | 0.9736% | 1.631% | 1.797% | 1.83%
| PDM3 |  | -11.1% | -4.016% | 1.311% | 1.351% | -3.664% | 4.951% | 8.348% | -5.484% | -3.421% | 0.7041% | -0.02392% | -4.634% | 3.737% | 6.612% | 3.082% | -11.19% | -3.628% | 2.16% | 0.8254% | 7.29% | 8.154%
| Z-Score (Absolute) |  | 1.81 | 0.73 | 0.07 | 0.08 | 0.68 | 0.62 | 1.14 | 0.96 | 0.64 | 0.02 | 0.13 | 0.83 | 0.44 | 0.88 | 0.34 | 1.82 | 0.67 | 0.2 | 0 | 0.98 | 1.11
| Analytical results for Mn in OREAS 237b (Certified Value 0.028 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02865 | 0.029037 | 0.0293 | 0.030292 | 0.0291 | 0.0289 | 0.0284 | 0.0318 | 0.0281 | 0.0286 | 0.0281 | 0.0264 | 0.0296 | 0.029661 | 0.0256 | 0.0273 | 0.028129 | 0.03237 | 0.0288 | 0.026 | 0.0333 |  |  |  | 1
| 1 | 2 | 0.02929 | 0.028651 | 0.0292 | 0.0304 | 0.0295 | 0.0278 | 0.028 | 0.0319 | 0.0278 | 0.0295 | 0.028 | 0.0271 | 0.0274 | 0.030109 | 0.0256 | 0.0276 | 0.027192 | 0.03231 | 0.0286 | 0.02595 | 0.0335 |  |  |  | 30
| 1 | 3 | 0.029 | 0.029207 | 0.0292 | 0.030486 | 0.029 | 0.0278 | 0.0289 | 0.0317 | 0.0272 | 0.029 | 0.0282 | 0.0269 | 0.0285 | 0.029845 | 0.0256 | 0.0276 | 0.026986 | 0.03207 | 0.0284 | 0.02659 | 0.0335 |  |  |  | 16
| 1 | 4 | 0.02916 | 0.028609 | 0.03 | 0.030203 | 0.0298 | 0.0287 | 0.028 | 0.0322 | 0.0283 | 0.029 | 0.028 | 0.0257 | 0.0278 | 0.030109 | 0.0263 | 0.0274 | 0.026484 | 0.03201 | 0.0273 | 0.02686 | 0.0331 |  |  |  | 0.02827
| 1 | 5 | 0.02887 | 0.029201 | 0.0307 | 0.030316 | 0.0295 | 0.0287 | 0.0288 | 0.0318 | 0.0275 | 0.0289 | 0.0285 | 0.0264 | 0.0303 | 0.030088 | 0.0263 | 0.0279 | 0.027058 | 0.03173 | 0.0282 | 0.02575 | 0.033 |  |  |  | 95
| 1 | 6 | 0.02928 | 0.02887 | 0.0294 | 0.030435 | 0.0299 | 0.0281 | 0.0281 | 0.0326 | 0.028 | 0.0281 | 0.0278 | 0.0265 | 0.0281 | 0.029913 | 0.0256 | 0.0275 | 0.026702 | 0.03241 | 0.0285 | 0.02698 | 0.0339
| Mean |  | 0.029042 | 0.028929 | 0.029633 | 0.030355 | 0.029467 | 0.028333 | 0.028367 | 0.032 | 0.027817 | 0.02885 | 0.0281 | 0.0265 | 0.028617 | 0.029954 | 0.025833 | 0.02755 | 0.027092 | 0.03215 | 0.0283 | 0.026355 | 0.033383
| Median |  | 0.02908 | 0.028953 | 0.02935 | 0.030358 | 0.0295 | 0.0284 | 0.02825 | 0.03185 | 0.0279 | 0.02895 | 0.02805 | 0.02645 | 0.0283 | 0.030001 | 0.0256 | 0.02755 | 0.027022 | 0.03219 | 0.02845 | 0.026295 | 0.0334
| Std Dev. |  | 0.00025183 | 0.00026305 | 0.00060222 | 0.00010399 | 0.00036148 | 0.00049261 | 0.00040332 | 0.00034059 | 0.00040702 | 0.00046797 | 0.00023664 | 0.0004858 | 0.0011161 | 0.00018196 | 0.00036148 | 0.00020736 | 0.00056913 | 0.0002623 | 0.00052915 | 0.00052095 | 0.00032506
| Rel.Std.Dev. |  | 0.8671% | 0.9093% | 2.032% | 0.3426% | 1.227% | 1.739% | 1.422% | 1.064% | 1.463% | 1.622% | 0.8421% | 1.833% | 3.9% | 0.6075% | 1.399% | 0.7527% | 2.101% | 0.8159% | 1.87% | 1.977% | 0.9737%
| PDM3 |  | 2.731% | 2.332% | 4.824% | 7.378% | 4.234% | 0.2251% | 0.343% | 13.2% | -1.603% | 2.053% | -0.6003% | -6.26% | 1.227% | 5.958% | -8.618% | -2.546% | -4.167% | 13.73% | 0.1071% | -6.773% | 18.09%
| Z-Score (Absolute) |  | 0.28 | 0.21 | 0.67 | 1.15 | 0.56 | 0.19 | 0.17 | 2.24 | 0.53 | 0.15 | 0.34 | 1.4 | 0 | 0.89 | 1.86 | 0.71 | 1.01 | 2.34 | 0.21 | 1.5 | 3.16
| Analytical results for Mo in OREAS 237b (Certified Value 0.88 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.86 | 0.91 | < 2 | 0.8 | 0.86 | 0.9 | 0.9 | 0.9 | 0.89 | 0.79 | 0.89 | 0.87 | 0.9316 | < 1 | 1.066 | 0.82 | < 1 | 0.89 | 0.7172 | 2.2848 |  |  |  |  | 1
| 1 | 2 | 0.91 | 0.84 | < 2 | 0.8 | 0.87 | 0.9 | 0.8 | 0.87 | 0.88 | 0.8 | 0.86 | 0.87 | 0.95978 | < 1 | 1.075 | 0.79 | < 1 | 0.91 | 0.6108 | 2.0684 |  |  |  |  | 31
| 1 | 3 | 0.9 | 0.98 | < 2 | 0.8 | 0.87 | 0.9 | 0.8 | 0.86 | 0.86 | 0.78 | 0.86 | 0.86 | 0.89096 | < 1 | 1.04 | 0.91 | < 1 | 0.84 | 0.7535 | 2.8812 |  |  |  |  | 16
| 1 | 4 | 0.86 | 0.87 | < 2 | 0.8 | 0.89 | 0.9 | 0.8 | 0.93 | 0.86 | 0.78 | 0.86 | 0.84 | 0.87296 | < 1 | 1.106 | 0.86 | < 1 | 0.87 | 0.7457 | 2.719 |  |  |  |  | 0.87621
| 1 | 5 | 0.87 | 0.84 | < 2 | 0.8 | 0.76 | 1 | 0.8 | 0.85 | 0.86 | 0.87 | 1 | 0.94 | 0.86156 | < 1 | 1.095 | 0.75 | < 1 | 0.92 | 0.6043 | 2.4483 |  |  |  |  | 96
| 1 | 6 | 0.91 | 0.88 | < 2 | 0.8 | 0.81 | 1 | 0.9 | 0.87 | 0.86 | 0.78 | 0.85 | 0.92 | 0.87593 | < 1 | 0.963 | 0.84 | < 1 | 0.93 | 0.5395 | 2.5569
| Mean |  | 0.885 | 0.88667 |  | 0.8 | 0.84333 | 0.93333 | 0.83333 | 0.88 | 0.86833 | 0.8 | 0.88667 | 0.88333 | 0.8988 |  | 1.0575 | 0.82833 |  | 0.89333 | 0.66183 | 2.4931
| Median |  | 0.885 | 0.875 |  | 0.8 | 0.865 | 0.9 | 0.8 | 0.87 | 0.86 | 0.785 | 0.86 | 0.87 | 0.88344 |  | 1.0705 | 0.83 |  | 0.9 | 0.664 | 2.5026
| Std Dev. |  | 0.02429 | 0.052789 |  | 1.2162e-16 | 0.048854 | 0.05164 | 0.05164 | 0.029665 | 0.013292 | 0.035214 | 0.057155 | 0.038297 | 0.038558 |  | 0.051718 | 0.055648 |  | 0.033862 | 0.088749 | 0.29376
| Rel.Std.Dev. |  | 2.745% | 5.954% |  | 1.52e-14% | 5.793% | 5.533% | 6.197% | 3.371% | 1.531% | 4.402% | 6.446% | 4.336% | 4.29% |  | 4.891% | 6.718% |  | 3.791% | 13.41% | 11.78%
| PDM3 |  | 1.004% | 1.194% |  | -8.697% | -3.752% | 6.52% | -4.893% | 0.4331% | -0.8984% | -8.697% | 1.194% | 0.8135% | 2.579% |  | 20.69% | -5.464% |  | 1.955% | -24.47% | 184.5%
| Z-Score (Absolute) |  | 0.04 | 0.07 | 1.6 | 1.1 | 0.52 | 0.7 | 0.65 | 0.02 | 0.18 | 1.1 | 0.07 | 0.02 | 0.23 | 5.15 | 2.37 | 0.72 | 5.15 | 0.16 | 2.96 | 21.73
| Analytical results for Na in OREAS 237b (Certified Value 0.082 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.082 | 0.13164 | 0.093 | 0.08 | 0.09 | 0.08 | 0.11 | 0.11 | 0.08 | 0.09 | 0.09 | 0.07 | 0.092 | 0.080317 | 0.074 | 0.08 | 0.090554 | 0.08 | 0.09 | 0.0728 | 0.0687 |  |  |  | 1
| 1 | 2 | 0.082 | 0.1306 | 0.092 | 0.073333 | 0.09 | 0.07 | 0.11 | 0.11 | 0.08 | 0.09 | 0.09 | 0.08 | 0.085 | 0.085787 | 0.074 | 0.08 | 0.086845 | 0.08 | 0.09 | 0.0711 | 0.0686 |  |  |  | #N/A
| 1 | 3 | 0.083 | 0.13021 | 0.094 | 0.073333 | 0.09 | 0.07 | 0.11 | 0.13 | 0.08 | 0.09 | 0.09 | 0.07 | 0.088 | 0.084994 | 0.074 | 0.08 | 0.087559 | 0.08 | 0.09 | 0.0734 | 0.0677 |  |  |  | 16
| 1 | 4 | 0.083 | 0.12883 | 0.105 | 0.076667 | 0.09 | 0.08 | 0.11 | 0.13 | 0.08 | 0.09 | 0.09 | 0.07 | 0.084 | 0.079932 | 0.074 | 0.08 | 0.084238 | 0.08 | 0.08 | 0.0746 | 0.0659 |  |  |  | 0.081824
| 1 | 5 | 0.082 | 0.13188 | 0.1 | 0.08 | 0.09 | 0.08 | 0.11 | 0.12 | 0.08 | 0.09 | 0.09 | 0.07 | 0.093 | 0.084322 | 0.074 | 0.08 | 0.086711 | 0.08 | 0.09 | 0.0711 | 0.0676 |  |  |  | 97
| 1 | 6 | 0.085 | 0.12903 | 0.096 | 0.08 | 0.09 | 0.08 | 0.11 | 0.12 | 0.08 | 0.08 | 0.09 | 0.07 | 0.087 | 0.081265 | 0.067 | 0.08 | 0.083571 | 0.08 | 0.09 | 0.0735 | 0.0689
| Mean |  | 0.082833 | 0.13036 | 0.096667 | 0.077222 | 0.09 | 0.076667 | 0.11 | 0.12 | 0.08 | 0.088333 | 0.09 | 0.071667 | 0.088167 | 0.082769 | 0.072833 | 0.08 | 0.08658 | 0.08 | 0.088333 | 0.07275 | 0.0679
| Median |  | 0.0825 | 0.1304 | 0.095 | 0.078333 | 0.09 | 0.08 | 0.11 | 0.12 | 0.08 | 0.09 | 0.09 | 0.07 | 0.0875 | 0.082794 | 0.074 | 0.08 | 0.086778 | 0.08 | 0.09 | 0.0731 | 0.06815
| Std Dev. |  | 0.001169 | 0.0012775 | 0.0049666 | 0.0032773 | 1.5202e-17 | 0.005164 | 0 | 0.0089443 | 0 | 0.0040825 | 1.5202e-17 | 0.0040825 | 0.003656 | 0.0025609 | 0.0028577 | 0 | 0.0025051 | 0 | 0.0040825 | 0.0014039 | 0.0011189
| Rel.Std.Dev. |  | 1.411% | 0.9799% | 5.138% | 4.244% | 1.689e-14% | 6.736% | 0% | 7.454% | 0% | 4.622% | 1.689e-14% | 5.696% | 4.147% | 3.094% | 3.924% | 0% | 2.893% | 0% | 4.622% | 1.93% | 1.648%
| PDM3 |  | 1.233% | 59.32% | 18.14% | -5.624% | 9.992% | -6.303% | 34.43% | 46.66% | -2.229% | 7.955% | 9.992% | -12.41% | 7.751% | 1.155% | -10.99% | -2.229% | 5.812% | -2.229% | 7.955% | -11.09% | -17.02%
| Z-Score (Absolute) |  | 0 | 5.2 | 1.51 | 0.61 | 0.78 | 0.67 | 2.97 | 4.06 | 0.31 | 0.6 | 0.78 | 1.22 | 0.58 | 0.01 | 1.08 | 0.31 | 0.41 | 0.31 | 0.6 | 1.1 | 1.63
| Analytical results for Nb in OREAS 237b (Certified Value 0.39 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.49 | 0.6 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.42 | 0.39 | 0.27 | 0.34 | 0.5303 | 0.48 | 0.33 | 0.3305 | [9 empty] | 1
| 1 | 2 | 0.5 | 0.5 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.4 | 0.45 | 0.3 | 0.36 | 0.54957 | 0.51 | 0.3 | 0.296 | [9 empty] | 32
| 1 | 3 | 0.51 | 0.3 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.36 | 0.35 | 0.29 | 0.37 | 0.53715 | 0.51 | 0.26 | 0.3038 | [9 empty] | 16
| 1 | 4 | 0.54 | 0.6 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.39 | 0.35 | 0.29 | 0.36 | 0.56648 | 0.46 | 0.27 | 0.3163 | [9 empty] | 0.38659
| 1 | 5 | 0.49 | 0.4 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.37 | 0.38 | 0.31 | 0.37 | 0.53063 | 0.59 | 0.76 | 0.2711 | [9 empty] | 98
| 1 | 6 | 0.51 | 0.5 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.36 | 0.36 | 0.28 | 0.36 | 0.56345 | 0.48 | 0.29 | 0.2491
| Mean |  | 0.50667 | 0.48333 |  | 0.7 |  | 0.4 | 0.2 | 0.38333 | 0.38 | 0.29 | 0.36 | 0.54626 | 0.505 | 0.36833 | 0.29447
| Median |  | 0.505 | 0.5 |  | 0.7 |  | 0.4 | 0.2 | 0.38 | 0.37 | 0.29 | 0.36 | 0.54336 | 0.495 | 0.295 | 0.2999
| Std Dev. |  | 0.018619 | 0.1169 |  | 1.2162e-16 |  | 6.0809e-17 | 3.0405e-17 | 0.024221 | 0.037947 | 0.014142 | 0.010954 | 0.016107 | 0.045935 | 0.19343 | 0.029895
| Rel.Std.Dev. |  | 3.675% | 24.19% |  | 1.737e-14% |  | 1.52e-14% | 1.52e-14% | 6.319% | 9.986% | 4.877% | 3.043% | 2.949% | 9.096% | 52.52% | 10.15%
| PDM3 |  | 31.06% | 25.03% |  | 81.07% |  | 3.469% | -48.27% | -0.842% | -1.704% | -24.98% | -6.878% | 41.3% | 30.63% | -4.722% | -23.83%
| Z-Score (Absolute) |  | 0.69 | 0.54 | 13.49 | 1.93 | 0.64 | 0 | 1.28 | 0.11 | 0.13 | 0.71 | 0.26 | 0.94 | 0.67 | 0.2 | 0.68
| Analytical results for Nd in OREAS 237b (Indicative Value 23.3 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 1
| 1 | 1 | 27.4408 | 20.1 | 23.58 | 21.573 | [20 empty] | 1
| 1 | 2 | 26.1093 | 23.8 | 22.94 | 17.9752 | [20 empty] | 9
| 1 | 3 | 27.3419 | 24.7 | 24.39 | 21.021 | [20 empty] | 16
| 1 | 4 | 27.2852 | 22 | 25.27 | 22.128 | [20 empty] | 23.2618
| 1 | 5 | 27.4614 | 20.9 | 23.58 | 18.2716 | [20 empty] | 15
| 1 | 6 | 27.1696 | 20.9 | 25.63 | 16.7154
| Mean |  | 27.1347 | 22.0667 | 24.2317 | 19.614
| Median |  | 27.3135 | 21.45 | 23.985 | 19.6463
| Std Dev. |  | 0.51354 | 1.8184 | 1.0561 | 2.2373
| Rel.Std.Dev. |  | 1.893% | 8.241% | 4.358% | 11.41%
| PDM3 |  | 16.65% | -5.138% | 4.17% | -15.68%
| Z-Score (Absolute) |  | 1.16 | 0.32 | 0.32 | 1.03
| Analytical results for Ni in OREAS 237b (Certified Value 60 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 57.7 | 52.8295 | 58.1 | 55.3367 | 58 | 64 | 65 | 63.7 | 58.8 | 62.6 | 62.2 | 57.7 | 63.4 | 62.0163 | 60 | 43.59 | 59.4 | 75.45 | 56.21 | 64 | 54.53 | 57.404 | 61.0686 |  | 1
| 1 | 2 | 57.8 | 53.5895 | 59.8 | 55.2233 | 59 | 62 | 64 | 62.8 | 59.7 | 62.4 | 61.8 | 58.1 | 58.5 | 64.7874 | 60 | 43.47 | 57.4 | 71.95 | 57.4 | 64.9 | 53.105 | 58.067 | 61.1617 |  | 34
| 1 | 3 | 57.8 | 54.0835 | 61.4 | 54.62 | 58 | 62 | 66 | 63.9 | 58 | 61.1 | 61.2 | 58.1 | 62.2 | 62.2102 | 60 | 42.75 | 60.9 | 73.05 | 57.32 | 62.6 | 55.765 | 57.2101 | 61.3445 |  | 16
| 1 | 4 | 57.7 | 53.2335 | 60.7 | 55.2867 | 59 | 64 | 64 | 63.7 | 60.1 | 61.3 | 60.8 | 57.6 | 60.7 | 63.2108 | 61 | 44.5 | 62.1 | 72.09 | 57.5 | 63.6 | 55.29 | 57.6678 | 59.9151 |  | 59.8278
| 1 | 5 | 57.8 | 52.8865 | 59.9 | 55.27 | 58 | 62 | 65 | 63.9 | 59.1 | 60.5 | 62.7 | 59.2 | 65.1 | 65.2869 | 61 | 44.42 | 60.2 | 68.32 | 57.27 | 63.9 | 53.865 | 59.9746 | 61.148 |  | 99
| 1 | 6 | 57.9 | 53.8745 | 59.6 | 55.4533 | 59 | 63 | 64 | 63.5 | 58.8 | 60.4 | 60.7 | 58.6 | 60.7 | 65.1593 | 58 | 43.35 | 57.4 | 69.64 | 56.21 | 66.1 | 55.48 | 58.7067 | 62.0344
| Mean |  | 57.7833 | 53.4162 | 59.9167 | 55.1983 | 58.5 | 62.8333 | 64.6667 | 63.5833 | 59.0833 | 61.3833 | 61.5667 | 58.2167 | 61.7667 | 63.7785 | 60 | 43.68 | 59.5667 | 71.75 | 56.985 | 64.1833 | 54.6725 | 58.1717 | 61.1121
| Median |  | 57.8 | 53.4115 | 59.85 | 55.2783 | 58.5 | 62.5 | 64.5 | 63.7 | 58.95 | 61.2 | 61.5 | 58.1 | 61.45 | 63.9991 | 60 | 43.53 | 59.8 | 72.02 | 57.295 | 63.95 | 54.91 | 57.8674 | 61.1549
| Std Dev. |  | 0.075277 | 0.51837 | 1.1161 | 0.294 | 0.54772 | 0.98319 | 0.8165 | 0.41191 | 0.7414 | 0.93256 | 0.80166 | 0.59805 | 2.3201 | 1.4891 | 1.0954 | 0.67022 | 1.8981 | 2.52 | 0.60533 | 1.1957 | 1.035 | 1.0311 | 0.68454
| Rel.Std.Dev. |  | 0.1303% | 0.9704% | 1.863% | 0.5326% | 0.9363% | 1.565% | 1.263% | 0.6478% | 1.255% | 1.519% | 1.302% | 1.027% | 3.756% | 2.335% | 1.826% | 1.534% | 3.186% | 3.512% | 1.062% | 1.863% | 1.893% | 1.772% | 1.12%
| PDM3 |  | -3.417% | -10.72% | 0.1485% | -7.738% | -2.219% | 5.024% | 8.088% | 6.277% | -1.244% | 2.6% | 2.906% | -2.693% | 3.241% | 6.603% | 0.2878% | -26.99% | -0.4365% | 19.93% | -4.752% | 7.28% | -8.617% | -2.768% | 2.147%
| Z-Score (Absolute) |  | 0.67 | 2.05 | 0 | 1.49 | 0.45 | 0.92 | 1.5 | 1.16 | 0.26 | 0.46 | 0.52 | 0.54 | 0.58 | 1.22 | 0.03 | 5.13 | 0.11 | 3.74 | 0.93 | 1.35 | 1.66 | 0.55 | 0.38
| Analytical results for P in OREAS 237b (Certified Value 0.058 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.0531 | 0.051 | 0.057572 | 0.055 | 0.0597 | 0.0575 | 0.059 | 0.058 | 0.059 | 0.058 | 0.057 | 0.06 | 0.059173 | 0.0567 | 0.05 | 0.057802 | 0.06 | 0.0582 | 0.0631 | 0.0641 |  |  |  |  | 1
| 1 | 2 | 0.0542 | 0.053 | 0.058004 | 0.055 | 0.0589 | 0.0568 | 0.059 | 0.058 | 0.059 | 0.058 | 0.059 | 0.055 | 0.062428 | 0.0572 | 0.05 | 0.056886 | 0.06 | 0.0582 | 0.0624 | 0.0643 |  |  |  |  | #N/A
| 1 | 3 | 0.0555 | 0.054 | 0.057483 | 0.057 | 0.0594 | 0.0584 | 0.06 | 0.057 | 0.06 | 0.059 | 0.058 | 0.058 | 0.06025 | 0.0541 | 0.05 | 0.056701 | 0.06 | 0.0573 | 0.0639 | 0.0642 |  |  |  |  | 16
| 1 | 4 | 0.0542 | 0.054 | 0.057597 | 0.056 | 0.06 | 0.058 | 0.06 | 0.058 | 0.06 | 0.059 | 0.055 | 0.056 | 0.062893 | 0.0589 | 0.05 | 0.057459 | 0.06 | 0.0554 | 0.0639 | 0.0637 |  |  |  |  | 0.057965
| 1 | 5 | 0.0556 | 0.053 | 0.058021 | 0.056 | 0.0594 | 0.058 | 0.059 | 0.057 | 0.059 | 0.059 | 0.057 | 0.062 | 0.05923 | 0.055 | 0.05 | 0.057852 | 0.06 | 0.0576 | 0.0623 | 0.0643 |  |  |  |  | 100
| 1 | 6 | 0.0551 | 0.052 | 0.057815 | 0.057 | 0.0592 | 0.0566 | 0.06 | 0.058 | 0.057 | 0.058 | 0.058 | 0.057 | 0.059166 | 0.0554 | 0.05 | 0.056884 | 0.06 | 0.0578 | 0.0646 | 0.0649
| Mean |  | 0.054617 | 0.052833 | 0.057749 | 0.056 | 0.059433 | 0.05755 | 0.0595 | 0.057667 | 0.059 | 0.0585 | 0.057333 | 0.058 | 0.060523 | 0.056217 | 0.05 | 0.057264 | 0.06 | 0.057417 | 0.063367 | 0.06425
| Median |  | 0.05465 | 0.053 | 0.057706 | 0.056 | 0.0594 | 0.05775 | 0.0595 | 0.058 | 0.059 | 0.0585 | 0.0575 | 0.0575 | 0.05974 | 0.05605 | 0.05 | 0.057173 | 0.06 | 0.0577 | 0.0635 | 0.06425
| Std Dev. |  | 0.00096212 | 0.001169 | 0.00023154 | 0.00089443 | 0.00038297 | 0.00072042 | 0.00054772 | 0.0005164 | 0.0010954 | 0.00054772 | 0.0013663 | 0.0026077 | 0.0017121 | 0.0017337 | 7.6012e-18 | 0.00050546 | 0 | 0.0010477 | 0.00092014 | 0.00038859
| Rel.Std.Dev. |  | 1.762% | 2.213% | 0.4009% | 1.597% | 0.6444% | 1.252% | 0.9205% | 0.8955% | 1.857% | 0.9363% | 2.383% | 4.496% | 2.829% | 3.084% | 1.52e-14% | 0.8827% | 0% | 1.825% | 1.452% | 0.6048%
| PDM3 |  | -5.777% | -8.854% | -0.3736% | -3.391% | 2.533% | -0.7165% | 2.648% | -0.5152% | 1.785% | 0.9224% | -1.09% | 0.05981% | 4.413% | -3.017% | -13.74% | -1.21% | 3.51% | -0.9465% | 9.318% | 10.84%
| Z-Score (Absolute) |  | 1.3 | 2.06 | 0.02 | 0.72 | 0.73 | 0.07 | 0.76 | 0.02 | 0.55 | 0.33 | 0.16 | 0.12 | 1.19 | 0.63 | 3.25 | 0.19 | 0.97 | 0.12 | 2.39 | 2.76
| Analytical results for Pb in OREAS 237b (Certified Value 9.97 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  | 3
| 1 | 1 | 9.9 | 10.6 | 9.8533 | 9.9 | 10 | 9.5 | 10.2 | 9.9 | 9.6 | 10.2 | 10.3 | 9.69 | 10.6429 | 8 | 26 | 10.1 | 1.98 | 8.7 | 9.2 | 11.5 | 11 | 8.8096 |  |  | 1
| 1 | 2 | 9.9 | 11 | 9.5267 | 10.2 | 8 | 9.6 | 9.9 | 10.1 | 9.4 | 10.3 | 10.2 | 9.3 | 11.2141 | 9 | 17.11 | 9.7 | 5.24 | 9 | 9 | 11.1 | 11 | 7.9469 |  |  | 35
| 1 | 3 | 10.1 | 11.1 | 10.5433 | 9.8 | 10 | 9.5 | 10.2 | 9.9 | 9.9 | 10 | 10.2 | 10.18 | 10.4505 | 8 | 20.08 | 10 | 2.82 | 8.9 | 9 | 12.2 | 12 | 8.6347 |  |  | 16
| 1 | 4 | 10.1 | 10.8 | 9.7867 | 9.9 | 8 | 9.3 | 11.7 | 10.3 | 9.8 | 10.6 | 10.1 | 11.09 | 11.6155 | 9 | 43.62 | 9.6 | 5.2 | 8.4 | 8.9 | 11.8 | 11 | 9.2833 |  |  | 9.9661
| 1 | 5 | 10.1 | 10.5 | 9.3867 | 10.1 | 10 | 9.5 | 11.9 | 9.7 | 9.9 | 10.3 | 10.2 | 10.61 | 11.1704 | 9 | 22.54 | 10 |  | 8.3 | 9.3 | 11 | 12 | 8.7346 |  |  | 101
| 1 | 6 | 10.1 | 10.3 | 9.6933 | 10.1 | 9 | 9.7 | 11.8 | 10.2 | 9.4 | 10 | 10.2 | 10.39 | 11.0891 | 9 | 14.47 | 10.1 |  | 8.7 | 9.2 | 11.3 | 11 | 9.187
| Mean |  | 10.0333 | 10.7167 | 9.7983 | 10 | 9.1667 | 9.5167 | 10.95 | 10.0167 | 9.6667 | 10.2333 | 10.2 | 10.21 | 11.0304 | 8.6667 | 23.97 | 9.9167 | 3.81 | 8.6667 | 9.1 | 11.4833 | 11.3333 | 8.766
| Median |  | 10.1 | 10.7 | 9.74 | 10 | 9.5 | 9.5 | 10.95 | 10 | 9.7 | 10.25 | 10.2 | 10.285 | 11.1298 | 9 | 21.31 | 10 | 4.01 | 8.7 | 9.1 | 11.4 | 11 | 8.7721
| Std Dev. |  | 0.10328 | 0.30605 | 0.40311 | 0.15492 | 0.98319 | 0.13292 | 0.93968 | 0.22286 | 0.23381 | 0.22509 | 0.063246 | 0.64284 | 0.42085 | 0.5164 | 10.4377 | 0.2137 | 1.6639 | 0.27325 | 0.15492 | 0.4535 | 0.5164 | 0.47663
| Rel.Std.Dev. |  | 1.029% | 2.856% | 4.114% | 1.549% | 10.73% | 1.397% | 8.582% | 2.225% | 2.419% | 2.2% | 0.6201% | 6.296% | 3.815% | 5.958% | 43.54% | 2.155% | 43.67% | 3.153% | 1.702% | 3.949% | 4.556% | 5.437%
| PDM3 |  | 0.6744% | 7.531% | -1.684% | 0.3399% | -8.022% | -4.51% | 9.872% | 0.5072% | -3.005% | 2.681% | 2.347% | 2.447% | 10.68% | -13.04% | 140.5% | -0.4962% | -61.77% | -13.04% | -8.691% | 15.22% | 13.72% | -12.04%
| Z-Score (Absolute) |  | 0.02 | 0.62 | 0.18 | 0.01 | 0.73 | 0.43 | 0.82 | 0.01 | 0.3 | 0.2 | 0.17 | 0.18 | 0.89 | 1.17 | 12.15 | 0.08 | 5.39 | 1.17 | 0.79 | 1.28 | 1.15 | 1.08
| Analytical results for Pd in OREAS 237b (Indicative Value 26.3 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 1
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 1
| 1 | 1 | < 1 | < 10 | 29.6 | [21 empty] | 1
| 1 | 2 | < 1 | < 10 | 24.2 | [21 empty] | 1
| 1 | 3 | < 1 | < 10 | 33.2 | [21 empty] | 16
| 1 | 4 | < 1 | < 10 | 27.8 | [21 empty] | 26.3167
| 1 | 5 | < 1 | < 10 | 14.4 | [21 empty] | 16
| 1 | 6 | < 1 | < 10 | 28.7
| Mean |  |  |  | 26.3167
| Median |  |  |  | 28.25
| Std Dev. |  |  |  | 6.5196
| Rel.Std.Dev. |  |  |  | 24.77%
| PDM3 |  |  |  | -1.332e-13%
| Z-Score (Absolute) |  | 0.67 | 0 | 3.19
| Analytical results for Pr in OREAS 237b (Indicative Value 6.24 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 7.2127 | 5.7 | 6.269 | 5.6678 | [20 empty] | 1
| 1 | 2 | 7.2755 | 6.2 | 6.351 | 4.8136 | [20 empty] | 11
| 1 | 3 | 7.205 | 6.5 | 6.568 | 5.5423 | [20 empty] | 16
| 1 | 4 | 7.3429 | 6.1 | 6.707 | 5.9012 | [20 empty] | 6.2427
| 1 | 5 | 7.4572 | 5.9 | 6.116 | 4.801 | [20 empty] | 17
| 1 | 6 | 7.3564 | 5.6 | 6.834 | 4.4033
| Mean |  | 7.3083 | 6 | 6.4742 | 5.1882
| Median |  | 7.3092 | 6 | 6.4595 | 5.178
| Std Dev. |  | 0.096486 | 0.33466 | 0.27497 | 0.59499
| Rel.Std.Dev. |  | 1.32% | 5.578% | 4.247% | 11.47%
| PDM3 |  | 17.07% | -3.887% | 3.708% | -16.89%
| Z-Score (Absolute) |  | 1.12 | 0.25 | 0.25 | 1.1
| Analytical results for Pt in OREAS 237b (Indicative Value 2.18 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 2
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 2
| 1 | 1 | 2 | < 5 | 1.1 | [21 empty] | 1
| 1 | 2 | 2 | < 5 | 2.4 | [21 empty] | 1
| 1 | 3 | 2 | < 5 | 3.6 | [21 empty] | 16
| 1 | 4 | 2 | < 5 | 4 | [21 empty] | 2.1833
| 1 | 5 | 2 | < 5 | 1.6 | [21 empty] | 18
| 1 | 6 | 2 | < 5 | 1.5
| Mean |  | 2 |  | 2.3667
| Median |  | 2 |  | 2
| Std Dev. |  | 0 |  | 1.1944
| Rel.Std.Dev. |  | 0% |  | 50.47%
| PDM3 |  | -8.397% |  | 8.397%
| Z-Score (Absolute) |  | 1.85 | 0.67 | 0
| Analytical results for Rb in OREAS 237b (Certified Value 85 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 84.47 | 72.0671 | 64.7 | 84.95 | 79.8 | 97.9 | 80.6 | 86.8 | 89.7 | 85.6 | 83.2 | 84.1802 | 71.3 | 86.7 | 87.6 | 91.8899 | 83.9425 | [7 empty] | 1
| 1 | 2 | 84.21 | 72.4936 | 65 | 86.87 | 76 | 94.1 | 80.2 | 87.7 | 88.3 | 85.9 | 84.9 | 85.1968 | 70.93 | 81.6 | 89.6 | 78.921 | 83.8016 | [7 empty] | 37
| 1 | 3 | 87.13 | 73.1473 | 66.1 | 84.34 | 80.6 | 99.8 | 81.3 | 85.7 | 85.1 | 85.7 | 82.2 | 84.1545 | 71.59 | 87.3 | 87 | 91.7069 | 88.271 | [7 empty] | 16
| 1 | 4 | 85.41 | 71.4657 | 65.7 | 84.29 | 82.8 | 96.8 | 80.1 | 87.8 | 85.2 | 85 | 83.2 | 85.2099 | 73.65 | 87.9 | 86.7 | 93.0614 | 84.8673 | [7 empty] | 84.9409
| 1 | 5 | 85.6 | 72.3222 | 67.2 | 84.72 | 81.8 | 95.9 | 81.6 | 85.1 | 84.7 | 85.7 | 86.7 | 85.601 | 74.7 | 80.4 | 90 | 77.733 | 83.0389 | [7 empty] | 102
| 1 | 6 | 85 | 72.4012 | 66.4 | 86.94 | 80.5 | 94.2 | 81.6 | 85.6 | 84.9 | 84.8 | 84.7 | 85.083 | 71.34 | 84.3 | 90.8 | 71.5392 | 86.0979
| Mean |  | 85.3033 | 72.3162 | 65.85 | 85.3517 | 80.25 | 96.45 | 80.9 | 86.45 | 86.3167 | 85.45 | 84.15 | 84.9042 | 72.2517 | 84.7 | 88.6167 | 84.1419 | 85.0032
| Median |  | 85.205 | 72.3617 | 65.9 | 84.835 | 80.55 | 96.35 | 80.95 | 86.25 | 85.15 | 85.65 | 83.95 | 85.1399 | 71.465 | 85.5 | 88.6 | 85.314 | 84.4049
| Std Dev. |  | 1.041 | 0.55042 | 0.9225 | 1.2279 | 2.339 | 2.2061 | 0.68702 | 1.1502 | 2.1321 | 0.44159 | 1.6109 | 0.59713 | 1.5408 | 3.1388 | 1.7302 | 9.2085 | 1.9143
| Rel.Std.Dev. |  | 1.22% | 0.7611% | 1.401% | 1.439% | 2.915% | 2.287% | 0.8492% | 1.331% | 2.47% | 0.5168% | 1.914% | 0.7033% | 2.133% | 3.706% | 1.952% | 10.94% | 2.252%
| PDM3 |  | 0.4267% | -14.86% | -22.48% | 0.4836% | -5.523% | 13.55% | -4.757% | 1.777% | 1.62% | 0.5994% | -0.9311% | -0.04314% | -14.94% | -0.2836% | 4.327% | -0.9406% | 0.07335%
| Z-Score (Absolute) |  | 0.19 | 6.01 | 9.1 | 0.21 | 2.22 | 5.51 | 1.91 | 0.74 | 0.67 | 0.26 | 0.36 | 0 | 6.04 | 0.1 | 1.77 | 0.36 | 0.05
| Analytical results for Re in OREAS 237b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0009 | [12 empty] | 1
| 1 | 2 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.001 | [12 empty] | 38
| 1 | 3 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0011 | [12 empty] | 16
| 1 | 4 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0006 | [12 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0005 | [12 empty] | 103
| 1 | 6 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.001
| Mean | [7 empty] | 0.001 |  |  |  |  | 0.00085
| Median | [7 empty] | 0.001 |  |  |  |  | 0.00095
| Std Dev. | [7 empty] | 0 |  |  |  |  | 0.0002429
| Rel.Std.Dev. | [7 empty] | 0% |  |  |  |  | 28.58%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 131.49 | 2.02 | 131.49 | 10.11 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 1.21
| Analytical results for Rh in OREAS 237b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Rh | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.0005 | [23 empty] | 1
| 1 | 2 | 0.0005 | [23 empty] | 13
| 1 | 3 | 0.0008 | [23 empty] | 16
| 1 | 4 | 0.0014 | [23 empty] | 0.00061667
| 1 | 5 | 0.0002 | [23 empty] | 19
| 1 | 6 | 0.0003
| Mean |  | 0.00061667
| Median |  | 0.0005
| Std Dev. |  | 0.00043551
| Rel.Std.Dev. |  | 70.62%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 237b (Indicative Value < 0.0037 ppm) | [25 empty] | Indicative
Ru | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | < 0.0037 | [23 empty] | 1
| 1 | 2 | 0.0019 | [23 empty] | 14
| 1 | 3 | < 0.0019 | [23 empty] | 16
| 1 | 4 | < 0.0047 | [23 empty] | < 0.0037
| 1 | 5 | < 0.0037 | [23 empty] | 20
| Mean |  | 0.0019
| Median |  | 0.0019
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 237b (Certified Value 0.192 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.18 | 0.20434 | 0.176 | 0.16267 | 0.19 | 0.18 | 0.19 | 0.21 | 0.22 | 0.21 | 0.2 | 0.19 | 0.21 | 0.165 | 0.18 | 0.006086 | 0.17 | 0.2 | 0.2231 | 0.2189 |  |  |  |  | 1
| 1 | 2 | 0.18 | 0.20297 | 0.175 | 0.16267 | 0.19 | 0.16 | 0.19 | 0.21 | 0.21 | 0.21 | 0.2 | 0.19 | 0.19 | 0.163 | 0.18 | 0.005691 | 0.18 | 0.2 | 0.2223 | 0.8119 |  |  |  |  | 23
| 1 | 3 | 0.19 | 0.20722 | 0.177 | 0.16267 | 0.19 | 0.17 | 0.2 | 0.21 | 0.21 | 0.21 | 0.2 | 0.18 | 0.2 | 0.154 | 0.18 | 0.005866 | 0.18 | 0.2 | 0.2266 | 0.1948 |  |  |  |  | 16
| 1 | 4 | 0.19 | 0.20983 | 0.18 | 0.16267 | 0.19 | 0.17 | 0.2 | 0.21 | 0.21 | 0.21 | 0.2 | 0.18 | 0.19 | 0.165 | 0.18 | 0.006076 | 0.18 | 0.19 | 0.2262 | 0.1931 |  |  |  |  | 0.19179
| 1 | 5 | 0.18 | 0.20596 | 0.175 | 0.16267 | 0.19 | 0.17 | 0.2 | 0.21 | 0.21 | 0.21 | 0.2 | 0.18 | 0.21 | 0.158 | 0.18 | 0.005934 | 0.18 | 0.2 | 0.2198 | 0.2164 |  |  |  |  | 104
| 1 | 6 | 0.18 | 0.2069 | 0.175 | 0.16267 | 0.2 | 0.17 | 0.19 | 0.21 | 0.21 | 0.2 | 0.2 | 0.18 | 0.2 | 0.157 | 0.18 | 0.005641 | 0.17 | 0.2 | 0.2274 | 0.2065
| Mean |  | 0.18333 | 0.2062 | 0.17633 | 0.16267 | 0.19167 | 0.17 | 0.195 | 0.21 | 0.21167 | 0.20833 | 0.2 | 0.18333 | 0.2 | 0.16033 | 0.18 | 0.0058823 | 0.17667 | 0.19833 | 0.22423 | 0.30693
| Median |  | 0.18 | 0.20643 | 0.1755 | 0.16267 | 0.19 | 0.17 | 0.195 | 0.21 | 0.21 | 0.21 | 0.2 | 0.18 | 0.2 | 0.1605 | 0.18 | 0.0059 | 0.18 | 0.2 | 0.22465 | 0.21145
| Std Dev. |  | 0.005164 | 0.0023926 | 0.0019664 | 0 | 0.0040825 | 0.0063246 | 0.0054772 | 0 | 0.0040825 | 0.0040825 | 3.0405e-17 | 0.005164 | 0.0089443 | 0.0046332 | 3.0405e-17 | 0.00018803 | 0.005164 | 0.0040825 | 0.0029723 | 0.24761
| Rel.Std.Dev. |  | 2.817% | 1.16% | 1.115% | 0% | 2.13% | 3.72% | 2.809% | 0% | 1.929% | 1.96% | 1.52e-14% | 2.817% | 4.472% | 2.89% | 1.689e-14% | 3.196% | 2.923% | 2.058% | 1.326% | 80.67%
| PDM3 |  | -4.41% | 7.514% | -8.06% | -15.19% | -0.0652% | -11.36% | 1.673% | 9.494% | 10.36% | 8.625% | 4.28% | -4.41% | 4.28% | -16.4% | -6.148% | -96.93% | -7.886% | 3.411% | 16.92% | 60.03%
| Z-Score (Absolute) |  | 0.43 | 0.55 | 0.72 | 1.31 | 0.07 | 0.99 | 0.07 | 0.71 | 0.78 | 0.64 | 0.28 | 0.43 | 0.28 | 1.41 | 0.57 | 7.98 | 0.71 | 0.21 | 1.32 | 4.84
| Analytical results for Sb in OREAS 237b (Certified Value 343 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 0
| 1 | 1 | 386.6 | 266 | 369.22 | 295.81 | 217.36 | 302 | 366.96 | 369 | 385 | 365 | 362 | 432.41 | 395.2478 | 279 | 353.7 | 287 | 422.01 | 406.4 | 309.74 | 252.2 | 165.5603 |  |  |  | 1
| 1 | 2 | 397.26 | 273 | 370.96 | 304.97 | 195.96 | 296 | 334.14 | 356 | 404 | 370 | 378 | 393.69 | 398.6528 | 285 | 363.4 | 276 | 409.52 | 407.7 | 298.23 | 265.6 | 161.6449 |  |  |  | 39
| 1 | 3 | 403.2 | 323 | 367.635 | 310.89 | 197.99 | 320 | 339.21 | 357 | 385 | 380 | 368 | 416.75 | 387.4824 | 280 | 375.2 | 286 | 430.04 | 412.9 | 284.77 | 262.6 | 162.2782 |  |  |  | 16
| 1 | 4 | 392.5 | 286 | 368.43 | 304.35 | 225.06 | 315 | 358.71 | 371 | 391 | 380 | 366 | 411.83 | 394.6494 | 284 | 400.4 | 286 | 412.56 | 410.2 | 297.44 | 255 | 161.9594 |  |  |  | 342.6233
| 1 | 5 | 397.5 | 291 | 367.8567 | 295.14 | 222.27 | 315 | 347.77 | 363 | 392 | 373 | 357 | 443.56 | 395.055 | 293 | 397.4 | 287 | 415.51 | 409.8 | 302.79 | 253.6 | 161.517 |  |  |  | 105
| 1 | 6 | 402.89 | 286 | 371.01 | 304.51 | 193.01 | 307 | 339.97 | 370 | 384 | 355 | 366 | 409.15 | 391.59 | 284 | 364.7 | 298 | 394.03 | 402.8 | 318.25 | 255.1 | 162.0447
| Mean |  | 396.6583 | 287.5 | 369.1853 | 302.6117 | 208.6083 | 309.1667 | 347.7933 | 364.3333 | 390.1667 | 370.5 | 366.1667 | 417.8983 | 393.7796 | 284.1667 | 375.8 | 286.6667 | 413.945 | 408.3 | 301.87 | 257.35 | 162.5007
| Median |  | 397.38 | 286 | 368.825 | 304.43 | 207.675 | 311 | 343.87 | 366 | 388 | 371.5 | 366 | 414.29 | 394.8522 | 284 | 369.95 | 286.5 | 414.035 | 408.75 | 300.51 | 255.05 | 162.002
| Std Dev. |  | 6.3429 | 19.7459 | 1.498 | 6.0466 | 14.4909 | 9.1086 | 12.7108 | 6.6833 | 7.5741 | 9.5656 | 6.9976 | 17.7191 | 3.8141 | 4.9565 | 19.1701 | 6.9761 | 12.2037 | 3.4997 | 11.4575 | 5.4184 | 1.5239
| Rel.Std.Dev. |  | 1.599% | 6.868% | 0.4057% | 1.998% | 6.946% | 2.946% | 3.655% | 1.834% | 1.941% | 2.582% | 1.911% | 4.24% | 0.9686% | 1.744% | 5.101% | 2.434% | 2.948% | 0.8571% | 3.796% | 2.105% | 0.9378%
| PDM3 |  | 15.77% | -16.09% | 7.753% | -11.68% | -39.11% | -9.765% | 1.509% | 6.336% | 13.88% | 8.136% | 6.871% | 21.97% | 14.93% | -17.06% | 9.683% | -16.33% | 20.82% | 19.17% | -11.89% | -24.89% | -52.57%
| Z-Score (Absolute) |  | 0.44 | 1.04 | 0.07 | 0.84 | 2.12 | 0.75 | 0.22 | 0 | 0.35 | 0.08 | 0.02 | 0.73 | 0.4 | 1.09 | 0.16 | 1.06 | 0.67 | 0.6 | 0.85 | 1.45 | 2.74
| Analytical results for Sc in OREAS 237b (Certified Value 7.29 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 7.21 | 6.4354 | 6.6 | 6.8533 | 7 | 8.5 | 8 | 7 | 8.2 | 8.3 | 7.7 | 7.8 | 7 | 7.8784 | 6 | 6.341 | 7.5 | 7.2 | 7.2 | 6.79 | 10.7029 | 7.2352 |  |  | 1
| 1 | 2 | 7.29 | 6.4415 | 6.7 | 6.88 | 7 | 7.9 | 8 | 7 | 8.3 | 8.3 | 7.4 | 7.8 | 6.6 | 7.7688 | 6 | 6.284 | 7.6 | 7.3 | 6.9 | 6.63 | 10.6629 | 7.1135 |  |  | 40
| 1 | 3 | 7.39 | 6.6053 | 7 | 6.7833 | 7 | 8.1 | 8 | 7.1 | 8.1 | 8.1 | 7.5 | 7.8 | 6.8 | 7.7549 | 6 | 6.303 | 7.8 | 7.2 | 7 | 6.88 | 10.6639 | 7.4441 |  |  | 16
| 1 | 4 | 7.28 | 6.3797 | 7.3 | 6.8567 | 7 | 8.3 | 8 | 7 | 8.3 | 8.3 | 7.5 | 7.8 | 6.6 | 7.5157 | 6 | 6.636 | 7.5 | 7.3 | 6.9 | 6.95 | 10.5417 | 7.2545 |  |  | 7.2922
| 1 | 5 | 7.36 | 6.5897 | 6.8 | 6.86 | 7 | 8.3 | 8 | 7 | 8 | 8 | 7.3 | 8 | 7.3 | 7.5335 | 6 | 6.439 | 7.5 | 7.1 | 8.6 | 6.69 | 10.4194 | 7.0649 |  |  | 106
| 1 | 6 | 7.33 | 6.6508 | 6.7 | 6.8967 | 7 | 8.4 | 8 | 7.1 | 8.1 | 8 | 7.7 | 7.9 | 6.8 | 7.9475 | 6 | 6.351 | 7.6 | 7.1 | 6.9 | 6.92 | 10.7493 | 7.3554
| Mean |  | 7.31 | 6.5171 | 6.85 | 6.855 | 7 | 8.25 | 8 | 7.0333 | 8.1667 | 8.1667 | 7.5167 | 7.85 | 6.85 | 7.7331 | 6 | 6.3923 | 7.5833 | 7.2 | 7.25 | 6.81 | 10.6233 | 7.2446
| Median |  | 7.31 | 6.5156 | 6.75 | 6.8583 | 7 | 8.3 | 8 | 7 | 8.15 | 8.2 | 7.5 | 7.8 | 6.8 | 7.7619 | 6 | 6.346 | 7.55 | 7.2 | 6.95 | 6.835 | 10.6634 | 7.2448
| Std Dev. |  | 0.064187 | 0.11153 | 0.25884 | 0.038801 | 0 | 0.21679 | 0 | 0.05164 | 0.12111 | 0.15055 | 0.16021 | 0.083666 | 0.26646 | 0.17659 | 0 | 0.13085 | 0.1169 | 0.089443 | 0.67157 | 0.12946 | 0.12138 | 0.14265
| Rel.Std.Dev. |  | 0.8781% | 1.711% | 3.779% | 0.566% | 0% | 2.628% | 0% | 0.7342% | 1.483% | 1.844% | 2.131% | 1.066% | 3.89% | 2.284% | 0% | 2.047% | 1.542% | 1.242% | 9.263% | 1.901% | 1.143% | 1.969%
| PDM3 |  | 0.2437% | -10.63% | -6.064% | -5.996% | -4.007% | 13.13% | 9.706% | -3.55% | 11.99% | 11.99% | 3.078% | 7.649% | -6.064% | 6.046% | -17.72% | -12.34% | 3.992% | -1.265% | -0.5791% | -6.613% | 45.68% | -0.653%
| Z-Score (Absolute) |  | 0.1 | 1.24 | 0.67 | 0.67 | NA | 1.69 | NA | 0.37 | 1.55 | 1.55 | 0.45 | 1.01 | 0.67 | 0.81 | NA | 1.45 | 0.56 | 0.08 | 0 | 0.74 | 5.69 | 0.01
Results from laboratories 1.08, 1.10 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 237b (Indicative Value 0.24 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.17 | 0.4 | < 5 | < 1 | 0.8 | < 1 | 2.5 | 0.4 | 0.2 | 0.3 | 0.3 | 0.2 | < 1 | 0.13 | < 1 | 0.2 | 0.229 | 2.8955 | 0.31318 | [5 empty] | 1
| 1 | 2 | 0.16 | 0.4 | < 5 | < 1 | 0.7 | < 1 | 2.4 | 0.3 | 0.2 | 0.2 | 0.3 | 0.1 | < 1 | 0.14 | < 1 | 0.2 | 0.2069 | 3.1372 | 0.32377 | [5 empty] | 41
| 1 | 3 | 0.15 | 0.3 | < 5 | < 1 | 0.7 | < 1 | 2 | 0.2 | 0.2 | 0.2 | 0.4 | 0.2 | < 1 | 0.15 | < 1 | 0.2 | 0.1959 | 3.2033 | 0.31222 | [5 empty] | 16
| 1 | 4 | 0.19 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 2.2 | 0.3 | 0.3 | 0.2 | 0.3 | 0.3 | < 1 | 0.16 | < 1 | 0.2 | 0.1821 | 2.8086 | 0.32592 | [5 empty] | 0.2439
| 1 | 5 | 0.19 | 0.4 | < 5 | < 1 | 0.9 | < 1 | 2.2 | 0.2 | 0.3 | 0.4 | 0.4 | 0.1 | < 1 | 0.17 | < 1 | 0.2 | 0.1904 | 2.4907 | 0.31179 | [5 empty] | 14
| 1 | 6 | 0.17 | 0.3 | < 5 | < 1 | 0.9 | < 1 | 2 | 0.4 | 0.3 | 0.2 | 0.3 | < 0.1 | < 1 | 0.18 | < 1 | 0.2 | 0.1393 | 3.0506 | 0.32686
| Mean |  | 0.17167 | 0.33333 |  |  | 0.78333 |  | 2.2167 | 0.3 | 0.25 | 0.25 | 0.33333 | 0.18 |  | 0.155 |  | 0.2 | 0.1906 | 2.931 | 0.31896
| Median |  | 0.17 | 0.35 |  |  | 0.75 |  | 2.2 | 0.3 | 0.25 | 0.2 | 0.3 | 0.2 |  | 0.155 |  | 0.2 | 0.19315 | 2.9731 | 0.31848
| Std Dev. |  | 0.016021 | 0.08165 |  |  | 0.098319 |  | 0.20412 | 0.089443 | 0.054772 | 0.083666 | 0.05164 | 0.083666 |  | 0.018708 |  | 3.0405e-17 | 0.029908 | 0.26118 | 0.0072696
| Rel.Std.Dev. |  | 9.333% | 24.49% |  |  | 12.55% |  | 9.209% | 29.81% | 21.91% | 33.47% | 15.49% | 46.48% |  | 12.07% |  | 1.52e-14% | 15.69% | 8.911% | 2.279%
| PDM3 |  | -29.62% | 36.67% |  |  | 221.2% |  | 808.8% | 23% | 2.501% | 2.501% | 36.67% | -26.2% |  | -36.45% |  | -18% | -21.85% | 1102% | 30.77%
| Z-Score (Absolute) |  | 0.65 | 0 | 8.77 | 0.67 | 1.82 | 0.67 | 7.62 | 0.13 | 0.34 | 0.34 | 0 | 0.71 | 0.67 | 0.72 | 0.67 | 0.54 | 0.58 | 10.51 | 0.06
| Analytical results for Si in OREAS 237b (Indicative Value 0.439 wt.%) | [25 empty] | Indicative
Si | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | wt.%
|  |  | AR*OES | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.4312 | [23 empty] | 1
| 1 | 2 | 0.4413 | [23 empty] | 9
| 1 | 3 | 0.4425 | [23 empty] | 16
| 1 | 4 | 0.4503 | [23 empty] | 0.43877
| 1 | 5 | 0.4458 | [23 empty] | 15
| 1 | 6 | 0.4215
| Mean |  | 0.43877
| Median |  | 0.4419
| Std Dev. |  | 0.010566
| Rel.Std.Dev. |  | 2.408%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 237b (Indicative Value 4.42 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 5.1695 | 4 | 4.392 | 4.1208 | [20 empty] | 1
| 1 | 2 | 5.062 | 4.2 | 4.29 | 3.5938 | [20 empty] | 10
| 1 | 3 | 5.3154 | 4.6 | 4.499 | 4.2224 | [20 empty] | 16
| 1 | 4 | 5.237 | 3.3 | 4.67 | 4.3551 | [20 empty] | 4.4191
| 1 | 5 | 5.355 | 4.7 | 4.366 | 3.558 | [20 empty] | 16
| 1 | 6 | 5.4 | 3.6 | 4.743 | 3.3095
| Mean |  | 5.2565 | 4.0667 | 4.4933 | 3.8599
| Median |  | 5.2762 | 4.1 | 4.4455 | 3.8573
| Std Dev. |  | 0.12614 | 0.55015 | 0.17967 | 0.42652
| Rel.Std.Dev. |  | 2.4% | 13.53% | 3.999% | 11.05%
| PDM3 |  | 18.95% | -7.975% | 1.68% | -12.65%
| Z-Score (Absolute) |  | 2.08 | 0.45 | 0.45 | 0.89
| Analytical results for Sn in OREAS 237b (Certified Value 1.67 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.72 | 2.7223 | 1.46 | < 10 | 1.6 | 1.7 | 1.8 | 1.6 | 1.6 | 1.7 | 1.7 | 1.7 | 1.7912 | < 1 | 1.7 | < 5 | 1.6 | 1.5197 | 2.3957 | [5 empty] | 1
| 1 | 2 | 1.72 | 2.595 | 1.58 | < 10 | 1.6 | 1.7 | 1.9 | 1.6 | 1.7 | 1.6 | 1.7 | 1.7 | 1.8479 | < 1 | 1.7 | < 5 | 1.7 | 1.2834 | 2.5344 | [5 empty] | 26
| 1 | 3 | 1.71 | 2.6314 | 1.67 | < 10 | 1.6 | 1.8 | 2.1 | 1.6 | 1.6 | 1.5 | 1.6 | 1.7 | 1.7788 | < 1 | 1.7 | < 5 | 1.6 | 1.5462 | 2.4043 | [5 empty] | 16
| 1 | 4 | 1.71 | 2.5002 | 1.64 | < 10 | 1.6 | 1.8 | 1.9 | 1.7 | 1.7 | 1.6 | 1.6 | 1.7 | 1.7396 | < 1 | 1.7 | < 5 | 1.6 | 1.4937 | 2.3794 | [5 empty] | 1.6654
| 1 | 5 | 1.75 | 2.69 | 1.63 | < 10 | 1.6 | 1.8 | 1.8 | 1.7 | 1.6 | 1.6 | 1.6 | 1.7 | 1.7897 | < 1 | 1.6 | < 5 | 1.7 | 1.2737 | 2.4485 | [5 empty] | 107
| 1 | 6 | 1.74 | 2.6057 | 1.58 | < 10 | 1.6 | 1.7 | 2.4 | 1.7 | 1.6 | 1.6 | 1.6 | 1.7 | 1.7513 | < 1 | 1.6 | < 5 | 1.7 | 1.1934 | 2.4722
| Mean |  | 1.725 | 2.6241 | 1.5933 |  | 1.6 | 1.75 | 1.9833 | 1.65 | 1.6333 | 1.6 | 1.6333 | 1.7 | 1.7831 |  | 1.6667 |  | 1.65 | 1.385 | 2.4391
| Median |  | 1.72 | 2.6186 | 1.605 |  | 1.6 | 1.75 | 1.9 | 1.65 | 1.6 | 1.6 | 1.6 | 1.7 | 1.7842 |  | 1.7 |  | 1.65 | 1.3885 | 2.4264
| Std Dev. |  | 0.016432 | 0.078205 | 0.074207 |  | 2.4324e-16 | 0.054772 | 0.23166 | 0.054772 | 0.05164 | 0.063246 | 0.05164 | 0 | 0.03801 |  | 0.05164 |  | 0.054772 | 0.1519 | 0.05818
| Rel.Std.Dev. |  | 0.9526% | 2.98% | 4.657% |  | 1.52e-14% | 3.13% | 11.68% | 3.32% | 3.162% | 3.953% | 3.162% | 0% | 2.132% |  | 3.098% |  | 3.32% | 10.97% | 2.385%
| PDM3 |  | 3.579% | 57.57% | -4.327% |  | -3.927% | 5.08% | 19.09% | -0.9244% | -1.925% | -3.927% | -1.925% | 2.078% | 7.066% |  | 0.07635% |  | -0.9244% | -16.84% | 46.46%
| Z-Score (Absolute) |  | 0.54 | 8.8 | 0.67 | 30.65 | 0.61 | 0.77 | 2.91 | 0.15 | 0.31 | 0.61 | 0.31 | 0.31 | 1.07 | 10.73 | 0 | 7.66 | 0.15 | 2.59 | 7.1
| Analytical results for Sr in OREAS 237b (Certified Value 26.2 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 25.81 | 27.0864 | 32.2 | 36.6467 | 26.3 | 28 | 33.2 | 25 | 26.5 | 27.2 | 25.8 | 26.7 | 28.1 | 27.1811 | 26 | 22.59 | 23.6 | 25.3 | 27.4 | 20.6 | 124.5915 | 25.4821 |  |  | 1
| 1 | 2 | 26.18 | 27.0536 | 29.8 | 36.8933 | 26.6 | 26 | 32.2 | 26 | 27 | 26.5 | 26 | 27.5 | 25.7 | 27.9218 | 25 | 22.22 | 23.9 | 24.7 | 27.6 | 20.3 | 122.6351 | 25.3134 |  |  | 27
| 1 | 3 | 26.51 | 26.5489 | 31.7 | 35.7967 | 26 | 26 | 34 | 26 | 26.5 | 25.7 | 25.4 | 27.6 | 26.6 | 27.4173 | 26 | 22.78 | 24.1 | 24.5 | 27.1 | 20.9 | 122.6521 | 25.9954 |  |  | 16
| 1 | 4 | 26.42 | 26.5566 | 33.1 | 36.2533 | 25.8 | 27 | 33 | 26 | 27 | 26.2 | 25.6 | 26.7 | 25.6 | 27.1349 | 26 | 23.16 | 24 | 24.6 | 26.2 | 21 | 121.8523 | 25.426 |  |  | 26.1532
| 1 | 5 | 26.09 | 26.5687 | 31.9 | 36.1667 | 26.1 | 27 | 32.7 | 26 | 26 | 25.1 | 26.3 | 27.4 | 28.5 | 26.9904 | 26 | 23.39 | 24.3 | 23.8 | 27.8 | 20.3 | 122.1551 | 24.6712 |  |  | 108
| 1 | 6 | 26.54 | 26.7322 | 33.6 | 37.06 | 26.6 | 27 | 32.3 | 26 | 26.2 | 25.3 | 25.2 | 27.6 | 26.3 | 27.759 | 24 | 22.4 | 24.1 | 23.4 | 28.1 | 21 | 122.7848 | 26.6181
| Mean |  | 26.2583 | 26.7577 | 32.05 | 36.4694 | 26.2333 | 26.8333 | 32.9 | 25.8333 | 26.5333 | 26 | 25.7167 | 27.25 | 26.8 | 27.4007 | 25.5 | 22.7567 | 24 | 24.3833 | 27.3667 | 20.6833 | 122.7785 | 25.5844
| Median |  | 26.3 | 26.6504 | 32.05 | 36.45 | 26.2 | 27 | 32.85 | 26 | 26.5 | 25.95 | 25.7 | 27.45 | 26.45 | 27.2992 | 26 | 22.685 | 24.05 | 24.55 | 27.5 | 20.75 | 122.6436 | 25.4541
| Std Dev. |  | 0.28435 | 0.25137 | 1.3217 | 0.47976 | 0.3266 | 0.75277 | 0.66332 | 0.40825 | 0.40825 | 0.78994 | 0.40208 | 0.43243 | 1.2264 | 0.37079 | 0.83666 | 0.44885 | 0.23664 | 0.67946 | 0.66533 | 0.33116 | 0.95604 | 0.66037
| Rel.Std.Dev. |  | 1.083% | 0.9394% | 4.124% | 1.316% | 1.245% | 2.805% | 2.016% | 1.58% | 1.539% | 3.038% | 1.563% | 1.587% | 4.576% | 1.353% | 3.281% | 1.972% | 0.986% | 2.787% | 2.431% | 1.601% | 0.7787% | 2.581%
| PDM3 |  | 0.402% | 2.312% | 22.55% | 39.45% | 0.3064% | 2.601% | 25.8% | -1.223% | 1.453% | -0.5858% | -1.669% | 4.194% | 2.473% | 4.77% | -2.498% | -12.99% | -8.233% | -6.767% | 4.64% | -20.91% | 369.5% | -2.175%
| Z-Score (Absolute) |  | 0.11 | 0.28 | 4.36 | 7.76 | 0.13 | 0.34 | 5.01 | 0.43 | 0.11 | 0.31 | 0.52 | 0.66 | 0.31 | 0.77 | 0.69 | 2.8 | 1.85 | 1.55 | 0.75 | 4.4 | 74.28 | 0.63
| Analytical results for Ta in OREAS 237b (Certified Value < 0.05 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0411 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.5677 | [10 empty] | 28
| 1 | 3 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0588 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.03 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.05 | 0.0186 | [10 empty] | < 0.05
| 1 | 5 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | 0.0781 | [10 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.03 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0716
| Mean | [5 empty] | 0.033333 |  |  | 0.01 | 0.01 |  |  |  | 0.035 | 0.13932
| Median | [5 empty] | 0.035 |  |  | 0.01 | 0.01 |  |  |  | 0.035 | 0.0652
| Std Dev. | [5 empty] | 0.008165 |  |  | 0 | 0 |  |  |  | 0.021213 | 0.21097
| Rel.Std.Dev. | [5 empty] | 24.49% |  |  | 0% | 0% |  |  |  | 60.61% | 151.4%
PDM3
| Z-Score (Absolute) |  | 0.87 | 0 | 108.25 | 0 | 0.36 | 0 | 1.09 | 0.69 | 0.66 | 0.87 | 0.87 | 0 | 0.44 | 5
| Analytical results for Tb in OREAS 237b (Indicative Value 0.47 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.56957 | 0.4 | 0.45 | 0.45 | 1.46 | [19 empty] | 1
| 1 | 2 | 0.56349 | 0.4 | 0.442 | 0.44 | 1.41 | [19 empty] | 11
| 1 | 3 | 0.56182 | 0.4 | 0.461 | 0.44 | 1.45 | [19 empty] | 16
| 1 | 4 | 0.54984 | 0.4 | 0.473 | 0.44 | 1.48 | [19 empty] | 0.46848
| 1 | 5 | 0.59039 | 0.4 | 0.451 | 0.45 | 1.42 | [19 empty] | 17
| 1 | 6 | 0.58947 | 0.4 | 0.482 | 0.44 | 1.44
| Mean |  | 0.57076 | 0.4 | 0.45983 | 0.44333 | 1.4433
| Median |  | 0.56653 | 0.4 | 0.456 | 0.44 | 1.445
| Std Dev. |  | 0.01617 | 6.0809e-17 | 0.015224 | 0.005164 | 0.02582
| Rel.Std.Dev. |  | 2.833% | 1.52e-14% | 3.311% | 1.165% | 1.789%
| PDM3 |  | 21.83% | -14.62% | -1.846% | -5.368% | 208.1%
| Z-Score (Absolute) |  | 1.25 | 0.67 | 0 | 0.19 | 11.08
| Analytical results for Te in OREAS 237b (Indicative Value 0.029 ppm) | [25 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.02 | 0.06 | 5.5533 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.06 | 0.0359 | [9 empty] | 1
| 1 | 2 | < 0.02 | 0.06 | 5.3963 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.01 | 0.02 | 0.02 | 0.01 | < 0.02 | < 0.05 | 0.04 | [10 empty] | 6
| 1 | 3 | < 0.02 | 0.05 | < 5 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.01 | 0.02 | < 0.05 | 0.05 | [10 empty] | 16
| 1 | 4 | < 0.02 | 0.04 | 6.2795 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03 | < 0.05 | 0.05 | [10 empty] | 0.029248
| 1 | 5 | < 0.02 | 0.05 | < 5 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.03 | 0.0239 | [9 empty] | 18
| 1 | 6 | < 0.02 | 0.03 | 5.4747 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.05 | 0.0239
| Mean |  |  | 0.048333 | 5.676 |  | 0.046667 |  |  | 0.018333 | 0.02 | 0.021667 | 0.011667 | 0.022 |  | 0.046667 | 0.0279
| Median |  |  | 0.05 | 5.514 |  | 0.045 |  |  | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 |  | 0.05 | 0.0239
| Std Dev. |  |  | 0.01169 | 0.40743 |  | 0.008165 |  |  | 0.0040825 | 0 | 0.0040825 | 0.0040825 | 0.0044721 |  | 0.010328 | 0.0069282
| Rel.Std.Dev. |  |  | 24.19% | 7.178% |  | 17.5% |  |  | 22.27% | 0% | 18.84% | 34.99% | 20.33% |  | 22.13% | 24.83%
| PDM3 |  |  | 65.25% | 19306% |  | 59.55% |  |  | -37.32% | -31.62% | -25.92% | -60.11% | -24.78% |  | 59.55% | -4.609%
| Z-Score (Absolute) |  | 1.52 | 2.36 | 464.49 | 0 | 2.19 | 0 | 2.53 | 0.67 | 0.51 | 0.34 | 1.35 | 0.51 | 0 | 2.19 | 0.29
| Analytical results for Th in OREAS 237b (Certified Value 12.3 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.676 | 12.2471 | 12.2 | 12.66 | 12.1 | 14.8 | 13 | 12.3 | 12 | 12.1 | 12.1 | 11.7 | 12.9804 | 11.68 | 12.2 | 13.09 | 15.3 | 12.7 | 9.9297 | 11.3129 |  |  |  |  | 1
| 1 | 2 | 11.645 | 11.9602 | 12.8 | 13.02 | 11.8 | 14.4 | 12.9 | 12.6 | 11.8 | 12.1 | 12.4 | 10.9 | 14.0096 | 11.52 | 11.8 | 14.87 | 15.3 | 12.8 | 7.8333 | 11.0941 |  |  |  |  | 31
| 1 | 3 | 11.649 | 12.4299 | 13.4 | 12.74 | 10.9 | 14.9 | 12.9 | 12.2 | 12.1 | 11.7 | 12.4 | 11.7 | 12.7399 | 11.93 | 12.3 | 15.47 | 15.6 | 12.5 | 9.9013 | 11.0072 |  |  |  |  | 16
| 1 | 4 | 12.105 | 12.3231 | 12.7 | 12.84 | 12.3 | 15 | 13 | 12.8 | 12 | 11.7 | 12.1 | 11.1 | 12.6695 | 12.08 | 11.7 | 13.63 | 16.2 | 12.3 | 10.3282 | 10.9643 |  |  |  |  | 12.2657
| 1 | 5 | 11.918 | 12.1597 | 12.6 | 12.79 | 12.3 | 15.3 | 13.2 | 12 | 12 | 12.3 | 12.3 | 12 | 13.3185 | 11.58 | 12.3 | 13.53 | 16.5 | 12.6 | 8.1563 | 10.9947 |  |  |  |  | 110
| 1 | 6 | 11.854 | 12.3703 | 12.5 | 13.01 | 12.2 | 15.1 | 12.8 | 12.2 | 11.5 | 11.9 | 12.2 | 11.8 | 13.3664 | 12.2 | 12.3 | 11.13 | 15 | 13.3 | 7.4168 | 11.0692
| Mean |  | 11.8078 | 12.2484 | 12.7 | 12.8433 | 11.9333 | 14.9167 | 12.9667 | 12.35 | 11.9 | 11.9667 | 12.25 | 11.5333 | 13.1807 | 11.8317 | 12.1 | 13.62 | 15.65 | 12.7 | 8.9276 | 11.0737
| Median |  | 11.765 | 12.2851 | 12.65 | 12.815 | 12.15 | 14.95 | 12.95 | 12.25 | 12 | 12 | 12.25 | 11.7 | 13.1495 | 11.805 | 12.25 | 13.58 | 15.45 | 12.65 | 9.0288 | 11.0382
| Std Dev. |  | 0.18531 | 0.16989 | 0.4 | 0.14569 | 0.53914 | 0.30605 | 0.13663 | 0.29496 | 0.21909 | 0.24221 | 0.13784 | 0.43205 | 0.49724 | 0.27946 | 0.27568 | 1.5149 | 0.58224 | 0.34059 | 1.264 | 0.12672
| Rel.Std.Dev. |  | 1.569% | 1.387% | 3.15% | 1.134% | 4.518% | 2.052% | 1.054% | 2.388% | 1.841% | 2.024% | 1.125% | 3.746% | 3.773% | 2.362% | 2.278% | 11.12% | 3.72% | 2.682% | 14.16% | 1.144%
| PDM3 |  | -3.733% | -0.1413% | 3.541% | 4.709% | -2.71% | 21.61% | 5.715% | 0.6871% | -2.982% | -2.438% | -0.1282% | -5.971% | 7.46% | -3.539% | -1.351% | 11.04% | 27.59% | 3.541% | -27.22% | -9.718%
| Z-Score (Absolute) |  | 0.66 | 0 | 0.67 | 0.89 | 0.47 | 3.99 | 1.07 | 0.15 | 0.52 | 0.42 | 0 | 1.07 | 1.39 | 0.62 | 0.22 | 2.05 | 5.09 | 0.67 | 4.97 | 1.76
| Analytical results for Ti in OREAS 237b (Certified Value 0.15 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1357 | 0.12 | 0.12867 | 0.1714 | 0.156 | 0.189 | 0.14 | 0.154 | 0.163 | 0.153 | 0.152 | 0.151 | 0.16525 | 0.144 | 0.13 | 0.14754 | 0.27 | 0.162 | 0.14361 | 0.2959 | 0.14863 |  |  |  | 1
| 1 | 2 | 0.1354 | 0.12 | 0.132 | 0.1665 | 0.152 | 0.185 | 0.139 | 0.152 | 0.167 | 0.153 | 0.157 | 0.139 | 0.16381 | 0.144 | 0.14 | 0.14365 | 0.26 | 0.16 | 0.13826 | 0.3034 | 0.14645 |  |  |  | 32
| 1 | 3 | 0.1365 | 0.12 | 0.12867 | 0.1651 | 0.145 | 0.19 | 0.14 | 0.151 | 0.166 | 0.154 | 0.155 | 0.146 | 0.1623 | 0.144 | 0.14 | 0.14819 | 0.27 | 0.159 | 0.14159 | 0.2986 | 0.15343 |  |  |  | 16
| 1 | 4 | 0.1384 | 0.13 | 0.12867 | 0.1655 | 0.153 | 0.187 | 0.139 | 0.155 | 0.164 | 0.153 | 0.147 | 0.141 | 0.16594 | 0.15 | 0.13 | 0.14315 | 0.27 | 0.153 | 0.14535 | 0.2997 | 0.14792 |  |  |  | 0.14975
| 1 | 5 | 0.1369 | 0.13 | 0.132 | 0.1666 | 0.151 | 0.191 | 0.142 | 0.151 | 0.166 | 0.155 | 0.152 | 0.155 | 0.16775 | 0.144 | 0.14 | 0.14443 | 0.28 | 0.157 | 0.14095 | 0.2936 | 0.14727 |  |  |  | 111
| 1 | 6 | 0.1363 | 0.12 | 0.132 | 0.1704 | 0.149 | 0.186 | 0.142 | 0.152 | 0.16 | 0.152 | 0.154 | 0.143 | 0.16715 | 0.138 | 0.13 | 0.13936 | 0.27 | 0.16 | 0.14527 | 0.3018 | 0.15437
| Mean |  | 0.13653 | 0.12333 | 0.13033 | 0.16758 | 0.151 | 0.188 | 0.14033 | 0.1525 | 0.16433 | 0.15333 | 0.15283 | 0.14583 | 0.16537 | 0.144 | 0.135 | 0.14438 | 0.27 | 0.1585 | 0.1425 | 0.29883 | 0.14968
| Median |  | 0.1364 | 0.12 | 0.13033 | 0.16655 | 0.1515 | 0.188 | 0.14 | 0.152 | 0.165 | 0.153 | 0.153 | 0.1445 | 0.1656 | 0.144 | 0.135 | 0.14404 | 0.27 | 0.1595 | 0.1426 | 0.29915 | 0.14828
| Std Dev. |  | 0.0010633 | 0.005164 | 0.0018257 | 0.0026514 | 0.0037417 | 0.0023664 | 0.0013663 | 0.0016432 | 0.002582 | 0.0010328 | 0.0034303 | 0.0061455 | 0.0020504 | 0.0037947 | 0.0054772 | 0.0032162 | 0.0063246 | 0.0031464 | 0.0027653 | 0.0036456 | 0.0033599
| Rel.Std.Dev. |  | 0.7788% | 4.187% | 1.401% | 1.582% | 2.478% | 1.259% | 0.9736% | 1.077% | 1.571% | 0.6736% | 2.244% | 4.214% | 1.24% | 2.635% | 4.057% | 2.227% | 2.342% | 1.985% | 1.941% | 1.22% | 2.245%
| PDM3 |  | -8.826% | -17.64% | -12.97% | 11.91% | 0.8343% | 25.54% | -6.289% | 1.836% | 9.738% | 2.392% | 2.059% | -2.616% | 10.43% | -3.84% | -9.85% | -3.583% | 80.3% | 5.843% | -4.838% | 99.55% | -0.04833%
| Z-Score (Absolute) |  | 0.91 | 1.75 | 1.31 | 1.05 | 0 | 2.34 | 0.67 | 0.09 | 0.84 | 0.15 | 0.12 | 0.33 | 0.91 | 0.45 | 1.01 | 0.42 | 7.52 | 0.47 | 0.54 | 9.35 | 0.08
| Analytical results for Tl in OREAS 237b (Certified Value 0.5 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.5 | 0.51 | < 10 | 0.53 | 0.5 | 0.54 | 0.53 | 0.49 | 0.47 | 0.49 | 0.53 | 0.5 | 0.5679 | 0.737 | 0.49 | < 5 | 0.5 | 0.4677 | 0.616 | [5 empty] | 1
| 1 | 2 | 0.5 | 0.52 | < 10 | 0.55 | 0.4 | 0.55 | 0.53 | 0.49 | 0.48 | 0.47 | 0.54 | 0.49 | 0.57686 | 0.766 | 0.46 | < 5 | 0.49 | 0.3961 | 0.6459 | [5 empty] | 33
| 1 | 3 | 0.49 | 0.53 | < 10 | 0.54 | 0.4 | 0.54 | 0.53 | 0.5 | 0.48 | 0.47 | 0.53 | 0.5 | 0.57629 | 0.847 | 0.48 | < 5 | 0.49 | 0.4731 | 0.6226 | [5 empty] | 16
| 1 | 4 | 0.5 | 0.52 | < 10 | 0.54 | 0.48 | 0.53 | 0.53 | 0.51 | 0.5 | 0.48 | 0.54 | 0.5 | 0.56888 | 0.771 | 0.45 | < 5 | 0.49 | 0.4863 | 0.6291 | [5 empty] | 0.50358
| 1 | 5 | 0.51 | 0.52 | < 10 | 0.54 | 0.47 | 0.53 | 0.54 | 0.49 | 0.51 | 0.48 | 0.53 | 0.51 | 0.57034 | 0.769 | 0.47 | < 5 | 0.51 | 0.3848 | 0.6418 | [5 empty] | 112
| 1 | 6 | 0.51 | 0.51 | < 10 | 0.53 | 0.42 | 0.54 | 0.54 | 0.5 | 0.48 | 0.48 | 0.54 | 0.51 | 0.59404 | 0.765 | 0.48 | < 5 | 0.49 | 0.3552 | 0.5938
| Mean |  | 0.50167 | 0.51833 |  | 0.53833 | 0.445 | 0.53833 | 0.53333 | 0.49667 | 0.48667 | 0.47833 | 0.535 | 0.50167 | 0.57572 | 0.77583 | 0.47167 |  | 0.495 | 0.4272 | 0.62487
| Median |  | 0.5 | 0.52 |  | 0.54 | 0.445 | 0.54 | 0.53 | 0.495 | 0.48 | 0.48 | 0.535 | 0.5 | 0.57332 | 0.7675 | 0.475 |  | 0.49 | 0.4319 | 0.62585
| Std Dev. |  | 0.0075277 | 0.0075277 |  | 0.0075277 | 0.043704 | 0.0075277 | 0.005164 | 0.008165 | 0.015055 | 0.0075277 | 0.0054772 | 0.0075277 | 0.0097365 | 0.037032 | 0.01472 |  | 0.0083666 | 0.055116 | 0.018952
| Rel.Std.Dev. |  | 1.501% | 1.452% |  | 1.398% | 9.821% | 1.398% | 0.9682% | 1.644% | 3.094% | 1.574% | 1.024% | 1.501% | 1.691% | 4.773% | 3.121% |  | 1.69% | 12.9% | 3.033%
| PDM3 |  | -0.3794% | 2.93% |  | 6.902% | -11.63% | 6.902% | 5.909% | -1.372% | -3.358% | -5.013% | 6.24% | -0.3794% | 14.33% | 54.06% | -6.337% |  | -1.703% | -15.17% | 24.09%
| Z-Score (Absolute) |  | 0.35 | 0 | 95.43 | 0.43 | 1.56 | 0.43 | 0.32 | 0.46 | 0.67 | 0.85 | 0.35 | 0.35 | 1.22 | 5.48 | 0.99 | 42.2 | 0.5 | 1.94 | 2.27
| Analytical results for Tm in OREAS 237b (Indicative Value 0.12 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.13468 | 0.1 | 0.115 | 0.1259 | [20 empty] | 1
| 1 | 2 | 0.12786 | 0.1 | 0.113 | 0.1024 | [20 empty] | 13
| 1 | 3 | 0.14159 | 0.1 | 0.118 | 0.1224 | [20 empty] | 16
| 1 | 4 | 0.14282 | 0.1 | 0.12 | 0.1283 | [20 empty] | 0.11783
| 1 | 5 | 0.14147 | 0.1 | 0.117 | 0.1038 | [20 empty] | 19
| 1 | 6 | 0.14193 | 0.1 | 0.124 | 0.0952
| Mean |  | 0.13839 | 0.1 | 0.11783 | 0.113
| Median |  | 0.14153 | 0.1 | 0.1175 | 0.1131
| Std Dev. |  | 0.0059433 | 1.5202e-17 | 0.0038687 | 0.014161
| Rel.Std.Dev. |  | 4.295% | 1.52e-14% | 3.283% | 12.53%
| PDM3 |  | 17.45% | -15.13% | 0.0002842% | -4.102%
| Z-Score (Absolute) |  | 1.74 | 1.17 | 0.18 | 0.18
| Analytical results for U in OREAS 237b (Certified Value 1.46 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.388 | 1.6179 | 1.4 | 1.61 | 1.41 | 1.94 | 1.5 | 1.36 | 1.42 | 1.32 | 1.5 | 1.43 | 1.5041 | 1 | 1.538 | 1.45 | 1.62 | 1.3098 | 3.8282 | [5 empty] | 1
| 1 | 2 | 1.355 | 1.526 | 1.5 | 1.64 | 1.33 | 1.9 | 1.5 | 1.38 | 1.38 | 1.32 | 1.53 | 1.28 | 1.4573 | 1 | 1.517 | 1.4 | 1.56 | 1.1256 | 3.7053 | [5 empty] | 35
| 1 | 3 | 1.367 | 1.608 | 1.5 | 1.6 | 1.27 | 1.93 | 1.5 | 1.35 | 1.41 | 1.3 | 1.53 | 1.47 | 1.5812 | 1 | 1.516 | 1.45 | 1.61 | 1.3009 | 3.4943 | [5 empty] | 16
| 1 | 4 | 1.428 | 1.6102 | 1.5 | 1.61 | 1.38 | 2 | 1.5 | 1.39 | 1.38 | 1.31 | 1.51 | 1.33 | 1.5096 | 1 | 1.543 | 1.44 | 1.54 | 1.3596 | 3.4946 | [5 empty] | 1.4593
| 1 | 5 | 1.404 | 1.7169 | 1.5 | 1.61 | 1.4 | 1.99 | 1.5 | 1.32 | 1.43 | 1.36 | 1.56 | 1.49 | 1.5783 | 1 | 1.532 | 1.46 | 1.6 | 1.1081 | 3.5636 | [5 empty] | 113
| 1 | 6 | 1.393 | 1.6857 | 1.5 | 1.64 | 1.34 | 1.98 | 1.5 | 1.35 | 1.34 | 1.34 | 1.54 | 1.47 | 1.6383 | 1 | 1.568 | 1.45 | 1.6 | 1.0482 | 3.5317
| Mean |  | 1.3892 | 1.6274 | 1.4833 | 1.6183 | 1.355 | 1.9567 | 1.5 | 1.3583 | 1.3933 | 1.325 | 1.5283 | 1.4117 | 1.5448 | 1 | 1.5357 | 1.4417 | 1.5883 | 1.2087 | 3.6029
| Median |  | 1.3905 | 1.614 | 1.5 | 1.61 | 1.36 | 1.96 | 1.5 | 1.355 | 1.395 | 1.32 | 1.53 | 1.45 | 1.544 | 1 | 1.535 | 1.45 | 1.6 | 1.2132 | 3.5476
| Std Dev. |  | 0.026088 | 0.067002 | 0.040825 | 0.017224 | 0.05244 | 0.039328 | 0 | 0.024833 | 0.033267 | 0.021679 | 0.02137 | 0.086352 | 0.065934 | 0 | 0.019253 | 0.02137 | 0.031252 | 0.12983 | 0.13519
| Rel.Std.Dev. |  | 1.878% | 4.117% | 2.752% | 1.064% | 3.87% | 2.01% | 0% | 1.828% | 2.388% | 1.636% | 1.398% | 6.117% | 4.268% | 0% | 1.254% | 1.482% | 1.968% | 10.74% | 3.752%
| PDM3 |  | -4.809% | 11.52% | 1.644% | 10.89% | -7.15% | 34.08% | 2.786% | -6.922% | -4.524% | -9.206% | 4.727% | -3.267% | 5.856% | -31.48% | 5.23% | -1.212% | 8.839% | -17.18% | 146.9%
| Z-Score (Absolute) |  | 0.69 | 0.91 | 0.06 | 0.85 | 0.92 | 3.12 | 0.06 | 0.9 | 0.66 | 1.12 | 0.25 | 0.54 | 0.36 | NA | 0.3 | 0.34 | 0.65 | 1.9 | 14.18
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 237b (Certified Value 69 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 71 | 75.03 | 55 | 67.7367 | 67 | 72 | 71 | 73 | 69 | 69 | 69 | 66 | 73 | 71.801 | 71 | 67 | 71.62 | 69.8 | 65 | 63.8 | 94.6766 | 70.8965 |  |  | 1
| 1 | 2 | 71.3 | 74.69 | 57 | 67.5067 | 65 | 69 | 71 | 72 | 68 | 71 | 69 | 67 | 67 | 69.5746 | 71 | 67 | 68.1 | 70.2 | 65 | 62.4 | 93.9711 | 70.3125 |  |  | 36
| 1 | 3 | 71 | 75.73 | 59 | 66.82 | 66 | 70 | 72 | 73 | 67 | 71 | 69 | 67 | 70 | 71.2929 | 71 | 67 | 69.47 | 69.4 | 64 | 64.5 | 92.8363 | 71.4309 |  |  | 16
| 1 | 4 | 71.7 | 76 | 58 | 67.12 | 65 | 72 | 70 | 73 | 69 | 70 | 70 | 64 | 69 | 68.0861 | 73 | 66 | 68.43 | 70.1 | 62 | 65 | 96.0638 | 69.8868 |  |  | 69.1929
| 1 | 5 | 71.7 | 75.16 | 58 | 67.2667 | 64 | 71 | 72 | 73 | 68 | 70 | 70 | 66 | 76 | 70.2366 | 72 | 66 | 69.25 | 70.2 | 64 | 62.7 | 95.2152 | 69.6414 |  |  | 114
| 1 | 6 | 71.4 | 74.16 | 56 | 67.8667 | 67 | 70 | 70 | 73 | 68 | 68 | 69 | 66 | 69 | 69.5085 | 74 | 67 | 68.51 | 69.8 | 64 | 65 | 95.3678 | 70.8758
| Mean |  | 71.35 | 75.1283 | 57.1667 | 67.3861 | 65.6667 | 70.6667 | 71 | 72.8333 | 68.1667 | 69.8333 | 69.3333 | 66 | 70.6667 | 70.0833 | 72 | 66.6667 | 69.23 | 69.9167 | 64 | 63.9 | 94.6885 | 70.5073
| Median |  | 71.35 | 75.095 | 57.5 | 67.3867 | 65.5 | 70.5 | 71 | 73 | 68 | 70 | 69 | 66 | 69.5 | 69.9056 | 71.5 | 67 | 68.88 | 69.95 | 64 | 64.15 | 94.9459 | 70.5941
| Std Dev. |  | 0.31464 | 0.67265 | 1.472 | 0.39343 | 1.2111 | 1.2111 | 0.89443 | 0.40825 | 0.75277 | 1.169 | 0.5164 | 1.0954 | 3.266 | 1.3428 | 1.2649 | 0.5164 | 1.281 | 0.31252 | 1.0954 | 1.1384 | 1.1468 | 0.68017
| Rel.Std.Dev. |  | 0.441% | 0.8953% | 2.575% | 0.5838% | 1.844% | 1.714% | 1.26% | 0.5605% | 1.104% | 1.674% | 0.7448% | 1.66% | 4.622% | 1.916% | 1.757% | 0.7746% | 1.85% | 0.447% | 1.712% | 1.782% | 1.211% | 0.9647%
| PDM3 |  | 3.118% | 8.578% | -17.38% | -2.611% | -5.096% | 2.13% | 2.612% | 5.261% | -1.483% | 0.9256% | 0.203% | -4.614% | 2.13% | 1.287% | 4.057% | -3.651% | 0.05369% | 1.046% | -7.505% | -7.649% | 36.85% | 1.9%
| Z-Score (Absolute) |  | 0.52 | 1.85 | 4.47 | 0.88 | 1.48 | 0.28 | 0.4 | 1.04 | 0.6 | 0.01 | 0.19 | 1.36 | 0.28 | 0.07 | 0.75 | 1.13 | 0.23 | 0.01 | 2.07 | 2.1 | 8.73 | 0.22
| Analytical results for W in OREAS 237b (Certified Value 0.49 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.49 | 0.3 | 7.96 | 0.6 | 0.44 | 0.5 | 0.5 | 0.56 | 0.58 | 0.47 | 0.55 | 0.4 | < 1 | < 1 | 0.4 | < 5 | 0.44 | 0.4372 | [6 empty] | 1
| 1 | 2 | 0.48 | 0.4 | 8.06 | 0.6 | 0.42 | 0.5 | 0.5 | 0.57 | 0.57 | 0.5 | 0.53 | 0.39 | < 1 | < 1 | 0.4 | < 5 | 0.44 | 0.4193 | [6 empty] | 37
| 1 | 3 | 0.46 | 0.4 | 6.272 | 0.6 | 0.34 | 0.6 | 0.5 | 0.57 | 0.55 | 0.47 | 0.52 | 0.41 | < 1 | < 1 | 0.4 | < 5 | 0.41 | 0.4425 | [6 empty] | 16
| 1 | 4 | 0.6 | 0.5 | 7.608 | 0.6 | 0.43 | 0.5 | 0.5 | 0.56 | 0.55 | 0.48 | 0.53 | 0.4 | < 1 | < 1 | 0.4 | < 5 | 0.44 | 0.4603 | [6 empty] | 0.49421
| 1 | 5 | 0.56 | 0.5 | 8.264 | 0.6 | 0.38 | 0.5 | 0.5 | 0.52 | 0.54 | 0.48 | 0.55 | 0.43 | < 1 | < 1 | 0.4 | < 5 | 0.48 | 0.4829 | [6 empty] | 115
| 1 | 6 | 0.53 | 0.3 | 8.38 | 0.6 | 0.36 | 0.5 | 0.5 | 0.55 | 0.51 | 0.47 | 0.54 | 0.42 | < 1 | < 1 | 0.4 | < 5 | 0.45 | 0.3492
| Mean |  | 0.52 | 0.4 | 7.7573 | 0.6 | 0.395 | 0.51667 | 0.5 | 0.555 | 0.55 | 0.47833 | 0.53667 | 0.40833 |  |  | 0.4 |  | 0.44333 | 0.4319
| Median |  | 0.51 | 0.4 | 8.01 | 0.6 | 0.4 | 0.5 | 0.5 | 0.56 | 0.55 | 0.475 | 0.535 | 0.405 |  |  | 0.4 |  | 0.44 | 0.43985
| Std Dev. |  | 0.053292 | 0.089443 | 0.77532 | 0 | 0.040866 | 0.040825 | 0 | 0.018708 | 0.024495 | 0.01169 | 0.012111 | 0.01472 |  |  | 6.0809e-17 |  | 0.022509 | 0.045926
| Rel.Std.Dev. |  | 10.25% | 22.36% | 9.995% | 0% | 10.35% | 7.902% | 0% | 3.371% | 4.454% | 2.444% | 2.257% | 3.605% |  |  | 1.52e-14% |  | 5.077% | 10.63%
| PDM3 |  | 5.218% | -19.06% | 1470% | 21.41% | -20.07% | 4.544% | 1.171% | 12.3% | 11.29% | -3.213% | 8.591% | -17.38% |  |  | -19.06% |  | -10.29% | -12.61%
| Z-Score (Absolute) |  | 0.12 | NA | 87.53 | 1.09 | 1.39 | NA | NA | 0.54 | 0.48 | 0.38 | 0.32 | 1.23 | 0.12 | 0.12 | NA | 24.03 | 0.81 | 0.94
Results from laboratories 1.06, 1.10, 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 237b (Certified Value 9.23 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 9.23 | 8.5279 | 8.02 | 8.8933 | 10.07 | 8.91 | 9.78 | 9.2 | 9.52 | 10.15 | 9.16 | 9.47 | 10.0707 | 9 | 8.071 | 9.47 | 9.76 | 8.0845 | 10.1168 | 9.0212 |  |  |  |  | 1
| 1 | 2 | 9.19 | 8.7444 | 8.39 | 8.9 | 10.19 | 7.87 | 9.64 | 9.1 | 9.72 | 9.78 | 9.32 | 9.71 | 9.8502 | 9 | 8.202 | 9 | 9.23 | 7.904 | 10.6521 | 9.1063 |  |  |  |  | 38
| 1 | 3 | 9.54 | 8.6147 | 8.5 | 8.8367 | 10.05 | 8.46 | 9.77 | 9.2 | 9.51 | 9.45 | 9.33 | 9.39 | 9.8282 | 9 | 8.424 | 9.42 | 9.33 | 8.189 | 10.3682 | 9.2039 |  |  |  |  | 16
| 1 | 4 | 9.39 | 8.5763 | 8.63 | 8.86 | 10.05 | 8.83 | 9.92 | 9.2 | 9.7 | 9.65 | 9.3 | 9.46 | 9.9612 | 9 | 8.516 | 9.23 | 8.99 | 8.3125 | 10.3336 | 9.1041 |  |  |  |  | 9.2314
| 1 | 5 | 9.4 | 8.5049 | 8.39 | 8.8867 | 10.03 | 8.83 | 9.77 | 9.3 | 9.23 | 9.6 | 9.27 | 9.79 | 9.9042 | 9 | 8.256 | 8.89 | 11.34 | 8.056 | 10.5841 | 8.7477 |  |  |  |  | 116
| 1 | 6 | 9.32 | 8.5796 | 8.02 | 8.9867 | 10.21 | 8.64 | 9.66 | 9.3 | 9.4 | 9.28 | 9.17 | 9.53 | 9.7689 | 9 | 8.656 | 9.02 | 9.33 | 8.2745 | 10.7277 | 9.4187
| Mean |  | 9.345 | 8.5913 | 8.325 | 8.8939 | 10.1 | 8.59 | 9.7567 | 9.2167 | 9.5133 | 9.6517 | 9.2583 | 9.5583 | 9.8972 | 9 | 8.3542 | 9.1717 | 9.6633 | 8.1367 | 10.4638 | 9.1003
| Median |  | 9.355 | 8.578 | 8.39 | 8.89 | 10.06 | 8.735 | 9.77 | 9.2 | 9.515 | 9.625 | 9.285 | 9.5 | 9.8772 | 9 | 8.34 | 9.125 | 9.33 | 8.1367 | 10.4762 | 9.1052
| Std Dev. |  | 0.12724 | 0.084688 | 0.25225 | 0.051226 | 0.07874 | 0.38869 | 0.10053 | 0.075277 | 0.18479 | 0.29876 | 0.075211 | 0.15702 | 0.1074 | 0 | 0.21686 | 0.23912 | 0.85836 | 0.15233 | 0.23068 | 0.22034
| Rel.Std.Dev. |  | 1.362% | 0.9857% | 3.03% | 0.576% | 0.7796% | 4.525% | 1.03% | 0.8168% | 1.942% | 3.095% | 0.8124% | 1.643% | 1.085% | 0% | 2.596% | 2.607% | 8.883% | 1.872% | 2.205% | 2.421%
| PDM3 |  | 1.231% | -6.934% | -9.818% | -3.656% | 9.41% | -6.948% | 5.69% | -0.1593% | 3.054% | 4.553% | 0.2921% | 3.542% | 7.213% | -2.506% | -9.502% | -0.6468% | 4.679% | -11.86% | 13.35% | -1.419%
| Z-Score (Absolute) |  | 0.14 | 1.11 | 1.55 | 0.61 | 1.4 | 1.11 | 0.83 | 0.07 | 0.42 | 0.65 | 0 | 0.5 | 1.06 | NA | 1.51 | 0.14 | 0.67 | 1.87 | 2.01 | 0.26
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 237b (Certified Value 0.84 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | 1.32 | [18 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | [18 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [18 empty] | 2
| 1 | 1 | 0.85909 | 0.8 | 0.729 | 0.8 | 1.01 | 1.1399 | [18 empty] | 1
| 1 | 2 | 0.84454 | 0.8 | 0.716 | 0.8 | 0.98 | 1.199 | [18 empty] | 39
| 1 | 3 | 0.86685 | 0.8 | 0.757 | 0.8 | 1.02 | 1.1696 | [18 empty] | 16
| 1 | 4 | 0.85324 | 0.8 | 0.76 | 0.8 | 1.03 | 1.1698 | [18 empty] | 0.84496
| 1 | 5 | 0.90446 | 0.8 | 0.735 | 0.8 | 1 | 1.1699 | [18 empty] | 117
| 1 | 6 | 0.87365 | 0.8 | 0.78 | 0.8 | 1.03 | 1.1995
| Mean |  | 0.86697 | 0.8 | 0.74617 | 0.8 | 1.0117 | 1.1746
| Median |  | 0.86297 | 0.8 | 0.746 | 0.8 | 1.015 | 1.1698
| Std Dev. |  | 0.020995 | 1.2162e-16 | 0.023592 | 1.2162e-16 | 0.019408 | 0.022314
| Rel.Std.Dev. |  | 2.422% | 1.52e-14% | 3.162% | 1.52e-14% | 1.918% | 1.9%
| PDM3 |  | 2.605% | -5.321% | -11.69% | -5.321% | 19.73% | 39.01%
| Z-Score (Absolute) |  | 0.37 | 0.37 | 0.97 | 0.37 | 1.99 | 3.81
| Analytical results for Zn in OREAS 237b (Certified Value 85 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 1
| 1 | 1 | 78.8 | 85.954 | 84.1 | 77.7233 | 81 | 86 | 88 | 84 | 85 | 82 | 85 | 83 | 89.5 | 89.7383 | 85 | 81.31 | 77 | 88.47 | 86.2 | 80 | 97.9776 | 90 | 80.7051 | 85.6342 | 1
| 1 | 2 | 79 | 86.261 | 84.6 | 78.0333 | 84 | 82 | 88 | 83 | 85 | 85 | 85 | 85 | 84.1 | 86.4972 | 86 | 78.62 | 77 | 85.44 | 87.6 | 81 | 95.3904 | 91 | 80.0173 | 86.8118 | 40
| 1 | 3 | 81.2 | 86.405 | 86.3 | 78.005 | 82 | 84 | 90 | 84 | 83 | 83 | 86 | 84 | 86.5 | 88.5571 | 85 | 84.1 | 78 | 85.59 | 86.9 | 79 | 98.1344 | 92 | 79.1018 | 86.9908 | 16
| 1 | 4 | 79.9 | 86.963 | 84.5 | 77.03 | 82 | 85 | 89 | 86 | 85 | 84 | 86 | 81 | 85.4 | 85.3571 | 88 | 90.81 | 77 | 88.01 | 87.4 | 76 | 99.4672 | 90 | 80.0066 | 85.5361 | 84.9363
| 1 | 5 | 81.4 | 86.702 | 93.8 | 77.02 | 82 | 83 | 89 | 86 | 84 | 83 | 86 | 83 | 92 | 85.1956 | 87 | 83.36 | 78 | 86.94 | 87.5 | 78 | 96.7344 | 91 | 84.325 | 86.109 | 118
| 1 | 6 | 79.8 | 87.101 | 87.5 | 78.03 | 84 | 85 | 88 | 88 | 85 | 81 | 85 | 84 | 85.5 | 86.3601 | 83 | 82.91 | 76 | 85.67 | 88.4 | 80 | 99.064 | 91 | 84.1594 | 86.7573
| Mean |  | 80.0167 | 86.5643 | 86.8 | 77.6403 | 82.5 | 84.1667 | 88.6667 | 85.1667 | 84.5 | 83 | 85.5 | 83.3333 | 87.1667 | 86.9509 | 85.6667 | 83.5183 | 77.1667 | 86.6867 | 87.3333 | 79 | 97.7947 | 90.8333 | 81.3859 | 86.3065
| Median |  | 79.85 | 86.5535 | 85.45 | 77.8642 | 82 | 84.5 | 88.5 | 85 | 85 | 83 | 85.5 | 83.5 | 86 | 86.4287 | 85.5 | 83.135 | 77 | 86.305 | 87.45 | 79.5 | 98.056 | 91 | 80.3612 | 86.4331
| Std Dev. |  | 1.0852 | 0.43722 | 3.665 | 0.49061 | 1.2247 | 1.472 | 0.8165 | 1.8348 | 0.83666 | 1.4142 | 0.54772 | 1.3663 | 2.9851 | 1.818 | 1.7512 | 4.0687 | 0.75277 | 1.3256 | 0.73666 | 1.7889 | 1.5141 | 0.75277 | 2.271 | 0.63424
| Rel.Std.Dev. |  | 1.356% | 0.5051% | 4.222% | 0.6319% | 1.485% | 1.749% | 0.9209% | 2.154% | 0.9901% | 1.704% | 0.6406% | 1.64% | 3.425% | 2.091% | 2.044% | 4.872% | 0.9755% | 1.529% | 0.8435% | 2.264% | 1.548% | 0.8287% | 2.79% | 0.7349%
| PDM3 |  | -5.792% | 1.917% | 2.194% | -8.59% | -2.868% | -0.9061% | 4.392% | 0.2712% | -0.5137% | -2.28% | 0.6637% | -1.887% | 2.626% | 2.372% | 0.8599% | -1.669% | -9.148% | 2.061% | 2.822% | -6.989% | 15.14% | 6.943% | -4.18% | 1.613%
| Z-Score (Absolute) |  | 1.87 | 0.43 | 0.52 | 2.71 | 1 | 0.41 | 1.17 | 0.06 | 0.29 | 0.82 | 0.06 | 0.7 | 0.64 | 0.57 | 0.12 | 0.64 | 2.87 | 0.48 | 0.7 | 2.23 | 4.38 | 1.93 | 1.39 | 0.34
| Analytical results for Zr in OREAS 237b (Certified Value 19.1 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 18.89 | 31.46 | 12.9 | 21.992 | 22.5 | 12.4 | 27.1 | 18.3 | 18.9 | 20.8 | 19.6 | 19.8 | 21.9586 | 12 | 15.8 | 22 | 22.4 | 13 | 96.4435 | 17.3592 |  |  |  |  | 1
| 1 | 2 | 18.65 | 30.92 | 19.7 | 21.9799 | 21.5 | 10.4 | 23.4 | 17.9 | 19.1 | 19.8 | 20.3 | 20.6 | 21.898 | 13 | 16.2 | 21.8 | 22.7 | 13.27 | 96.6283 | 17.3327 |  |  |  |  | 41
| 1 | 3 | 18.87 | 32.42 | 18.1 | 21.7342 | 21.8 | 10.3 | 26.1 | 18.2 | 19 | 19.1 | 19.7 | 20.8 | 21.7579 | 12 | 16.5 | 20.9 | 21.9 | 13.64 | 96.4666 | 17.8503 |  |  |  |  | 16
| 1 | 4 | 19.19 | 31.36 | 13.8 | 21.7583 | 20.9 | 11.6 | 25.4 | 18.1 | 19.1 | 19.2 | 19.3 | 20 | 22.103 | 12 | 16.3 | 22.8 | 20.9 | 13.54 | 96.3964 | 17.873 |  |  |  |  | 19.1101
| 1 | 5 | 18.71 | 32.19 | 10.1 | 21.9824 | 20.7 | 12.8 | 26.8 | 18.3 | 18.2 | 18.8 | 20.1 | 21 | 22.0996 | 13 | 17 | 21.7 | 22.6 | 13 | 96.3262 | 16.5584 |  |  |  |  | 119
| 1 | 6 | 19.8 | 31.26 | 12.5 | 22.3446 | 22 | 11.8 | 26.7 | 18.6 | 18.7 | 18.8 | 19.3 | 20.2 | 22.4194 | 12 | 16.3 | 21.3 | 22.6 | 13.14 | 96.7989 | 19.0238
| Mean |  | 19.0183 | 31.6017 | 14.5167 | 21.9652 | 21.5667 | 11.55 | 25.9167 | 18.2333 | 18.8333 | 19.4167 | 19.7167 | 20.4 | 22.0394 | 12.3333 | 16.35 | 21.75 | 22.1833 | 13.265 | 96.51 | 17.6662
| Median |  | 18.88 | 31.41 | 13.35 | 21.9811 | 21.65 | 11.7 | 26.4 | 18.25 | 18.95 | 19.15 | 19.65 | 20.4 | 22.0291 | 12 | 16.3 | 21.75 | 22.5 | 13.205 | 96.455 | 17.6048
| Std Dev. |  | 0.42654 | 0.57891 | 3.6444 | 0.21973 | 0.6802 | 1.0232 | 1.3732 | 0.23381 | 0.34448 | 0.7705 | 0.41191 | 0.47329 | 0.22711 | 0.5164 | 0.3937 | 0.6473 | 0.69113 | 0.27289 | 0.17347 | 0.81882
| Rel.Std.Dev. |  | 2.243% | 1.832% | 25.1% | 1% | 3.154% | 8.859% | 5.299% | 1.282% | 1.829% | 3.968% | 2.089% | 2.32% | 1.03% | 4.187% | 2.408% | 2.976% | 3.116% | 2.057% | 0.1797% | 4.635%
| PDM3 |  | -0.4799% | 65.37% | -24.04% | 14.94% | 12.86% | -39.56% | 35.62% | -4.588% | -1.448% | 1.604% | 3.174% | 6.75% | 15.33% | -35.46% | -14.44% | 13.81% | 16.08% | -30.59% | 405% | -7.555%
| Z-Score (Absolute) |  | 0.21 | 3.56 | 1.56 | 0.67 | 0.55 | 2.45 | 1.86 | 0.44 | 0.26 | 0.09 | 0 | 0.2 | 0.7 | NA | 1.01 | 0.61 | 0.74 | 1.93 | 23.03 | 0.61
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C725: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C743: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C923: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C959: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C977: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C995: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1013: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1031: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1140: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1158: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1177: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1196: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1214: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1232: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 237b (Indicative Value 14.35 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.36 | 1
| 1 | 2 | 14.34 | 16
| Mean |  | 14.35 | 16
| Median |  | 14.35 | 14.35
| Std Dev. |  | 0.014142 | 22
| Rel.Std.Dev. |  | 0.09855%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 237b (Indicative Value 1.25 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.25 | 17
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 237b (Indicative Value 5.8 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.79 | 18
| Mean |  | 5.795 | 16
| Median |  | 5.795 | 5.795
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 237b (Indicative Value 3.27 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.26 | 1
| 1 | 2 | 3.27 | 19
| Mean |  | 3.265 | 16
| Median |  | 3.265 | 3.265
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.2166%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 237b (Indicative Value 2.85 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.85 | 1
| 1 | 2 | 2.85 | 16
| Mean |  | 2.85 | 16
| Median |  | 2.85 | 2.85
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 237b (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 17
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 237b (Indicative Value 1.03 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.02 | 18
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 237b (Indicative Value 0.151 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.151 | 1
| 1 | 2 | 0.15 | 19
| Mean |  | 0.1505 | 16
| Median |  | 0.1505 | 0.1505
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 0.4698%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 237b (Indicative Value 68.18 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 68.22 | 1
| 1 | 2 | 68.14 | 16
| Mean |  | 68.18 | 16
| Median |  | 68.18 | 68.18
| Std Dev. |  | 0.056569 | 22
| Rel.Std.Dev. |  | 0.08297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 237b (Indicative Value 0.482 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.481 | 1
| 1 | 2 | 0.483 | 17
| Mean |  | 0.482 | 16
| Median |  | 0.482 | 0.482
| Std Dev. |  | 0.0014142 | 23
| Rel.Std.Dev. |  | 0.2934%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 237b (Indicative Value 0.815 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.82 | 1
| 1 | 2 | 0.81 | 18
| Mean |  | 0.815 | 16
| Median |  | 0.815 | 0.815
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.8676%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 237b (Indicative Value 2.09 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.09 | 1
| 1 | 2 | 2.08 | 21
| Mean |  | 2.085 | 16
| Median |  | 2.085 | 2.085
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.3391%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 237b (Indicative Value 0.08 wt.%) |  |  |  | Indicative
C | Round | Replicate | BV Geo |  | 1
| No | No | Perth |  | wt.%
|  |  | IRC |  | 3
3
| 1 | 1 | 0.07 |  | 1
| 1 | 2 | 0.09 |  | 23
| Mean |  | 0.08 |  | 16
| Median |  | 0.08 |  | 0.08
| Std Dev. |  | 0.014142 |  | 29
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 237b (Indicative Value 0.185 wt.%) |  |  |  | Indicative
S | Round | Replicate | Inspectorate | BV Geo | 1
| No | No | Gandhidham | Perth | wt.%
|  |  | IRC | IRC | 3
3
| 1 | 1 | 0.19 | 0.18 | 1
| 1 | 2 | 0.19 | 0.18 | 23
| 1 | 3 | 0.19 |  | 16
| 1 | 4 | 0.19 |  | 0.185
| 1 | 5 | 0.19 |  | 29
| 1 | 6 | 0.19
| Mean |  | 0.19 | 0.18
| Median |  | 0.19 | 0.18
| Std Dev. |  | 3.0405e-17 | 0
| Rel.Std.Dev. |  | 1.6e-14% | 0%
| PDM3 |  | 2.703% | -2.703%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 237b (Indicative Value 0.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 25
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 237b (Indicative Value 598 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 598 | 1
| 1 | 2 | 598 | 26
| Mean |  | 598 | 16
| Median |  | 598 | 598
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 237b (Indicative Value 700 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 697 | 1
| 1 | 2 | 702 | 27
| Mean |  | 699.5 | 16
| Median |  | 699.5 | 699.5
| Std Dev. |  | 3.5355 | 33
| Rel.Std.Dev. |  | 0.5054%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 237b (Indicative Value 2.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3 | 28
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 237b (Indicative Value 0.39 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.4 | 29
| Mean |  | 0.39 | 16
| Median |  | 0.39 | 0.39
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 237b (Indicative Value 0.25 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.2 | 5
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0.070711 | 36
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 237b (Indicative Value 80 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 79.5 | 1
| 1 | 2 | 79.6 | 31
| Mean |  | 79.55 | 16
| Median |  | 79.55 | 79.55
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.08889%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 237b (Indicative Value 18.2 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.2 | 1
| 1 | 2 | 18.2 | 32
| Mean |  | 18.2 | 16
| Median |  | 18.2 | 18.2
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 237b (Indicative Value 151 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 154 | 1
| 1 | 2 | 148 | 33
| Mean |  | 151 | 16
| Median |  | 151 | 151
| Std Dev. |  | 4.2426 | 39
| Rel.Std.Dev. |  | 2.81%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 237b (Indicative Value 9.46 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.47 | 1
| 1 | 2 | 9.45 | 34
| Mean |  | 9.46 | 16
| Median |  | 9.46 | 9.46
| Std Dev. |  | 0.014142 | 40
| Rel.Std.Dev. |  | 0.1495%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 237b (Indicative Value 30 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 35
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 237b (Indicative Value 5.74 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.72 | 1
| 1 | 2 | 5.76 | 9
| Mean |  | 5.74 | 16
| Median |  | 5.74 | 5.74
| Std Dev. |  | 0.028284 | 42
| Rel.Std.Dev. |  | 0.4928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 237b (Indicative Value 3.45 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.55 | 1
| 1 | 2 | 3.35 | 10
| Mean |  | 3.45 | 16
| Median |  | 3.45 | 3.45
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 4.099%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 237b (Indicative Value 1.36 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.37 | 1
| 1 | 2 | 1.35 | 11
| Mean |  | 1.36 | 16
| Median |  | 1.36 | 1.36
| Std Dev. |  | 0.014142 | 44
| Rel.Std.Dev. |  | 1.04%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 237b (Indicative Value 18.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.1 | 1
| 1 | 2 | 18.5 | 39
| Mean |  | 18.8 | 16
| Median |  | 18.8 | 18.8
| Std Dev. |  | 0.42426 | 45
| Rel.Std.Dev. |  | 2.257%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 237b (Indicative Value 6.02 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.93 | 1
| 1 | 2 | 6.1 | 12
| Mean |  | 6.015 | 16
| Median |  | 6.015 | 6.015
| Std Dev. |  | 0.12021 | 46
| Rel.Std.Dev. |  | 1.998%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 237b (Indicative Value 1.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.55 | 1
| 1 | 2 | 1.55 | 6
| Mean |  | 1.55 | 16
| Median |  | 1.55 | 1.55
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 237b (Indicative Value 6.82 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.77 | 1
| 1 | 2 | 6.87 | 25
| Mean |  | 6.82 | 16
| Median |  | 6.82 | 6.82
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 1.037%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 237b (Indicative Value 1.16 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.16 | 1
| 1 | 2 | 1.15 | 14
| Mean |  | 1.155 | 16
| Median |  | 1.155 | 1.155
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.6122%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 237b (Indicative Value < 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | < 0.05 | 27
| Mean |  |  | 16
| Median |  |  | < 0.05
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 237b (Indicative Value 40.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.5 | 1
| 1 | 2 | 40.6 | 28
| Mean |  | 40.55 | 16
| Median |  | 40.55 | 40.55
| Std Dev. |  | 0.070711 | 34
| Rel.Std.Dev. |  | 0.1744%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 237b (Indicative Value 0.47 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.47 | 1
| 1 | 2 | 0.46 | 29
| Mean |  | 0.465 | 16
| Median |  | 0.465 | 0.465
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 237b (Indicative Value 0.044 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0437 | 1
| 1 | 2 | 0.0434 | 30
| Mean |  | 0.04355 | 16
| Median |  | 0.04355 | 0.04355
| Std Dev. |  | 0.00021213 | 36
| Rel.Std.Dev. |  | 0.4871%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 237b (Indicative Value 1.2 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 31
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 237b (Indicative Value 16.1 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.1 | 1
| 1 | 2 | 16 | 32
| Mean |  | 16.05 | 16
| Median |  | 16.05 | 16.05
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.4406%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 237b (Indicative Value 37 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.7 | 1
| 1 | 2 | 37.2 | 9
| Mean |  | 36.95 | 16
| Median |  | 36.95 | 36.95
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 0.9568%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 237b (Indicative Value 69 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 68 | 1
| 1 | 2 | 70 | 34
| Mean |  | 69 | 16
| Median |  | 69 | 69
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 237b (Indicative Value 21.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21 | 1
| 1 | 2 | 22 | 35
| Mean |  | 21.5 | 16
| Median |  | 21.5 | 21.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 237b (Indicative Value 9.65 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.6 | 1
| 1 | 2 | 9.7 | 11
| Mean |  | 9.65 | 16
| Median |  | 9.65 | 9.65
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 237b (Indicative Value 160 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 160 | 37
| Mean |  | 160 | 16
| Median |  | 160 | 160
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 237b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 237b (Indicative Value 479 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 480 | 1
| 1 | 2 | 478 | 39
| Mean |  | 479 | 16
| Median |  | 479 | 479
| Std Dev. |  | 1.4142 | 45
| Rel.Std.Dev. |  | 0.2952%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 237b (Indicative Value 14.6 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.4 | 1
| 1 | 2 | 14.8 | 40
| Mean |  | 14.6 | 16
| Median |  | 14.6 | 14.6
| Std Dev. |  | 0.28284 | 46
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 237b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 237b (Indicative Value 7.25 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.37 | 1
| 1 | 2 | 7.13 | 10
| Mean |  | 7.25 | 16
| Median |  | 7.25 | 7.25
| Std Dev. |  | 0.16971 | 31
| Rel.Std.Dev. |  | 2.341%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 237b (Indicative Value 3.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.6 | 26
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 237b (Indicative Value 122 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 123 | 1
| 1 | 2 | 121 | 27
| Mean |  | 122 | 16
| Median |  | 122 | 122
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.159%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 237b (Indicative Value 1.26 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.26 | 28
| Mean |  | 1.255 | 16
| Median |  | 1.255 | 1.255
| Std Dev. |  | 0.0070711 | 34
| Rel.Std.Dev. |  | 0.5634%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 237b (Indicative Value 0.97 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.98 | 1
| 1 | 2 | 0.95 | 11
| Mean |  | 0.965 | 16
| Median |  | 0.965 | 0.965
| Std Dev. |  | 0.021213 | 35
| Rel.Std.Dev. |  | 2.198%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 237b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 6
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 237b (Indicative Value 15.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.2 | 1
| 1 | 2 | 15.1 | 31
| Mean |  | 15.15 | 16
| Median |  | 15.15 | 15.15
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.4667%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 237b (Indicative Value 0.494 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.492 | 1
| 1 | 2 | 0.496 | 32
| Mean |  | 0.494 | 16
| Median |  | 0.494 | 0.494
| Std Dev. |  | 0.0028284 | 38
| Rel.Std.Dev. |  | 0.5726%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 237b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 237b (Indicative Value 0.53 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.53 | 1
| 1 | 2 | 0.52 | 13
| Mean |  | 0.525 | 16
| Median |  | 0.525 | 0.525
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 237b (Indicative Value 3.29 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.25 | 1
| 1 | 2 | 3.32 | 35
| Mean |  | 3.285 | 16
| Median |  | 3.285 | 3.285
| Std Dev. |  | 0.049497 | 41
| Rel.Std.Dev. |  | 1.507%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 237b (Indicative Value 111 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 112 | 1
| 1 | 2 | 110 | 36
| Mean |  | 111 | 16
| Median |  | 111 | 111
| Std Dev. |  | 1.4142 | 42
| Rel.Std.Dev. |  | 1.274%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 237b (Indicative Value 3.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.5 | 37
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 237b (Indicative Value 31.2 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.2 | 1
| 1 | 2 | 31.1 | 38
| Mean |  | 31.15 | 16
| Median |  | 31.15 | 31.15
| Std Dev. |  | 0.070711 | 44
| Rel.Std.Dev. |  | 0.227%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 237b (Indicative Value 3.37 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.36 | 1
| 1 | 2 | 3.38 | 39
| Mean |  | 3.37 | 16
| Median |  | 3.37 | 3.37
| Std Dev. |  | 0.014142 | 45
| Rel.Std.Dev. |  | 0.4196%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 237b (Indicative Value 95 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 90 | 40
| Mean |  | 95 | 16
| Median |  | 95 | 95
| Std Dev. |  | 7.0711 | 46
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 237b (Indicative Value 240 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 236 | 1
| 1 | 2 | 244 | 41
| Mean |  | 240 | 16
| Median |  | 240 | 240
| Std Dev. |  | 5.6569 | 47
| Rel.Std.Dev. |  | 2.357%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.