## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 267
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 13.2103 | 13.0617 | 13.3588 | 13.175 | 13.2455
PhotonAssay
Au, Gold (ppm) | 13.5614 | 13.4585 | 13.6643 | 13.5252 | 13.5976
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 13.4171 | 13.1447 | 13.6896 | 13.3779 | 13.4564
Cyanide Leach
Au, Gold (ppm) | 12.9572 | 12.6539 | 13.2605 | 12.9438 | 12.9706
4-Acid Digestion
Ag, Silver (ppm) | 5.5404 | 5.3186 | 5.7622 | 5.4054 | 5.6754
Al, Aluminium (wt.%) | 5.046 | 4.914 | 5.178 | 4.9561 | 5.1359
As, Arsenic (ppm) | 136.9254 | 131.4598 | 142.391 | 133.9439 | 139.9069
Ba, Barium (ppm) | 381.1845 | 369.0815 | 393.2875 | 374.7355 | 387.6335
Be, Beryllium (ppm) | 1.8045 | 1.7285 | 1.8805 | 1.74 | 1.8689
Bi, Bismuth (ppm) | 1.281 | 1.1677 | 1.3942 | 1.2075 | 1.3545
Ca, Calcium (wt.%) | 1.1324 | 1.1041 | 1.1607 | 1.1149 | 1.1498
Ce, Cerium (ppm) | 58.2847 | 55.064 | 61.5053 | 56.6968 | 59.8726
Co, Cobalt (ppm) | 11.0372 | 10.5585 | 11.516 | 10.7169 | 11.3576
Cr, Chromium (ppm) | 97.8292 | 91.9181 | 103.7403 | 95.2133 | 100.4451
Cs, Caesium (ppm) | 5.4953 | 5.2456 | 5.745 | 5.3402 | 5.6505
Cu, Copper (ppm) | 23.8994 | 22.5521 | 25.2467 | 22.7801 | 25.0187
Dy, Dysprosium (ppm) | 3.0711 | 2.8702 | 3.272 | 2.9585 | 3.1836
Er, Erbium (ppm) | 1.499 | 1.332 | 1.6661 | 1.4348 | 1.5633
Eu, Europium (ppm) | 1.0285 | 0.93029 | 1.1267 | 0.97416 | 1.0828
Fe, Iron (wt.%) | 2.8626 | 2.7989 | 2.9262 | 2.824 | 2.9011
Ga, Gallium (ppm) | 14.5213 | 13.7476 | 15.295 | 14.0461 | 14.9965
Gd, Gadolinium (ppm) | 4.102 | 3.5162 | 4.6878 | 3.9877 | 4.2163
Hf, Hafnium (ppm) | 2.9642 | 2.7542 | 3.1742 | 2.8182 | 3.1102
Ho, Holmium (ppm) | 0.56362 | 0.51009 | 0.61714 | 0.51953 | 0.6077
In, Indium (ppm) | 0.060833 | 0.05219 | 0.069477 | 0.055173 | 0.066494
K, Potassium (wt.%) | 1.2467 | 1.2207 | 1.2727 | 1.2272 | 1.2662
La, Lanthanum (ppm) | 30.9554 | 29.67 | 32.2408 | 30.0512 | 31.8596
Li, Lithium (ppm) | 55.367 | 53.3977 | 57.3364 | 54.4619 | 56.2722
Lu, Lutetium (ppm) | 0.20256 | 0.18422 | 0.2209 | 0.18188 | 0.22325
Mg, Magnesium (wt.%) | 0.92851 | 0.90362 | 0.95341 | 0.91346 | 0.94357
Mn, Manganese (wt.%) | 0.027838 | 0.027284 | 0.028392 | 0.027408 | 0.028268
Mo, Molybdenum (ppm) | 6.9371 | 6.5535 | 7.3207 | 6.6983 | 7.176
Na, Sodium (wt.%) | 0.50469 | 0.48926 | 0.52013 | 0.49524 | 0.51415
Nb, Niobium (ppm) | 13.8542 | 13.0385 | 14.67 | 13.502 | 14.2065
Nd, Neodymium (ppm) | 26.6693 | 25.098 | 28.2407 | 25.7421 | 27.5966
Ni, Nickel (ppm) | 48.9097 | 46.7184 | 51.101 | 47.7151 | 50.1044
P, Phosphorus (wt.%) | 0.0476 | 0.046031 | 0.049169 | 0.04658 | 0.04862
Pb, Lead (ppm) | 15.2088 | 13.9367 | 16.4809 | 14.5838 | 15.8338
Pr, Praseodymium (ppm) | 6.9529 | 6.4513 | 7.4546 | 6.7697 | 7.1362
Rb, Rubidium (ppm) | 71.0678 | 68.2329 | 73.9027 | 69.0928 | 73.0428
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.024426 | 0.02251 | 0.026341 | 0.023875 | 0.024976
Sb, Antimony (ppm) | 5.1463 | 4.1496 | 6.143 | 4.6674 | 5.6252
Sc, Scandium (ppm) | 9.2516 | 8.7204 | 9.7828 | 9.0482 | 9.455
Sm, Samarium (ppm) | 5.1092 | 4.7238 | 5.4946 | 4.8886 | 5.3299
Sn, Tin (ppm) | 5.9531 | 5.5634 | 6.3428 | 5.6908 | 6.2153
Sr, Strontium (ppm) | 158.9316 | 154.5336 | 163.3297 | 156.4166 | 161.4467
Ta, Tantalum (ppm) | 0.93065 | 0.8722 | 0.98911 | 0.8898 | 0.9715
Tb, Terbium (ppm) | 0.54324 | 0.5164 | 0.57007 | 0.52772 | 0.55876
Te, Tellurium (ppm) | 0.067733 | 0.037676 | 0.097791 | IND | IND
Th, Thorium (ppm) | 9.1332 | 8.7355 | 9.5309 | 8.9163 | 9.35
Ti, Titanium (wt.%) | 0.37288 | 0.36124 | 0.38453 | 0.36547 | 0.3803
Tl, Thallium (ppm) | 0.39265 | 0.36599 | 0.41931 | 0.37537 | 0.40992
U, Uranium (ppm) | 1.9291 | 1.8099 | 2.0483 | 1.8536 | 2.0046
V, Vanadium (ppm) | 68.0536 | 65.9313 | 70.1759 | 66.3292 | 69.778
W, Tungsten (ppm) | 7.3236 | 6.9928 | 7.6544 | 7.0235 | 7.6237
Y, Yttrium (ppm) | 14.0028 | 13.2191 | 14.7865 | 13.6034 | 14.4022
Yb, Ytterbium (ppm) | 1.3787 | 1.263 | 1.4944 | 1.2906 | 1.4668
Zn, Zinc (ppm) | 43.4263 | 41.7913 | 45.0613 | 42.1122 | 44.7405
Zr, Zirconium (ppm) | 112.226 | 106.9124 | 117.5396 | 109.0085 | 115.4435
Aqua Regia Digestion
Ag, Silver (ppm) | 5.3495 | 5.158 | 5.5409 | 5.2161 | 5.4828
Al, Aluminium (wt.%) | 0.8788 | 0.81956 | 0.93803 | 0.84987 | 0.90772
As, Arsenic (ppm) | 111.617 | 107.1955 | 116.0385 | 108.4621 | 114.7719
Ba, Barium (ppm) | 71.4725 | 66.9539 | 75.9911 | 68.5218 | 74.4232
Be, Beryllium (ppm) | 0.52094 | 0.49373 | 0.54815 | 0.50178 | 0.5401
Bi, Bismuth (ppm) | 1.0349 | 0.93849 | 1.1312 | 0.95951 | 1.1102
Ca, Calcium (wt.%) | 0.27926 | 0.26223 | 0.2963 | 0.26903 | 0.28949
Ce, Cerium (ppm) | 31.8835 | 29.7381 | 34.0289 | 30.5521 | 33.2148
Co, Cobalt (ppm) | 7.0171 | 6.5347 | 7.4994 | 6.7384 | 7.2957
Cr, Chromium (ppm) | 46.7753 | 44.8245 | 48.7261 | 45.2167 | 48.3338
Cs, Caesium (ppm) | 1.1847 | 1.1256 | 1.2438 | 1.1413 | 1.228
Cu, Copper (ppm) | 19.7317 | 18.7406 | 20.7228 | 18.8594 | 20.604
Fe, Iron (wt.%) | 2.0245 | 1.9403 | 2.1088 | 1.9745 | 2.0746
Ga, Gallium (ppm) | 3.7107 | 3.4051 | 4.0162 | 3.5004 | 3.9209
Hf, Hafnium (ppm) | 0.48259 | 0.43587 | 0.5293 | 0.45506 | 0.51011
Hg, Mercury (ppm) | 0.042833 | 0.030374 | 0.055293 | IND | IND
In, Indium (ppm) | 0.020233 | 0.0166 | 0.023866 | 0.016414 | 0.024052
K, Potassium (wt.%) | 0.15821 | 0.14444 | 0.17199 | 0.15315 | 0.16328
La, Lanthanum (ppm) | 15.1627 | 14.081 | 16.2443 | 14.655 | 15.6704
Li, Lithium (ppm) | 5.5561 | 5.1006 | 6.0116 | 5.3355 | 5.7767
Mg, Magnesium (wt.%) | 0.37136 | 0.35341 | 0.3893 | 0.35604 | 0.38667
Mn, Manganese (wt.%) | 0.014759 | 0.014052 | 0.015465 | 0.014325 | 0.015192
Mo, Molybdenum (ppm) | 5.8875 | 5.686 | 6.089 | 5.6965 | 6.0785
Na, Sodium (wt.%) | 0.082544 | 0.072269 | 0.09282 | 0.079449 | 0.08564
Nb, Niobium (ppm) | 0.34868 | 0.28752 | 0.40985 | 0.31412 | 0.38325
Ni, Nickel (ppm) | 39.6436 | 37.3301 | 41.9571 | 38.2265 | 41.0607
P, Phosphorus (wt.%) | 0.031842 | 0.030476 | 0.033207 | 0.03088 | 0.032804
Pb, Lead (ppm) | 8.6041 | 7.907 | 9.3012 | 8.2571 | 8.9511
Rb, Rubidium (ppm) | 11.2477 | 10.3593 | 12.1361 | 10.7978 | 11.6976
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.022195 | 0.019925 | 0.024464 | 0.020857 | 0.023533
Sc, Scandium (ppm) | 2.2298 | 1.9858 | 2.4739 | 2.052 | 2.4077
Sn, Tin (ppm) | 1.1762 | 1.0089 | 1.3434 | 0.93633 | 1.4161
Sr, Strontium (ppm) | 30.6498 | 28.232 | 33.0675 | 29.6195 | 31.6801
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Te, Tellurium (ppm) | 0.053499 | 0.032173 | 0.074825 | IND | IND
Th, Thorium (ppm) | 5.9634 | 5.6452 | 6.2815 | 5.82 | 6.1067
Ti, Titanium (wt.%) | 0.059906 | 0.053141 | 0.06667 | 0.057574 | 0.062237
Tl, Thallium (ppm) | 0.090064 | 0.078917 | 0.10121 | 0.076279 | 0.10385
U, Uranium (ppm) | 0.89894 | 0.84948 | 0.94841 | 0.86863 | 0.92926
V, Vanadium (ppm) | 20.8081 | 19.7332 | 21.8829 | 20.2095 | 21.4066
W, Tungsten (ppm) | 1.8138 | 1.5717 | 2.056 | 1.685 | 1.9426
Y, Yttrium (ppm) | 6.2889 | 5.9184 | 6.6593 | 6.1093 | 6.4684
Zn, Zinc (ppm) | 23.5071 | 21.7676 | 25.2465 | 22.3353 | 24.6788
Zr, Zirconium (ppm) | 21.2834 | 20.2321 | 22.3346 | 20.386 | 22.1808
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 267
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Cd | ppm | 0.032833 | Hg | ppm | 0.0675 | Tm | ppm | 0.21176
Ge | ppm | 0.26021 | Se | ppm | 0.75478
Aqua Regia Digestion
B | ppm | 5.2932 | Ge | ppm | 0.0725 | Se | ppm | 0.26976
Cd | ppm | 0.025796 | Ho | ppm | 0.28012 | Sm | ppm | 3.0023
Dy | ppm | 1.6076 | Lu | ppm | 0.062021 | Tb | ppm | 0.32045
Er | ppm | 0.65586 | Nd | ppm | 15.5104 | Tm | ppm | < 0.1
Eu | ppm | 0.5728 | Pr | ppm | 4.2042 | Yb | ppm | 0.48429
Gd | ppm | 2.4256 | Sb | ppm | 1.658
Borate Fusion XRF
Al2O3 | wt.% | 10.095 | MgO | wt.% | 1.6 | SiO2 | wt.% | 77.06
CaO | wt.% | 1.585 | MnO | wt.% | 0.03 | TiO2 | wt.% | 0.7
Fe2O3 | wt.% | 4.245 | P2O5 | wt.% | 0.1085
K2O | wt.% | 1.49 | S | wt.% | 0.020823
Thermogravimetry
LOI1000 | wt.% | 2.145
Infrared Combustion
C | wt.% | 0.06 | S | wt.% | 0.015
Laser Ablation ICP-MS
Ag | ppm | 7.2 | Hf | ppm | 4.81 | Sn | ppm | 8.7
As | ppm | 157.5 | Ho | ppm | 0.745 | Sr | ppm | 151
Ba | ppm | 388 | In | ppm | 0.05 | Ta | ppm | 1.685
Be | ppm | 1.9 | La | ppm | 31.25 | Tb | ppm | 0.66
Bi | ppm | 1.63 | Lu | ppm | 0.275 | Te | ppm | < 0.2
Cd | ppm | < 0.1 | Mn | wt.% | 0.0288 | Th | ppm | 9.55
Ce | ppm | 60.05 | Mo | ppm | 6.6 | Ti | wt.% | 0.415
Co | ppm | 11.85 | Nb | ppm | 16.3 | Tl | ppm | 0.3
Cr | ppm | 112 | Nd | ppm | 27.35 | Tm | ppm | 0.3
Cs | ppm | 6.305 | Ni | ppm | 57 | U | ppm | 2.07
Cu | ppm | 31 | Pb | ppm | 18 | V | ppm | 71.95
Dy | ppm | 3.785 | Pr | ppm | 7.29 | W | ppm | 8.375
Er | ppm | 2.035 | Rb | ppm | 77.3 | Y | ppm | 19.25
Eu | ppm | 1.045 | Re | ppm | < 0.01 | Yb | ppm | 2.025
Ga | ppm | 16.95 | Sb | ppm | 7.9 | Zn | ppm | 50
Gd | ppm | 4.345 | Sc | ppm | 10.25 | Zr | ppm | 172.5
Ge | ppm | 1.35 | Sm | ppm | 5.24
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 267
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 13.2103 | 0.40043 | 12.4094 | 14.0111 | 12.009 | 14.4116 | 3.031% | 6.062% | 9.094% | 12.5498 | 13.8708
PhotonAssay
Au, ppm | 13.5614 | 0.28426 | 12.9928 | 14.1299 | 12.7086 | 14.4141 | 2.096% | 4.192% | 6.288% | 12.8833 | 14.2394
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 13.4171 | 0.52628 | 12.3646 | 14.4697 | 11.8383 | 14.996 | 3.922% | 7.845% | 11.77% | 12.7463 | 14.088
Cyanide Leach
Au, ppm | 12.9572 | 0.64283 | 11.6715 | 14.2429 | 11.0287 | 14.8857 | 4.961% | 9.922% | 14.88% | 12.3093 | 13.6051
4-Acid Digestion
Ag, ppm | 5.5404 | 0.21496 | 5.1105 | 5.9703 | 4.8955 | 6.1853 | 3.88% | 7.76% | 11.64% | 5.2634 | 5.8174
Al, wt.% | 5.046 | 0.18633 | 4.6733 | 5.4186 | 4.487 | 5.605 | 3.693% | 7.385% | 11.08% | 4.7937 | 5.2983
As, ppm | 136.9254 | 5.409 | 126.1074 | 147.7434 | 120.6985 | 153.1524 | 3.95% | 7.901% | 11.85% | 130.0791 | 143.7717
Ba, ppm | 381.1845 | 12.2222 | 356.74 | 405.629 | 344.5178 | 417.8512 | 3.206% | 6.413% | 9.619% | 362.1253 | 400.2437
Be, ppm | 1.8045 | 0.072176 | 1.6601 | 1.9488 | 1.588 | 2.021 | 4% | 8% | 12% | 1.7143 | 1.8947
Bi, ppm | 1.281 | 0.12258 | 1.0358 | 1.5261 | 0.91325 | 1.6487 | 9.569% | 19.14% | 28.71% | 1.2169 | 1.345
Ca, wt.% | 1.1324 | 0.030595 | 1.0712 | 1.1936 | 1.0406 | 1.2242 | 2.702% | 5.404% | 8.106% | 1.0758 | 1.189
Ce, ppm | 58.2847 | 3.2313 | 51.8221 | 64.7473 | 48.5908 | 67.9786 | 5.544% | 11.09% | 16.63% | 55.3704 | 61.1989
Co, ppm | 11.0372 | 0.25759 | 10.522 | 11.5524 | 10.2645 | 11.81 | 2.334% | 4.668% | 7.001% | 10.4854 | 11.5891
Cr, ppm | 97.8292 | 7.0662 | 83.6968 | 111.9616 | 76.6306 | 119.0278 | 7.223% | 14.45% | 21.67% | 92.9378 | 102.7207
Cs, ppm | 5.4953 | 0.15403 | 5.1873 | 5.8034 | 5.0333 | 5.9574 | 2.803% | 5.606% | 8.409% | 5.2206 | 5.7701
Cu, ppm | 23.8994 | 1.3905 | 21.1183 | 26.6805 | 19.7278 | 28.071 | 5.818% | 11.64% | 17.45% | 22.7044 | 25.0944
Dy, ppm | 3.0711 | 0.092192 | 2.8867 | 3.2555 | 2.7945 | 3.3477 | 3.002% | 6.004% | 9.006% | 2.9175 | 3.2246
Er, ppm | 1.499 | 0.092801 | 1.3134 | 1.6846 | 1.2206 | 1.7774 | 6.191% | 12.38% | 18.57% | 1.4241 | 1.574
Eu, ppm | 1.0285 | 0.038989 | 0.95051 | 1.1065 | 0.91153 | 1.1455 | 3.791% | 7.582% | 11.37% | 0.97707 | 1.0799
Fe, wt.% | 2.8626 | 0.068486 | 2.7256 | 2.9995 | 2.6571 | 3.068 | 2.392% | 4.785% | 7.177% | 2.7194 | 3.0057
Ga, ppm | 14.5213 | 0.74329 | 13.0347 | 16.0079 | 12.2914 | 16.7512 | 5.119% | 10.24% | 15.36% | 13.7952 | 15.2474
Gd, ppm | 4.102 | 0.32425 | 3.4535 | 4.7505 | 3.1292 | 5.0748 | 7.905% | 15.81% | 23.71% | 3.8969 | 4.3071
Hf, ppm | 2.9642 | 0.19494 | 2.5743 | 3.3541 | 2.3794 | 3.549 | 6.577% | 13.15% | 19.73% | 2.816 | 3.1124
Ho, ppm | 0.56362 | 0.029215 | 0.50519 | 0.62205 | 0.47597 | 0.65126 | 5.183% | 10.37% | 15.55% | 0.53544 | 0.5918
In, ppm | 0.060833 | 0.0051276 | 0.050578 | 0.071089 | 0.04545 | 0.076216 | 8.429% | 16.86% | 25.29% | 0.057792 | 0.063875
K, wt.% | 1.2467 | 0.022816 | 1.2011 | 1.2923 | 1.1782 | 1.3151 | 1.83% | 3.66% | 5.49% | 1.1844 | 1.309
La, ppm | 30.9554 | 1.2769 | 28.4017 | 33.5092 | 27.1248 | 34.786 | 4.125% | 8.25% | 12.37% | 29.4076 | 32.5032
Li, ppm | 55.367 | 2.1037 | 51.1596 | 59.5745 | 49.0558 | 61.6783 | 3.8% | 7.599% | 11.4% | 52.5987 | 58.1354
Lu, ppm | 0.20256 | 0.006964 | 0.18863 | 0.21649 | 0.18167 | 0.22345 | 3.438% | 6.876% | 10.31% | 0.19243 | 0.21269
Mg, wt.% | 0.92851 | 0.033744 | 0.86103 | 0.996 | 0.82728 | 1.0297 | 3.634% | 7.268% | 10.9% | 0.88209 | 0.97494
Mn, wt.% | 0.027838 | 0.0007299 | 0.026378 | 0.029298 | 0.025648 | 0.030028 | 2.622% | 5.244% | 7.866% | 0.026446 | 0.02923
Mo, ppm | 6.9371 | 0.34482 | 6.2475 | 7.6268 | 5.9027 | 7.9716 | 4.971% | 9.941% | 14.91% | 6.5903 | 7.284
Na, wt.% | 0.50469 | 0.021898 | 0.4609 | 0.54849 | 0.439 | 0.57039 | 4.339% | 8.678% | 13.02% | 0.47946 | 0.52993
Nb, ppm | 13.8542 | 0.98022 | 11.8938 | 15.8147 | 10.9136 | 16.7949 | 7.075% | 14.15% | 21.23% | 13.1615 | 14.547
Nd, ppm | 26.6693 | 0.84611 | 24.9771 | 28.3616 | 24.131 | 29.2077 | 3.173% | 6.345% | 9.518% | 25.3359 | 28.0028
Ni, ppm | 48.9097 | 2.3863 | 44.1371 | 53.6823 | 41.7508 | 56.0687 | 4.879% | 9.758% | 14.64% | 46.4642 | 51.3552
P, wt.% | 0.0476 | 0.0011999 | 0.0452 | 0.05 | 0.044 | 0.0512 | 2.521% | 5.042% | 7.563% | 0.04522 | 0.04998
Pb, ppm | 15.2088 | 1.5012 | 12.2064 | 18.2112 | 10.7052 | 19.7124 | 9.871% | 19.74% | 29.61% | 14.4484 | 15.9692
Pr, ppm | 6.9529 | 0.3408 | 6.2713 | 7.6345 | 5.9305 | 7.9753 | 4.902% | 9.803% | 14.7% | 6.6053 | 7.3006
Rb, ppm | 71.0678 | 2.5713 | 65.9251 | 76.2104 | 63.3538 | 78.7818 | 3.618% | 7.236% | 10.85% | 67.5144 | 74.6212
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.024426 | 0.0039123 | 0.016601 | 0.03225 | 0.012689 | 0.036163 | 16.02% | 32.03% | 48.05% | 0.023204 | 0.025647
Sb, ppm | 5.1463 | 1.5917 | 1.9629 | 8.3297 | 0.37124 | 9.9214 | 30.93% | 61.86% | 92.79% | 4.889 | 5.4037
Sc, ppm | 9.2516 | 0.60183 | 8.0479 | 10.4553 | 7.4461 | 11.0571 | 6.505% | 13.01% | 19.52% | 8.789 | 9.7142
Sm, ppm | 5.1092 | 0.18614 | 4.7369 | 5.4815 | 4.5508 | 5.6676 | 3.643% | 7.286% | 10.93% | 4.8538 | 5.3647
Sn, ppm | 5.9531 | 0.28764 | 5.3778 | 6.5283 | 5.0901 | 6.816 | 4.832% | 9.664% | 14.5% | 5.6554 | 6.2507
Sr, ppm | 158.9316 | 3.5166 | 151.8985 | 165.9648 | 148.3819 | 169.4813 | 2.213% | 4.425% | 6.638% | 150.9851 | 166.8782
Ta, ppm | 0.93065 | 0.037402 | 0.85585 | 1.0055 | 0.81845 | 1.0429 | 4.019% | 8.038% | 12.06% | 0.88412 | 0.97718
Tb, ppm | 0.54324 | 0.015599 | 0.51204 | 0.57443 | 0.49644 | 0.59003 | 2.871% | 5.743% | 8.614% | 0.51607 | 0.5704
Te, ppm | 0.067733 | 0.011339 | 0.045055 | 0.090411 | 0.033717 | 0.10175 | 16.74% | 33.48% | 50.22% | 0.064347 | 0.07112
Th, ppm | 9.1332 | 0.34015 | 8.4529 | 9.8135 | 8.1127 | 10.1536 | 3.724% | 7.449% | 11.17% | 8.6765 | 9.5898
Ti, wt.% | 0.37288 | 0.01322 | 0.34645 | 0.39932 | 0.33323 | 0.41254 | 3.545% | 7.09% | 10.64% | 0.35424 | 0.39153
Tl, ppm | 0.39265 | 0.025343 | 0.34196 | 0.44333 | 0.31662 | 0.46867 | 6.454% | 12.91% | 19.36% | 0.37301 | 0.41228
U, ppm | 1.9291 | 0.1209 | 1.6873 | 2.1709 | 1.5664 | 2.2918 | 6.267% | 12.53% | 18.8% | 1.8326 | 2.0256
V, ppm | 68.0536 | 1.8917 | 64.2701 | 71.837 | 62.3784 | 73.7287 | 2.78% | 5.559% | 8.339% | 64.6509 | 71.4563
W, ppm | 7.3236 | 0.35965 | 6.6043 | 8.0429 | 6.2447 | 8.4025 | 4.911% | 9.822% | 14.73% | 6.9574 | 7.6898
Y, ppm | 14.0028 | 0.69749 | 12.6078 | 15.3978 | 11.9103 | 16.0953 | 4.981% | 9.962% | 14.94% | 13.3027 | 14.7029
Yb, ppm | 1.3787 | 0.070056 | 1.2386 | 1.5188 | 1.1685 | 1.5889 | 5.081% | 10.16% | 15.24% | 1.3098 | 1.4476
Zn, ppm | 43.4263 | 1.5991 | 40.2282 | 46.6244 | 38.6291 | 48.2235 | 3.682% | 7.364% | 11.05% | 41.255 | 45.5976
Zr, ppm | 112.226 | 2.7366 | 106.7527 | 117.6992 | 104.0161 | 120.4359 | 2.438% | 4.877% | 7.315% | 106.6147 | 117.8373
Aqua Regia Digestion
Ag, ppm | 5.3495 | 0.17971 | 4.99 | 5.7089 | 4.8103 | 5.8886 | 3.359% | 6.719% | 10.08% | 5.082 | 5.6169
Al, wt.% | 0.8788 | 0.08019 | 0.71842 | 1.0392 | 0.63823 | 1.1194 | 9.125% | 18.25% | 27.37% | 0.83486 | 0.92274
As, ppm | 111.617 | 4.2551 | 103.1068 | 120.1272 | 98.8517 | 124.3823 | 3.812% | 7.624% | 11.44% | 106.0362 | 117.1979
Ba, ppm | 71.4725 | 8.5212 | 54.4301 | 88.5149 | 45.9089 | 97.0361 | 11.92% | 23.84% | 35.77% | 67.8989 | 75.0461
Be, ppm | 0.52094 | 0.018079 | 0.48478 | 0.5571 | 0.4667 | 0.57518 | 3.47% | 6.941% | 10.41% | 0.49489 | 0.54699
Bi, ppm | 1.0349 | 0.084757 | 0.86534 | 1.2044 | 0.78058 | 1.2891 | 8.19% | 16.38% | 24.57% | 0.98311 | 1.0866
Ca, wt.% | 0.27926 | 0.023747 | 0.23177 | 0.32676 | 0.20802 | 0.3505 | 8.503% | 17.01% | 25.51% | 0.2653 | 0.29323
Ce, ppm | 31.8835 | 2.6099 | 26.6636 | 37.1033 | 24.0537 | 39.7133 | 8.186% | 16.37% | 24.56% | 30.2893 | 33.4777
Co, ppm | 7.0171 | 0.58196 | 5.8531 | 8.181 | 5.2712 | 8.7629 | 8.294% | 16.59% | 24.88% | 6.6662 | 7.3679
Cr, ppm | 46.7753 | 2.0937 | 42.588 | 50.9626 | 40.4943 | 53.0563 | 4.476% | 8.952% | 13.43% | 44.4365 | 49.114
Cs, ppm | 1.1847 | 0.057488 | 1.0697 | 1.2996 | 1.0122 | 1.3571 | 4.853% | 9.705% | 14.56% | 1.1254 | 1.2439
Cu, ppm | 19.7317 | 1.1024 | 17.5268 | 21.9365 | 16.4244 | 23.039 | 5.587% | 11.17% | 16.76% | 18.7451 | 20.7183
Fe, wt.% | 2.0245 | 0.093688 | 1.8372 | 2.2119 | 1.7435 | 2.3056 | 4.628% | 9.255% | 13.88% | 1.9233 | 2.1258
Ga, ppm | 3.7107 | 0.43475 | 2.8412 | 4.5802 | 2.4064 | 5.0149 | 11.72% | 23.43% | 35.15% | 3.5251 | 3.8962
Hf, ppm | 0.48259 | 0.052428 | 0.37773 | 0.58744 | 0.3253 | 0.63987 | 10.86% | 21.73% | 32.59% | 0.45846 | 0.50672
Hg, ppm | 0.042833 | 0.0080535 | 0.026726 | 0.05894 | 0.018673 | 0.066994 | 18.8% | 37.6% | 56.41% | 0.040692 | 0.044975
In, ppm | 0.020233 | 0.0016491 | 0.016935 | 0.023531 | 0.015286 | 0.02518 | 8.15% | 16.3% | 24.45% | 0.019221 | 0.021245
K, wt.% | 0.15821 | 0.019879 | 0.11846 | 0.19797 | 0.098577 | 0.21785 | 12.56% | 25.13% | 37.69% | 0.1503 | 0.16613
La, ppm | 15.1627 | 1.2975 | 12.5677 | 17.7576 | 11.2703 | 19.0551 | 8.557% | 17.11% | 25.67% | 14.4045 | 15.9208
Li, ppm | 5.5561 | 0.60343 | 4.3492 | 6.7629 | 3.7458 | 7.3664 | 10.86% | 21.72% | 32.58% | 5.2783 | 5.8339
Mg, wt.% | 0.37136 | 0.021671 | 0.32801 | 0.4147 | 0.30634 | 0.43637 | 5.836% | 11.67% | 17.51% | 0.35279 | 0.38992
Mn, wt.% | 0.014759 | 0.00093155 | 0.012895 | 0.016622 | 0.011964 | 0.017553 | 6.312% | 12.62% | 18.94% | 0.014021 | 0.015496
Mo, ppm | 5.8875 | 0.23494 | 5.4176 | 6.3574 | 5.1827 | 6.5923 | 3.991% | 7.981% | 11.97% | 5.5931 | 6.1819
Na, wt.% | 0.082544 | 0.012271 | 0.058003 | 0.10709 | 0.045732 | 0.11936 | 14.87% | 29.73% | 44.6% | 0.078417 | 0.086672
Nb, ppm | 0.34868 | 0.066548 | 0.21559 | 0.48178 | 0.14904 | 0.54833 | 19.09% | 38.17% | 57.26% | 0.33125 | 0.36612
Ni, ppm | 39.6436 | 2.5896 | 34.4643 | 44.8229 | 31.8747 | 47.4125 | 6.532% | 13.06% | 19.6% | 37.6614 | 41.6258
P, wt.% | 0.031842 | 0.0012414 | 0.029359 | 0.034325 | 0.028118 | 0.035566 | 3.899% | 7.797% | 11.7% | 0.03025 | 0.033434
Pb, ppm | 8.6041 | 0.91891 | 6.7663 | 10.4419 | 5.8474 | 11.3608 | 10.68% | 21.36% | 32.04% | 8.1739 | 9.0343
Rb, ppm | 11.2477 | 0.91417 | 9.4194 | 13.076 | 8.5052 | 13.9902 | 8.128% | 16.26% | 24.38% | 10.6853 | 11.8101
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.022195 | 0.0035083 | 0.015178 | 0.029211 | 0.01167 | 0.03272 | 15.81% | 31.61% | 47.42% | 0.021085 | 0.023305
Sc, ppm | 2.2298 | 0.19823 | 1.8334 | 2.6263 | 1.6351 | 2.8245 | 8.89% | 17.78% | 26.67% | 2.1183 | 2.3413
Sn, ppm | 1.1762 | 0.15649 | 0.86322 | 1.4892 | 0.70673 | 1.6457 | 13.3% | 26.61% | 39.91% | 1.1174 | 1.235
Sr, ppm | 30.6498 | 3.0893 | 24.4712 | 36.8283 | 21.3819 | 39.9176 | 10.08% | 20.16% | 30.24% | 29.1173 | 32.1822
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Te, ppm | 0.053499 | 0.011453 | 0.030594 | 0.076404 | 0.019141 | 0.087857 | 21.41% | 42.81% | 64.22% | 0.050824 | 0.056174
Th, ppm | 5.9634 | 0.32823 | 5.3069 | 6.6198 | 4.9787 | 6.9481 | 5.504% | 11.01% | 16.51% | 5.6652 | 6.2615
Ti, wt.% | 0.059906 | 0.010541 | 0.038823 | 0.080988 | 0.028282 | 0.091529 | 17.6% | 35.19% | 52.79% | 0.05691 | 0.062901
Tl, ppm | 0.090064 | 0.012886 | 0.064291 | 0.11584 | 0.051405 | 0.12872 | 14.31% | 28.62% | 42.92% | 0.085561 | 0.094567
U, ppm | 0.89894 | 0.034699 | 0.82954 | 0.96834 | 0.79485 | 1.003 | 3.86% | 7.72% | 11.58% | 0.854 | 0.94389
V, ppm | 20.8081 | 0.7545 | 19.2991 | 22.3171 | 18.5446 | 23.0716 | 3.626% | 7.252% | 10.88% | 19.7677 | 21.8485
W, ppm | 1.8138 | 0.25444 | 1.305 | 2.3227 | 1.0505 | 2.5772 | 14.03% | 28.06% | 42.08% | 1.7231 | 1.9045
Y, ppm | 6.2889 | 0.46297 | 5.3629 | 7.2148 | 4.9 | 7.6778 | 7.362% | 14.72% | 22.09% | 5.9744 | 6.6033
Zn, ppm | 23.5071 | 2.3423 | 18.8224 | 28.1917 | 16.48 | 30.5341 | 9.964% | 19.93% | 29.89% | 22.3317 | 24.6824
Zr, ppm | 21.2834 | 1.0343 | 19.2147 | 23.352 | 18.1804 | 24.3863 | 4.86% | 9.719% | 14.58% | 20.2192 | 22.3475
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 267
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PAAU02 | PhotonAssay (Au) two-cycle analysis
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 267
Lab Slot | Laboratory Name and Location
Accura Geomysore Labs Private Limited, Jonnagiri, Kurnool, India
Accura Gold Minerals Testing, Namakkal, Tamilnadu, India
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
Paragon Geochemical Laboratories, Surrey, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Indo Mineral Research, Bungursari, West Java, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 267 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 8 | 8 | 290562 | 0.083873 | 13.9847 | 14.1671
2 | 1 | ANSLu | 1 | 19 | 11 | 290563 | 0.082183 | 14.0918 | 14.1728
2 | 1 | ANSLu | 1 | 18 | 5 | 290564 | 0.083671 | 14.506 | 14.1948
2 | 1 | ANSLu | 1 | 4 | 7 | 290565 | 0.08737 | 14.1874 | 14.1779
2 | 1 | ANSLu | 1 | 7 | 8 | 290566 | 0.082486 | 14.1257 | 14.1746
2 | 1 | ANSLu | 1 | 16 | 10 | 290567 | 0.08635 | 13.8588 | 14.1604
2 | 1 | ANSLu | 1 | 17 | 5 | 290568 | 0.083945 | 13.722 | 14.1531
2 | 1 | ANSLu | 1 | 1 | 1 | 290569 | 0.085469 | 14.3817 | 14.1882
2 | 1 | ANSLu | 1 | 5 | 2 | 290570 | 0.084928 | 14.0682 | 14.1715
2 | 1 | ANSLu | 1 | 11 | 9 | 290571 | 0.086482 | 14.3988 | 14.1891
2 | 1 | ANSLu | 1 | 12 | 9 | 290572 | 0.083781 | 14.2269 | 14.18
2 | 1 | ANSLu | 1 | 13 | 4 | 290573 | 0.085638 | 14.2541 | 14.1814
2 | 1 | ANSLu | 1 | 15 | 10 | 290574 | 0.084712 | 14.0428 | 14.1702
2 | 1 | ANSLu | 1 | 10 | 3 | 290575 | 0.084509 | 14.2911 | 14.1834
2 | 1 | ANSLu | 1 | 6 | 2 | 290576 | 0.082334 | 14.2979 | 14.1838
2 | 1 | ANSLu | 1 | 9 | 3 | 290577 | 0.084913 | 14.5274 | 14.196
2 | 1 | ANSLu | 1 | 14 | 4 | 290578 | 0.086334 | 13.9656 | 14.1661
2 | 1 | ANSLu | 1 | 20 | 11 | 290579 | 0.087314 | 13.8987 | 14.1625
2 | 1 | ANSLu | 1 | 3 | 7 | 290580 | 0.083884 | 14.2454 | 14.181
2 | 1 | ANSLu | 1 | 2 | 1 | 290581 | 0.087691 | 14.472 | 14.193
Mean | [6 empty] | 0.084893 | 14.1773 | 14.1773
Median | [7 empty] | 14.2071 | 14.1789
SD | [7 empty] | 0.22444 | 0.011939
RSD | [7 empty] | 1.583% | 0.08421%
RSD Upscale Factor | [7 empty] | 0.053196
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 267 (Certified Value 13.21 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.30 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*OES | FA*GRAV | FA*AAS | FA*OES | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 25g | 30g | 50g | 30g | 30g | 30g | 50g | 40g | 30g | 30g | 30g | 30g | 30g | 50g | 40g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 3
| 1 | 1 | 14.3817 | 13.739 | 13.435 | 13.015 | 12.663 | 12.8718 | 12.7 | 12.82 | 13.66 | 13.45 | 13.3 | 13.3 | 12.85 | 13.5665 | 13.0767 | 15.0562 | 13.3 | 13.3 | 13.301 | 13.4 | 12.58 | 12.7 | > 5 | 13.1 | 13.9 | 13.625 | 1
| 1 | 2 | 14.472 | 13.614 | 13.419 | 13.047 | 12.85 | 12.8572 | 13.3 | 11.99 | 13.68 | 13.4 | 13.55 | 13.35 | 13 | 13.674 | 12.13 | 13.9736 | 13 | 13.34 | 12.75 | 13.1 | 12.46 | 12.8 | > 5 | 11.6 | 13.2 | 13.9932 | #N/A
| 1 | 3 | 14.2454 | 13.572 | 13.412 | 13.137 | 12.762 | 13.0824 | 13 | 12.74 | 13.67 | 13.55 | 13.75 | 13.3 | 13 | 13.537 | 12.2833 | 14.7307 | 13.3 | 13.26 | 13.333 | 13.4 | 12.689 | 12.1 | > 5 | 13.7 | 13.1 | 13.6825 | 16
| 1 | 4 | 14.1874 | 13.526 | 13.418 | 13.242 | 12.258 | 13.0649 | 13.2 | 12.16 | 13.67 | 13.25 | 13.65 | 13.45 | 13.15 | 13.5675 | 12.9567 | 14.7897 | 13.3 | 13.21 | 12.99 | 13.4 | 12.652 | 12.9 | > 5 | 13.1 | 13.4 | 13.731 | 13.2103
| 1 | 5 | 14.0682 | 13.793 | 13.427 | 13.001 | 12.988 | 13.081 | 13.9 | 12.9 | 13.68 | 13.35 | 13.55 | 13.4 | 13.05 | 13.486 | 12.93 | 14.5698 | 13.3 | 13.39 | 13.701 | 13.3 | 12.658 | 12.2 | > 5 | 13.2 | 12.8 | 13.415 | 7
| 1 | 6 | 14.2979 | 13.854 | 13.425 | 13.208 | 12.931 | 13.0309 | 13.2 | 13.67 | 13.69 | 13.6 | 13.5 | 13.4 | 13 | 13.4725 | 13.07 | 14.6851 | 13.3 | 13.25 | 13.522 | 13.3 | 12.549 | 12.7 | > 5 | 14.8 | 13.4 | 13.5515
|  | 7 | 14.1257
|  | 8 | 13.9847
|  | 9 | 14.5274
|  | 10 | 14.2911
|  | 11 | 14.3988
|  | 12 | 14.2269
|  | 13 | 14.2541
|  | 14 | 13.9656
|  | 15 | 14.0428
|  | 16 | 13.8588
|  | 17 | 13.722
|  | 18 | 14.506
|  | 19 | 14.0918
|  | 20 | 13.8987
| Mean |  | 14.1773 | 13.683 | 13.4227 | 13.1083 | 12.742 | 12.998 | 13.2167 | 12.7133 | 13.675 | 13.4333 | 13.55 | 13.3667 | 13.0083 | 13.5506 | 12.7411 | 14.6342 | 13.25 | 13.2917 | 13.2662 | 13.3167 | 12.598 | 12.5667 |  | 13.25 | 13.3 | 13.6664
| Median |  | 14.2071 | 13.6765 | 13.422 | 13.092 | 12.806 | 13.0479 | 13.2 | 12.78 | 13.675 | 13.425 | 13.55 | 13.375 | 13 | 13.5518 | 12.9433 | 14.7079 | 13.3 | 13.28 | 13.317 | 13.35 | 12.616 | 12.7 |  | 13.15 | 13.3 | 13.6537
| Std Dev. |  | 0.22444 | 0.13131 | 0.0080664 | 0.10258 | 0.26412 | 0.10517 | 0.39707 | 0.59859 | 0.010488 | 0.1291 | 0.15166 | 0.060553 | 0.09704 | 0.072421 | 0.42093 | 0.36185 | 0.12247 | 0.065549 | 0.34725 | 0.1169 | 0.085517 | 0.33267 |  | 1.0368 | 0.36878 | 0.19461
| Rel.Std.Dev. |  | 1.583% | 0.9596% | 0.0601% | 0.7826% | 2.073% | 0.8092% | 3.004% | 4.708% | 0.0767% | 0.961% | 1.119% | 0.453% | 0.746% | 0.5344% | 3.304% | 2.473% | 0.9243% | 0.4932% | 2.618% | 0.8779% | 0.6788% | 2.647% |  | 7.825% | 2.773% | 1.424%
| PDM3 |  | 7.321% | 3.578% | 1.608% | -0.7717% | -3.545% | -1.607% | 0.04836% | -3.762% | 3.518% | 1.688% | 2.572% | 1.184% | -1.529% | 2.576% | -3.552% | 10.78% | 0.3007% | 0.6161% | 0.4231% | 0.8053% | -4.635% | -4.872% |  | 0.3007% | 0.6792% | 3.453%
| Z-Score (Absolute) |  | NA | 1.01 | 0.36 | 0.42 | 1.34 | 0.7 | 0.15 | 1.41 | 0.99 | 0.38 | 0.67 | 0.22 | 0.67 | 0.68 | 1.34 | 3.37 | 0.07 | 0.03 | 0.03 | 0.09 | 1.7 | 1.77 | NA | 0.07 | 0.05 | 0.96
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 267 (Certified Value 13.56 ppm) | [14 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | ppm
|  |  | INAA | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | 2
|  |  | 0.085g | [12 empty] | 3
| 1 | 1 | 14.3817 | 13.62 | 13.56 | 13.953 | 13.3 | 13.46 | 13.88 | 14.223 | 13.522 | 13.695 | 12.941 | 13.345 | 13.416 | 1
| 1 | 2 | 14.472 | 13.51 | 13.34 | 13.041 | 13.3 | 13.65 | 13.33 | 13.322 | 13.446 | 13.475 | 13.177 | 13.481 | 13.393 | #N/A
| 1 | 3 | 14.2454 | 13.89 | 13.18 | 13.764 | 13.5 | 13.54 | 13.32 | 14.335 | 13.507 | 13.717 | 13.551 | 13.298 | 13.437 | 16
| 1 | 4 | 14.1874 | 13.35 | 13.37 | 14.112 | 13.5 | 13.69 | 12.97 | 13.656 | 13.678 | 13.466 | 13.783 | 13.471 | 13.722 | 13.5614
| 1 | 5 | 14.0682 | 13.23 | 13.52 | 14.125 | 13.5 | 13.45 | 13.84 | 14.36 | 13.625 | 13.763 | 13.738 | 13.334 | 13.67 | 9
| 1 | 6 | 14.2979 | 13.9 | 13.4 | 13.496 | 13.3 | 13.6 | 13.22 | 13.568 | 13.631 | 13.898 | 13.746 | 13.521 | 13.796
|  | 7 | 14.1257
|  | 8 | 13.9847
|  | 9 | 14.5274
|  | 10 | 14.2911
|  | 11 | 14.3988
|  | 12 | 14.2269
|  | 13 | 14.2541
|  | 14 | 13.9656
|  | 15 | 14.0428
|  | 16 | 13.8588
|  | 17 | 13.722
|  | 18 | 14.506
|  | 19 | 14.0918
|  | 20 | 13.8987
| Mean |  | 14.1773 | 13.5833 | 13.395 | 13.7485 | 13.4 | 13.565 | 13.4267 | 13.9107 | 13.5682 | 13.669 | 13.4893 | 13.4083 | 13.5723
| Median |  | 14.2071 | 13.565 | 13.385 | 13.8585 | 13.4 | 13.57 | 13.325 | 13.9395 | 13.5735 | 13.706 | 13.6445 | 13.408 | 13.5535
| Std Dev. |  | 0.22444 | 0.2758 | 0.1362 | 0.41958 | 0.10954 | 0.098944 | 0.36009 | 0.44907 | 0.089486 | 0.16918 | 0.35099 | 0.093393 | 0.17713
| Rel.Std.Dev. |  | 1.583% | 2.03% | 1.017% | 3.052% | 0.8175% | 0.7294% | 2.682% | 3.228% | 0.6595% | 1.238% | 2.602% | 0.6965% | 1.305%
| PDM3 |  | 4.542% | 0.162% | -1.227% | 1.38% | -1.19% | 0.02683% | -0.9932% | 2.576% | 0.05018% | 0.7937% | -0.5311% | -1.128% | 0.08091%
| Z-Score (Absolute) |  | NA | 0.09 | 0.95 | 1.01 | 0.93 | 0.01 | 0.78 | 1.91 | 0.01 | 0.57 | 0.43 | 0.88 | 0.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 267 (Certified Value 13.42 ppm) | [20 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.30 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | 2
|  |  | 0.085g | 25g | 30g | 30g | 50g | 30g | 25g | 25g | 25g | 25g | 25g | 30g | 50g | 25g | 30g | 25g | 10g | 15g | 30g | 3
| 1 | 1 | 14.3817 | 13.24 | 13.371 | 13.1 | 13.91 | 13.54 | 13.15 | 14.55 | 13.65 | 12.3 | 13.4 | 13.415 | 12.7775 | 13.2552 | 12.8 | > 0.5 | 10.24 | > 10 | 13.4875 | 1
| 1 | 2 | 14.472 | 13.31 | 13.352 | 13.5 | 14.26 | 13.59 | 13.25 | 14.8 | 13.75 | 12.75 | 13.25 | 13.483 | 13.045 | 13.6136 | 12.9 | > 0.5 | 12.43 | > 10 | 13.2823 | #N/A
| 1 | 3 | 14.2454 | 13.25 | 13.335 | 13.6 | 14.62 | 13.61 | 13.2 | 13.9 | 13.5 | 12.5 | 14.2 | 13.545 | 12.6337 | 13.8768 | 12.8 | > 0.5 | 11.68 | > 10 | 13.3997 | 16
| 1 | 4 | 14.1874 | 13.27 | 13.377 | 13.7 | 13.61 | 13.59 | 13.3 | 13.9 | 13.55 | 12.6 | 13.85 | 13.389 | 12.54 | 14.1736 | 12.7 | > 0.5 | 10.59 | > 10 | 13.2845 | 13.4171
| 1 | 5 | 14.0682 | 13.22 | 13.388 | 13.5 | 14.2 | 13.55 | 13.1 | 14 | 13.85 | 12.6 | 14.6 | 13.425 | 12.585 | 13.5296 | 12.8 | > 0.5 | 11.03 | > 10 | 13.2036 | 11
| 1 | 6 | 14.2979 | 13.01 | 13.363 | 13.4 | 13.5 | 13.65 | 13.5 | 14.65 | 13.85 | 12.2 | 13.7 | 13.658 | 12.5738 | 13.8488 | 12.8 | > 0.5 | 10.29 | > 10 | 13.4015
|  | 7 | 14.1257
|  | 8 | 13.9847
|  | 9 | 14.5274
|  | 10 | 14.2911
|  | 11 | 14.3988
|  | 12 | 14.2269
|  | 13 | 14.2541
|  | 14 | 13.9656
|  | 15 | 14.0428
|  | 16 | 13.8588
|  | 17 | 13.722
|  | 18 | 14.506
|  | 19 | 14.0918
|  | 20 | 13.8987
| Mean |  | 14.1773 | 13.2167 | 13.3643 | 13.4667 | 14.0167 | 13.5883 | 13.25 | 14.3 | 13.6917 | 12.4917 | 13.8333 | 13.4858 | 12.6925 | 13.7163 | 12.8 |  | 11.0433 |  | 13.3432
| Median |  | 14.2071 | 13.245 | 13.367 | 13.5 | 14.055 | 13.59 | 13.225 | 14.275 | 13.7 | 12.55 | 13.775 | 13.454 | 12.6094 | 13.7312 | 12.8 |  | 10.81 |  | 13.3421
| Std Dev. |  | 0.22444 | 0.10577 | 0.018864 | 0.20656 | 0.42439 | 0.040208 | 0.14142 | 0.4111 | 0.14972 | 0.20595 | 0.50365 | 0.10124 | 0.19175 | 0.31952 | 0.063246 |  | 0.86565 |  | 0.10397
| Rel.Std.Dev. |  | 1.583% | 0.8003% | 0.1412% | 1.534% | 3.028% | 0.2959% | 1.067% | 2.875% | 1.094% | 1.649% | 3.641% | 0.7507% | 1.511% | 2.329% | 0.4941% |  | 7.839% |  | 0.7792%
| PDM3 |  | 5.666% | -1.494% | -0.3936% | 0.3691% | 4.468% | 1.276% | -1.246% | 6.58% | 2.046% | -6.898% | 3.102% | 0.512% | -5.401% | 2.229% | -4.6% |  | -17.69% |  | -0.5511%
| Z-Score (Absolute) |  | NA | 0.46 | 0.12 | 0.12 | 1.41 | 0.4 | 0.39 | 2.07 | 0.65 | 2.16 | 0.98 | 0.16 | 1.69 | 0.7 | 1.44 | NA | 5.55 | NA | 0.17
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 267 (Certified Value 12.96 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*MS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 200g | 50g | 200g | 50g | 30g | 30g | 30g | 100g | 50g | 10g | 05g | 05g | 15g | 50g | 15g | 100g | 3
| 1 | 1 | 14.3817 | 12.81 | 12.7782 | 12.8144 | 11.687 | 13.38 | 14.2795 | 11.8 | 12.17 | 13.2 | 14.1 | 12.9 | 12.9 | 13.076 | 12.7688 | 14.368 | 13.1 | 13.522 | 13.05 | 6.497 | 12 | 13.839 | 1
| 1 | 2 | 14.472 | 13.34 | 12.8708 | 12.836 | 12.976 | 13.39 | 13.8077 | 11.8 | 11.75 | 13.25 | 14.05 | 13.1 | 12.8 | 13.16 | 12.578 | 14.972 | 13.1 | 13.556 | 12.87 | 6.472 | 11.44 | 13.061 | #N/A
| 1 | 3 | 14.2454 | 13.29 | 12.8155 | 12.824 | 12.716 | 13.33 | 14.0675 | 12 | 12.31 | 13.22 | 14 | 13 | 13 | 12.982 | 12.4788 | 14.664 | 13.2 | 13.493 | 12.87 | 6.518 | 11.45 | 13.647 | 16
| 1 | 4 | 14.1874 | 13.14 | 12.7493 | 12.8128 | 12.372 | 13.68 | 13.8557 | 12.4 | 12.07 | 13.28 | 14.2 | 13.1 | 13 | 12.892 | 12.91 | 14.28 | 13.2 | 13.475 | 13.14 | 6.493 | 11.84 | 13.433 | 12.9572
| 1 | 5 | 14.0682 | 13.15 | 12.6732 | 12.8216 | 13.138 | 13.38 | 13.8996 | 12.1 | 11.92 | 13.2 | 14 | 13.05 | 13.15 | 13.032 | 12.7696 | 14.724 | 13.2 | 13.433 | 12.96 | 6.482 | 11.36 | 13.187 | 13
| 1 | 6 | 14.2979 | 13.06 | 12.8342 | 12.8016 | 12.267 | 13.23 | 14.1189 | 11.9 | 11.95 | 13.17 | 13.9 | 13.15 | 12.9 | 13.005 | 12.0512 | 14.256 | 13.2 | 13.538 | 12.6 | 6.272 | 11.52 | 13.346
|  | 7 | 14.1257
|  | 8 | 13.9847
|  | 9 | 14.5274
|  | 10 | 14.2911
|  | 11 | 14.3988
|  | 12 | 14.2269
|  | 13 | 14.2541
|  | 14 | 13.9656
|  | 15 | 14.0428
|  | 16 | 13.8588
|  | 17 | 13.722
|  | 18 | 14.506
|  | 19 | 14.0918
|  | 20 | 13.8987
| Mean |  | 14.1773 | 13.1317 | 12.7869 | 12.8184 | 12.526 | 13.3983 | 14.0048 | 12 | 12.0283 | 13.22 | 14.0417 | 13.05 | 12.9583 | 13.0245 | 12.5927 | 14.544 | 13.1667 | 13.5028 | 12.915 | 6.4557 | 11.6017 | 13.4188
| Median |  | 14.2071 | 13.145 | 12.7968 | 12.818 | 12.544 | 13.38 | 13.9835 | 11.95 | 12.01 | 13.21 | 14.025 | 13.075 | 12.95 | 13.0185 | 12.6734 | 14.516 | 13.2 | 13.5075 | 12.915 | 6.4875 | 11.485 | 13.3895
| Std Dev. |  | 0.22444 | 0.18841 | 0.069991 | 0.01167 | 0.53063 | 0.15039 | 0.1814 | 0.22804 | 0.19823 | 0.039497 | 0.10206 | 0.089443 | 0.12007 | 0.090345 | 0.30639 | 0.28761 | 0.05164 | 0.045119 | 0.18663 | 0.091303 | 0.25678 | 0.28841
| Rel.Std.Dev. |  | 1.583% | 1.435% | 0.5474% | 0.09104% | 4.236% | 1.122% | 1.295% | 1.9% | 1.648% | 0.2988% | 0.7269% | 0.6854% | 0.9266% | 0.6937% | 2.433% | 1.978% | 0.3922% | 0.3341% | 1.445% | 1.414% | 2.213% | 2.149%
| PDM3 |  | 9.417% | 1.347% | -1.314% | -1.071% | -3.328% | 3.405% | 8.085% | -7.387% | -7.169% | 2.028% | 8.37% | 0.7163% | 0.008807% | 0.5195% | -2.813% | 12.25% | 1.617% | 4.211% | -0.3256% | -50.18% | -10.46% | 3.563%
| Z-Score (Absolute) |  | NA | 0.18 | 0.41 | 0.35 | 0.85 | 0.64 | 1.68 | 1.75 | 1.7 | 0.33 | 1.74 | 0.04 | 0.11 | 0 | 0.74 | 2.6 | 0.24 | 0.82 | 0.19 | 11.23 | 2.43 | 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 267 (Certified Value 5.54 ppm) | [18 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 5.72 | 5.17 | 5.3195 | 5.7 | 5.51 | 5.81 | 5.48 | 5.3 | 5.89 | 5.59 | 0.96579 | 5.4 | 5.17 | 5.73 | 5.6 | 6.61 | 5.4 | 1
| 1 | 2 | 5.77 | 5.29 | 5.2369 | 5.7 | 5.41 | 5.79 | 5.98 | 5.41 | 5.66 | 5.58 | 0.95917 | 5.4 | 5.31 | 5.61 | 5.6 | 6.2 | 5.4 | 20
| 1 | 3 | 5.63 | 5.37 | 5.5662 | 5.7 | 5.41 | 5.5 | 5.71 | 5.43 | 6.02 | 5.34 | 0.95732 | 5.4 | 5.24 | 5.74 | 5.7 | 6.27 | 5.3 | 16
| 1 | 4 | 5.71 | 5.29 | 5.3516 | 5.7 | 5.47 | 5.78 | 5.71 | 5.44 | 5.97 | 5.57 | 0.9182 | 5.4 | 5.39 | 5.8 | 5.6 | 6.78 | 5.4 | 5.5404
| 1 | 5 | 5.76 | 5.34 | 5.4228 | 5.5 | 5.18 | 5.77 | 5.64 | 5.45 | 6.02 | 5.41 | 0.91418 | 5.4 | 5.35 | 5.61 | 5.6 | 6.3 | 5.3 | 15
| 1 | 6 | 5.73 | 5.15 | 5.3803 | 5.8 | 5.47 | 5.78 | 5.7 | 5.47 | 6.05 | 5.54 | 0.9963 | 5.3 | 5.19 | 5.75 | 5.6 | 6.3 | 5.4
| Mean |  | 5.72 | 5.2683 | 5.3796 | 5.6833 | 5.4083 | 5.7383 | 5.7033 | 5.4167 | 5.935 | 5.505 | 0.95183 | 5.3833 | 5.275 | 5.7067 | 5.6167 | 6.41 | 5.3667
| Median |  | 5.725 | 5.29 | 5.366 | 5.7 | 5.44 | 5.78 | 5.705 | 5.435 | 5.995 | 5.555 | 0.95824 | 5.4 | 5.275 | 5.735 | 5.6 | 6.3 | 5.4
| Std Dev. |  | 0.0498 | 0.089536 | 0.11085 | 0.098319 | 0.11839 | 0.11754 | 0.16158 | 0.060553 | 0.14598 | 0.10445 | 0.031 | 0.040825 | 0.088938 | 0.078655 | 0.040825 | 0.23013 | 0.05164
| Rel.Std.Dev. |  | 0.8706% | 1.7% | 2.061% | 1.73% | 2.189% | 2.048% | 2.833% | 1.118% | 2.46% | 1.897% | 3.257% | 0.7584% | 1.686% | 1.378% | 0.7269% | 3.59% | 0.9622%
| PDM3 |  | 3.241% | -4.911% | -2.903% | 2.58% | -2.384% | 3.572% | 2.941% | -2.234% | 7.122% | -0.6392% | -82.82% | -2.835% | -4.791% | 3.001% | 1.376% | 15.7% | -3.136%
| Z-Score (Absolute) |  | 0.64 | 1.17 | 0.73 | 0.49 | 0.61 | 0.71 | 0.57 | 0.58 | 1.5 | 0.22 | NA | 0.71 | 1.14 | 0.58 | 0.22 | 3.4 | 0.78
| Analytical results for Al in OREAS 267 (Certified Value 5.05 wt.%) | [18 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 5.2055 | 5.2011 | 5.2502 | 4.72 | 5.15 | 5.23 | 5.15 | 5.08 | 5.02 | 5.11 | 6.2312 | 4.99 | 4.72 | 5.0072 | 4.8712 | 4.83 | 4.96 | 1
| 1 | 2 | 5.1828 | 5.2134 | 5.1863 | 4.5 | 5.23 | 5.24 | 5.44 | 5.24 | 5.14 | 5.08 | 6.2054 | 4.92 | 4.87 | 5.0392 | 4.8274 | 4.92 | 5 | #N/A
| 1 | 3 | 5.1907 | 5.2232 | 5.5395 | 4.84 | 5.19 | 5.11 | 5.31 | 5.16 | 5.09 | 5.16 | 5.8661 | 4.92 | 4.68 | 5.0254 | 4.9129 | 4.96 | 5.03 | 16
| 1 | 4 | 5.1902 | 5.2181 | 5.1799 | 4.52 | 5.08 | 5.21 | 5.22 | 5.2 | 5.05 | 5.02 | 5.6193 | 4.87 | 4.82 | 5.0411 | 4.8401 | 4.94 | 4.98 | 5.046
| 1 | 5 | 5.2962 | 5.206 | 5.2425 | 4.57 | 5.15 | 5.21 | 5.25 | 5.14 | 5.1 | 5.07 | 6.1751 | 4.96 | 4.8 | 5.107 | 4.8953 | 4.83 | 4.98 | 16
| 1 | 6 | 5.2585 | 5.2398 | 5.2101 | 4.67 | 5.15 | 5.19 | 5.23 | 5.16 | 5.14 | 5.08 | 5.8372 | 4.89 | 4.78 | 4.9802 | 4.8696 | 4.87 | 4.97
| Mean |  | 5.2207 | 5.217 | 5.2681 | 4.6367 | 5.1583 | 5.1983 | 5.2667 | 5.1633 | 5.09 | 5.0867 | 5.989 | 4.925 | 4.7783 | 5.0333 | 4.8694 | 4.8917 | 4.9867
| Median |  | 5.1981 | 5.2158 | 5.2263 | 4.62 | 5.15 | 5.21 | 5.24 | 5.16 | 5.095 | 5.08 | 6.0206 | 4.92 | 4.79 | 5.0323 | 4.8704 | 4.895 | 4.98
| Std Dev. |  | 0.046102 | 0.013767 | 0.13597 | 0.13125 | 0.049967 | 0.046655 | 0.099331 | 0.054283 | 0.048166 | 0.046332 | 0.25095 | 0.044159 | 0.068823 | 0.042639 | 0.032214 | 0.056362 | 0.025033
| Rel.Std.Dev. |  | 0.8831% | 0.2639% | 2.581% | 2.831% | 0.9687% | 0.8975% | 1.886% | 1.051% | 0.9463% | 0.9109% | 4.19% | 0.8966% | 1.44% | 0.8471% | 0.6616% | 1.152% | 0.502%
| PDM3 |  | 3.461% | 3.388% | 4.401% | -8.112% | 2.226% | 3.019% | 4.373% | 2.325% | 0.8722% | 0.8061% | 18.69% | -2.398% | -5.304% | -0.2505% | -3.499% | -3.058% | -1.176%
| Z-Score (Absolute) |  | 0.67 | 0.66 | 0.92 | 2.34 | 0.35 | 0.56 | 0.91 | 0.38 | 0 | 0.02 | 4.64 | 0.85 | 1.61 | 0.29 | 1.14 | 1.02 | 0.53
| Analytical results for As in OREAS 267 (Certified Value 137 ppm) | [18 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 0
|  |  |  | 0.25g | [15 empty] | 0
| 1 | 1 | 135.5 | 114.709 | 134.0691 | 143 | 113 | 131.5 | 138.5 | 127.5 | 134.5 | 138.5 | 113.1457 | 133 | 131 | 137 | 149.5 | 107.2 | 140.5 | 1
| 1 | 2 | 133.8 | 116.722 | 135.4143 | 140 | 115 | 135 | 148.5 | 130 | 133 | 137 | 109.3102 | 133 | 131 | 136 | 147.3 | 113.3 | 139.1 | 21
| 1 | 3 | 135.3 | 117.723 | 144.9971 | 139 | 111 | 129 | 145 | 131 | 137 | 139 | 114.7183 | 133 | 130 | 136 | 152.1 | 116.1 | 138.9 | 16
| 1 | 4 | 135.4 | 114.302 | 133.9965 | 139 | 113 | 133 | 141.5 | 129 | 135 | 137.5 | 113.1599 | 133 | 133 | 139 | 146.4 | 109.2 | 135.7 | 136.9254
| 1 | 5 | 138.6 | 113.356 | 136.3602 | 136 | 110 | 134 | 142 | 130.5 | 137.5 | 137 | 118.035 | 132 | 131 | 138 | 152.4 | 114.3 | 137.3 | 17
| 1 | 6 | 136.1 | 117.591 | 137.8443 | 140 | 112 | 134 | 141 | 130.5 | 138 | 140 | 119.0794 | 132 | 129 | 135 | 146.9 | 113.3 | 142.2
| Mean |  | 135.7833 | 115.7338 | 137.1136 | 139.5 | 112.3333 | 132.75 | 142.75 | 129.75 | 135.8333 | 138.1667 | 114.5748 | 132.6667 | 130.8333 | 136.8333 | 149.1 | 112.2333 | 138.95
| Median |  | 135.45 | 115.7155 | 135.8873 | 139.5 | 112.5 | 133.5 | 141.75 | 130.25 | 136 | 138 | 113.9391 | 133 | 131 | 136.5 | 148.4 | 113.3 | 139
| Std Dev. |  | 1.5766 | 1.8513 | 4.1261 | 2.2583 | 1.7512 | 2.1852 | 3.5036 | 1.2942 | 1.9664 | 1.2111 | 3.579 | 0.5164 | 1.3292 | 1.472 | 2.6616 | 3.3476 | 2.2924
| Rel.Std.Dev. |  | 1.161% | 1.6% | 3.009% | 1.619% | 1.559% | 1.646% | 2.454% | 0.9975% | 1.448% | 0.8765% | 3.124% | 0.3892% | 1.016% | 1.076% | 1.785% | 2.983% | 1.65%
| PDM3 |  | -0.8341% | -15.48% | 0.1374% | 1.88% | -17.96% | -3.049% | 4.254% | -5.24% | -0.7976% | 0.9065% | -16.32% | -3.11% | -4.449% | -0.06724% | 8.891% | -18.03% | 1.479%
| Z-Score (Absolute) |  | 0 | 3.64 | 0.24 | 0.67 | 4.25 | 0.55 | 1.26 | 1.09 | 0.01 | 0.43 | 3.85 | 0.57 | 0.9 | 0.19 | 2.42 | 4.27 | 0.57
| Analytical results for Ba in OREAS 267 (Certified Value 381 ppm) | [18 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 381.1 | 351.8 | 371.1761 | 388 | 397 | 370 | 380 | 380 | 380 | 370 | 316.5029 | 370 | 385 | 384.4 | 345.4 | 394 |  | 1
| 1 | 2 | 376.8 | 352.7 | 374.6763 | 387 | 401 | 400 | 410 | 380 | 390 | 370 | 312.2936 | 368 | 388 | 379 | 349.7 | 395 |  | 22
| 1 | 3 | 380.6 | 353.73 | 392.3885 | 393 | 390 | 370 | 400 | 380 | 380 | 380 | 323.3036 | 371 | 391 | 386.3 | 350.1 | 396 |  | 16
| 1 | 4 | 376.5 | 354.87 | 376.6572 | 381 | 394 | 380 | 390 | 380 | 380 | 360 | 321.4565 | 374 | 386 | 387.6 | 347.3 | 395 |  | 381.1845
| 1 | 5 | 383.6 | 352.57 | 379.5803 | 376 | 402 | 380 | 390 | 380 | 380 | 370 | 337.2964 | 364 | 394 | 383.4 | 345.8 | 396 |  | 18
| 1 | 6 | 376.8 | 353.84 | 379.1109 | 383 | 396 | 380 | 390 | 380 | 390 | 370 | 354.0599 | 376 | 381 | 382.3 | 348.9 | 396
| Mean |  | 379.2333 | 353.2517 | 378.9316 | 384.6667 | 396.6667 | 380 | 393.3333 | 380 | 383.3333 | 370 | 327.4855 | 370.5 | 387.5 | 383.8333 | 347.8667 | 395.3333
| Median |  | 378.7 | 353.215 | 377.8841 | 385 | 396.5 | 380 | 390 | 380 | 380 | 370 | 322.38 | 370.5 | 387 | 383.9 | 348.1 | 395.5
| Std Dev. |  | 2.9575 | 1.1018 | 7.2797 | 5.9554 | 4.4572 | 10.9545 | 10.328 | 0 | 5.164 | 6.3246 | 15.5415 | 4.2778 | 4.5935 | 3.0494 | 2.0047 | 0.8165
| Rel.Std.Dev. |  | 0.7799% | 0.3119% | 1.921% | 1.548% | 1.124% | 2.883% | 2.626% | 0% | 1.347% | 1.709% | 4.746% | 1.155% | 1.185% | 0.7945% | 0.5763% | 0.2065%
| PDM3 |  | -0.5119% | -7.328% | -0.591% | 0.9135% | 4.062% | -0.3107% | 3.187% | -0.3107% | 0.5637% | -2.934% | -14.09% | -2.803% | 1.657% | 0.6949% | -8.741% | 3.712%
| Z-Score (Absolute) |  | 0.06 | 2.12 | 0.08 | 0.37 | 1.32 | 0 | 1.06 | 0 | 0.26 | 0.79 | 4.17 | 0.75 | 0.59 | 0.3 | 2.55 | 1.22
| Analytical results for Be in OREAS 267 (Certified Value 1.8 ppm) | [18 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.84 | 1.2414 | 1.7946 | 2 | 1.93 | 1.75 | 1.78 | 1.77 | 1.81 | 1.78 | 1.4414 | 1.8 | 1.7 | 1.91 | 1.62 | 2 |  | 1
| 1 | 2 | 1.74 | 1.2543 | 1.8478 | 2 | 1.94 | 1.76 | 1.89 | 1.8 | 1.84 | 1.79 | 1.445 | 1.8 | 1.8 | 1.89 | 1.68 | 2 |  | 23
| 1 | 3 | 1.72 | 1.673 | 1.8735 | 2.1 | 1.98 | 1.74 | 1.84 | 1.8 | 1.82 | 1.8 | 1.5113 | 1.8 | 1.7 | 1.95 | 1.65 | 2 |  | 16
| 1 | 4 | 1.78 | 1.6937 | 1.8378 | 2.1 | 1.89 | 1.74 | 1.8 | 1.8 | 1.83 | 1.75 | 1.3718 | 1.8 | 1.9 | 1.8 | 1.7 | 2 |  | 1.8045
| 1 | 5 | 1.78 | 1.6003 | 1.8214 | 1.9 | 1.93 | 1.76 | 1.82 | 1.78 | 1.83 | 1.78 | 1.4957 | 1.8 | 1.8 | 1.85 | 1.71 | 2 |  | 19
| 1 | 6 | 1.79 | 1.5823 | 1.8375 | 2 | 1.99 | 1.79 | 1.81 | 1.78 | 1.86 | 1.76 | 1.451 | 1.8 | 1.7 | 1.85 | 1.73 | 2
| Mean |  | 1.775 | 1.5075 | 1.8354 | 2.0167 | 1.9433 | 1.7567 | 1.8233 | 1.7883 | 1.8317 | 1.7767 | 1.4527 | 1.8 | 1.7667 | 1.875 | 1.6817 | 2
| Median |  | 1.78 | 1.5913 | 1.8376 | 2 | 1.935 | 1.755 | 1.815 | 1.79 | 1.83 | 1.78 | 1.448 | 1.8 | 1.75 | 1.87 | 1.69 | 2
| Std Dev. |  | 0.041833 | 0.20554 | 0.026389 | 0.075277 | 0.036697 | 0.018619 | 0.038297 | 0.013292 | 0.017224 | 0.018619 | 0.049048 | 0 | 0.08165 | 0.05282 | 0.040702 | 0
| Rel.Std.Dev. |  | 2.357% | 13.63% | 1.438% | 3.733% | 1.888% | 1.06% | 2.1% | 0.7432% | 0.9403% | 1.048% | 3.376% | 0% | 4.622% | 2.817% | 2.42% | 0%
| PDM3 |  | -1.634% | -16.46% | 1.715% | 11.76% | 7.695% | -2.65% | 1.045% | -0.8948% | 1.507% | -1.541% | -19.5% | -0.2483% | -2.096% | 3.908% | -6.806% | 10.84%
| Z-Score (Absolute) |  | 0.21 | 4.37 | 0.73 | 3.55 | 2.41 | 0.49 | 0.54 | 0 | 0.67 | 0.18 | 5.22 | 0.18 | 0.34 | 1.35 | 1.66 | NA
| Analytical results for Bi in OREAS 267 (Certified Value 1.28 ppm) | [18 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 1.42 | 0.87158 | 1.3152 | 1.32 | 2.57 | 1.22 | 1.16 | 1.24 | 1.33 | 1.28 | 1.0545 | 1.23 | 1.24 | 1.33 | 1.22 | 1.17 | 1.5 | 1
| 1 | 2 | 1.34 | 0.87948 | 1.2629 | 1.38 | 2.73 | 1.22 | 1.31 | 1.3 | 1.41 | 1.3 | 1.0073 | 1.22 | 1.37 | 1.43 | 0.99 | 1.15 | 1.4 | 24
| 1 | 3 | 1.39 | 0.85991 | 1.3928 | 1.34 | 2.43 | 1.17 | 1.45 | 1.27 | 1.35 | 1.21 | 1.0045 | 1.21 | 1.34 | 1.46 | 1.11 | 1.22 | 1.4 | 16
| 1 | 4 | 1.34 | 0.80929 | 1.353 | 1.36 | 2.62 | 1.19 | 1.2 | 1.42 | 1.4 | 1.17 | 1.043 | 1.21 | 1.35 | 1.62 | 1.26 | 1.1 | 1.4 | 1.281
| 1 | 5 | 1.33 | 0.85522 | 1.3283 | 1.24 | 2.41 | 1.28 | 1.21 | 1.34 | 1.39 | 1.26 | 1.0639 | 1.23 | 1.38 | 1.67 | 1.11 | 1.17 | 1.3 | 20
| 1 | 6 | 1.37 | 0.86929 | 1.4236 | 1.36 | 2.48 | 1.24 | 1.18 | 1.26 | 1.5 | 1.28 | 1.0315 | 1.21 | 1.35 | 1.4 | 1.21 | 1.14 | 1.4
| Mean |  | 1.365 | 0.85746 | 1.346 | 1.3333 | 2.54 | 1.22 | 1.2517 | 1.305 | 1.3967 | 1.25 | 1.0341 | 1.2183 | 1.3383 | 1.485 | 1.15 | 1.1583 | 1.4
| Median |  | 1.355 | 0.8646 | 1.3406 | 1.35 | 2.525 | 1.22 | 1.205 | 1.285 | 1.395 | 1.27 | 1.0372 | 1.215 | 1.35 | 1.445 | 1.16 | 1.16 | 1.4
| Std Dev. |  | 0.035071 | 0.025121 | 0.057304 | 0.050067 | 0.12329 | 0.038471 | 0.11017 | 0.066257 | 0.059217 | 0.0498 | 0.024424 | 0.0098319 | 0.050365 | 0.13217 | 0.099398 | 0.039707 | 0.063246
| Rel.Std.Dev. |  | 2.569% | 2.93% | 4.258% | 3.755% | 4.854% | 3.153% | 8.802% | 5.077% | 4.24% | 3.984% | 2.362% | 0.807% | 3.763% | 8.901% | 8.643% | 3.428% | 4.518%
| PDM3 |  | 6.559% | -33.06% | 5.072% | 4.087% | 98.29% | -4.761% | -2.289% | 1.875% | 9.031% | -2.419% | -19.27% | -4.891% | 4.477% | 15.93% | -10.23% | -9.575% | 9.291%
| Z-Score (Absolute) |  | 0.47 | 3.48 | 0.32 | 0.22 | 9.61 | 0.66 | 0.41 | 0 | 0.71 | 0.43 | 2.11 | 0.67 | 0.26 | 1.4 | 1.21 | 1.14 | 0.74
| Analytical results for Ca in OREAS 267 (Certified Value 1.13 wt.%) | [18 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.1369 | 1.1015 | 1.1281 | 1.17 | 1.12 | 1.13 | 1.19 | 1.1 | 1.13 | 1.13 | 1.1581 | 1.11 | 1.05 | 1.1317 | 1.1386 | 1.02 | 1.17 | 1
| 1 | 2 | 1.1269 | 1.1062 | 1.1246 | 1.14 | 1.14 | 1.16 | 1.26 | 1.12 | 1.16 | 1.14 | 1.1639 | 1.11 | 1.09 | 1.1188 | 1.1264 | 1.03 | 1.18 | #N/A
| 1 | 3 | 1.1291 | 1.1066 | 1.197 | 1.16 | 1.14 | 1.12 | 1.22 | 1.09 | 1.14 | 1.14 | 1.1688 | 1.1 | 1.08 | 1.1395 | 1.1481 | 1.05 | 1.17 | 16
| 1 | 4 | 1.1348 | 1.11 | 1.1301 | 1.15 | 1.11 | 1.13 | 1.2 | 1.1 | 1.14 | 1.12 | 1.248 | 1.09 | 1.06 | 1.1486 | 1.1315 | 1.04 | 1.18 | 1.1324
| 1 | 5 | 1.1619 | 1.1079 | 1.1401 | 1.12 | 1.13 | 1.15 | 1.22 | 1.1 | 1.15 | 1.13 | 1.2269 | 1.11 | 1.09 | 1.1395 | 1.1336 | 1.01 | 1.17 | 21
| 1 | 6 | 1.141 | 1.1096 | 1.1367 | 1.17 | 1.13 | 1.15 | 1.21 | 1.1 | 1.17 | 1.13 | 1.2271 | 1.1 | 1.07 | 1.1123 | 1.1268 | 1.02 | 1.18
| Mean |  | 1.1384 | 1.107 | 1.1428 | 1.1517 | 1.1283 | 1.14 | 1.2167 | 1.1017 | 1.1483 | 1.1317 | 1.1988 | 1.1033 | 1.0733 | 1.1317 | 1.1342 | 1.0283 | 1.175
| Median |  | 1.1359 | 1.1073 | 1.1334 | 1.155 | 1.13 | 1.14 | 1.215 | 1.1 | 1.145 | 1.13 | 1.1979 | 1.105 | 1.075 | 1.1356 | 1.1326 | 1.025 | 1.175
| Std Dev. |  | 0.012593 | 0.0030821 | 0.027144 | 0.019408 | 0.01169 | 0.015492 | 0.024221 | 0.0098319 | 0.01472 | 0.0075277 | 0.039461 | 0.008165 | 0.01633 | 0.013784 | 0.0081953 | 0.01472 | 0.0054772
| Rel.Std.Dev. |  | 1.106% | 0.2784% | 2.375% | 1.685% | 1.036% | 1.359% | 1.991% | 0.8925% | 1.282% | 0.6652% | 3.292% | 0.74% | 1.521% | 1.218% | 0.7226% | 1.431% | 0.4661%
| PDM3 |  | 0.5355% | -2.242% | 0.9188% | 1.704% | -0.3565% | 0.6738% | 7.444% | -2.711% | 1.41% | -0.0621% | 5.867% | -2.564% | -5.214% | -0.05622% | 0.1587% | -9.188% | 3.765%
| Z-Score (Absolute) |  | 0.16 | 1.05 | 0.33 | 0.67 | 0.22 | 0.22 | 3.18 | 1.25 | 0.55 | 0.1 | 2.49 | 1.19 | 2.34 | 0.09 | 0 | 4.08 | 1.57
| Analytical results for Cd in OREAS 267 (Indicative Value 0.033 ppm) | [18 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.04 | 1.14 | < 0.2 | < 0.3 | 0.08 | < 0.02 | 0.02 | < 0.02 | 0.03 | 0.03 | 0.69197 | < 0.05 | 0.04 | 0.03 | < 0.02 | < 0.2 | 0.1 | 1
| 1 | 2 | 0.03 | 1.13 | < 0.2 | < 0.3 | 0.07 | 0.04 | 0.02 | 0.02 | 0.03 | 0.02 | 0.71461 | < 0.05 | 0.05 | 0.03 | 0.25 | < 0.2 | 0.1 | 1
| 1 | 3 | 0.05 | 1.07 | < 0.2 | < 0.3 | 0.07 | 0.02 | 0.03 | 0.02 | 0.04 | 0.04 | 0.71204 | < 0.05 | 0.06 | 0.04 | 0.13 | < 0.2 | 0.1 | 16
| 1 | 4 | 0.03 | 1.1 | < 0.2 | < 0.3 | 0.08 | < 0.02 | 0.02 | 0.03 | 0.03 | 0.02 | 0.67499 | < 0.05 | 0.04 | 0.03 | < 0.02 | < 0.2 | 0.2 | 0.032833
| 1 | 5 | 0.04 | 1.12 | < 0.2 | < 0.3 | 0.08 | < 0.02 | 0.04 | 0.03 | 0.04 | 0.03 | 0.6894 | < 0.05 | 0.06 | 0.03 | < 0.02 | < 0.2 | < 0.1 | 7
| 1 | 6 | 0.04 | 1.15 | < 0.2 | < 0.3 | 0.07 | 0.02 | 0.03 | < 0.02 | < 0.02 | 0.03 | 0.73263 | < 0.05 | 0.06 | 0.06 | < 0.02 | < 0.2 | < 0.1
| Mean |  | 0.038333 | 1.1183 |  |  | 0.075 | 0.026667 | 0.026667 | 0.025 | 0.034 | 0.028333 | 0.70261 |  | 0.051667 | 0.036667 | 0.19 |  | 0.125
| Median |  | 0.04 | 1.125 |  |  | 0.075 | 0.02 | 0.025 | 0.025 | 0.03 | 0.03 | 0.70201 |  | 0.055 | 0.03 | 0.19 |  | 0.1
| Std Dev. |  | 0.0075277 | 0.029269 |  |  | 0.0054772 | 0.011547 | 0.008165 | 0.0057735 | 0.0054772 | 0.0075277 | 0.020902 |  | 0.0098319 | 0.012111 | 0.084853 |  | 0.05
| Rel.Std.Dev. |  | 19.64% | 2.617% |  |  | 7.303% | 43.3% | 30.62% | 23.09% | 16.11% | 26.57% | 2.975% |  | 19.03% | 33.03% | 44.66% |  | 40%
| PDM3 |  | 16.75% | 3306% |  |  | 128.4% | -18.78% | -18.78% | -23.86% | 3.553% | -13.71% | 2040% |  | 57.36% | 11.68% | 478.7% |  | 280.7%
| Z-Score (Absolute) |  | 0.18 | 28.95 | 1.48 | 2.83 | 0.81 | 0.72 | 0.49 | 0.67 | 0.4 | 0.45 | 17.74 | 0.54 | 0.18 | 0.22 | 0.67 | 1.48 | NA
| Analytical results for Ce in OREAS 267 (Certified Value 58 ppm) | [18 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 59.97 | 56.6267 | 57.665 | 56.5 | 58 | 60.7 | 60.7 | 58.6 | 61.9 | 62 | 52.0864 | 53.71 | 56.4 | 54.65 | 59 |  |  | 1
| 1 | 2 | 59.55 | 56.5678 | 56.9167 | 57.5 | 58 | 63.5 | 65.1 | 59.9 | 58.1 | 62.3 | 50.3278 | 54.76 | 57.28 | 54.67 | 57 |  |  | 26
| 1 | 3 | 61.5 | 56.8893 | 60.7526 | 56.9 | 56 | 59.9 | 63.6 | 62 | 64.4 | 59 | 51.5438 | 53.73 | 58.6 | 55.67 | 56 |  |  | 16
| 1 | 4 | 59.19 | 56.5514 | 56.8981 | 57.3 | 57 | 61.7 | 61 | 61.3 | 62.9 | 61.1 | 51.1883 | 52.65 | 57.6 | 56.22 | 57 |  |  | 58.2847
| 1 | 5 | 61.86 | 56.2606 | 60.0189 | 55.1 | 57 | 59 | 61.9 | 60.4 | 61.7 | 61.2 | 55.7141 | 53.22 | 57.15 | 54.13 | 59 |  |  | 22
| 1 | 6 | 60.72 | 56.937 | 60.136 | 57.2 | 57 | 63.7 | 62.2 | 60.4 | 64.3 | 62.5 | 55.5505 | 55.01 | 56.21 | 55.06 | 57
| Mean |  | 60.465 | 56.6388 | 58.7312 | 56.75 | 57.1667 | 61.4167 | 62.4167 | 60.4333 | 62.2167 | 61.35 | 52.7351 | 53.8467 | 57.2067 | 55.0667 | 57.5
| Median |  | 60.345 | 56.5972 | 58.842 | 57.05 | 57 | 61.2 | 62.05 | 60.4 | 62.4 | 61.6 | 51.8151 | 53.72 | 57.215 | 54.865 | 57
| Std Dev. |  | 1.0761 | 0.248 | 1.7611 | 0.88034 | 0.75277 | 1.9125 | 1.6654 | 1.1708 | 2.319 | 1.2849 | 2.3164 | 0.89939 | 0.86588 | 0.76191 | 1.2247
| Rel.Std.Dev. |  | 1.78% | 0.4379% | 2.999% | 1.551% | 1.317% | 3.114% | 2.668% | 1.937% | 3.727% | 2.094% | 4.392% | 1.67% | 1.514% | 1.384% | 2.13%
| PDM3 |  | 3.741% | -2.824% | 0.7661% | -2.633% | -1.918% | 5.374% | 7.089% | 3.686% | 6.746% | 5.259% | -9.521% | -7.614% | -1.85% | -5.521% | -1.346%
| Z-Score (Absolute) |  | 0.68 | 0.2 | 0.28 | 0.17 | 0.08 | 0.9 | 1.13 | 0.67 | 1.08 | 0.89 | 1.1 | 0.84 | 0.07 | 0.56 | 0
| Analytical results for Co in OREAS 267 (Certified Value 11 ppm) | [18 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
|  |  |  | 0.25g | [15 empty] | 2
| 1 | 1 | 11.4 | 10.85 | 10.8856 | 11 | 10.9 | 10 | 11.2 | 10.6 | 12 | 10.2 | 9.6204 | 11 | 10.7 | 10.9 | 12 | 14.14 | 11.1 | 1
| 1 | 2 | 11.2 | 10.91 | 10.7833 | 10 | 11.3 | 10.4 | 11.9 | 11 | 11.5 | 9.9 | 9.6388 | 11 | 11.4 | 10.8 | 11.9 | 14.54 | 11 | 27
| 1 | 3 | 11.2 | 10.1 | 11.6039 | 10 | 10.8 | 9.6 | 12 | 11 | 12.3 | 9.8 | 9.8632 | 11 | 10.9 | 11 | 12.2 | 13.87 | 10.9 | 16
| 1 | 4 | 11.1 | 10.94 | 10.9831 | 10 | 10.9 | 9.9 | 11.2 | 11.2 | 12.3 | 10.1 | 10.1305 | 11 | 11.4 | 11 | 11.9 | 13.99 | 10.9 | 11.0372
| 1 | 5 | 11.2 | 10.9 | 11.3075 | 10 | 10.6 | 9.9 | 11 | 11.2 | 12.1 | 10.1 | 10.1331 | 11 | 11.3 | 10.9 | 12.1 | 12.5 | 10.7 | 23
| 1 | 6 | 11.4 | 10.44 | 11.2824 | 10 | 11 | 10.2 | 11.4 | 11.2 | 12.1 | 10.2 | 10.3519 | 11 | 10.8 | 10.8 | 12 | 13.26 | 10.8
| Mean |  | 11.25 | 10.69 | 11.141 | 10.1667 | 10.9167 | 10 | 11.45 | 11.0333 | 12.05 | 10.05 | 9.9563 | 11 | 11.0833 | 10.9 | 12.0167 | 13.7167 | 10.9
| Median |  | 11.2 | 10.875 | 11.1327 | 10 | 10.9 | 9.95 | 11.3 | 11.1 | 12.1 | 10.1 | 9.9968 | 11 | 11.1 | 10.9 | 12 | 13.93 | 10.9
| Std Dev. |  | 0.12247 | 0.34386 | 0.30987 | 0.40825 | 0.23166 | 0.27568 | 0.40866 | 0.23381 | 0.29496 | 0.16432 | 0.29677 | 0 | 0.31885 | 0.089443 | 0.1169 | 0.7272 | 0.14142
| Rel.Std.Dev. |  | 1.089% | 3.217% | 2.781% | 4.016% | 2.122% | 2.757% | 3.569% | 2.119% | 2.448% | 1.635% | 2.981% | 0% | 2.877% | 0.8206% | 0.9729% | 5.302% | 1.297%
| PDM3 |  | 1.928% | -3.146% | 0.9398% | -7.888% | -1.092% | -9.398% | 3.74% | -0.03529% | 9.176% | -8.945% | -9.793% | -0.3373% | 0.4177% | -1.243% | 8.874% | 24.28% | -1.243%
| Z-Score (Absolute) |  | 0.67 | 0.84 | 0.38 | NA | 0.22 | 2.7 | 1.21 | 0.09 | 2.83 | 2.56 | 2.82 | 0 | 0.22 | 0.27 | 2.74 | NA | 0.27
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 267 (Certified Value 98 ppm) | [18 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 112 | 85.2505 | 95.9736 | 63 | 82 | 97 | 93 | 94 | 105 | 96 | 129.4846 | 73 | 99 | 105 | 101 | 102 |  | 1
| 1 | 2 | 111 | 87.8955 | 100.5841 | 67 | 87 | 98 | 92 | 96 | 100 | 96 | 132.6938 | 75 | 100 | 105 | 96.28 | 101 |  | 28
| 1 | 3 | 111 | 86.3775 | 104.9387 | 63 | 89 | 95 | 94 | 94 | 103 | 94 | 116.5175 | 81 | 98 | 108 | 99.22 | 102 |  | 16
| 1 | 4 | 112 | 89.197 | 95.6556 | 65 | 84 | 96 | 92 | 94 | 98 | 94 | 120.2645 | 79 | 103 | 105 | 96.04 | 101 |  | 97.8292
| 1 | 5 | 113 | 85.1355 | 100.7074 | 64 | 88 | 99 | 93 | 95 | 97 | 97 | 121.3024 | 79 | 103 | 107 | 95.69 | 102 |  | 24
| 1 | 6 | 113 | 86.0835 | 100.1503 | 75 | 90 | 99 | 92 | 95 | 103 | 93 | 124.2747 | 81 | 99 | 106 | 101.5 | 101
| Mean |  | 112 | 86.6566 | 99.6683 | 66.1667 | 86.6667 | 97.3333 | 92.6667 | 94.6667 | 101 | 95 | 124.0896 | 78 | 100.3333 | 106 | 98.2883 | 101.5
| Median |  | 112 | 86.2305 | 100.3672 | 64.5 | 87.5 | 97.5 | 92.5 | 94.5 | 101.5 | 95 | 122.7885 | 79 | 99.5 | 105.5 | 97.75 | 101.5
| Std Dev. |  | 0.89443 | 1.5928 | 3.4549 | 4.5789 | 3.0768 | 1.633 | 0.8165 | 0.8165 | 3.1623 | 1.5492 | 6.0483 | 3.2863 | 2.1602 | 1.2649 | 2.622 | 0.54772
| Rel.Std.Dev. |  | 0.7986% | 1.838% | 3.466% | 6.92% | 3.55% | 1.678% | 0.8811% | 0.8625% | 3.131% | 1.631% | 4.874% | 4.213% | 2.153% | 1.193% | 2.668% | 0.5396%
| PDM3 |  | 14.49% | -11.42% | 1.88% | -32.37% | -11.41% | -0.5069% | -5.277% | -3.233% | 3.241% | -2.892% | 26.84% | -20.27% | 2.56% | 8.352% | 0.4693% | 3.752%
| Z-Score (Absolute) |  | 2.17 | 1.7 | 0.28 | 4.83 | 1.7 | 0.07 | 0.79 | 0.48 | 0.49 | 0.43 | 4.01 | 3.02 | 0.39 | 1.25 | 0.07 | 0.56
| Analytical results for Cs in OREAS 267 (Certified Value 5.5 ppm) | [18 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.46 | 5.45 | 5.4074 | 5.96 | 5.63 | 5.52 | 5.23 | 5.34 | 5.75 | 5.64 | 4.4545 | 5.3 | 5.42 | 5.56 | 6.4 |  |  | 1
| 1 | 2 | 5.45 | 5.3317 | 5.4134 | 5.62 | 5.65 | 5.73 | 5.62 | 5.5 | 5.64 | 5.47 | 4.3606 | 5.2 | 5.57 | 5.51 | 6.3 |  |  | 29
| 1 | 3 | 5.5 | 5.6481 | 5.7644 | 6.06 | 5.46 | 5.39 | 5.53 | 5.56 | 6.35 | 5.36 | 4.4223 | 5.2 | 5.31 | 5.57 | 6.2 |  |  | 16
| 1 | 4 | 5.42 | 5.5245 | 5.4138 | 5.66 | 5.47 | 5.63 | 5.26 | 5.49 | 5.96 | 5.47 | 4.3374 | 5.3 | 5.36 | 5.63 | 6.2 |  |  | 5.4953
| 1 | 5 | 5.44 | 5.9162 | 5.6723 | 5.65 | 5.4 | 5.55 | 5.4 | 5.45 | 5.85 | 5.5 | 4.7887 | 5.3 | 5.36 | 5.59 | 6.5 |  |  | 25
| 1 | 6 | 5.44 | 5.5806 | 5.4983 | 5.78 | 5.31 | 5.63 | 5.37 | 5.39 | 5.95 | 5.62 | 4.7097 | 5.3 | 5.37 | 5.51 | 6.2
| Mean |  | 5.4517 | 5.5752 | 5.5283 | 5.7883 | 5.4867 | 5.575 | 5.4017 | 5.455 | 5.9167 | 5.51 | 4.5122 | 5.2667 | 5.3983 | 5.5617 | 6.3
| Median |  | 5.445 | 5.5526 | 5.456 | 5.72 | 5.465 | 5.59 | 5.385 | 5.47 | 5.9 | 5.485 | 4.4384 | 5.3 | 5.365 | 5.565 | 6.25
| Std Dev. |  | 0.027142 | 0.1996 | 0.15383 | 0.18291 | 0.13186 | 0.11658 | 0.15145 | 0.079687 | 0.24476 | 0.10469 | 0.18995 | 0.05164 | 0.091086 | 0.046655 | 0.12649
| Rel.Std.Dev. |  | 0.4979% | 3.58% | 2.783% | 3.16% | 2.403% | 2.091% | 2.804% | 1.461% | 4.137% | 1.9% | 4.21% | 0.9805% | 1.687% | 0.8389% | 2.008%
| PDM3 |  | -0.7948% | 1.453% | 0.5989% | 5.332% | -0.1579% | 1.45% | -1.705% | -0.7341% | 7.667% | 0.2667% | -17.89% | -4.161% | -1.765% | 1.207% | 14.64%
| Z-Score (Absolute) |  | 0.6 | 0.67 | 0.19 | 2.88 | 0.24 | 0.67 | 1.12 | 0.57 | 4.21 | 0 | 10.32 | 2.52 | 1.16 | 0.53 | 8.17
| Analytical results for Cu in OREAS 267 (Certified Value 23.9 ppm) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 25.1 | 23.47 | 24.2585 | 22 | 29 | 23.9 | 24 | 23.1 | 23.5 | 23.8 | 31.0067 | 25 | 26.1 | 24.5 | 19.6 | 21.72 | 21.2 | 1
| 1 | 2 | 24.6 | 23.67 | 23.8269 | 21 | 29.2 | 24.9 | 25.3 | 24.3 | 22.7 | 23.2 | 32.5887 | 25 | 25.7 | 24.1 | 20.9 | 25.75 | 21.8 | 30
| 1 | 3 | 25.2 | 23.23 | 25.3559 | 22 | 27.4 | 23.3 | 24.9 | 23.9 | 24 | 22.7 | 29.4481 | 25 | 25.8 | 24.5 | 20.7 | 22.91 | 21.1 | 16
| 1 | 4 | 25 | 23.54 | 23.8088 | 21 | 27.7 | 24.4 | 24.3 | 24.6 | 28 | 23.5 | 29.6523 | 26 | 26.7 | 25 | 18.9 | 21.09 | 21.9 | 23.8994
| 1 | 5 | 25.8 | 23.15 | 24.7248 | 22 | 26.7 | 24 | 24.6 | 23.8 | 24.3 | 23.6 | 30.6354 | 24 | 27 | 24.8 | 22.3 | 21.48 | 21.7 | 26
| 1 | 6 | 25 | 25.81 | 24.5046 | 22 | 26.5 | 24.4 | 24.2 | 23.6 | 24.1 | 23.5 | 29.9959 | 24 | 25.4 | 24.4 | 23.7 | 23.93 | 21.8
| Mean |  | 25.1167 | 23.8117 | 24.4133 | 21.6667 | 27.75 | 24.15 | 24.55 | 23.8833 | 24.4333 | 23.3833 | 30.5545 | 24.8333 | 26.1167 | 24.55 | 21.0167 | 22.8133 | 21.5833
| Median |  | 25.05 | 23.505 | 24.3816 | 22 | 27.55 | 24.2 | 24.45 | 23.85 | 24.05 | 23.5 | 30.3157 | 25 | 25.95 | 24.5 | 20.8 | 22.315 | 21.75
| Std Dev. |  | 0.392 | 0.99801 | 0.58772 | 0.5164 | 1.1362 | 0.54681 | 0.48477 | 0.52694 | 1.8392 | 0.38687 | 1.1575 | 0.75277 | 0.61779 | 0.31464 | 1.7577 | 1.779 | 0.34303
| Rel.Std.Dev. |  | 1.561% | 4.191% | 2.407% | 2.383% | 4.094% | 2.264% | 1.975% | 2.206% | 7.527% | 1.654% | 3.788% | 3.031% | 2.366% | 1.282% | 8.364% | 7.798% | 1.589%
| PDM3 |  | 5.093% | -0.3671% | 2.15% | -9.342% | 16.11% | 1.049% | 2.722% | -0.06723% | 2.234% | -2.159% | 27.85% | 3.908% | 9.277% | 2.722% | -12.06% | -4.544% | -9.691%
| Z-Score (Absolute) |  | 0.67 | 0.58 | 0 | 2.63 | 3.2 | 0.25 | 0.13 | 0.51 | 0.02 | 0.99 | 5.89 | 0.4 | 1.63 | 0.13 | 3.26 | 1.53 | 2.71
| Analytical results for Dy in OREAS 267 (Certified Value 3.07 ppm) | [18 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 3.1396 | 2.9326 | 3.1 | 3 | 1.8531 | 3.03 | [11 empty] | 1
| 1 | 2 | 3.1206 | 2.959 | 3.1 | 3.11 | 1.8069 | 3.05 | [11 empty] | 6
| 1 | 3 | 3.1554 | 3.1655 | 3.2 | 2.9 | 1.8486 | 3.04 | [11 empty] | 16
| 1 | 4 | 3.171 | 2.9656 | 3.1 | 2.92 | 1.8338 | 3.01 | [11 empty] | 3.0711
| 1 | 5 | 3.1754 | 3.1548 | 3.6 | 3.08 | 1.9821 | 2.94 | [11 empty] | 27
| 1 | 6 | 3.1717 | 3.1111 | 3.2 | 3.03 | 1.9448 | 2.96
| Mean |  | 3.1556 | 3.0481 | 3.2167 | 3.0067 | 1.8782 | 3.005
| Median |  | 3.1632 | 3.0384 | 3.15 | 3.015 | 1.8508 | 3.02
| Std Dev. |  | 0.021745 | 0.10697 | 0.19408 | 0.084301 | 0.068977 | 0.045056
| Rel.Std.Dev. |  | 0.6891% | 3.509% | 6.034% | 2.804% | 3.673% | 1.499%
| PDM3 |  | 2.753% | -0.7479% | 4.741% | -2.097% | -38.84% | -2.152%
| Z-Score (Absolute) |  | 1.15 | 0.19 | 1.69 | 0.19 | 10.29 | 0.2
| Analytical results for Er in OREAS 267 (Certified Value 1.5 ppm) | [18 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.5917 | 1.5391 | 1.5 | 1.47 | 0.81603 | 1.44 | [11 empty] | 1
| 1 | 2 | 1.575 | 1.5095 | 1.4 | 1.32 | 0.8185 | 1.45 | [11 empty] | 7
| 1 | 3 | 1.5912 | 1.6525 | 1.4 | 1.35 | 0.86731 | 1.51 | [11 empty] | 16
| 1 | 4 | 1.5761 | 1.5244 | 1.5 | 1.39 | 0.84443 | 1.51 | [11 empty] | 1.499
| 1 | 5 | 1.5783 | 1.6683 | 1.5 | 1.39 | 0.8529 | 1.52 | [11 empty] | 28
| 1 | 6 | 1.6057 | 1.6233 | 1.4 | 1.4 | 0.88391 | 2.14
| Mean |  | 1.5863 | 1.5862 | 1.45 | 1.3867 | 0.84718 | 1.595
| Median |  | 1.5848 | 1.5812 | 1.45 | 1.39 | 0.84866 | 1.51
| Std Dev. |  | 0.012056 | 0.069881 | 0.054772 | 0.050859 | 0.026786 | 0.26913
| Rel.Std.Dev. |  | 0.76% | 4.406% | 3.777% | 3.668% | 3.162% | 16.87%
| PDM3 |  | 5.824% | 5.813% | -3.271% | -7.496% | -43.49% | 6.401%
| Z-Score (Absolute) |  | 0.63 | 0.63 | 0.63 | 1.22 | 6.23 | 0.71
| Analytical results for Eu in OREAS 267 (Certified Value 1.03 ppm) | [18 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.0441 | 1.0171 | 1.09 | 0.99 | 2.2484 | 1 | [11 empty] | 1
| 1 | 2 | 1.0446 | 0.9963 | 1.05 | 1.03 | 2.215 | 0.99 | [11 empty] | 8
| 1 | 3 | 1.0364 | 1.0504 | 1.09 | 1.14 | 2.0954 | 1 | [11 empty] | 16
| 1 | 4 | 1.0038 | 1.0181 | 1.02 | 1.09 | 2.0359 | 1 | [11 empty] | 1.0285
| 1 | 5 | 1.0199 | 1.0232 | 0.96 | 1.03 | 2.1737 | 1.01 | [11 empty] | 29
| 1 | 6 | 1.0005 | 1.0803 | 1.01 | 1.04 | 2.2244 | 0.98
| Mean |  | 1.0249 | 1.0309 | 1.0367 | 1.0533 | 2.1655 | 0.99667
| Median |  | 1.0282 | 1.0207 | 1.035 | 1.035 | 2.1944 | 1
| Std Dev. |  | 0.019781 | 0.029778 | 0.050465 | 0.053166 | 0.083149 | 0.010328
| Rel.Std.Dev. |  | 1.93% | 2.889% | 4.868% | 5.047% | 3.84% | 1.036%
| PDM3 |  | -0.3503% | 0.2346% | 0.7948% | 2.415% | 110.6% | -3.094%
| Z-Score (Absolute) |  | 0.42 | 0.14 | 0.14 | 0.93 | 53.66 | 1.76
| Analytical results for Fe in OREAS 267 (Certified Value 2.86 wt.%) | [18 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.92 | 2.943 | 2.8659 | 2.84 | 2.84 | 2.89 | 2.89 | 2.77 | 2.82 | 2.84 | 3.2663 | 2.86 | 2.71 | 2.8939 | 2.7504 | 2.82 | 2.93 | 1
| 1 | 2 | 2.89 | 2.9447 | 2.8568 | 2.91 | 2.9 | 2.93 | 3.05 | 2.82 | 2.87 | 2.86 | 3.2214 | 2.82 | 2.81 | 2.8953 | 2.7133 | 2.82 | 2.94 | #N/A
| 1 | 3 | 2.9 | 2.9524 | 3.0455 | 2.83 | 2.87 | 2.86 | 2.98 | 2.77 | 2.85 | 2.85 | 3.1492 | 2.82 | 2.75 | 2.8981 | 2.7691 | 2.88 | 2.92 | 16
| 1 | 4 | 2.9 | 2.958 | 2.879 | 2.89 | 2.8 | 2.9 | 2.94 | 2.8 | 2.83 | 2.8 | 3.0363 | 2.81 | 2.78 | 2.8845 | 2.7572 | 2.87 | 2.91 | 2.8626
| 1 | 5 | 2.96 | 2.9511 | 2.914 | 2.84 | 2.85 | 2.95 | 2.97 | 2.79 | 2.85 | 2.81 | 3.1873 | 2.86 | 2.79 | 2.9104 | 2.7255 | 2.8 | 2.91 | 30
| 1 | 6 | 2.94 | 2.9375 | 2.8927 | 2.91 | 2.86 | 2.93 | 2.95 | 2.79 | 2.9 | 2.83 | 3.2768 | 2.85 | 2.73 | 2.8382 | 2.6927 | 2.83 | 2.94
| Mean |  | 2.9183 | 2.9478 | 2.909 | 2.87 | 2.8533 | 2.91 | 2.9633 | 2.79 | 2.8533 | 2.8317 | 3.1896 | 2.8367 | 2.7617 | 2.8867 | 2.7347 | 2.8367 | 2.925
| Median |  | 2.91 | 2.9479 | 2.8859 | 2.865 | 2.855 | 2.915 | 2.96 | 2.79 | 2.85 | 2.835 | 3.2043 | 2.835 | 2.765 | 2.8946 | 2.738 | 2.825 | 2.925
| Std Dev. |  | 0.027142 | 0.0074254 | 0.069879 | 0.037417 | 0.033267 | 0.032863 | 0.052789 | 0.018974 | 0.028752 | 0.023166 | 0.089035 | 0.022509 | 0.038166 | 0.025197 | 0.02913 | 0.031411 | 0.013784
| Rel.Std.Dev. |  | 0.93% | 0.2519% | 2.402% | 1.304% | 1.166% | 1.129% | 1.781% | 0.6801% | 1.008% | 0.8181% | 2.791% | 0.7935% | 1.382% | 0.8729% | 1.065% | 1.107% | 0.4712%
| PDM3 |  | 1.948% | 2.978% | 1.622% | 0.26% | -0.3222% | 1.657% | 3.521% | -2.535% | -0.3222% | -1.079% | 11.42% | -0.9044% | -3.524% | 0.8446% | -4.467% | -0.9044% | 2.181%
| Z-Score (Absolute) |  | 0.8 | 1.21 | 0.67 | 0.12 | 0.12 | 0.68 | 1.43 | 1.01 | 0.12 | 0.42 | NA | 0.35 | 1.41 | 0.35 | 1.79 | 0.35 | 0.89
| Analytical results for Ga in OREAS 267 (Certified Value 14.5 ppm) | [18 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 15.01 | 11.324 | 14.234 | 15.9 | 15.6 | 14.3 | 14.45 | 13.9 | 13.85 | 14.15 | 14.0305 | 14.7 | 15.1 | 14.21 | 14 | 12.89 | 16 | 1
| 1 | 2 | 14.36 | 11.3045 | 14.4254 | 15.6 | 15.3 | 14.85 | 15.25 | 13.9 | 13.25 | 13.85 | 13.853 | 14.6 | 15.9 | 14.24 | 14 | 13.32 | 17 | 34
| 1 | 3 | 14.76 | 11.479 | 14.8891 | 15.4 | 15.5 | 13.75 | 14.85 | 14.25 | 14.75 | 13.6 | 14.2604 | 14.7 | 15.1 | 14.16 | 14 | 12.64 | 15 | 16
| 1 | 4 | 14.69 | 10.9405 | 15.0825 | 15.8 | 15.3 | 14.15 | 14.6 | 14.05 | 14.1 | 13.9 | 14.2318 | 14.6 | 15.8 | 13.94 | 14 | 12.81 | 16 | 14.5213
| 1 | 5 | 15.05 | 12.0975 | 14.8795 | 14.7 | 15.8 | 14.35 | 14.7 | 13.95 | 14.1 | 13.95 | 14.6562 | 14.4 | 16 | 13.96 | 14 | 13.55 | 16 | 31
| 1 | 6 | 14.66 | 10.8495 | 15.1839 | 16.2 | 15.5 | 14.55 | 14.6 | 14.1 | 14.35 | 14.25 | 14.7927 | 14.2 | 15.2 | 14.09 | 14 | 13.38 | 16
| Mean |  | 14.755 | 11.3325 | 14.7824 | 15.6 | 15.5 | 14.325 | 14.7417 | 14.025 | 14.0667 | 13.95 | 14.3041 | 14.5333 | 15.5167 | 14.1 | 14 | 13.0983 | 16
| Median |  | 14.725 | 11.3143 | 14.8843 | 15.7 | 15.5 | 14.325 | 14.65 | 14 | 14.1 | 13.925 | 14.2461 | 14.6 | 15.5 | 14.125 | 14 | 13.105 | 16
| Std Dev. |  | 0.25352 | 0.44598 | 0.37424 | 0.51769 | 0.18974 | 0.37115 | 0.28181 | 0.13693 | 0.50266 | 0.23022 | 0.36004 | 0.19664 | 0.42622 | 0.12696 | 0 | 0.36581 | 0.63246
| Rel.Std.Dev. |  | 1.718% | 3.935% | 2.532% | 3.319% | 1.224% | 2.591% | 1.912% | 0.9763% | 3.573% | 1.65% | 2.517% | 1.353% | 2.747% | 0.9005% | 0% | 2.793% | 3.953%
| PDM3 |  | 1.609% | -21.96% | 1.798% | 7.428% | 6.74% | -1.352% | 1.518% | -3.418% | -3.131% | -3.934% | -1.496% | 0.08289% | 6.855% | -2.901% | -3.59% | -9.799% | 10.18%
| Z-Score (Absolute) |  | 0.7 | 4.84 | 0.74 | 2.06 | 1.9 | 0 | 0.67 | 0.49 | 0.42 | 0.61 | 0.03 | 0.34 | 1.93 | 0.36 | NA | 1.99 | NA
Results from laboratories 21 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 267 (Certified Value 4.1 ppm) | [18 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 4.6577 | 3.9093 | 3.8 | 3.71 | 1.368 | 3.94 | [11 empty] | 1
| 1 | 2 | 4.6614 | 3.8082 | 3.7 | 4.35 | 1.3502 | 4.29 | [11 empty] | 9
| 1 | 3 | 4.6863 | 4.1965 | 3.8 | 4.01 | 1.4016 | 3.99 | [11 empty] | 16
| 1 | 4 | 4.6288 | 3.95 | 3.9 | 4.04 | 1.3686 | 3.99 | [11 empty] | 4.102
| 1 | 5 | 4.6273 | 3.9981 | 4.2 | 3.77 | 1.5031 | 3.91 | [11 empty] | 32
| 1 | 6 | 4.6324 | 4.324 | 3.9 | 3.81 | 1.4557 | 3.87
| Mean |  | 4.649 | 4.031 | 3.8833 | 3.9483 | 1.4079 | 3.9983
| Median |  | 4.645 | 3.9741 | 3.85 | 3.91 | 1.3851 | 3.965
| Std Dev. |  | 0.023576 | 0.19256 | 0.17224 | 0.23718 | 0.059753 | 0.15026
| Rel.Std.Dev. |  | 0.5071% | 4.777% | 4.435% | 6.007% | 4.244% | 3.758%
| PDM3 |  | 13.33% | -1.73% | -5.331% | -3.746% | -65.68% | -2.527%
| Z-Score (Absolute) |  | 6.17 | 0.53 | 0.82 | 0.23 | 23.43 | 0.23
| Analytical results for Ge in OREAS 267 (Indicative Value 0.26 ppm) | [18 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.4 | < 1 | 1 | 0.15 | 0.06 | 0.14 | 0.11 | 0.12 | < 0.05 | 1.2 | 0.4 | 1.3 | 3.3 |  |  |  |  | 1
| 1 | 2 | 1.4 | < 1 | 1.1 | 0.13 | 0.05 | 0.15 | 0.13 | 0.1 | 0.06 | 1.2 | 0.4 | 1.3 | 3.2 |  |  |  |  | 2
| 1 | 3 | 1.4 | < 1 | 1 | 0.14 | 0.06 | 0.13 | 0.12 | 0.11 | 0.08 | 1.2 | 0.4 | 1.4 | 3.5 |  |  |  |  | 16
| 1 | 4 | 1.4 | < 1 | 1 | 0.15 | 0.06 | 0.14 | 0.11 | 0.12 | 0.07 | 1.3 | 0.4 | 1.3 | 3.9 |  |  |  |  | 0.26021
| 1 | 5 | 1.4 | < 1 | 1 | 0.16 | 0.06 | 0.16 | 0.12 | 0.12 | 0.06 | 1.2 | 0.4 | 1.3 | 3.1 |  |  |  |  | 8
| 1 | 6 | 1.3 | < 1 | 1.1 | 0.14 | 0.06 | 0.14 | 0.1 | 0.13 | 0.08 | 1.2 | 0.4 | 1.3 | 3.8
| Mean |  | 1.3833 |  | 1.0333 | 0.145 | 0.058333 | 0.14333 | 0.115 | 0.11667 | 0.07 | 1.2167 | 0.4 | 1.3167 | 3.4667
| Median |  | 1.4 |  | 1 | 0.145 | 0.06 | 0.14 | 0.115 | 0.12 | 0.07 | 1.2 | 0.4 | 1.3 | 3.4
| Std Dev. |  | 0.040825 |  | 0.05164 | 0.010488 | 0.0040825 | 0.010328 | 0.010488 | 0.010328 | 0.01 | 0.040825 | 6.0809e-17 | 0.040825 | 0.3266
| Rel.Std.Dev. |  | 2.951% |  | 4.997% | 7.233% | 6.999% | 7.206% | 9.12% | 8.853% | 14.29% | 3.355% | 1.52e-14% | 3.101% | 9.421%
| PDM3 |  | 431.6% |  | 297.1% | -44.28% | -77.58% | -44.92% | -55.8% | -55.16% | -73.1% | 367.6% | 53.72% | 406% | 1232%
| Z-Score (Absolute) |  | 1.96 | 0.2 | 1.27 | 0.51 | 0.68 | 0.51 | 0.57 | 0.57 | 0.67 | 1.63 | 0 | 1.83 | 6.13
| Analytical results for Hf in OREAS 267 (Certified Value 2.96 ppm) | [18 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.2 | 2.1791 | 2.981 | 2.9 | 3.01 | 3 | 2.9 | 2.8 | 3 | 2.8 | 2.7921 | 3.2 | 3.08 | 3.08 | 3.24 | 3.1 |  | 1
| 1 | 2 | 3.11 | 2.216 | 2.9104 | 2.9 | 2.68 | 2.8 | 3 | 2.9 | 2.8 | 2.8 | 2.6632 | 3.1 | 3.23 | 3.02 | 3.31 | 2.9 |  | 20
| 1 | 3 | 3.21 | 2.1615 | 3.0682 | 3.2 | 2.74 | 2.8 | 2.8 | 2.8 | 3 | 2.7 | 2.6746 | 3.3 | 3.15 | 3.2 | 3.36 | 2.9 |  | 16
| 1 | 4 | 3.05 | 2.2189 | 3.0135 | 2.9 | 2.79 | 2.8 | 2.8 | 2.9 | 3.2 | 2.8 | 2.4845 | 3.2 | 3.17 | 3.06 | 3.62 | 2.8 |  | 2.9642
| 1 | 5 | 3.25 | 2.1746 | 3.0405 | 3 | 2.51 | 2.9 | 3 | 2.9 | 3.1 | 2.7 | 2.6701 | 3.2 | 3.19 | 3.13 | 3.17 | 3 |  | 33
| 1 | 6 | 3.08 | 2.1448 | 3.0429 | 2.6 | 2.71 | 2.8 | 2.8 | 2.9 | 3.1 | 2.7 | 2.7665 | 3.1 | 3.1 | 2.96 | 3.22 | 2.8
| Mean |  | 3.15 | 2.1825 | 3.0094 | 2.9167 | 2.74 | 2.85 | 2.8833 | 2.8667 | 3.0333 | 2.75 | 2.6752 | 3.1833 | 3.1533 | 3.075 | 3.32 | 2.9167
| Median |  | 3.155 | 2.1768 | 3.027 | 2.9 | 2.725 | 2.8 | 2.85 | 2.9 | 3.05 | 2.75 | 2.6724 | 3.2 | 3.16 | 3.07 | 3.275 | 2.9
| Std Dev. |  | 0.080747 | 0.029601 | 0.056868 | 0.19408 | 0.16297 | 0.083666 | 0.098319 | 0.05164 | 0.13663 | 0.054772 | 0.10819 | 0.075277 | 0.056095 | 0.083845 | 0.16162 | 0.1169
| Rel.Std.Dev. |  | 2.563% | 1.356% | 1.89% | 6.654% | 5.948% | 2.936% | 3.41% | 1.801% | 4.504% | 1.992% | 4.044% | 2.365% | 1.779% | 2.727% | 4.868% | 4.008%
| PDM3 |  | 6.268% | -26.37% | 1.526% | -1.603% | -7.563% | -3.853% | -2.728% | -3.29% | 2.332% | -7.226% | -9.75% | 7.393% | 6.381% | 3.738% | 12% | -1.603%
| Z-Score (Absolute) |  | 0.97 | 3.05 | 0.38 | 0 | 0.73 | 0.28 | 0.14 | 0.21 | 0.48 | 0.69 | 1 | 1.11 | 0.98 | 0.66 | 1.67 | 0
| Analytical results for Hg in OREAS 267 (Indicative Value 0.068 ppm) | [18 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 1.06 | 1.21 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | [15 empty] | 3
3
| 1 | 1 | < 1 | 0.08 | [15 empty] | 1
| 1 | 2 | < 1 | 0.08 | [15 empty] | 1
| 1 | 3 | < 1 | < 0.01 | [15 empty] | 16
| 1 | 4 | < 1 | 0.08 | [15 empty] | 0.0675
| 1 | 5 | < 1 | 0.08 | [15 empty] | 7
| 1 | 6 | < 1 | 0.08
| Mean |  |  | 0.08
| Median |  |  | 0.08
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 18.52%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 267 (Certified Value 0.56 ppm) | [18 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.55125 | 0.55398 | 0.6 | 0.54 | 0.35383 | 0.55 | [11 empty] | 1
| 1 | 2 | 0.56742 | 0.5931 | 0.6 | 0.55 | 0.36087 | 0.54 | [11 empty] | 5
| 1 | 3 | 0.58691 | 0.59792 | 0.6 | 0.5 | 0.33634 | 0.56 | [11 empty] | 16
| 1 | 4 | 0.56841 | 0.53869 | 0.6 | 0.51 | 0.33613 | 0.55 | [11 empty] | 0.56362
| 1 | 5 | 0.57311 | 0.5714 | 0.6 | 0.52 | 0.36771 | 0.54 | [11 empty] | 34
| 1 | 6 | 0.57996 | 0.58632 | 0.6 | 0.55 | 0.36259 | 0.53
| Mean |  | 0.57118 | 0.57357 | 0.6 | 0.52833 | 0.35291 | 0.545
| Median |  | 0.57076 | 0.57886 | 0.6 | 0.53 | 0.35735 | 0.545
| Std Dev. |  | 0.012225 | 0.023419 | 0 | 0.02137 | 0.013663 | 0.010488
| Rel.Std.Dev. |  | 2.14% | 4.083% | 0% | 4.045% | 3.871% | 1.924%
| PDM3 |  | 1.341% | 1.766% | 6.455% | -6.26% | -37.38% | -3.303%
| Z-Score (Absolute) |  | 0.39 | 0.46 | 1.25 | 0.89 | 6.12 | 0.39
| Analytical results for In in OREAS 267 (Certified Value 0.061 ppm) | [18 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.06 | < 0.2 | < 0.1 | 0.19 | 0.06 | 0.06 | 0.057 | 0.065 | 0.061 | 0.03503 | 0.05 | 0.06 | 0.06 | < 1 | 0.09 |  |  | 1
| 1 | 2 | 0.07 | < 0.2 | < 0.1 | 0.19 | 0.062 | 0.066 | 0.056 | 0.062 | 0.057 | 0.031091 | 0.05 | 0.07 | 0.06 | < 1 | 0.09 |  |  | 22
| 1 | 3 | 0.06 | < 0.2 | < 0.1 | 0.19 | 0.059 | 0.063 | 0.056 | 0.064 | 0.059 | 0.027655 | 0.06 | 0.06 | 0.06 | < 1 | 0.08 |  |  | 16
| 1 | 4 | 0.06 | < 0.2 | < 0.1 | 0.18 | 0.067 | 0.059 | 0.058 | 0.066 | 0.06 | 0.031342 | 0.05 | 0.06 | 0.06 | < 1 | 0.09 |  |  | 0.060833
| 1 | 5 | 0.07 | < 0.2 | < 0.1 | 0.19 | 0.061 | 0.059 | 0.059 | 0.066 | 0.058 | 0.029666 | 0.05 | 0.07 | 0.06 | < 1 | 0.09 |  |  | 35
| 1 | 6 | 0.07 | < 0.2 | < 0.1 | 0.2 | 0.066 | 0.063 | 0.051 | 0.065 | 0.06 | 0.029834 | 0.06 | 0.06 | 0.07 | < 1 | 0.08
| Mean |  | 0.065 |  |  | 0.19 | 0.0625 | 0.061667 | 0.056167 | 0.064667 | 0.059167 | 0.03077 | 0.053333 | 0.063333 | 0.061667 |  | 0.086667
| Median |  | 0.065 |  |  | 0.19 | 0.0615 | 0.0615 | 0.0565 | 0.065 | 0.0595 | 0.030463 | 0.05 | 0.06 | 0.06 |  | 0.09
| Std Dev. |  | 0.0054772 |  |  | 0.0063246 | 0.0032711 | 0.0028048 | 0.0027869 | 0.0015055 | 0.001472 | 0.0024648 | 0.005164 | 0.005164 | 0.0040825 |  | 0.005164
| Rel.Std.Dev. |  | 8.427% |  |  | 3.329% | 5.234% | 4.548% | 4.962% | 2.328% | 2.488% | 8.01% | 9.682% | 8.154% | 6.62% |  | 5.958%
| PDM3 |  | 6.849% |  |  | 212.3% | 2.74% | 1.37% | -7.671% | 6.301% | -2.74% | -49.42% | -12.33% | 4.11% | 1.37% |  | 42.47%
| Z-Score (Absolute) |  | 0.27 | 3.99 | 1.33 | 13.57 | 0 | 0.09 | 0.67 | 0.23 | 0.35 | 3.38 | 0.98 | 0.09 | 0.09 | 46.58 | 2.57
| Analytical results for K in OREAS 267 (Certified Value 1.25 wt.%) | [18 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.253 | 1.3245 | 1.2721 | 1.31 | 1.26 | 1.23 | 1.25 | 1.23 | 1.21 | 1.25 | 1.1856 | 1.22 | 1.16 | 1.2402 | 1.24 | 1.25 |  | 1
| 1 | 2 | 1.2469 | 1.3218 | 1.2521 | 1.29 | 1.28 | 1.22 | 1.32 | 1.25 | 1.25 | 1.25 | 1.2134 | 1.22 | 1.2 | 1.2395 | 1.26 | 1.26 |  | #N/A
| 1 | 3 | 1.2533 | 1.3263 | 1.2755 | 1.3 | 1.27 | 1.21 | 1.29 | 1.25 | 1.24 | 1.26 | 1.2539 | 1.21 | 1.21 | 1.2471 | 1.28 | 1.25 |  | 16
| 1 | 4 | 1.2512 | 1.3283 | 1.2531 | 1.3 | 1.24 | 1.23 | 1.26 | 1.25 | 1.23 | 1.23 | 1.1904 | 1.2 | 1.2 | 1.2443 | 1.28 | 1.26 |  | 1.2467
| 1 | 5 | 1.2712 | 1.327 | 1.2666 | 1.25 | 1.26 | 1.24 | 1.28 | 1.25 | 1.24 | 1.24 | 1.2759 | 1.23 | 1.23 | 1.2482 | 1.23 | 1.25 |  | 36
| 1 | 6 | 1.2665 | 1.3224 | 1.2605 | 1.29 | 1.26 | 1.24 | 1.28 | 1.24 | 1.25 | 1.24 | 1.3255 | 1.21 | 1.2 | 1.2278 | 1.25 | 1.25
| Mean |  | 1.257 | 1.325 | 1.2633 | 1.29 | 1.2617 | 1.2283 | 1.28 | 1.245 | 1.2367 | 1.245 | 1.2408 | 1.215 | 1.2 | 1.2412 | 1.2567 | 1.2533
| Median |  | 1.2531 | 1.3254 | 1.2636 | 1.295 | 1.26 | 1.23 | 1.28 | 1.25 | 1.24 | 1.245 | 1.2337 | 1.215 | 1.2 | 1.2422 | 1.255 | 1.25
| Std Dev. |  | 0.0095627 | 0.0025958 | 0.0097444 | 0.020976 | 0.013292 | 0.01169 | 0.024495 | 0.0083666 | 0.015055 | 0.010488 | 0.054626 | 0.010488 | 0.022804 | 0.0074413 | 0.020656 | 0.005164
| Rel.Std.Dev. |  | 0.7607% | 0.1959% | 0.7713% | 1.626% | 1.053% | 0.9517% | 1.914% | 0.672% | 1.217% | 0.8424% | 4.403% | 0.8632% | 1.9% | 0.5995% | 1.644% | 0.412%
| PDM3 |  | 0.8281% | 6.284% | 1.333% | 3.474% | 1.201% | -1.473% | 2.672% | -0.1358% | -0.8043% | -0.1358% | -0.4743% | -2.542% | -3.745% | -0.442% | 0.8% | 0.5326%
| Z-Score (Absolute) |  | 0.42 | 4.09 | 0.76 | 2.2 | 0.67 | 1.12 | 1.66 | 0.22 | 0.67 | 0.22 | 0.45 | 1.84 | 2.65 | 0.43 | 0.4 | 0.22
| Analytical results for La in OREAS 267 (Certified Value 31 ppm) | [18 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 30.65 | 30.044 | 29.3204 | 30.1 | 31.1 | 32.4 | 29.6 | 30.9 | 31 | 31.9 | 25.744 | 26.3 | 28.17 | 32.49 | 41 |  |  | 1
| 1 | 2 | 30.04 | 30.8653 | 29.2617 | 30.4 | 31.1 | 32.9 | 32.4 | 30.7 | 29.8 | 32.1 | 25.5423 | 27.7 | 28.41 | 32.97 | 41 |  |  | 23
| 1 | 3 | 30.89 | 30.4706 | 30.4699 | 30.8 | 30.2 | 31.9 | 31.9 | 32.3 | 32.2 | 30.7 | 25.5879 | 26.5 | 28.8 | 32.23 | 41 |  |  | 16
| 1 | 4 | 29.83 | 30.3991 | 29.0766 | 30.6 | 31.1 | 32 | 30.1 | 31.9 | 31.6 | 31.9 | 25.5054 | 26.4 | 28.46 | 32.01 | 41 |  |  | 30.9554
| 1 | 5 | 31.23 | 30.9711 | 30.5363 | 29.4 | 31 | 33.1 | 31.2 | 31.8 | 32.3 | 31.2 | 27.8361 | 26.6 | 28.47 | 33.46 | 40 |  |  | 37
| 1 | 6 | 30.47 | 30.2133 | 30.4017 | 30.6 | 30.4 | 32.6 | 31 | 31.6 | 32.7 | 32.4 | 27.2481 | 27.1 | 28.21 | 33.07 | 41
| Mean |  | 30.5183 | 30.4939 | 29.8445 | 30.3167 | 30.8167 | 32.4833 | 31.0333 | 31.5333 | 31.6 | 31.7 | 26.244 | 26.7667 | 28.42 | 32.705 | 40.8333
| Median |  | 30.56 | 30.4349 | 29.8611 | 30.5 | 31.05 | 32.5 | 31.1 | 31.7 | 31.9 | 31.9 | 25.666 | 26.55 | 28.435 | 32.73 | 41
| Std Dev. |  | 0.52285 | 0.3623 | 0.69054 | 0.50761 | 0.40702 | 0.47924 | 1.0558 | 0.61536 | 1.064 | 0.62929 | 1.0258 | 0.53541 | 0.22592 | 0.55287 | 0.40825
| Rel.Std.Dev. |  | 1.713% | 1.188% | 2.314% | 1.674% | 1.321% | 1.475% | 3.402% | 1.951% | 3.367% | 1.985% | 3.909% | 2% | 0.7949% | 1.69% | 0.9998%
| PDM3 |  | -1.412% | -1.491% | -3.589% | -2.063% | -0.4482% | 4.936% | 0.2517% | 1.867% | 2.082% | 2.405% | -15.22% | -13.53% | -8.191% | 5.652% | 31.91%
| Z-Score (Absolute) |  | 0.23 | 0.25 | 0.74 | 0.38 | 0 | 1.27 | 0.17 | 0.55 | 0.6 | 0.67 | 3.49 | 3.09 | 1.83 | 1.44 | 7.65
| Analytical results for Li in OREAS 267 (Certified Value 55 ppm) | [18 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 53.3 | 81.6867 | 53.5681 | 56 | 56 | 55.6 | 56.9 | 56 | 54.3 | 56.8 | 45.0998 | 53.2 | 50 | 56.2 | 56.59 | 57.2 |  | 1
| 1 | 2 | 52.9 | 81.0414 | 53.4705 | 53 | 55 | 56 | 60.4 | 56 | 55.9 | 56.5 | 45.7347 | 53.6 | 51 | 55.8 | 58.11 | 57.1 |  | #N/A
| 1 | 3 | 52.2 | 81.0888 | 56.1833 | 58 | 56 | 55.2 | 58.7 | 56.3 | 55.3 | 57.1 | 48.629 | 54 | 50 | 55.9 | 58.35 | 56.4 |  | 16
| 1 | 4 | 52.4 | 81.0853 | 55.2461 | 53 | 54 | 55.5 | 57.5 | 56.4 | 55.2 | 55.2 | 45.9411 | 54.6 | 51 | 56 | 58.23 | 57.7 |  | 55.367
| 1 | 5 | 52.3 | 81.8765 | 53.6049 | 54 | 55 | 56.5 | 58.1 | 56.6 | 55.1 | 55.8 | 47.0752 | 52.2 | 53 | 55.7 | 56.87 | 58.5 |  | 38
| 1 | 6 | 52.8 | 81.2186 | 54.179 | 56 | 56 | 57.2 | 58 | 56.1 | 55.7 | 56.2 | 47.1021 | 53 | 52 | 55.8 | 57.53 | 59
| Mean |  | 52.65 | 81.3329 | 54.3753 | 55 | 55.3333 | 56 | 58.2667 | 56.2333 | 55.25 | 56.2667 | 46.597 | 53.4333 | 51.1667 | 55.9 | 57.6133 | 57.65
| Median |  | 52.6 | 81.1537 | 53.892 | 55 | 55.5 | 55.8 | 58.05 | 56.2 | 55.25 | 56.35 | 46.5082 | 53.4 | 51 | 55.85 | 57.82 | 57.45
| Std Dev. |  | 0.42308 | 0.35768 | 1.1072 | 2 | 0.8165 | 0.74027 | 1.2078 | 0.24221 | 0.55767 | 0.69186 | 1.2671 | 0.83347 | 1.169 | 0.17889 | 0.74516 | 0.96073
| Rel.Std.Dev. |  | 0.8036% | 0.4398% | 2.036% | 3.636% | 1.476% | 1.322% | 2.073% | 0.4307% | 1.009% | 1.23% | 2.719% | 1.56% | 2.285% | 0.32% | 1.293% | 1.666%
| PDM3 |  | -4.907% | 46.9% | -1.791% | -0.6629% | -0.06089% | 1.143% | 5.237% | 1.565% | -0.2114% | 1.625% | -15.84% | -3.493% | -7.586% | 0.9626% | 4.057% | 4.123%
| Z-Score (Absolute) |  | 1.24 | 10.71 | 0.52 | 0.26 | 0.12 | 0.16 | 1.1 | 0.26 | 0.15 | 0.27 | 3.76 | 0.91 | 1.85 | 0.12 | 0.83 | 0.85
| Analytical results for Lu in OREAS 267 (Certified Value 0.2 ppm) | [18 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.06 | 1.07 | 1.09 | 1.17 | 1.19 | 1.20 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.20496 | 0.2 | 0.18 | 0.10406 | 0.19 | 0.21 | < 1 | [10 empty] | 1
| 1 | 2 | 0.20104 | 0.2 | 0.21 | 0.098413 | 0.21 | 0.2 | < 1 | [10 empty] | 6
| 1 | 3 | 0.19616 | 0.2 | 0.21 | 0.10228 | 0.21 | 0.21 | < 1 | [10 empty] | 16
| 1 | 4 | 0.20053 | 0.2 | 0.2 | 0.10698 | 0.2 | 0.21 | < 1 | [10 empty] | 0.20256
| 1 | 5 | 0.21287 | 0.2 | 0.2 | 0.10742 | 0.2 | 0.26 | < 1 | [10 empty] | 39
| 1 | 6 | 0.20632 | 0.2 | 0.2 | 0.10577 | 0.2 | 0.23 | < 1
| Mean |  | 0.20364 | 0.2 | 0.2 | 0.10415 | 0.20167 | 0.22
| Median |  | 0.203 | 0.2 | 0.2 | 0.10491 | 0.2 | 0.21
| Std Dev. |  | 0.0057696 | 3.0405e-17 | 0.010954 | 0.0033973 | 0.0075277 | 0.021909
| Rel.Std.Dev. |  | 2.833% | 1.52e-14% | 5.477% | 3.262% | 3.733% | 9.959%
| PDM3 |  | 0.5344% | -1.265% | -1.265% | -48.58% | -0.4421% | 8.609%
| Z-Score (Absolute) |  | 0.67 | 0.57 | 0.57 | 33.24 | 0 | 6.25 | 101.7
| Analytical results for Mg in OREAS 267 (Certified Value 0.929 wt.%) | [18 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.9162 | 0.96967 | 0.94248 | 0.98 | 0.94 | 0.93 | 0.94 | 0.88 | 0.91 | 0.91 | 1.0559 | 0.919 | 0.85 | 0.9382 | 0.9342 | 0.9 | 0.98 | 1
| 1 | 2 | 0.8969 | 0.96551 | 0.94978 | 0.97 | 0.87 | 0.95 | 1 | 0.9 | 0.92 | 0.9 | 1.0554 | 0.909 | 0.87 | 0.9279 | 0.9318 | 0.9 | 0.97 | #N/A
| 1 | 3 | 0.9143 | 0.9692 | 0.95294 | 0.96 | 0.97 | 0.93 | 0.97 | 0.88 | 0.92 | 0.91 | 1.0122 | 0.908 | 0.85 | 0.9414 | 0.9437 | 0.91 | 0.98 | 16
| 1 | 4 | 0.9059 | 0.96859 | 0.96254 | 0.97 | 0.88 | 0.93 | 0.95 | 0.88 | 0.91 | 0.89 | 0.98855 | 0.92 | 0.88 | 0.9365 | 0.9254 | 0.91 | 0.98 | 0.92851
| 1 | 5 | 0.9302 | 0.96657 | 0.97067 | 0.95 | 0.92 | 0.94 | 0.97 | 0.89 | 0.92 | 0.9 | 1.0447 | 0.914 | 0.87 | 0.9346 | 0.9362 | 0.89 | 0.97 | 40
| 1 | 6 | 0.9266 | 0.96768 | 0.96208 | 0.98 | 0.92 | 0.96 | 0.96 | 0.89 | 0.93 | 0.9 | 1.0526 | 0.903 | 0.87 | 0.9097 | 0.9369 | 0.9 | 0.97
| Mean |  | 0.91502 | 0.96787 | 0.95675 | 0.96833 | 0.91667 | 0.94 | 0.965 | 0.88667 | 0.91833 | 0.90167 | 1.0349 | 0.91217 | 0.865 | 0.93138 | 0.9347 | 0.90167 | 0.975
| Median |  | 0.91525 | 0.96814 | 0.95751 | 0.97 | 0.92 | 0.935 | 0.965 | 0.885 | 0.92 | 0.9 | 1.0487 | 0.9115 | 0.87 | 0.93555 | 0.9352 | 0.9 | 0.975
| Std Dev. |  | 0.012474 | 0.0016021 | 0.010222 | 0.01169 | 0.037238 | 0.012649 | 0.020736 | 0.008165 | 0.0075277 | 0.0075277 | 0.02804 | 0.0066758 | 0.012247 | 0.011538 | 0.0060538 | 0.0075277 | 0.0054772
| Rel.Std.Dev. |  | 1.363% | 0.1655% | 1.068% | 1.207% | 4.062% | 1.346% | 2.149% | 0.9209% | 0.8197% | 0.8349% | 2.709% | 0.7319% | 1.416% | 1.239% | 0.6477% | 0.8349% | 0.5618%
| PDM3 |  | -1.454% | 4.239% | 3.041% | 4.289% | -1.276% | 1.237% | 3.93% | -4.507% | -1.096% | -2.891% | 11.46% | -1.761% | -6.84% | 0.3091% | 0.6663% | -2.891% | 5.007%
| Z-Score (Absolute) |  | 0.37 | 0.83 | 0.58 | 0.84 | 0.33 | 0.2 | 0.76 | 1.01 | 0.3 | 0.67 | 2.35 | 0.44 | 1.51 | 0 | 0.08 | 0.67 | 0.99
| Analytical results for Mn in OREAS 267 (Certified Value 0.028 wt.%) | [18 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0278 | 0.027508 | 0.028187 | 0.0287 | 0.0283 | 0.0271 | 0.0277 | 0.0274 | 0.0272 | 0.0272 | 0.026748 | 0.0275 | 0.0286 | 0.0278 | 0.0272 | 0.02493 | 0.0296 | 1
| 1 | 2 | 0.0272 | 0.027549 | 0.02845 | 0.0283 | 0.0286 | 0.027 | 0.0295 | 0.0278 | 0.0273 | 0.0276 | 0.026575 | 0.0274 | 0.0282 | 0.0277 | 0.0267 | 0.02454 | 0.0297 | 25
| 1 | 3 | 0.0278 | 0.02773 | 0.028496 | 0.028 | 0.0278 | 0.0268 | 0.0286 | 0.0275 | 0.0274 | 0.0278 | 0.027261 | 0.027 | 0.0276 | 0.0278 | 0.0273 | 0.02512 | 0.0296 | 16
| 1 | 4 | 0.0275 | 0.027726 | 0.02832 | 0.0283 | 0.0282 | 0.027 | 0.0277 | 0.0279 | 0.0274 | 0.0269 | 0.027243 | 0.0268 | 0.0285 | 0.0276 | 0.0271 | 0.02492 | 0.0295 | 0.027838
| 1 | 5 | 0.028 | 0.027743 | 0.028855 | 0.0272 | 0.028 | 0.0271 | 0.0283 | 0.0277 | 0.0275 | 0.0275 | 0.028055 | 0.0273 | 0.0291 | 0.028 | 0.027 | 0.02467 | 0.0295 | 41
| 1 | 6 | 0.028 | 0.027804 | 0.028527 | 0.0282 | 0.0269 | 0.0275 | 0.0282 | 0.028 | 0.0276 | 0.0273 | 0.028983 | 0.0272 | 0.0276 | 0.0277 | 0.0271 | 0.0249 | 0.0296
| Mean |  | 0.027717 | 0.027677 | 0.028473 | 0.028117 | 0.027967 | 0.027083 | 0.028333 | 0.027717 | 0.0274 | 0.027383 | 0.027478 | 0.0272 | 0.028267 | 0.027767 | 0.027067 | 0.024847 | 0.029583
| Median |  | 0.0278 | 0.027728 | 0.028473 | 0.02825 | 0.0281 | 0.02705 | 0.02825 | 0.02775 | 0.0274 | 0.0274 | 0.027252 | 0.02725 | 0.02835 | 0.02775 | 0.0271 | 0.02491 | 0.0296
| Std Dev. |  | 0.00031252 | 0.00011901 | 0.00022593 | 0.00050365 | 0.00058878 | 0.00023166 | 0.00067132 | 0.00023166 | 0.00014142 | 0.00031885 | 0.00089983 | 0.00026077 | 0.00059217 | 0.00013663 | 0.00020656 | 0.00020743 | 7.5277e-05
| Rel.Std.Dev. |  | 1.128% | 0.43% | 0.7935% | 1.791% | 2.105% | 0.8554% | 2.369% | 0.8358% | 0.5161% | 1.164% | 3.275% | 0.9587% | 2.095% | 0.4921% | 0.7631% | 0.8348% | 0.2545%
| PDM3 |  | -0.4363% | -0.5805% | 2.279% | 1.001% | 0.4618% | -2.711% | 1.779% | -0.4363% | -1.574% | -1.634% | -1.295% | -2.292% | 1.539% | -0.2567% | -2.771% | -10.75% | 6.269%
| Z-Score (Absolute) |  | 0 | 0.07 | 1.27 | 0.67 | 0.42 | 1.07 | 1.04 | 0 | 0.53 | 0.56 | 0.4 | 0.87 | 0.93 | 0.08 | 1.1 | 4.84 | 3.15
| Analytical results for Mo in OREAS 267 (Certified Value 6.94 ppm) | [18 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 7.2 | < 10 | < 50 | 6 | 5.93 | 6.3 | 6.97 | 6.78 | 6.57 | 6.74 | 8.0943 | 6.7 | 6.98 | 7.1 | 7 | 3.88 | 7.8 | 1
| 1 | 2 | 7.1 | < 10 | < 50 | 7 | 5.63 | 6.57 | 7.37 | 6.88 | 6.35 | 6.51 | 8.078 | 6.8 | 7.05 | 7 | 6.6 | 4.75 | 7.6 | 26
| 1 | 3 | 7.3 | < 10 | < 50 | 7 | 5.65 | 6.06 | 7.33 | 6.91 | 6.83 | 6.33 | 8.1128 | 6.9 | 7.07 | 7.2 | 6.8 | 4.7 | 7.5 | 16
| 1 | 4 | 7.2 | < 10 | < 50 | 6 | 5.64 | 6.62 | 6.98 | 7.02 | 6.69 | 6.67 | 8.1854 | 6.8 | 7.25 | 7 | 7 | 5.03 | 7.7 | 6.9371
| 1 | 5 | 7.3 | < 10 | < 50 | 7 | 5.39 | 6.42 | 7.19 | 6.92 | 6.71 | 6.81 | 8.2829 | 6.8 | 7.03 | 7.1 | 6.9 | 4.11 | 7.8 | 42
| 1 | 6 | 7.1 | < 10 | < 50 | 6 | 5.45 | 6.37 | 7.11 | 6.99 | 6.7 | 6.7 | 8.3353 | 6.9 | 6.97 | 7 | 6.8 | 3.81 | 7.1
| Mean |  | 7.2 |  |  | 6.5 | 5.615 | 6.39 | 7.1583 | 6.9167 | 6.6417 | 6.6267 | 8.1814 | 6.8167 | 7.0583 | 7.0667 | 6.85 | 4.38 | 7.5833
| Median |  | 7.2 |  |  | 6.5 | 5.635 | 6.395 | 7.15 | 6.915 | 6.695 | 6.685 | 8.1491 | 6.8 | 7.04 | 7.05 | 6.85 | 4.405 | 7.65
| Std Dev. |  | 0.089443 |  |  | 0.54772 | 0.18929 | 0.20179 | 0.17023 | 0.08501 | 0.16497 | 0.17626 | 0.10675 | 0.075277 | 0.10167 | 0.08165 | 0.15166 | 0.51178 | 0.26394
| Rel.Std.Dev. |  | 1.242% |  |  | 8.427% | 3.371% | 3.158% | 2.378% | 1.229% | 2.484% | 2.66% | 1.305% | 1.104% | 1.44% | 1.155% | 2.214% | 11.68% | 3.481%
| PDM3 |  | 3.789% |  |  | -6.301% | -19.06% | -7.887% | 3.189% | -0.2949% | -4.259% | -4.475% | 17.94% | -1.736% | 1.747% | 1.867% | -1.256% | -36.86% | 9.315%
| Z-Score (Absolute) |  | 0.72 | 4.29 | 41.29 | NA | 2.89 | 1.12 | 0.63 | 0.08 | 0.55 | 0.59 | 2.96 | 0.15 | 0.4 | 0.42 | 0.08 | 5.71 | 1.6
Results from laboratory 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 267 (Certified Value 0.505 wt.%) | [18 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.5019 | 0.54291 | 0.53237 | 0.53 | 0.51 | 0.51 | 0.51 | 0.5 | 0.49 | 0.5 | 0.59474 | 0.497 | 0.62 | 0.5084 | 0.4643 | 0.45 | 0.51 | 1
| 1 | 2 | 0.4989 | 0.54717 | 0.52047 | 0.52 | 0.52 | 0.51 | 0.53 | 0.5 | 0.5 | 0.5 | 0.59866 | 0.494 | 0.65 | 0.5028 | 0.4601 | 0.48 | 0.51 | #N/A
| 1 | 3 | 0.5027 | 0.54574 | 0.5466 | 0.53 | 0.51 | 0.5 | 0.52 | 0.5 | 0.49 | 0.51 | 0.57517 | 0.493 | 0.66 | 0.5126 | 0.4669 | 0.45 | 0.51 | 16
| 1 | 4 | 0.5031 | 0.54923 | 0.52872 | 0.52 | 0.5 | 0.51 | 0.51 | 0.51 | 0.49 | 0.5 | 0.57677 | 0.487 | 0.66 | 0.5093 | 0.4612 | 0.46 | 0.51 | 0.50469
| 1 | 5 | 0.5121 | 0.54763 | 0.52851 | 0.51 | 0.5 | 0.51 | 0.52 | 0.51 | 0.49 | 0.5 | 0.58823 | 0.498 | 0.66 | 0.5136 | 0.4634 | 0.47 | 0.51 | 43
| 1 | 6 | 0.5089 | 0.54787 | 0.52101 | 0.52 | 0.51 | 0.51 | 0.51 | 0.5 | 0.5 | 0.5 | 0.59551 | 0.49 | 0.63 | 0.5045 | 0.4606 | 0.47 | 0.5
| Mean |  | 0.5046 | 0.54676 | 0.52961 | 0.52167 | 0.50833 | 0.50833 | 0.51667 | 0.50333 | 0.49333 | 0.50167 | 0.58818 | 0.49317 | 0.64667 | 0.50853 | 0.46275 | 0.46333 | 0.50833
| Median |  | 0.5029 | 0.5474 | 0.52861 | 0.52 | 0.51 | 0.51 | 0.515 | 0.5 | 0.49 | 0.5 | 0.59148 | 0.4935 | 0.655 | 0.50885 | 0.4623 | 0.465 | 0.51
| Std Dev. |  | 0.0049067 | 0.002198 | 0.0095474 | 0.0075277 | 0.0075277 | 0.0040825 | 0.008165 | 0.005164 | 0.005164 | 0.0040825 | 0.010058 | 0.0041673 | 0.017512 | 0.0042884 | 0.0026113 | 0.012111 | 0.0040825
| Rel.Std.Dev. |  | 0.9724% | 0.402% | 1.803% | 1.443% | 1.481% | 0.8031% | 1.58% | 1.026% | 1.047% | 0.8138% | 1.71% | 0.845% | 2.708% | 0.8433% | 0.5643% | 2.614% | 0.8031%
| PDM3 |  | -0.01877% | 8.334% | 4.938% | 3.363% | 0.721% | 0.721% | 2.372% | -0.2697% | -2.251% | -0.6% | 16.54% | -2.284% | 28.13% | 0.7606% | -8.311% | -8.195% | 0.721%
| Z-Score (Absolute) |  | 0.19 | 1.94 | 1.08 | 0.67 | 0 | 0 | 0.42 | 0.25 | 0.76 | 0.34 | 4.04 | 0.77 | 7 | 0.01 | 2.31 | 2.28 | 0
| Analytical results for Nb in OREAS 267 (Certified Value 13.9 ppm) | [18 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 14.81 | 13.8413 | 13.3 | 12 | 14 | 14.9 | 13.5 | 14 | 14.7 | 20.4625 | 12.4 | 12.3 | 13.88 | 15.17 | 14.2 |  |  | 1
| 1 | 2 | 14.68 | 13.6766 | 13.6 | 12.8 | 14.8 | 15.9 | 13.5 | 13.4 | 14.4 | 20.2631 | 12.3 | 12.7 | 13.5 | 15.18 | 13.9 |  |  | 27
| 1 | 3 | 14.65 | 14.8045 | 13.6 | 11.3 | 13.3 | 15.7 | 13.7 | 14.1 | 14.2 | 20.3352 | 12.3 | 12.4 | 14.28 | 15.28 | 14 |  |  | 16
| 1 | 4 | 14.68 | 13.6213 | 13.5 | 13.6 | 13.8 | 15.1 | 13.8 | 14.1 | 14.8 | 18.82 | 12.5 | 12.8 | 14.05 | 15.51 | 14 |  |  | 13.8542
| 1 | 5 | 14.83 | 14.0313 | 13.4 | 12.4 | 14 | 15.1 | 13.6 | 14.2 | 14.8 | 20.3937 | 12.3 | 13 | 13.68 | 14.86 | 13.8 |  |  | 44
| 1 | 6 | 14.72 | 14.0818 | 13.8 | 11.8 | 14.5 | 15 | 13.3 | 14.4 | 14.8 | 20.6655 | 12.3 | 12.2 | 13.51 | 15.13 | 13.4
| Mean |  | 14.7283 | 14.0095 | 13.5333 | 12.3167 | 14.0667 | 15.2833 | 13.5667 | 14.0333 | 14.6167 | 20.1566 | 12.35 | 12.5667 | 13.8167 | 15.1883 | 13.8833
| Median |  | 14.7 | 13.9363 | 13.55 | 12.2 | 14 | 15.1 | 13.55 | 14.1 | 14.75 | 20.3644 | 12.3 | 12.55 | 13.78 | 15.175 | 13.95
| Std Dev. |  | 0.074677 | 0.43069 | 0.17512 | 0.81097 | 0.52789 | 0.41191 | 0.17512 | 0.33862 | 0.25626 | 0.6691 | 0.083666 | 0.31411 | 0.31181 | 0.21142 | 0.27142
| Rel.Std.Dev. |  | 0.507% | 3.074% | 1.294% | 6.584% | 3.753% | 2.695% | 1.291% | 2.413% | 1.753% | 3.32% | 0.6775% | 2.5% | 2.257% | 1.392% | 1.955%
| PDM3 |  | 6.309% | 1.12% | -2.316% | -11.1% | 1.533% | 10.32% | -2.076% | 1.293% | 5.503% | 45.49% | -10.86% | -9.294% | -0.2713% | 9.629% | 0.2099%
| Z-Score (Absolute) |  | 0.8 | 0 | 0.53 | 1.88 | 0.06 | 1.41 | 0.49 | 0.03 | 0.67 | 6.83 | 1.84 | 1.6 | 0.21 | 1.31 | 0.14
| Analytical results for Nd in OREAS 267 (Certified Value 26.7 ppm) | [18 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | 1.21 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 1
2
| 1 | 1 | 27.5835 | 26.1151 | 25 | 27.1 | 20.1443 | 25.87 | 23 | [10 empty] | 1
| 1 | 2 | 27.5991 | 26.3398 | 25.9 | 27.1 | 19.1847 | 26.56 | 19 | [10 empty] | 7
| 1 | 3 | 27.5745 | 27.7571 | 26.1 | 25.9 | 20.1682 | 26.75 | 19 | [10 empty] | 16
| 1 | 4 | 27.6058 | 26.7948 | 25.3 | 27.6 | 20.0053 | 26.51 | 20 | [10 empty] | 26.6693
| 1 | 5 | 27.6356 | 27.2756 | 25 | 27 | 21.6822 | 27.12 | 18 | [10 empty] | 45
| 1 | 6 | 27.5726 | 26.9869 | 25.4 | 27.1 | 21.1786 | 25.93 | 21
| Mean |  | 27.5952 | 26.8782 | 25.45 | 26.9667 | 20.3939 | 26.4567 | 20
| Median |  | 27.5913 | 26.8908 | 25.35 | 27.1 | 20.1563 | 26.535 | 19.5
| Std Dev. |  | 0.023847 | 0.60354 | 0.45935 | 0.56451 | 0.89465 | 0.4819 | 1.7889
| Rel.Std.Dev. |  | 0.08642% | 2.245% | 1.805% | 2.093% | 4.387% | 1.821% | 8.944%
| PDM3 |  | 3.472% | 0.7831% | -4.572% | 1.115% | -23.53% | -0.7975% | -25.01%
| Z-Score (Absolute) |  | 0.76 | 0.28 | 0.67 | 0.34 | 4.06 | 0 | 4.32
| Analytical results for Ni in OREAS 267 (Certified Value 48.9 ppm) | [18 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 49.9 | 44.59 | 48.2857 | 49 | 47.4 | 50.2 | 50.1 | 50.4 | 50.3 | 48.6 | 55.2258 | 50 | 47.8 | 49.6 | 51.8 | 44.95 | 51.2 | 1
| 1 | 2 | 49.4 | 44.55 | 49.605 | 45 | 46.9 | 52.1 | 52.8 | 51 | 48.8 | 47 | 57.1389 | 49 | 48.8 | 49.1 | 51.7 | 43.38 | 49.6 | 29
| 1 | 3 | 50 | 45.29 | 50.9103 | 49 | 46.7 | 48.4 | 53 | 51.3 | 51.9 | 46.6 | 57.4275 | 49 | 48 | 49.8 | 52.3 | 43.26 | 50.8 | 16
| 1 | 4 | 49.2 | 43.74 | 49.4465 | 47 | 47.4 | 50.2 | 50.8 | 52.3 | 51.7 | 48.6 | 55.2083 | 49 | 48.3 | 49.5 | 51.3 | 44.82 | 49.5 | 48.9097
| 1 | 5 | 50.5 | 46.05 | 49.7387 | 47 | 45.8 | 49.8 | 50 | 51.8 | 51.2 | 47.4 | 62.6516 | 49 | 47.5 | 50.6 | 52.1 | 42.86 | 48.2 | 46
| 1 | 6 | 50.2 | 44.32 | 50.4667 | 49 | 45.5 | 51 | 50.9 | 51.1 | 51 | 48.3 | 56.2298 | 49 | 47.8 | 50.2 | 51.8 | 44.77 | 48.5
| Mean |  | 49.8667 | 44.7567 | 49.7422 | 47.6667 | 46.6167 | 50.2833 | 51.2667 | 51.3167 | 50.8167 | 47.75 | 57.3136 | 49.1667 | 48.0333 | 49.8 | 51.8333 | 44.0067 | 49.6333
| Median |  | 49.95 | 44.57 | 49.6718 | 48 | 46.8 | 50.2 | 50.85 | 51.2 | 51.1 | 47.85 | 56.6843 | 49 | 47.9 | 49.7 | 51.8 | 44.075 | 49.55
| Std Dev. |  | 0.48854 | 0.80647 | 0.90742 | 1.633 | 0.80353 | 1.2336 | 1.3171 | 0.66156 | 1.1374 | 0.8666 | 2.775 | 0.40825 | 0.45898 | 0.53292 | 0.34448 | 0.93799 | 1.1978
| Rel.Std.Dev. |  | 0.9797% | 1.802% | 1.824% | 3.426% | 1.724% | 2.453% | 2.569% | 1.289% | 2.238% | 1.815% | 4.842% | 0.8303% | 0.9556% | 1.07% | 0.6646% | 2.131% | 2.413%
| PDM3 |  | 1.957% | -8.491% | 1.702% | -2.542% | -4.688% | 2.808% | 4.819% | 4.921% | 3.899% | -2.371% | 17.18% | 0.5254% | -1.792% | 1.82% | 5.978% | -10.02% | 1.479%
| Z-Score (Absolute) |  | 0.05 | 2.14 | 0 | 0.89 | 1.34 | 0.23 | 0.65 | 0.67 | 0.46 | 0.85 | 3.24 | 0.25 | 0.73 | 0.02 | 0.9 | 2.46 | 0.05
| Analytical results for P in OREAS 267 (Certified Value 0.048 wt.%) | [18 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0469 | 0.051276 | 0.047272 | 0.043 | 0.0493 | 0.047 | 0.048 | 0.046 | 0.047 | 0.047 | 0.058543 | 0.048 | 0.0446 | 0.0483 | 0.0485 | 0.045 | 0.047 | 1
| 1 | 2 | 0.0463 | 0.05197 | 0.048033 | 0.043 | 0.0486 | 0.048 | 0.051 | 0.047 | 0.048 | 0.048 | 0.0569 | 0.048 | 0.0467 | 0.0479 | 0.0484 | 0.05 | 0.047 | #N/A
| 1 | 3 | 0.0464 | 0.052123 | 0.050282 | 0.046 | 0.0494 | 0.047 | 0.049 | 0.047 | 0.047 | 0.048 | 0.057082 | 0.047 | 0.0442 | 0.048 | 0.0484 | 0.048 | 0.047 | 16
| 1 | 4 | 0.0466 | 0.051443 | 0.048396 | 0.041 | 0.0485 | 0.047 | 0.048 | 0.047 | 0.047 | 0.046 | 0.058156 | 0.048 | 0.0468 | 0.0487 | 0.0479 | 0.046 | 0.048 | 0.0476
| 1 | 5 | 0.0474 | 0.051879 | 0.049372 | 0.043 | 0.0488 | 0.048 | 0.049 | 0.047 | 0.047 | 0.048 | 0.060233 | 0.047 | 0.0471 | 0.0489 | 0.048 | 0.046 | 0.047 | 47
| 1 | 6 | 0.0471 | 0.051752 | 0.049744 | 0.043 | 0.0492 | 0.049 | 0.048 | 0.046 | 0.048 | 0.047 | 0.059305 | 0.047 | 0.0454 | 0.0476 | 0.0484 | 0.047 | 0.047
| Mean |  | 0.046783 | 0.051741 | 0.04885 | 0.043167 | 0.048967 | 0.047667 | 0.048833 | 0.046667 | 0.047333 | 0.047333 | 0.05837 | 0.0475 | 0.0458 | 0.048233 | 0.048267 | 0.047 | 0.047167
| Median |  | 0.04675 | 0.051816 | 0.048884 | 0.043 | 0.049 | 0.0475 | 0.0485 | 0.047 | 0.047 | 0.0475 | 0.05835 | 0.0475 | 0.04605 | 0.04815 | 0.0484 | 0.0465 | 0.047
| Std Dev. |  | 0.00042622 | 0.00032357 | 0.001139 | 0.0016021 | 0.00038297 | 0.0008165 | 0.001169 | 0.0005164 | 0.0005164 | 0.0008165 | 0.0012836 | 0.00054772 | 0.0012377 | 0.00049666 | 0.00025033 | 0.0017889 | 0.00040825
| Rel.Std.Dev. |  | 0.9111% | 0.6254% | 2.332% | 3.711% | 0.7821% | 1.713% | 2.394% | 1.107% | 1.091% | 1.725% | 2.199% | 1.153% | 2.702% | 1.03% | 0.5186% | 3.806% | 0.8655%
| PDM3 |  | -1.716% | 8.699% | 2.626% | -9.314% | 2.871% | 0.1401% | 2.591% | -1.961% | -0.5602% | -0.5602% | 22.63% | -0.2101% | -3.781% | 1.331% | 1.401% | -1.26% | -0.9103%
| Z-Score (Absolute) |  | 0.63 | 3.73 | 1.19 | 3.81 | 1.29 | 0.15 | 1.17 | 0.73 | 0.15 | 0.15 | 9.56 | 0 | 1.5 | 0.64 | 0.67 | 0.44 | 0.29
| Analytical results for Pb in OREAS 267 (Certified Value 15.2 ppm) | [18 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 15.1 | 10.031 | 16.768 | 13 | 19.2 | 14.8 | 14.7 | 14.9 | 15.4 | 14.4 | 16.288 | 15 | 16 | 16.7 | 19 | 19 | 11.3 | 1
| 1 | 2 | 15.1 | 9.837 | 17.5011 | 14 | 19.7 | 14.6 | 15.4 | 15 | 14.9 | 14.2 | 16.0209 | 15 | 16.9 | 16.5 | 20 | 20.3 | 11.7 | 30
| 1 | 3 | 15.4 | 8.8025 | 17.1142 | 14 | 20.7 | 14.5 | 15 | 15.2 | 16 | 14 | 16.5943 | 15 | 16.2 | 16.7 | 23 | 21.2 | 12.1 | 16
| 1 | 4 | 15.1 | 10.976 | 17.3351 | 12 | 19.7 | 14.6 | 14.5 | 15.5 | 15.5 | 13.9 | 16.2282 | 15 | 16.2 | 16.7 | 19 | 22.8 | 11.3 | 15.2088
| 1 | 5 | 15.7 | 9.308 | 16.9056 | 14 | 19 | 15 | 15 | 15 | 15.4 | 14.1 | 14.8585 | 15 | 16.6 | 16.2 | 19 | 22.4 | 11.5 | 48
| 1 | 6 | 15.1 | 9.596 | 18.1917 | 13 | 19.5 | 14.9 | 14.6 | 14.7 | 16.4 | 13.8 | 15.4155 | 15 | 16.8 | 16.4 | 18 | 22.7 | 11.9
| Mean |  | 15.25 | 9.7584 | 17.3026 | 13.3333 | 19.6333 | 14.7333 | 14.8667 | 15.05 | 15.6 | 14.0667 | 15.9009 | 15 | 16.45 | 16.5333 | 19.6667 | 21.4 | 11.6333
| Median |  | 15.1 | 9.7165 | 17.2247 | 13.5 | 19.6 | 14.7 | 14.85 | 15 | 15.45 | 14.05 | 16.1245 | 15 | 16.4 | 16.6 | 19 | 21.8 | 11.6
| Std Dev. |  | 0.251 | 0.73567 | 0.51183 | 0.8165 | 0.59217 | 0.19664 | 0.33267 | 0.27386 | 0.52536 | 0.21602 | 0.64411 | 0 | 0.36742 | 0.20656 | 1.7512 | 1.5271 | 0.3266
| Rel.Std.Dev. |  | 1.646% | 7.539% | 2.958% | 6.124% | 3.016% | 1.335% | 2.238% | 1.82% | 3.368% | 1.536% | 4.051% | 0% | 2.234% | 1.249% | 8.904% | 7.136% | 2.807%
| PDM3 |  | 0.2709% | -35.84% | 13.77% | -12.33% | 29.09% | -3.126% | -2.25% | -1.044% | 2.572% | -7.51% | 4.551% | -1.373% | 8.161% | 8.709% | 29.31% | 40.71% | -23.51%
| Z-Score (Absolute) |  | 0.11 | 3.58 | 1.19 | NA | 2.66 | 0.44 | 0.35 | 0.24 | 0.11 | 0.86 | 0.3 | NA | 0.65 | 0.7 | NA | 3.78 | 2.4
Results from laboratories 7, 18 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 267 (Certified Value 6.95 ppm) | [18 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 6.3696 | 6.9951 | 7.5 | 6.88 | 5.3461 | 7.1 | [11 empty] | 1
| 1 | 2 | 6.3516 | 6.919 | 7.2 | 6.94 | 5.1502 | 7.19 | [11 empty] | 8
| 1 | 3 | 6.3702 | 7.2863 | 7.5 | 6.82 | 5.5992 | 7.25 | [11 empty] | 16
| 1 | 4 | 6.3543 | 6.9181 | 7 | 6.89 | 5.2904 | 7.26 | [11 empty] | 6.9529
| 1 | 5 | 6.4048 | 7.1256 | 6.9 | 6.99 | 5.6796 | 7.1 | [11 empty] | 49
| 1 | 6 | 6.4093 | 7.1642 | 7.4 | 7.04 | 5.6308 | 6.96
| Mean |  | 6.3766 | 7.068 | 7.25 | 6.9267 | 5.4494 | 7.1433
| Median |  | 6.3699 | 7.0604 | 7.3 | 6.915 | 5.4727 | 7.145
| Std Dev. |  | 0.024802 | 0.14849 | 0.25884 | 0.079917 | 0.21626 | 0.11361
| Rel.Std.Dev. |  | 0.389% | 2.101% | 3.57% | 1.154% | 3.969% | 1.59%
| PDM3 |  | -8.289% | 1.656% | 4.273% | -0.3778% | -21.62% | 2.738%
| Z-Score (Absolute) |  | 2.1 | 0.24 | 0.85 | 0.24 | 5.24 | 0.49
| Analytical results for Rb in OREAS 267 (Certified Value 71 ppm) | [18 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 71.62 | 66.4958 | 68.2187 | 71.5 | 74 | 69.4 | 71 | 72.3 | 75.1 | 72.5 | 74.3892 | 67.3 | 70.7 | 71.54 | 68.36 | 84.9 |  | 1
| 1 | 2 | 70.34 | 66.5208 | 70.3828 | 71.5 | 73 | 71 | 76.7 | 73.5 | 71 | 70 | 74.199 | 67.2 | 73.4 | 70.24 | 69.87 | 86.2 |  | 32
| 1 | 3 | 71.98 | 67.9718 | 73.0703 | 71.9 | 71 | 66.4 | 73.1 | 73.7 | 75.4 | 69.4 | 68.7519 | 68.3 | 70.6 | 70.87 | 68.06 | 86.1 |  | 16
| 1 | 4 | 70.86 | 67.8439 | 69.7069 | 72.9 | 73 | 69.9 | 72.3 | 73.8 | 76.9 | 70.9 | 63.9069 | 69.4 | 71 | 71.31 | 68.47 | 85.4 |  | 71.0678
| 1 | 5 | 72.61 | 66.3683 | 70.4266 | 69.8 | 73 | 68.4 | 73.7 | 73.6 | 74.7 | 70.7 | 75.8523 | 66.6 | 73 | 71.18 | 70.41 | 84.3 |  | 50
| 1 | 6 | 72.11 | 67.8642 | 72.2902 | 71.2 | 71 | 70.6 | 72.1 | 73.2 | 76.2 | 71.9 | 72.4624 | 66.3 | 70.8 | 70.77 | 70.98 | 85
| Mean |  | 71.5867 | 67.1775 | 70.6826 | 71.4667 | 72.5 | 69.2833 | 73.15 | 73.35 | 74.8833 | 70.9 | 71.5936 | 67.5167 | 71.5833 | 70.985 | 69.3583 | 85.3167
| Median |  | 71.8 | 67.1823 | 70.4047 | 71.5 | 73 | 69.65 | 72.7 | 73.55 | 75.25 | 70.8 | 73.3307 | 67.25 | 70.9 | 71.025 | 69.17 | 85.2
| Std Dev. |  | 0.8437 | 0.78706 | 1.7586 | 1.0093 | 1.2247 | 1.6833 | 1.9675 | 0.55408 | 2.0605 | 1.1541 | 4.4868 | 1.1514 | 1.2656 | 0.46177 | 1.2222 | 0.73598
| Rel.Std.Dev. |  | 1.179% | 1.172% | 2.488% | 1.412% | 1.689% | 2.43% | 2.69% | 0.7554% | 2.752% | 1.628% | 6.267% | 1.705% | 1.768% | 0.6505% | 1.762% | 0.8626%
| PDM3 |  | 0.7301% | -5.474% | -0.5421% | 0.5613% | 2.015% | -2.511% | 2.93% | 3.211% | 5.369% | -0.2361% | 0.7399% | -4.997% | 0.7254% | -0.1165% | -2.405% | 20.05%
| Z-Score (Absolute) |  | 0.03 | 2.26 | 0.44 | 0.03 | 0.51 | 1.16 | 0.84 | 0.95 | 1.74 | 0.32 | 0.04 | 2.08 | 0.03 | 0.28 | 1.12 | 7.15
| Analytical results for Re in OREAS 267 (Certified Value < 0.002 ppm) | [18 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.21 | 1.24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | < 0.002 | < 0.01 | < 0.001 | 0.008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.001544 | < 0.05 | < 0.003 | < 0.005 |  |  |  |  | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.001 | 0.008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.00051462 | < 0.05 | < 0.003 | < 0.005 |  |  |  |  | 33
| 1 | 3 | < 0.002 | < 0.01 | < 0.001 | 0.008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.00064327 | < 0.05 | < 0.003 | < 0.005 |  |  |  |  | 16
| 1 | 4 | < 0.002 | < 0.01 | < 0.001 | 0.008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.00038596 | < 0.05 | < 0.003 | < 0.005 |  |  |  |  | < 0.002
| 1 | 5 | < 0.002 | < 0.01 | < 0.001 | 0.008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.0011579 | < 0.05 | < 0.003 | < 0.005 |  |  |  |  | 51
| 1 | 6 | < 0.002 | < 0.01 | < 0.001 | 0.011 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.00064327 | < 0.05 | < 0.003 | < 0.005
| Mean |  |  |  | 0.001 | 0.0085 | [5 empty] | 0.00081483
| Median |  |  |  | 0.001 | 0.008 | [5 empty] | 0.00064327
| Std Dev. |  |  |  |  | 0.0012247 | [5 empty] | 0.00044323
| Rel.Std.Dev. |  |  |  |  | 14.41% | [5 empty] | 54.4%
PDM3
| Z-Score (Absolute) |  | 0 | 14.57 | 1.52 | 27.31 | 0 | 0 | 0 | 0 | 0 | 0.67 | 87.4 | 1.82 | 5.46
| Analytical results for S in OREAS 267 (Certified Value 0.024 wt.%) | [18 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0236 | 0.022963 | 0.025063 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.03 | 0.023 | 0.03 | 0.0238 | 0.0251 | 0.02 | < 0.05 |  | 1
| 1 | 2 | 0.0237 | 0.022808 | 0.024695 | 0.02 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.03 | 0.023 | 0.04 | 0.0238 | 0.0253 | 0.02 | < 0.05 |  | 14
| 1 | 3 | 0.0237 | 0.022646 | 0.026582 | 0.02 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.03 | 0.023 | 0.03 | 0.0236 | 0.0246 | 0.02 | < 0.05 |  | 16
| 1 | 4 | 0.0238 | 0.023114 | 0.024501 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 | 0.023 | 0.04 | 0.0244 | 0.0241 | 0.02 | < 0.05 |  | 0.024426
| 1 | 5 | 0.0237 | 0.023033 | 0.024781 | 0.02 | 0.03 | 0.02 | 0.03 | 0.02 | 0.03 | 0.03 | 0.023 | 0.04 | 0.025 | 0.032 | 0.02 | < 0.05 |  | 52
| 1 | 6 | 0.0235 | 0.022911 | 0.025043 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.023 | 0.04 | 0.0237 | 0.0268 | 0.01 | < 0.05
| Mean |  | 0.023667 | 0.022912 | 0.025111 | 0.02 | 0.03 | 0.026667 | 0.025 | 0.02 | 0.03 | 0.026667 | 0.023 | 0.036667 | 0.02405 | 0.026317 | 0.018333
| Median |  | 0.0237 | 0.022937 | 0.024912 | 0.02 | 0.03 | 0.03 | 0.025 | 0.02 | 0.03 | 0.03 | 0.023 | 0.04 | 0.0238 | 0.0252 | 0.02
| Std Dev. |  | 0.00010328 | 0.00016721 | 0.00075165 | 0 | 0 | 0.005164 | 0.0054772 | 0 | 0 | 0.005164 | 3.8006e-18 | 0.005164 | 0.00054314 | 0.0029294 | 0.0040825
| Rel.Std.Dev. |  | 0.4364% | 0.7298% | 2.993% | 0% | 0% | 19.36% | 21.91% | 0% | 0% | 19.36% | 1.652e-14% | 14.08% | 2.258% | 11.13% | 22.27%
| PDM3 |  | -3.107% | -6.195% | 2.806% | -18.12% | 22.82% | 9.175% | 2.351% | -18.12% | 22.82% | 9.175% | -5.837% | 50.12% | -1.538% | 7.742% | -24.94%
| Z-Score (Absolute) |  | 0.49 | 0.77 | 0.04 | 1.84 | 1.84 | 0.61 | 0 | 1.84 | 1.84 | 0.61 | 0.74 | 4.29 | 0.35 | 0.48 | 2.45 | 0
| Analytical results for Sb in OREAS 267 (Certified Value 5.15 ppm) | [18 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 2
| 1 | 1 | 8.3 | 7.7709 | < 10 | 7.9 | 3.39 | 4.04 | 8.83 | 3.98 | 3.77 | 4.32 | 9.7211 | 5.2 | 4.81 | 5.1 | 4.32 | 7.6 | 4.7 | 1
| 1 | 2 | 7.96 | 7.8184 | < 10 | 9.4 | 3.32 | 4.19 | 6.67 | 3.97 | 3.54 | 4.24 | 9.4346 | 4.5 | 4.67 | 5.06 | 4.43 | 7.08 | 4.2 | 9
| 1 | 3 | 8.31 | 7.7749 | < 10 | 6.8 | 3.4 | 4.07 | 16.6 | 4.03 | 3.81 | 4.1 | 9.1766 | 4.3 | 4.77 | 7.94 | 4.25 | 6.02 | 4.7 | 16
| 1 | 4 | 10.42 | 7.7939 | < 10 | 8.8 | 3.24 | 4.19 | 6.49 | 4.19 | 3.81 | 4.26 | 9.4598 | 4.5 | 4.87 | 45.82 | 4.81 | 6.49 | 4.6 | 5.1463
| 1 | 5 | 8.43 | 7.8058 | < 10 | 9.1 | 3.33 | 4.03 | 22.9 | 4.01 | 3.85 | 4.21 | 8.8975 | 5.1 | 4.8 | 15.1 | 4.62 | 6.86 | 6.2 | 53
| 1 | 6 | 9.12 | 7.8256 | < 10 | 7.3 | 3.36 | 4.12 | 6.06 | 4.09 | 3.79 | 4.26 | 9.4422 | 4.6 | 4.84 | 6.03 | 4.09 | 6.39 | 3.9
| Mean |  | 8.7567 | 7.7983 |  | 8.2167 | 3.34 | 4.1067 | 11.2583 | 4.045 | 3.7617 | 4.2317 | 9.3553 | 4.7 | 4.7933 | 14.175 | 4.42 | 6.74 | 4.7167
| Median |  | 8.37 | 7.7999 |  | 8.35 | 3.345 | 4.095 | 7.75 | 4.02 | 3.8 | 4.25 | 9.4384 | 4.55 | 4.805 | 6.985 | 4.375 | 6.675 | 4.65
| Std Dev. |  | 0.89979 | 0.022475 |  | 1.0458 | 0.05831 | 0.07174 | 6.9393 | 0.082885 | 0.11179 | 0.073869 | 0.28286 | 0.36332 | 0.069474 | 15.9557 | 0.26077 | 0.56089 | 0.79352
| Rel.Std.Dev. |  | 10.28% | 0.2882% |  | 12.73% | 1.746% | 1.747% | 61.64% | 2.049% | 2.972% | 1.746% | 3.024% | 7.73% | 1.449% | 112.6% | 5.9% | 8.322% | 16.82%
| PDM3 |  | 70.15% | 51.53% |  | 59.66% | -35.1% | -20.2% | 118.8% | -21.4% | -26.91% | -17.77% | 81.79% | -8.673% | -6.859% | 175.4% | -14.11% | 30.97% | -8.349%
| Z-Score (Absolute) |  | 2.59 | 1.96 | 0.14 | 2.24 | 0.95 | 0.45 | 4.23 | 0.49 | 0.67 | 0.37 | 2.98 | 0.06 | 0 | 6.13 | 0.24 | 1.27 | 0.05
| Analytical results for Sc in OREAS 267 (Certified Value 9.25 ppm) | [18 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 9.4 | 9.702 | 9.2861 | 9 | 8.5 | 8.7 | 9.4 | 9.1 | 9.4 | 8.7 | 88.1544 | 9 | 8.5 | 9.2 | 8.4 | 10.6 |  | 1
| 1 | 2 | 9.4 | 9.691 | 9.2499 | 9 | 9 | 9.3 | 10.1 | 9.3 | 9.2 | 8.5 | 92.2629 | 9 | 8.9 | 9.1 | 8.44 | 10.7 |  | 35
| 1 | 3 | 9.4 | 9.768 | 9.9597 | 10 | 8.6 | 8.6 | 9.9 | 9.4 | 9.6 | 8.5 | 92.8231 | 9 | 8.3 | 9.3 | 8.52 | 10.6 |  | 16
| 1 | 4 | 9.2 | 9.779 | 9.2143 | 9 | 8.8 | 9 | 9.6 | 9.3 | 10 | 8.7 | 90.5533 | 9 | 8.1 | 9.3 | 8.4 | 10.6 |  | 9.2516
| 1 | 5 | 9.4 | 9.691 | 9.5297 | 9 | 9.3 | 8.9 | 9.5 | 9.4 | 9.6 | 8.7 | 102.1865 | 9 | 8.6 | 9.2 | 8.37 | 10.5 |  | 54
| 1 | 6 | 9.4 | 10.263 | 9.3596 | 10 | 9.2 | 9 | 9.8 | 9.5 | 9.6 | 8.9 | 100.982 | 9 | 8.9 | 9 | 8.37 | 10.6
| Mean |  | 9.3667 | 9.8157 | 9.4332 | 9.3333 | 8.9 | 8.9167 | 9.7167 | 9.3333 | 9.5667 | 8.6667 | 94.4937 | 9 | 8.55 | 9.1833 | 8.4167 | 10.6
| Median |  | 9.4 | 9.735 | 9.3228 | 9 | 8.9 | 8.95 | 9.7 | 9.35 | 9.6 | 8.7 | 92.543 | 9 | 8.55 | 9.2 | 8.4 | 10.6
| Std Dev. |  | 0.08165 | 0.22259 | 0.2811 | 0.5164 | 0.32249 | 0.24833 | 0.26394 | 0.13663 | 0.26583 | 0.15055 | 5.7404 | 0 | 0.32094 | 0.1169 | 0.056804 | 0.063246
| Rel.Std.Dev. |  | 0.8717% | 2.268% | 2.98% | 5.533% | 3.623% | 2.785% | 2.716% | 1.464% | 2.779% | 1.737% | 6.075% | 0% | 3.754% | 1.273% | 0.6749% | 0.5967%
| PDM3 |  | 1.244% | 6.097% | 1.963% | 0.8835% | -3.8% | -3.62% | 5.027% | 0.8835% | 3.406% | -6.322% | 921.4% | -2.72% | -7.584% | -0.7379% | -9.025% | 14.57%
| Z-Score (Absolute) |  | 0.05 | 0.78 | 0.16 | NA | 0.7 | 0.67 | 0.62 | 0 | 0.38 | 1.08 | NA | NA | 1.27 | 0.24 | 1.48 | 2.05
Results from laboratories 1.07 and 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 267 (Indicative Value 0.75 ppm) | [18 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | < 0.5 | < 5 | 0.3 | 1.11 | 1 | < 1 | 2 | < 1 | < 1 | 1.8276 | < 1 | < 2 | < 0.05 | < 0.5 |  |  |  | 1
| 1 | 2 | < 0.5 | < 5 | 0.4 | 1.16 | 1 | < 1 | 2 | < 1 | < 1 | 2.0003 | < 1 | < 2 | < 0.05 | < 0.5 |  |  |  | 2
| 1 | 3 | < 0.5 | < 5 | 0.2 | 1.19 | 1 | < 1 | 3 | 1 | < 1 | 1.8124 | < 1 | < 2 | < 0.05 | < 0.5 |  |  |  | 16
| 1 | 4 | < 0.5 | < 5 | 0.3 | 1.09 | 1 | < 1 | 2 | 1 | < 1 | 1.812 | < 1 | < 2 | 0.82 | < 0.5 |  |  |  | 0.75478
| 1 | 5 | 0.6 | < 5 | 0.4 | 1.15 | 1 | 1 | 2 | < 1 | < 1 | 1.7494 | < 1 | < 2 | < 0.05 | < 0.5 |  |  |  | 8
| 1 | 6 | < 0.5 | < 5 | 0.4 | 1.06 | 1 | < 1 | 2 | < 1 | < 1 | 1.7398 | < 1 | < 2 | 0.82 | < 0.5
| Mean |  | 0.6 |  | 0.33333 | 1.1267 | 1 | 1 | 2.1667 | 1 |  | 1.8236 |  |  | 0.82
| Median |  | 0.6 |  | 0.35 | 1.13 | 1 | 1 | 2 | 1 |  | 1.8122 |  |  | 0.82
| Std Dev. |  |  |  | 0.08165 | 0.048442 | 0 |  | 0.40825 | 0 |  | 0.093829 |  |  | 0
| Rel.Std.Dev. |  |  |  | 24.49% | 4.3% | 0% |  | 18.84% | 0% |  | 5.145% |  |  | 0%
| PDM3 |  | -20.51% |  | -55.84% | 49.27% | 32.49% | 32.49% | 187.1% | 32.49% |  | 141.6% |  |  | 8.641%
| Z-Score (Absolute) |  | 0.6 | 3.56 | 0.55 | 0.95 | 0.71 | 0.08 | 2.93 | 0.08 | 0.24 | 2.28 | 0.24 | 0.71 | 0.64 | 0.71
| Analytical results for Sm in OREAS 267 (Certified Value 5.11 ppm) | [18 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 4.9207 | 5.1139 | 5.2 | 5.25 | 3.8143 | 4.98 | [11 empty] | 1
| 1 | 2 | 4.8868 | 5.0942 | 4.8 | 5.59 | 3.7444 | 5.06 | [11 empty] | 36
| 1 | 3 | 5.0588 | 5.2725 | 5.1 | 5.28 | 3.882 | 4.94 | [11 empty] | 16
| 1 | 4 | 4.9215 | 4.9208 | 5.2 | 5.28 | 3.9229 | 4.98 | [11 empty] | 5.1092
| 1 | 5 | 5.0572 | 5.1902 | 5.3 | 5.46 | 4.1321 | 4.93 | [11 empty] | 55
| 1 | 6 | 5.123 | 5.1489 | 4.9 | 5.34 | 4.1164 | 5.48
| Mean |  | 4.9947 | 5.1234 | 5.0833 | 5.3667 | 3.9354 | 5.0617
| Median |  | 4.9894 | 5.1314 | 5.15 | 5.31 | 3.9024 | 4.98
| Std Dev. |  | 0.096912 | 0.11776 | 0.19408 | 0.13261 | 0.15849 | 0.20999
| Rel.Std.Dev. |  | 1.94% | 2.298% | 3.818% | 2.471% | 4.027% | 4.149%
| PDM3 |  | -2.242% | 0.2778% | -0.5067% | 5.039% | -22.98% | -0.9308%
| Z-Score (Absolute) |  | 0.82 | 0.53 | 0.11 | 3.08 | 11.91 | 0.11
| Analytical results for Sn in OREAS 267 (Certified Value 5.95 ppm) | [18 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.3 | 5.9037 | 6 | 6.1 | 5.6 | 5.7 | 5.7 | 5.9 | 5.8 | 3.0272 | 6 | 5.8 | 6 | 8.2 | 6.5 |  |  | 1
| 1 | 2 | 6.3 | 5.986 | 6 | 5.8 | 5.7 | 6 | 5.6 | 5.7 | 5.8 | 3.7667 | 5.6 | 6.2 | 6 | 7.95 | 6.6 |  |  | 20
| 1 | 3 | 6.4 | 6.3577 | 6 | 6.1 | 5.5 | 6 | 5.8 | 6.1 | 5.5 | 3.4789 | 6 | 5.9 | 6 | 7.19 | 6.1 |  |  | 16
| 1 | 4 | 6.5 | 5.8714 | 6 | 6 | 5.6 | 5.8 | 5.7 | 6.1 | 5.8 | 3.6074 | 5.8 | 6.1 | 7 | 7.6 | 6.3 |  |  | 5.9531
| 1 | 5 | 6.4 | 6.2794 | 6 | 5.8 | 5.5 | 5.7 | 5.8 | 6.1 | 5.8 | 2.895 | 5.8 | 6.3 | 6 | 7.46 | 6.6 |  |  | 56
| 1 | 6 | 6.4 | 6.2035 | 6 | 5.7 | 5.6 | 5.7 | 5.8 | 6.3 | 5.8 | 3.6788 | 5.7 | 5.8 | 5 | 8.12 | 5.9
| Mean |  | 6.3833 | 6.1003 | 6 | 5.9167 | 5.5833 | 5.8167 | 5.7333 | 6.0333 | 5.75 | 3.409 | 5.8167 | 6.0167 | 6 | 7.7533 | 6.3333
| Median |  | 6.4 | 6.0948 | 6 | 5.9 | 5.6 | 5.75 | 5.75 | 6.1 | 5.8 | 3.5432 | 5.8 | 6 | 6 | 7.775 | 6.4
| Std Dev. |  | 0.075277 | 0.20644 | 0 | 0.17224 | 0.075277 | 0.1472 | 0.08165 | 0.20656 | 0.12247 | 0.36191 | 0.16021 | 0.2137 | 0.63246 | 0.39988 | 0.28752
| Rel.Std.Dev. |  | 1.179% | 3.384% | 0% | 2.911% | 1.348% | 2.531% | 1.424% | 3.424% | 2.13% | 10.62% | 2.754% | 3.552% | 10.54% | 5.158% | 4.54%
| PDM3 |  | 7.228% | 2.473% | 0.7886% | -0.6113% | -6.211% | -2.291% | -3.691% | 1.349% | -3.411% | -42.74% | -2.291% | 1.069% | 0.7886% | 30.24% | 6.388%
| Z-Score (Absolute) |  | 1.72 | 0.68 | NA | 0 | 1.23 | 0.37 | 0.67 | 0.43 | 0.61 | 9.22 | 0.37 | 0.37 | NA | 6.76 | 1.53
Results from laboratories 1.07 and 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 267 (Certified Value 159 ppm) | [18 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 156.13 | 163.528 | 152.6869 | 161 | 163 | 157 | 160 | 157.5 | 153.5 | 159 | 194.4244 | 155 | 152.8 | 158 | 146.4 | 162 |  | 1
| 1 | 2 | 155.57 | 164.639 | 156.3539 | 161 | 164 | 161 | 167 | 156.5 | 152.5 | 158 | 192.2185 | 154 | 160.5 | 158 | 147.5 | 163 |  | 21
| 1 | 3 | 158.3 | 164.573 | 165.5782 | 159 | 161 | 156 | 165.5 | 158 | 157 | 157.5 | 183.2504 | 156 | 154.3 | 161 | 149.5 | 163 |  | 16
| 1 | 4 | 155.97 | 164.133 | 155.8431 | 160 | 165 | 159 | 159.5 | 158 | 156.5 | 156.5 | 181.1605 | 155 | 155 | 161 | 148.8 | 163 |  | 158.9316
| 1 | 5 | 160.39 | 164.221 | 159.6682 | 155 | 162 | 160.5 | 163.5 | 159 | 156.5 | 157.5 | 191.399 | 153 | 158.6 | 158 | 145.5 | 162 |  | 57
| 1 | 6 | 157.8 | 162.934 | 155.673 | 165 | 161 | 160.5 | 161 | 157 | 157.5 | 158.5 | 193.6546 | 155 | 154.6 | 155 | 146.6 | 164
| Mean |  | 157.36 | 164.0047 | 157.6339 | 160.1667 | 162.6667 | 159 | 162.75 | 157.6667 | 155.5833 | 157.8333 | 189.3512 | 154.6667 | 155.9667 | 158.5 | 147.3833 | 162.8333
| Median |  | 156.965 | 164.177 | 156.0985 | 160.5 | 162.5 | 159.75 | 162.25 | 157.75 | 156.5 | 157.75 | 191.8088 | 155 | 154.8 | 158 | 147.05 | 163
| Std Dev. |  | 1.8389 | 0.6576 | 4.4814 | 3.2506 | 1.633 | 2.0736 | 3.0781 | 0.8756 | 2.0595 | 0.8756 | 5.6741 | 1.0328 | 2.9358 | 2.2583 | 1.525 | 0.75277
| Rel.Std.Dev. |  | 1.169% | 0.401% | 2.843% | 2.03% | 1.004% | 1.304% | 1.891% | 0.5553% | 1.324% | 0.5548% | 2.997% | 0.6678% | 1.882% | 1.425% | 1.035% | 0.4623%
| PDM3 |  | -0.9889% | 3.192% | -0.8166% | 0.7771% | 2.35% | 0.04301% | 2.403% | -0.7959% | -2.107% | -0.6911% | 19.14% | -2.684% | -1.866% | -0.2716% | -7.266% | 2.455%
| Z-Score (Absolute) |  | 0.23 | 1.65 | 0.15 | 0.56 | 1.27 | 0.23 | 1.29 | 0.14 | 0.73 | 0.09 | 8.79 | 0.99 | 0.62 | 0.09 | 3.04 | 1.32
| Analytical results for Ta in OREAS 267 (Certified Value 0.93 ppm) | [18 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.07 | 0.92147 | 1 | 1.26 | 0.9 | 0.96 | 0.95 | 0.92 | 0.94 | 0.87337 | 0.83 | 0.9 | 0.94 | 0.88 | 1 |  |  | 1
| 1 | 2 | 1.09 | 0.92248 | 1 | 1.31 | 0.91 | 0.98 | 0.96 | 0.87 | 0.92 | 0.72847 | 0.81 | 0.98 | 0.92 | 0.88 | 1 |  |  | 22
| 1 | 3 | 1.15 | 0.95893 | 1 | 1.25 | 0.89 | 0.96 | 0.94 | 0.96 | 0.93 | 0.91236 | 0.81 | 1 | 1.02 | 0.9 | 1.1 |  |  | 16
| 1 | 4 | 1.09 | 0.96494 | 0.9 | 1.24 | 0.89 | 0.96 | 0.97 | 0.93 | 0.92 | 0.83046 | 0.82 | 0.9 | 0.95 | 0.89 | 1.1 |  |  | 0.93065
| 1 | 5 | 1.12 | 0.99025 | 1 | 1.39 | 0.92 | 0.96 | 0.94 | 0.92 | 0.93 | 0.9927 | 0.81 | 1.01 | 0.96 | 0.91 | 1 |  |  | 58
| 1 | 6 | 1.08 | 0.95867 | 1 | 1.35 | 0.92 | 0.93 | 0.93 | 0.92 | 0.9 | 0.90972 | 0.8 | 0.94 | 0.92 | 0.87 | 1
| Mean |  | 1.1 | 0.95279 | 0.98333 | 1.3 | 0.905 | 0.95833 | 0.94833 | 0.92 | 0.92333 | 0.87451 | 0.81333 | 0.955 | 0.95167 | 0.88833 | 1.0333
| Median |  | 1.09 | 0.9588 | 1 | 1.285 | 0.905 | 0.96 | 0.945 | 0.92 | 0.925 | 0.89155 | 0.81 | 0.96 | 0.945 | 0.885 | 1
| Std Dev. |  | 0.029665 | 0.026543 | 0.040825 | 0.060663 | 0.013784 | 0.016021 | 0.01472 | 0.028983 | 0.013663 | 0.089343 | 0.010328 | 0.048888 | 0.037103 | 0.01472 | 0.05164
| Rel.Std.Dev. |  | 2.697% | 2.786% | 4.152% | 4.666% | 1.523% | 1.672% | 1.552% | 3.15% | 1.48% | 10.22% | 1.27% | 5.119% | 3.899% | 1.657% | 4.997%
| PDM3 |  | 18.2% | 2.379% | 5.661% | 39.69% | -2.756% | 2.974% | 1.9% | -1.145% | -0.7863% | -6.032% | -12.61% | 2.616% | 2.258% | -4.547% | 11.03%
| Z-Score (Absolute) |  | 3.61 | 0.11 | NA | 8.37 | 1.03 | 0.24 | 0 | 0.67 | 0.59 | 1.76 | 3.21 | 0.16 | 0.08 | 1.43 | NA
Results from laboratories 1.07 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 267 (Certified Value 0.54 ppm) | [18 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.19 | 1.20 | 1.22 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.52838 | 0.54537 | 0.6 | 0.56 | < 0.005 | 0.55 | 0.54 | 0.62 | [9 empty] | 1
| 1 | 2 | 0.53616 | 0.54056 | 0.6 | 0.59 | < 0.005 | 0.57 | 0.55 | 0.63 | [9 empty] | 6
| 1 | 3 | 0.51052 | 0.54537 | 0.6 | 0.54 | < 0.005 | 0.56 | 0.53 | 0.62 | [9 empty] | 16
| 1 | 4 | 0.5289 | 0.53794 | 0.6 | 0.53 | < 0.005 | 0.55 | 0.55 | 0.62 | [9 empty] | 0.54324
| 1 | 5 | 0.54259 | 0.56296 | 0.6 | 0.56 | < 0.005 | 0.55 | 0.53 | 0.63 | [9 empty] | 59
| 1 | 6 | 0.52234 | 0.53757 | 0.5 | 0.54 | < 0.005 | 0.55 | 0.53 | 0.63
| Mean |  | 0.52815 | 0.54496 | 0.58333 | 0.55333 |  | 0.555 | 0.53833 | 0.625
| Median |  | 0.52864 | 0.54296 | 0.6 | 0.55 |  | 0.55 | 0.535 | 0.625
| Std Dev. |  | 0.011118 | 0.0094602 | 0.040825 | 0.021602 |  | 0.0083666 | 0.0098319 | 0.0054772
| Rel.Std.Dev. |  | 2.105% | 1.736% | 6.999% | 3.904% |  | 1.507% | 1.826% | 0.8764%
| PDM3 |  | -2.777% | 0.3176% | 7.381% | 1.859% |  | 2.166% | -0.9024% | 15.05%
| Z-Score (Absolute) |  | 1.13 | 0 | NA | 0.56 | 36.44 | 0.67 | 0.45 | 5.38
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 267 (Certified Value 0.068 ppm) | [18 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | < 0.2 | < 1 | < 0.1 | 0.28 | 0.05 | 0.07 | 0.06 | 0.11 | 0.09 | 0.51385 | < 0.1 | 0.12 | < 0.03 | < 0.5 | < 0.1 |  |  | 1
| 1 | 2 | < 0.2 | < 1 | 0.2 | 0.28 | 0.08 | 0.08 | 0.05 | 0.06 | 0.07 | 0.57684 | < 0.1 | 0.1 | < 0.03 | < 0.5 | < 0.1 |  |  | 24
| 1 | 3 | < 0.2 | < 1 | 0.1 | 0.27 | 0.06 | 0.06 | 0.09 | 0.08 | 0.08 | 0.547 | < 0.1 | 0.09 | < 0.03 | < 0.5 | < 0.1 |  |  | 16
| 1 | 4 | < 0.2 | < 1 | 0.1 | 0.25 | 0.06 | 0.07 | < 0.05 | 0.07 | 0.07 | 0.44754 | < 0.1 | 0.11 | < 0.03 | < 0.5 | < 0.1 |  |  | 0.067733
| 1 | 5 | < 0.2 | < 1 | < 0.1 | 0.3 | 0.08 | 0.07 | 0.06 | 0.06 | 0.06 | 0.51054 | < 0.1 | 0.12 | < 0.03 | < 0.5 | < 0.1 |  |  | 60
| 1 | 6 | < 0.2 | < 1 | 0.1 | 0.29 | 0.06 | 0.05 | 0.07 | 0.06 | 0.08 | 0.54368 | < 0.1 | 0.11 | < 0.03 | < 0.5 | 0.1
| Mean |  |  |  | 0.125 | 0.27833 | 0.065 | 0.066667 | 0.066 | 0.073333 | 0.075 | 0.52324 |  | 0.10833 |  |  | 0.1
| Median |  |  |  | 0.1 | 0.28 | 0.06 | 0.07 | 0.06 | 0.065 | 0.075 | 0.52877 |  | 0.11 |  |  | 0.1
| Std Dev. |  |  |  | 0.05 | 0.017224 | 0.012247 | 0.010328 | 0.015166 | 0.019664 | 0.010488 | 0.04435 |  | 0.01169
| Rel.Std.Dev. |  |  |  | 40% | 6.188% | 18.84% | 15.49% | 22.98% | 26.81% | 13.98% | 8.476% |  | 10.79%
| PDM3 |  |  |  | 84.55% | 310.9% | -4.035% | -1.575% | -2.559% | 8.268% | 10.73% | 672.5% |  | 59.94% |  |  | 47.64%
| Z-Score (Absolute) |  | 0.7 | 11.49 | NA | 5.51 | 0.25 | 0.2 | 0.4 | 0.02 | 0.02 | NA | 0.65 | 0.92 | 1.6 | 4.74 | NA
Results from laboratories 1.07 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 267 (Certified Value 9.13 ppm) | [18 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 9.33 | 8.7223 | 9.123 | 8.9 | 9.62 | 9.03 | 8.77 | 9.3 | 9.33 | 9.32 | 9.0404 | 8.78 | 9.4 | 9.45 | 9.03 | 11.7 |  | 1
| 1 | 2 | 9.13 | 8.6559 | 8.9311 | 8.8 | 9.19 | 9.01 | 9.36 | 9.54 | 8.94 | 9.02 | 8.7396 | 8.72 | 9.7 | 9.39 | 9.2 | 11.6 |  | 25
| 1 | 3 | 9.4 | 8.7591 | 9.5969 | 9 | 9.58 | 8.82 | 8.97 | 9.44 | 9.93 | 8.91 | 8.4551 | 8.77 | 9.2 | 9.52 | 8.89 | 11.6 |  | 16
| 1 | 4 | 9.15 | 8.6997 | 9.0286 | 9.2 | 9.49 | 8.79 | 8.73 | 9.73 | 9.6 | 9.53 | 8.2355 | 8.73 | 9.5 | 9.43 | 8.91 | 11.5 |  | 9.1332
| 1 | 5 | 9.31 | 8.6859 | 9.432 | 8.3 | 9.19 | 9.18 | 9.27 | 9.38 | 9.48 | 9.13 | 8.958 | 8.63 | 9.3 | 9.25 | 9.22 | 11.6 |  | 61
| 1 | 6 | 9.26 | 8.7221 | 9.3154 | 8.9 | 9.59 | 9 | 8.85 | 9.4 | 10 | 9.11 | 8.7146 | 8.56 | 9.3 | 9.19 | 9.23 | 11.6
| Mean |  | 9.2633 | 8.7075 | 9.2378 | 8.85 | 9.4433 | 8.9717 | 8.9917 | 9.465 | 9.5467 | 9.17 | 8.6906 | 8.6983 | 9.4 | 9.3717 | 9.08 | 11.6
| Median |  | 9.285 | 8.7109 | 9.2192 | 8.9 | 9.535 | 9.005 | 8.91 | 9.42 | 9.54 | 9.12 | 8.7271 | 8.725 | 9.35 | 9.41 | 9.115 | 11.6
| Std Dev. |  | 0.10577 | 0.035466 | 0.25444 | 0.30166 | 0.20096 | 0.14497 | 0.26506 | 0.15176 | 0.39363 | 0.22244 | 0.30311 | 0.08612 | 0.17889 | 0.12624 | 0.15748 | 0.063246
| Rel.Std.Dev. |  | 1.142% | 0.4073% | 2.754% | 3.409% | 2.128% | 1.616% | 2.948% | 1.603% | 4.123% | 2.426% | 3.488% | 0.9901% | 1.903% | 1.347% | 1.734% | 0.5452%
| PDM3 |  | 1.425% | -4.661% | 1.146% | -3.1% | 3.396% | -1.768% | -1.549% | 3.633% | 4.527% | 0.4033% | -4.846% | -4.761% | 2.922% | 2.611% | -0.5821% | 27.01%
| Z-Score (Absolute) |  | 0.17 | 1.42 | 0.1 | 1.01 | 0.68 | 0.66 | 0.61 | 0.75 | 0.98 | 0.1 | 1.47 | 1.45 | 0.56 | 0.48 | 0.35 | 6.85
| Analytical results for Ti in OREAS 267 (Certified Value 0.373 wt.%) | [18 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.3926 | 0.38376 | 0.37927 | 0.37 | 0.37 | 0.377 | 0.385 | 0.366 | 0.372 | 0.369 | 0.28629 | 0.357 | 0.33 | 0.3668 | 0.36 | 0.373 |  | 1
| 1 | 2 | 0.3916 | 0.38363 | 0.37878 | 0.344 | 0.38 | 0.379 | 0.398 | 0.375 | 0.381 | 0.377 | 0.28361 | 0.345 | 0.33 | 0.3634 | 0.36 | 0.373 |  | 26
| 1 | 3 | 0.3918 | 0.38283 | 0.40084 | 0.367 | 0.37 | 0.371 | 0.392 | 0.37 | 0.372 | 0.378 | 0.30193 | 0.355 | 0.32 | 0.3742 | 0.37 | 0.367 |  | 16
| 1 | 4 | 0.3927 | 0.38272 | 0.37838 | 0.338 | 0.37 | 0.375 | 0.384 | 0.37 | 0.373 | 0.367 | 0.29349 | 0.346 | 0.33 | 0.3672 | 0.37 | 0.375 |  | 0.37288
| 1 | 5 | 0.4025 | 0.38332 | 0.38402 | 0.35 | 0.38 | 0.383 | 0.393 | 0.371 | 0.377 | 0.371 | 0.30107 | 0.347 | 0.34 | 0.3636 | 0.36 | 0.368 |  | 62
| 1 | 6 | 0.3947 | 0.38439 | 0.38317 | 0.354 | 0.37 | 0.382 | 0.387 | 0.372 | 0.383 | 0.375 | 0.31837 | 0.346 | 0.32 | 0.3572 | 0.36 | 0.359
| Mean |  | 0.39432 | 0.38344 | 0.38408 | 0.35383 | 0.37333 | 0.37783 | 0.38983 | 0.37067 | 0.37633 | 0.37283 | 0.29746 | 0.34933 | 0.32833 | 0.3654 | 0.36333 | 0.36917
| Median |  | 0.39265 | 0.38347 | 0.38122 | 0.352 | 0.37 | 0.378 | 0.3895 | 0.3705 | 0.375 | 0.373 | 0.29728 | 0.3465 | 0.33 | 0.3652 | 0.36 | 0.3705
| Std Dev. |  | 0.0041566 | 0.00062489 | 0.0085479 | 0.012624 | 0.005164 | 0.0044907 | 0.0054191 | 0.0029439 | 0.0048028 | 0.0044907 | 0.012669 | 0.0052409 | 0.0075277 | 0.0056057 | 0.005164 | 0.0058793
| Rel.Std.Dev. |  | 1.054% | 0.163% | 2.226% | 3.568% | 1.383% | 1.189% | 1.39% | 0.7942% | 1.276% | 1.204% | 4.259% | 1.5% | 2.293% | 1.534% | 1.421% | 1.593%
| PDM3 |  | 5.748% | 2.831% | 3.001% | -5.109% | 0.1204% | 1.327% | 4.545% | -0.5947% | 0.925% | -0.01366% | -20.23% | -6.316% | -11.95% | -2.007% | -2.561% | -0.997%
| Z-Score (Absolute) |  | 1.51 | 0.78 | 0.83 | 1.2 | 0.11 | 0.41 | 1.21 | 0.07 | 0.31 | 0.07 | 4.98 | 1.5 | 2.91 | 0.43 | 0.56 | 0.17
| Analytical results for Tl in OREAS 267 (Certified Value 0.39 ppm) | [18 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.43 | 0.42022 | 0.41 | 0.35 | 0.38 | 0.37 | 0.39 | 0.42 | 0.39 | 0.086957 | 0.4 | 0.37 | 0.41 | 0.363 | < 0.5 | 0.5 |  | 1
| 1 | 2 | 0.41 | 0.40039 | 0.41 | 0.36 | 0.36 | 0.39 | 0.37 | 0.42 | 0.38 | 0.1044 | 0.4 | 0.41 | 0.41 | 0.347 | < 0.5 | 0.5 |  | 27
| 1 | 3 | 0.43 | 0.43992 | 0.4 | 0.36 | 0.36 | 0.38 | 0.39 | 0.43 | 0.38 | 0.086524 | 0.39 | 0.39 | 0.42 | 0.341 | < 0.5 | 0.5 |  | 16
| 1 | 4 | 0.42 | 0.39984 | 0.42 | 0.37 | 0.36 | 0.41 | 0.4 | 0.43 | 0.39 | 0.097969 | 0.39 | 0.39 | 0.42 | 0.348 | < 0.5 | 0.5 |  | 0.39265
| 1 | 5 | 0.44 | 0.4155 | 0.39 | 0.36 | 0.37 | 0.4 | 0.37 | 0.42 | 0.37 | 0.11443 | 0.39 | 0.41 | 0.41 | 0.284 | < 0.5 | 0.5 |  | 63
| 1 | 6 | 0.42 | 0.41975 | 0.39 | 0.35 | 0.37 | 0.41 | 0.38 | 0.44 | 0.38 | 0.093076 | 0.39 | 0.37 | 0.42 | 0.36 | < 0.5 | 0.4
| Mean |  | 0.425 | 0.41594 | 0.40333 | 0.35833 | 0.36667 | 0.39333 | 0.38333 | 0.42667 | 0.38167 | 0.097226 | 0.39333 | 0.39 | 0.415 | 0.3405 |  | 0.48333
| Median |  | 0.425 | 0.41762 | 0.405 | 0.36 | 0.365 | 0.395 | 0.385 | 0.425 | 0.38 | 0.095522 | 0.39 | 0.39 | 0.415 | 0.3475 |  | 0.5
| Std Dev. |  | 0.010488 | 0.014893 | 0.012111 | 0.0075277 | 0.008165 | 0.01633 | 0.012111 | 0.008165 | 0.0075277 | 0.010818 | 0.005164 | 0.017889 | 0.0054772 | 0.028905 |  | 0.040825
| Rel.Std.Dev. |  | 2.468% | 3.581% | 3.003% | 2.101% | 2.227% | 4.152% | 3.159% | 1.914% | 1.972% | 11.13% | 1.313% | 4.587% | 1.32% | 8.489% |  | 8.447%
| PDM3 |  | 8.24% | 5.931% | 2.722% | -8.739% | -6.617% | 0.175% | -2.372% | 8.664% | -2.796% | -75.24% | 0.175% | -0.674% | 5.693% | -13.28% |  | 23.1%
| Z-Score (Absolute) |  | 0.94 | 0.7 | 0.36 | 0.85 | 0.63 | 0.09 | 0.18 | 0.99 | 0.22 | 7.9 | 0.09 | 0 | 0.67 | 1.34 | 3.78 | NA
| Analytical results for Tm in OREAS 267 (Indicative Value 0.21 ppm) | [18 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 0.22429 | 0.2 | 0.19 | 0.1307 | 0.21 | [12 empty] | 1
| 1 | 2 | 0.20234 | 0.3 | 0.2 | 0.12077 | 0.24 | [12 empty] | 1
| 1 | 3 | 0.22268 | 0.2 | 0.19 | 0.13028 | 0.22 | [12 empty] | 16
| 1 | 4 | 0.22021 | 0.2 | 0.2 | 0.1275 | 0.22 | [12 empty] | 0.21176
| 1 | 5 | 0.22145 | 0.2 | 0.2 | 0.13327 | 0.22 | [12 empty] | 7
| 1 | 6 | 0.21077 | 0.2 | 0.21 | 0.13181 | 0.2
| Mean |  | 0.21696 | 0.21667 | 0.19833 | 0.12906 | 0.21833
| Median |  | 0.22083 | 0.2 | 0.2 | 0.13049 | 0.22
| Std Dev. |  | 0.0085916 | 0.040825 | 0.0075277 | 0.0044884 | 0.013292
| Rel.Std.Dev. |  | 3.96% | 18.84% | 3.795% | 3.478% | 6.088%
| PDM3 |  | 2.452% | 2.316% | -6.342% | -39.06% | 3.103%
| Z-Score (Absolute) |  | 0.12 | 0 | 7.42 | 35.45 | 0.67
| Analytical results for U in OREAS 267 (Certified Value 1.93 ppm) | [18 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 1.95 | 0.91629 | < 5 | 1.9 | 1.92 | 1.8 | 1.7 | 2 | 1.9 | 1.9 | 1.2766 | 1.89 | 1.9 | 2 | 1.56 | 2.06 | 2.1 | 1
| 1 | 2 | 1.91 | 0.91329 | < 5 | 2 | 1.93 | 1.8 | 1.7 | 2.1 | 1.8 | 1.8 | 1.2698 | 1.89 | 2 | 2.03 | 1.29 | 2.12 | 2.1 | 29
| 1 | 3 | 1.96 | 0.97109 | < 5 | 1.9 | 1.96 | 1.7 | 1.7 | 2 | 1.9 | 1.8 | 1.2777 | 1.93 | 1.9 | 2.07 | 1.31 | 2.05 | 2.1 | 16
| 1 | 4 | 1.9 | 0.94942 | < 5 | 2 | 1.92 | 1.9 | 1.6 | 2.1 | 1.9 | 1.8 | 1.2946 | 1.93 | 2 | 1.96 | 1.24 | 2.04 | 2.1 | 1.9291
| 1 | 5 | 1.95 | 0.93391 | < 5 | 1.8 | 1.9 | 1.8 | 1.7 | 2 | 1.9 | 1.8 | 1.3535 | 1.9 | 2 | 1.98 | 1.35 | 2.11 | 2.1 | 64
| 1 | 6 | 1.92 | 0.98929 | < 5 | 2 | 1.97 | 1.8 | 1.7 | 2.1 | 2 | 1.8 | 1.3756 | 1.9 | 1.9 | 1.91 | 1.55 | 2.11 | 2.1
| Mean |  | 1.9317 | 0.94555 |  | 1.9333 | 1.9333 | 1.8 | 1.6833 | 2.05 | 1.9 | 1.8167 | 1.3079 | 1.9067 | 1.95 | 1.9917 | 1.3833 | 2.0817 | 2.1
| Median |  | 1.935 | 0.94166 |  | 1.95 | 1.925 | 1.8 | 1.7 | 2.05 | 1.9 | 1.8 | 1.2861 | 1.9 | 1.95 | 1.99 | 1.33 | 2.085 | 2.1
| Std Dev. |  | 0.024833 | 0.030375 |  | 0.08165 | 0.026583 | 0.063246 | 0.040825 | 0.054772 | 0.063246 | 0.040825 | 0.045123 | 0.018619 | 0.054772 | 0.055648 | 0.13765 | 0.035449 | 0
| Rel.Std.Dev. |  | 1.286% | 3.212% |  | 4.223% | 1.375% | 3.514% | 2.425% | 2.672% | 3.329% | 2.247% | 3.45% | 0.9765% | 2.809% | 2.794% | 9.95% | 1.703% | 0%
| PDM3 |  | 0.1329% | -50.99% |  | 0.2193% | 0.2193% | -6.692% | -12.74% | 6.267% | -1.509% | -5.828% | -32.2% | -1.163% | 1.083% | 3.243% | -28.29% | 7.909% | 8.859%
| Z-Score (Absolute) |  | 0 | 5.62 | 3.24 | 0.01 | 0.01 | 0.75 | 1.42 | 0.67 | 0.18 | 0.66 | 3.55 | 0.14 | 0.1 | 0.34 | 3.12 | 0.85 | 0.96
| Analytical results for V in OREAS 267 (Certified Value 68 ppm) | [18 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 0
1
| 1 | 1 | 69 | 61.4284 | 64.6047 | 69 | 71 | 67 | 67 | 66 | 66 | 67 | 64.1548 | 66 | 65 | 68 | 71 | 64.66 | 71 | 1
| 1 | 2 | 69 | 60.4368 | 65.2916 | 67 | 72 | 68 | 70 | 65 | 69 | 68 | 60.2446 | 66 | 67 | 67 | 69 | 69.54 | 71 | 30
| 1 | 3 | 68 | 60.6832 | 70.0338 | 68 | 70 | 67 | 69 | 66 | 68 | 69 | 65.1209 | 66 | 65 | 66 | 71 | 68.3 | 70 | 16
| 1 | 4 | 68 | 63.6208 | 66.0664 | 69 | 71 | 67 | 68 | 65 | 68 | 67 | 64.5832 | 65 | 68 | 69 | 71 | 69.58 | 71 | 68.0536
| 1 | 5 | 70 | 62.5624 | 66.4038 | 64 | 71 | 69 | 69 | 67 | 68 | 67 | 67.2341 | 66 | 69 | 69 | 70 | 69.91 | 70 | 65
| 1 | 6 | 69 | 62.012 | 66.766 | 69 | 69 | 69 | 68 | 67 | 70 | 66 | 69.175 | 66 | 66 | 68 | 70 | 69.55 | 70
| Mean |  | 68.8333 | 61.7906 | 66.5277 | 67.6667 | 70.6667 | 67.8333 | 68.5 | 66 | 68.1667 | 67.3333 | 65.0854 | 65.8333 | 66.6667 | 67.8333 | 70.3333 | 68.59 | 70.5
| Median |  | 69 | 61.7202 | 66.2351 | 68.5 | 71 | 67.5 | 68.5 | 66 | 68 | 67 | 64.852 | 66 | 66.5 | 68 | 70.5 | 69.545 | 70.5
| Std Dev. |  | 0.75277 | 1.1987 | 1.887 | 1.9664 | 1.0328 | 0.98319 | 1.0488 | 0.89443 | 1.3292 | 1.0328 | 3.0297 | 0.40825 | 1.633 | 1.169 | 0.8165 | 2.0038 | 0.54772
| Rel.Std.Dev. |  | 1.094% | 1.94% | 2.836% | 2.906% | 1.462% | 1.449% | 1.531% | 1.355% | 1.95% | 1.534% | 4.655% | 0.6201% | 2.449% | 1.723% | 1.161% | 2.921% | 0.7769%
| PDM3 |  | 1.146% | -9.203% | -2.242% | -0.5685% | 3.84% | -0.3236% | 0.656% | -3.018% | 0.1662% | -1.058% | -4.362% | -3.263% | -2.038% | -0.3236% | 3.35% | 0.7882% | 3.595%
| Z-Score (Absolute) |  | 0.58 | 3.49 | 0.75 | 0.1 | 1.64 | 0 | 0.39 | 1.06 | 0.19 | 0.29 | 1.59 | 1.16 | 0.67 | 0 | 1.44 | 0.44 | 1.54
| Analytical results for W in OREAS 267 (Certified Value 7.32 ppm) | [18 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.8 | 7.4331 | 7.2 | 7.2 | 7.4 | 7.1 | 7 | 7.1 | 7.1 | 7.2382 | 7 | 6.8 | 7.5 | 7.65 | 7.3 |  |  | 1
| 1 | 2 | 7.6 | 7.5423 | 7.2 | 7.2 | 7.4 | 7.5 | 7.2 | 6.9 | 7 | 7.2065 | 6.8 | 7.1 | 7.4 | 7.49 | 7.3 |  |  | 31
| 1 | 3 | 7.6 | 7.8123 | 7.2 | 7.6 | 7.3 | 7.2 | 7 | 7.4 | 7 | 8.2899 | 6.8 | 6.9 | 7.7 | 7.86 | 7.3 |  |  | 16
| 1 | 4 | 7.7 | 8.0069 | 7.3 | 8.2 | 7.2 | 7.2 | 7.3 | 7.2 | 6.9 | 6.903 | 6.9 | 7.3 | 7.7 | 7.55 | 7.1 |  |  | 7.3236
| 1 | 5 | 7.8 | 7.8538 | 7.3 | 8.3 | 7.6 | 7.4 | 7 | 7.1 | 6.9 | 7.883 | 6.8 | 7.2 | 7.6 | 7.36 | 7.2 |  |  | 66
| 1 | 6 | 7.9 | 7.6381 | 7.5 | 7.5 | 7.6 | 7.3 | 7 | 7.1 | 6.9 | 6.4866 | 6.7 | 6.8 | 7.4 | 7.62 | 7.3
| Mean |  | 7.7333 | 7.7144 | 7.2833 | 7.6667 | 7.4167 | 7.2833 | 7.0833 | 7.1333 | 6.9667 | 7.3345 | 6.8333 | 7.0167 | 7.55 | 7.5883 | 7.25
| Median |  | 7.75 | 7.7252 | 7.25 | 7.55 | 7.4 | 7.25 | 7 | 7.1 | 6.95 | 7.2224 | 6.8 | 7 | 7.55 | 7.585 | 7.3
| Std Dev. |  | 0.12111 | 0.21408 | 0.1169 | 0.48028 | 0.16021 | 0.1472 | 0.13292 | 0.1633 | 0.08165 | 0.6549 | 0.10328 | 0.2137 | 0.13784 | 0.16845 | 0.083666
| Rel.Std.Dev. |  | 1.566% | 2.775% | 1.605% | 6.264% | 2.16% | 2.021% | 1.876% | 2.289% | 1.172% | 8.929% | 1.511% | 3.046% | 1.826% | 2.22% | 1.154%
| PDM3 |  | 5.595% | 5.336% | -0.5498% | 4.684% | 1.271% | -0.5498% | -3.281% | -2.598% | -4.874% | 0.1493% | -6.694% | -4.191% | 3.091% | 3.615% | -1.005%
| Z-Score (Absolute) |  | 1.14 | 1.09 | 0 | 0.97 | 0.34 | 0 | 0.51 | 0.38 | 0.8 | 0.13 | 1.14 | 0.67 | 0.67 | 0.77 | 0.08
| Analytical results for Y in OREAS 267 (Certified Value 14 ppm) | [18 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 13.67 | 13.3062 | 13.9364 | 14.1 | 12.9 | 14.1 | 14.3 | 13.4 | 19 | 14.6 | 11.4199 | 13.2 | 13.5 | 15.74 | 17.06 | 16.3 |  | 1
| 1 | 2 | 13.52 | 13.769 | 14.1439 | 14 | 12.9 | 15 | 15.4 | 13.6 | 14 | 14.4 | 11.2093 | 13.3 | 13.9 | 13.56 | 17.25 | 15.9 |  | 32
| 1 | 3 | 13.8 | 13.8214 | 14.763 | 14.7 | 12.6 | 13.6 | 14.8 | 13.7 | 15.2 | 14.2 | 11.764 | 13.3 | 13.7 | 14.05 | 17.31 | 15.7 |  | 16
| 1 | 4 | 14.61 | 13.8089 | 14.308 | 14.6 | 12.8 | 14.3 | 14.7 | 13.8 | 15.8 | 14.9 | 11.7122 | 13.4 | 13.3 | 13.6 | 17.53 | 15.9 |  | 14.0028
| 1 | 5 | 13.93 | 13.6342 | 14.2524 | 14.3 | 12.8 | 26.3 | 15.3 | 13.5 | 14.8 | 14.9 | 12.1148 | 13.2 | 13.4 | 13.81 | 17.39 | 16.4 |  | 67
| 1 | 6 | 13.5 | 13.8363 | 14.6483 | 14.9 | 12.9 | 14.3 | 14.4 | 13.6 | 15.4 | 14.8 | 11.9512 | 13.1 | 13.8 | 13.53 | 17.31 | 16.6
| Mean |  | 13.8383 | 13.696 | 14.342 | 14.4333 | 12.8167 | 16.2667 | 14.8167 | 13.6 | 15.7 | 14.6333 | 11.6953 | 13.25 | 13.6 | 14.0483 | 17.3083 | 16.1333
| Median |  | 13.735 | 13.789 | 14.2802 | 14.45 | 12.85 | 14.3 | 14.75 | 13.6 | 15.3 | 14.7 | 11.7381 | 13.25 | 13.6 | 13.705 | 17.31 | 16.1
| Std Dev. |  | 0.41219 | 0.20458 | 0.31112 | 0.3559 | 0.1169 | 4.9359 | 0.4535 | 0.14142 | 1.7286 | 0.28752 | 0.33423 | 0.10488 | 0.23664 | 0.85167 | 0.15549 | 0.35024
| Rel.Std.Dev. |  | 2.979% | 1.494% | 2.169% | 2.466% | 0.9121% | 30.34% | 3.061% | 1.04% | 11.01% | 1.965% | 2.858% | 0.7916% | 1.74% | 6.062% | 0.8983% | 2.171%
| PDM3 |  | -1.174% | -2.191% | 2.422% | 3.075% | -8.471% | 16.17% | 5.812% | -2.877% | 12.12% | 4.503% | -16.48% | -5.376% | -2.877% | 0.3252% | 23.61% | 15.22%
| Z-Score (Absolute) |  | 0.4 | 0.55 | 0.16 | 0.26 | 1.53 | 2.3 | 0.69 | 0.66 | 1.67 | 0.49 | 2.77 | 1.05 | 0.66 | 0.16 | 3.45 | 2.15
| Analytical results for Yb in OREAS 267 (Certified Value 1.38 ppm) | [18 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.19 | 1.20 | 1.22 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 1.392 | 1.4032 | 1.4 | 1.29 | 0.77071 | 1.4 | 1.9 | 1.64 | [9 empty] | 1
| 1 | 2 | 1.439 | 1.3457 | 1.2 | 1.33 | 0.75064 | 1.5 | 1.43 | 1.6 | [9 empty] | 8
| 1 | 3 | 1.3901 | 1.3804 | 1.5 | 1.27 | 0.76668 | 1.3 | 1.42 | 1.65 | [9 empty] | 16
| 1 | 4 | 1.42 | 1.3303 | 1.3 | 1.39 | 0.76899 | 1.4 | 1.48 | 1.67 | [9 empty] | 1.3787
| 1 | 5 | 1.4088 | 1.3804 | 1.3 | 1.26 | 0.79047 | 1.4 | 1.41 | 1.55 | [9 empty] | 68
| 1 | 6 | 1.4052 | 1.428 | 1.4 | 1.26 | 0.80882 | 1.4 | 1.75 | 1.62
| Mean |  | 1.4092 | 1.378 | 1.35 | 1.3 | 0.77605 | 1.4 | 1.565 | 1.6217
| Median |  | 1.407 | 1.3804 | 1.35 | 1.28 | 0.76985 | 1.4 | 1.455 | 1.63
| Std Dev. |  | 0.018319 | 0.035965 | 0.10488 | 0.051381 | 0.020465 | 0.063246 | 0.2083 | 0.042622
| Rel.Std.Dev. |  | 1.3% | 2.61% | 7.769% | 3.952% | 2.637% | 4.518% | 13.31% | 2.628%
| PDM3 |  | 2.212% | -0.05164% | -2.081% | -5.708% | -43.71% | 1.545% | 13.51% | 17.62%
| Z-Score (Absolute) |  | 0.12 | 0.3 | 0.67 | 1.35 | NA | 0 | 2.23 | 2.99
| Analytical results for Zn in OREAS 267 (Certified Value 43.4 ppm) | [18 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 45 | 29.8952 | 44.1529 | 42 | 49 | 43 | 44 | 42 | 43 | 43 | 37.1817 | 44 | 42 | 43 | 44 | 48.09 | 41 | 1
| 1 | 2 | 45 | 30.8888 | 43.8736 | 40 | 47 | 44 | 47 | 43 | 44 | 43 | 39.1986 | 43 | 41 | 43 | 43 | 48.99 | 41 | 34
| 1 | 3 | 45 | 31.3256 | 47.5985 | 42 | 48 | 43 | 45 | 43 | 44 | 44 | 34.9555 | 43 | 45 | 44 | 43 | 46.99 | 39 | 16
| 1 | 4 | 45 | 32.3136 | 44.1256 | 41 | 46 | 43 | 44 | 43 | 44 | 43 | 38.7088 | 44 | 44 | 44 | 43 | 46.77 | 40 | 43.4263
| 1 | 5 | 46 | 27.8624 | 44.9415 | 42 | 45 | 44 | 45 | 43 | 44 | 43 | 35.2691 | 43 | 43 | 44 | 44 | 42.75 | 39 | 69
| 1 | 6 | 44 | 29.0168 | 44.1648 | 42 | 48 | 44 | 45 | 43 | 45 | 43 | 35.9918 | 43 | 42 | 42 | 43 | 43.15 | 42
| Mean |  | 45 | 30.2171 | 44.8095 | 41.5 | 47.1667 | 43.5 | 45 | 42.8333 | 44 | 43.1667 | 36.8843 | 43.3333 | 42.8333 | 43.3333 | 43.3333 | 46.1233 | 40.3333
| Median |  | 45 | 30.392 | 44.1588 | 42 | 47.5 | 43.5 | 45 | 43 | 44 | 43 | 36.5868 | 43 | 42.5 | 43.5 | 43 | 46.88 | 40.5
| Std Dev. |  | 0.63246 | 1.6226 | 1.4133 | 0.83666 | 1.472 | 0.54772 | 1.0954 | 0.40825 | 0.63246 | 0.40825 | 1.7832 | 0.5164 | 1.472 | 0.8165 | 0.5164 | 2.5875 | 1.2111
| Rel.Std.Dev. |  | 1.405% | 5.37% | 3.154% | 2.016% | 3.121% | 1.259% | 2.434% | 0.9531% | 1.437% | 0.9457% | 4.835% | 1.192% | 3.436% | 1.884% | 1.192% | 5.61% | 3.003%
| PDM3 |  | 3.624% | -30.42% | 3.185% | -4.436% | 8.613% | 0.1697% | 3.624% | -1.365% | 1.321% | -0.5979% | -15.06% | -0.2141% | -1.365% | -0.2141% | -0.2141% | 6.211% | -7.122%
| Z-Score (Absolute) |  | 0.76 | 5.99 | 0.67 | 0.84 | 1.75 | 0.08 | 0.76 | 0.23 | 0.3 | 0.08 | 2.95 | 0 | 0.23 | 0 | 0 | 1.27 | 1.37
| Analytical results for Zr in OREAS 267 (Certified Value 112 ppm) | [18 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 113.7 | 94.4315 | 108.3141 | 110 | 106 | 111 | 112 | 111 | 110 | 103.5 | 117.3278 | 111 | 109.5 | 112 | 121.66 | 104 |  | 1
| 1 | 2 | 109.8 | 95.3765 | 109.1187 | 112 | 102 | 112.5 | 118.5 | 112.5 | 104.5 | 105.5 | 115.1534 | 112 | 113.3 | 107.3 | 127.65 | 105.3 |  | 35
| 1 | 3 | 111.3 | 95.702 | 115.7407 | 116 | 108 | 110.5 | 117.5 | 112.5 | 110.5 | 102 | 112.0011 | 110 | 113.6 | 118.3 | 126.8 | 107 |  | 16
| 1 | 4 | 111.4 | 96.4475 | 110.9416 | 113 | 108 | 110.5 | 112.5 | 115.5 | 110 | 103.5 | 107.9098 | 110 | 111.1 | 109.5 | 129.49 | 107.1 |  | 112.226
| 1 | 5 | 114.8 | 95.2295 | 111.8818 | 108 | 98 | 112 | 115.5 | 111.5 | 113 | 105.5 | 118.0151 | 112 | 113.1 | 111.2 | 124.79 | 106.1 |  | 70
| 1 | 6 | 110.9 | 95.5025 | 111.3254 | 115 | 107 | 113 | 114 | 111.5 | 110.5 | 112.5 | 119.586 | 112 | 111.4 | 107.6 | 124.22 | 103.6
| Mean |  | 111.9833 | 95.4483 | 111.2204 | 112.3333 | 104.8333 | 111.5833 | 115 | 112.4167 | 109.75 | 105.4167 | 114.9989 | 111.1667 | 112 | 110.9833 | 125.7683 | 105.5167
| Median |  | 111.35 | 95.4395 | 111.1335 | 112.5 | 106.5 | 111.5 | 114.75 | 112 | 110.25 | 104.5 | 116.2406 | 111.5 | 112.25 | 110.35 | 125.795 | 105.7
| Std Dev. |  | 1.8777 | 0.65614 | 2.6009 | 3.0111 | 4.0208 | 1.0685 | 2.6458 | 1.6253 | 2.8062 | 3.7204 | 4.355 | 0.98319 | 1.605 | 4.0455 | 2.7817 | 1.4878
| Rel.Std.Dev. |  | 1.677% | 0.6874% | 2.338% | 2.68% | 3.835% | 0.9576% | 2.301% | 1.446% | 2.557% | 3.529% | 3.787% | 0.8844% | 1.433% | 3.645% | 2.212% | 1.41%
| PDM3 |  | -0.2162% | -14.95% | -0.8961% | 0.09565% | -6.587% | -0.5726% | 2.472% | 0.1699% | -2.206% | -6.068% | 2.471% | -0.9439% | -0.2014% | -1.107% | 12.07% | -5.978%
| Z-Score (Absolute) |  | 0.29 | 8.07 | 0.09 | 0.47 | 3.32 | 0.09 | 1.82 | 0.51 | 0.84 | 3.03 | 1.82 | 0.12 | 0.3 | 0.21 | 7.27 | 2.98
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C851: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C869: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 267 (Certified Value 5.35 ppm) | [16 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 5.498 | 5.3561 | 4.95 | 5.49 | 5.35 | 5.3 | 5.32 | 5.52 | 0.22769 | 5.27 | 5.38 | 5.2 | 6.61 | 5.1 | 5.37 | 1
| 1 | 2 | 5.653 | 5.2623 | 5.6 | 5.57 | 5.44 | 5.34 | 5.4 | 5.34 | 0.3192 | 5.12 | 5.27 | 5.2 | 6.2 | 5.1 | 5.34 | 20
| 1 | 3 | 5.546 | 5.2885 | 5.23 | 5.35 | 5.35 | 6.01 | 5.3 | 5.53 | 0.2215 | 5.05 | 5.44 | 5.1 | 6.27 | 5.1 | 5.33 | 16
| 1 | 4 | 5.834 | 5.3406 | 5.52 | 5.47 | 5.38 | 5.62 | 5.31 | 5.47 | 0.23074 | 5.24 | 5.29 | 5.1 | 6.38 | 5.1 | 5.26 | 5.3495
| 1 | 5 | 5.755 | 5.3122 | 5.33 | 5.46 | 5.45 | 5.37 | 5.47 | 5.7 | 0.22747 | 5.05 | 5.51 | 5.2 | 6.3 | 5.1 | 5.38 | 72
| 1 | 6 | 5.721 | 5.2371 | 5.02 | 5.51 | 5.43 | 5.41 | 5.31 | 5.5 | 0.18885 | 5.26 | 5.34 | 5.2 | 6.3 | 5.2 | 5.56
| Mean |  | 5.6678 | 5.2995 | 5.275 | 5.475 | 5.4 | 5.5083 | 5.3517 | 5.51 | 0.23591 | 5.165 | 5.3717 | 5.1667 | 6.3433 | 5.1167 | 5.3733
| Median |  | 5.687 | 5.3004 | 5.28 | 5.48 | 5.405 | 5.39 | 5.315 | 5.51 | 0.22758 | 5.18 | 5.36 | 5.2 | 6.3 | 5.1 | 5.355
| Std Dev. |  | 0.12802 | 0.045714 | 0.26129 | 0.072595 | 0.045607 | 0.27007 | 0.068532 | 0.11593 | 0.043646 | 0.10407 | 0.091524 | 0.05164 | 0.14292 | 0.040825 | 0.10073
| Rel.Std.Dev. |  | 2.259% | 0.8626% | 4.953% | 1.326% | 0.8446% | 4.903% | 1.281% | 2.104% | 18.5% | 2.015% | 1.704% | 0.9995% | 2.253% | 0.7979% | 1.875%
| PDM3 |  | 5.951% | -0.9342% | -1.392% | 2.347% | 0.9448% | 2.97% | 0.04124% | 3.001% | -95.59% | -3.448% | 0.4151% | -3.417% | 18.58% | -4.352% | 0.4463%
| Z-Score (Absolute) |  | 1.67 | 0.41 | 0.55 | 0.58 | 0.16 | 0.77 | 0.12 | 0.78 | NA | 1.17 | 0 | 1.16 | 5.49 | 1.45 | 0
Results from laboratory 17 were removed due to their 1E-09 ppm reading resolution.
| Analytical results for Al in OREAS 267 (Certified Value 0.879 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.90311 | 0.91487 | 1.29 | 0.92 | 0.86 | 0.87 | 0.95 | 0.98 | 0.97496 | 1.1 | 0.77 | 1.13 | 0.84 | 0.73 |  | 1
| 1 | 2 | 0.92413 | 0.92354 | 1.19 | 0.92 | 0.85 | 0.89 | 0.95 | 0.98 | 0.89058 | 1.09 | 0.77 | 1.15 | 0.86 | 0.72 |  | #N/A
| 1 | 3 | 0.89395 | 0.91301 | 1.36 | 0.89 | 0.87 | 0.89 | 0.93 | 0.97 | 0.74115 | 1.1 | 0.75 | 1.17 | 0.85 | 0.73 |  | 16
| 1 | 4 | 0.913 | 0.89883 | 1.2 | 0.92 | 0.85 | 0.89 | 0.93 | 1 | 0.94049 | 1.09 | 0.78 | 1.18 | 0.85 | 0.71 |  | 0.8788
| 1 | 5 | 0.88838 | 0.92695 | 1.13 | 0.9 | 0.86 | 0.91 | 0.99 | 1.01 | 1.0052 | 1.1 | 0.79 | 1.16 | 0.86 | 0.72 |  | 73
| 1 | 6 | 0.89075 | 0.93179 | 1.24 | 0.9 | 0.86 | 0.91 | 0.97 | 1.02 | 0.77585 | 1.09 | 0.78 | 1.16 | 0.87 | 0.73
| Mean |  | 0.90222 | 0.91817 | 1.235 | 0.90833 | 0.85833 | 0.89333 | 0.95333 | 0.99333 | 0.88803 | 1.095 | 0.77333 | 1.1583 | 0.855 | 0.72333
| Median |  | 0.89853 | 0.9192 | 1.22 | 0.91 | 0.86 | 0.89 | 0.95 | 0.99 | 0.91554 | 1.095 | 0.775 | 1.16 | 0.855 | 0.725
| Std Dev. |  | 0.01406 | 0.011859 | 0.081179 | 0.013292 | 0.0075277 | 0.015055 | 0.023381 | 0.019664 | 0.10788 | 0.0054772 | 0.013663 | 0.017224 | 0.010488 | 0.008165
| Rel.Std.Dev. |  | 1.558% | 1.292% | 6.573% | 1.463% | 0.877% | 1.685% | 2.453% | 1.98% | 12.15% | 0.5002% | 1.767% | 1.487% | 1.227% | 1.129%
| PDM3 |  | 2.666% | 4.48% | 40.53% | 3.361% | -2.328% | 1.654% | 8.482% | 13.03% | 1.051% | 24.6% | -12% | 31.81% | -2.708% | -17.69%
| Z-Score (Absolute) |  | 0.04 | 0.18 | 4.52 | 0.04 | 0.64 | 0.16 | 0.66 | 1.21 | 0.24 | 2.6 | 1.81 | 3.47 | 0.69 | 2.5
| Analytical results for As in OREAS 267 (Certified Value 112 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 0
| 1 | 1 | 106.869 | 111.7195 | 109 | 115 | 112.5 | 111.5 | 109.5 | 106.5 | 99.6665 | 110 | 116 | 107.6 | 38.7 | 92.7 | 106 | 1
| 1 | 2 | 107.604 | 111.285 | 123 | 116 | 116 | 112 | 111 | 102.5 | 90.6478 | 112 | 119 | 108.4 | 44.8 | 93.1 | 106 | 21
| 1 | 3 | 110.42 | 109.7618 | 116 | 112.5 | 112.5 | 115 | 106.5 | 109 | 76.9535 | 113 | 120 | 106.8 | 45.4 | 92.4 | 106 | 16
| 1 | 4 | 110.85 | 109.9245 | 121 | 113.5 | 113.5 | 112 | 108.5 | 108 | 91.3647 | 113 | 122 | 108.2 | 44 | 92.3 | 104 | 111.617
| 1 | 5 | 112.542 | 110.8869 | 118 | 116.5 | 115 | 112 | 116 | 108.5 | 100.0404 | 112 | 118 | 108.5 | 38 | 93.3 | 106 | 74
| 1 | 6 | 109.896 | 111.1657 | 111 | 115 | 114 | 112.5 | 108.5 | 106.5 | 78.5295 | 111 | 116 | 110 | 38 | 93.6 | 108
| Mean |  | 109.6968 | 110.7906 | 116.3333 | 114.75 | 113.9167 | 112.5 | 110 | 106.8333 | 89.5337 | 111.8333 | 118.5 | 108.25 | 41.4833 | 92.9 | 106
| Median |  | 110.158 | 111.0263 | 117 | 115 | 113.75 | 112 | 109 | 107.25 | 91.0062 | 112 | 118.5 | 108.3 | 41.35 | 92.9 | 106
| Std Dev. |  | 2.1149 | 0.78306 | 5.5015 | 1.5083 | 1.3934 | 1.2649 | 3.2863 | 2.3594 | 9.97 | 1.169 | 2.3452 | 1.0654 | 3.5969 | 0.51769 | 1.2649
| Rel.Std.Dev. |  | 1.928% | 0.7068% | 4.729% | 1.314% | 1.223% | 1.124% | 2.988% | 2.208% | 11.14% | 1.045% | 1.979% | 0.9842% | 8.671% | 0.5573% | 1.193%
| PDM3 |  | -1.72% | -0.7404% | 4.225% | 2.807% | 2.06% | 0.7911% | -1.449% | -4.286% | -19.78% | 0.1938% | 6.167% | -3.017% | -62.83% | -16.77% | -5.032%
| Z-Score (Absolute) |  | 0.05 | 0.14 | 1.09 | 0.82 | 0.67 | 0.43 | 0 | 0.55 | 3.52 | 0.32 | 1.46 | 0.3 | 11.8 | 2.94 | 0.69
| Analytical results for B in OREAS 267 (Indicative Value 5.29 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | 30.062 | < 10 | 3 | < 10 | < 10 | 10 | < 10 | < 10 | 1.0304 | < 10 | < 5 |  |  |  |  | 1
| 1 | 2 | 29.381 | < 10 | 9 | < 10 | < 10 | 10 | < 10 | < 10 | 1.7233 | < 10 | < 5 |  |  |  |  | 4
| 1 | 3 | 129.066 | < 10 | 6 | < 10 | < 10 | < 10 | < 10 | < 10 | 1.1611 | < 10 | < 5 |  |  |  |  | 16
| 1 | 4 | 118.949 | < 10 | 5 | < 10 | < 10 | 10 | < 10 | < 10 | 1.5406 | < 10 | < 5 |  |  |  |  | 5.2932
| 1 | 5 | 130.202 | < 10 | 5 | < 10 | < 10 | 10 | < 10 | < 10 | 1.8921 | < 10 | < 5 |  |  |  |  | 10
| 1 | 6 | 35.97 | < 10 | 4 | < 10 | < 10 | 10 | < 10 | < 10 | 0.93013 | < 10 | < 5
| Mean |  | 78.9383 |  | 5.3333 |  |  | 10 |  |  | 1.3796
| Median |  | 77.4595 |  | 5 |  |  | 10 |  |  | 1.3508
| Std Dev. |  | 51.8318 |  | 2.0656 |  |  | 0 |  |  | 0.39456
| Rel.Std.Dev. |  | 65.66% |  | 38.73% |  |  | 0% |  |  | 28.6%
| PDM3 |  | 1391% |  | 0.7582% |  |  | 88.92% |  |  | -73.94%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 267 (Certified Value 71 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 0
| 1 | 1 | 66.327 | 73.429 | 84 | 70 | 70 | 70 | 80 | 80 | 90.2688 | 83 | 69 | 74.2 | 73 | 57 |  | 1
| 1 | 2 | 68.788 | 72.8229 | 99.1 | 70 | 70 | 80 | 80 | 70 | 66.6255 | 83 | 67 | 76.5 | 73 | 58 |  | 22
| 1 | 3 | 67.901 | 72.5179 | 89.3 | 70 | 70 | 50 | 80 | 70 | 49.9219 | 84 | 68 | 76.3 | 69 | 58 |  | 16
| 1 | 4 | 67.721 | 73.0921 | 94.8 | 70 | 70 | 50 | 80 | 80 | 68.502 | 84 | 67 | 79 | 71 | 59 |  | 71.4725
| 1 | 5 | 67.234 | 72.9538 | 91.9 | 70 | 70 | 80 | 80 | 80 | 66.3795 | 84 | 70 | 77.7 | 71 | 57 |  | 75
| 1 | 6 | 68.098 | 74.2533 | 87 | 70 | 70 | 80 | 80 | 80 | 50.3194 | 82 | 67 | 78 | 74 | 58
| Mean |  | 67.6782 | 73.1782 | 91.0167 | 70 | 70 | 68.3333 | 80 | 76.6667 | 65.3362 | 83.3333 | 68 | 76.95 | 71.8333 | 57.8333
| Median |  | 67.811 | 73.0229 | 90.6 | 70 | 70 | 75 | 80 | 80 | 66.5025 | 83.5 | 67.5 | 77.1 | 72 | 58
| Std Dev. |  | 0.83456 | 0.60643 | 5.4543 | 0 | 0 | 14.7196 | 0 | 5.164 | 14.8156 | 0.8165 | 1.2649 | 1.6766 | 1.8348 | 0.75277
| Rel.Std.Dev. |  | 1.233% | 0.8287% | 5.993% | 0% | 0% | 21.54% | 0% | 6.736% | 22.68% | 0.9798% | 1.86% | 2.179% | 2.554% | 1.302%
| PDM3 |  | -5.309% | 2.386% | 27.35% | -2.06% | -2.06% | -4.392% | 11.93% | 7.267% | -8.586% | 16.59% | -4.859% | 7.664% | 0.5049% | -19.08%
| Z-Score (Absolute) |  | 0.5 | 0.35 | 3.07 | 0.14 | 0.14 | 0.4 | 1.39 | 0.88 | 0.85 | 1.9 | 0.45 | 0.92 | 0.14 | 2
| Analytical results for Be in OREAS 267 (Certified Value 0.52 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.606 | 0.52189 | 0.7 | 0.51 | 0.5 | 0.52 | 0.55 | 0.53 | 0.6 | 0.8 | 1.7 | < 1 | 0.36 |  |  | 1
| 1 | 2 | 0.61 | 0.52326 | 0.6 | 0.51 | 0.5 | 0.52 | 0.55 | 0.52 | 0.6 | 0.8 | 1.7 | < 1 | 0.37 |  |  | 23
| 1 | 3 | 0.547 | 0.51849 | 0.7 | 0.49 | 0.51 | 0.51 | 0.54 | 0.52 | 0.6 | 0.8 | 1.7 | < 1 | 0.43 |  |  | 16
| 1 | 4 | 0.637 | 0.52385 | 0.6 | 0.5 | 0.5 | 0.5 | 0.54 | 0.54 | 0.6 | 0.8 | 1.7 | < 1 | 0.5 |  |  | 0.52094
| 1 | 5 | 0.594 | 0.52636 | 0.6 | 0.5 | 0.5 | 0.53 | 0.56 | 0.53 | 0.6 | 0.8 | 1.7 | < 1 | 0.43 |  |  | 76
| 1 | 6 | 0.603 | 0.52003 | 0.6 | 0.51 | 0.5 | 0.53 | 0.55 | 0.55 | 0.6 | 0.9 | 2 | < 1 | 0.45
| Mean |  | 0.5995 | 0.52231 | 0.63333 | 0.50333 | 0.50167 | 0.51833 | 0.54833 | 0.53167 | 0.6 | 0.81667 | 1.75 |  | 0.42333
| Median |  | 0.6045 | 0.52258 | 0.6 | 0.505 | 0.5 | 0.52 | 0.55 | 0.53 | 0.6 | 0.8 | 1.7 |  | 0.43
| Std Dev. |  | 0.029521 | 0.0028162 | 0.05164 | 0.008165 | 0.0040825 | 0.01169 | 0.0075277 | 0.01169 | 0 | 0.040825 | 0.12247 |  | 0.052026
| Rel.Std.Dev. |  | 4.924% | 0.5392% | 8.154% | 1.622% | 0.8138% | 2.255% | 1.373% | 2.199% | 0% | 4.999% | 6.999% |  | 12.29%
| PDM3 |  | 15.08% | 0.2634% | 21.57% | -3.38% | -3.7% | -0.5006% | 5.258% | 2.059% | 15.18% | 56.77% | 235.9% |  | -18.74%
| Z-Score (Absolute) |  | 3.28 | 0.16 | NA | 0.61 | 0.67 | 0 | 1.21 | 0.54 | NA | NA | NA | 0.74 | 3.84
Results from laboratories 7, 18, 19 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 267 (Certified Value 1.03 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 1.097 | 1.1098 | 1.11 | 0.97 | 0.94 | 0.98 | 1.07 | 1.08 | 1.0104 | 1.02 | 1 | 0.99 | 1.1 | 0.9 | 1.03 | 1
| 1 | 2 | 0.967 | 1.0506 | 1.23 | 1.07 | 1.03 | 0.98 | 1.15 | 1.06 | 1.031 | 1.07 | 0.99 | 0.95 | 1.2 | 1 | 1.11 | 24
| 1 | 3 | 1.264 | 1.0606 | 1.2 | 0.97 | 1.08 | 1.14 | 1.09 | 1.03 | 0.80939 | 1.05 | 1.01 | 0.93 | 1.1 | 0.9 | 1 | 16
| 1 | 4 | 1.026 | 1.0493 | 1.24 | 0.93 | 0.98 | 1.04 | 1.11 | 1.06 | 0.99661 | 1.04 | 1 | 0.97 | 1.1 | 0.9 | 1.14 | 1.0349
| 1 | 5 | 0.978 | 1.078 | 1.15 | 0.91 | 1.03 | 1.04 | 1.09 | 1.12 | 0.97335 | 1.08 | 1.05 | 1.01 | 1.1 | 0.9 | 1.09 | 77
| 1 | 6 | 0.973 | 1.0501 | 1.16 | 0.94 | 0.99 | 0.98 | 1.51 | 1.09 | 0.80647 | 1.1 | 0.97 | 1.01 | 1.2 | 0.9 | 1.01
| Mean |  | 1.0508 | 1.0664 | 1.1817 | 0.965 | 1.0083 | 1.0267 | 1.17 | 1.0733 | 0.93786 | 1.06 | 1.0033 | 0.97667 | 1.1333 | 0.91667 | 1.0633
| Median |  | 1.002 | 1.0556 | 1.18 | 0.955 | 1.01 | 1.01 | 1.1 | 1.07 | 0.98498 | 1.06 | 1 | 0.98 | 1.1 | 0.9 | 1.06
| Std Dev. |  | 0.11539 | 0.02392 | 0.050365 | 0.05648 | 0.048751 | 0.062823 | 0.16876 | 0.030768 | 0.10238 | 0.028983 | 0.026583 | 0.03266 | 0.05164 | 0.040825 | 0.05785
| Rel.Std.Dev. |  | 10.98% | 2.243% | 4.262% | 5.853% | 4.835% | 6.119% | 14.42% | 2.867% | 10.92% | 2.734% | 2.649% | 3.344% | 4.556% | 4.454% | 5.44%
| PDM3 |  | 1.544% | 3.05% | 14.19% | -6.75% | -2.563% | -0.7912% | 13.06% | 3.718% | -9.372% | 2.43% | -3.046% | -5.623% | 9.516% | -11.42% | 2.752%
| Z-Score (Absolute) |  | 0 | 0.22 | 1.86 | 1.22 | 0.6 | 0.34 | 1.69 | 0.32 | 1.6 | 0.13 | 0.67 | 1.05 | 1.17 | 1.9 | 0.18
| Analytical results for Ca in OREAS 267 (Certified Value 0.279 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.24923 | 0.30507 | 0.4 | 0.3 | 0.29 | 0.29 | 0.32 | 0.26 | 0.28014 | 0.3 | 0.26 | 0.2413 | 0.28 | 0.27 | 0.27 | 1
| 1 | 2 | 0.26374 | 0.30583 | 0.43 | 0.3 | 0.29 | 0.3 | 0.31 | 0.25 | 0.25293 | 0.31 | 0.26 | 0.2434 | 0.29 | 0.28 | 0.27 | #N/A
| 1 | 3 | 0.24793 | 0.29573 | 0.41 | 0.28 | 0.29 | 0.3 | 0.31 | 0.26 | 0.21785 | 0.31 | 0.25 | 0.2396 | 0.29 | 0.28 | 0.27 | 16
| 1 | 4 | 0.2613 | 0.2986 | 0.4 | 0.29 | 0.29 | 0.3 | 0.31 | 0.27 | 0.26693 | 0.3 | 0.26 | 0.2428 | 0.29 | 0.28 | 0.26 | 0.27926
| 1 | 5 | 0.24864 | 0.30586 | 0.4 | 0.29 | 0.29 | 0.3 | 0.33 | 0.27 | 0.2846 | 0.31 | 0.26 | 0.243 | 0.29 | 0.28 | 0.26 | 78
| 1 | 6 | 0.2797 | 0.30656 | 0.4 | 0.29 | 0.29 | 0.3 | 0.32 | 0.25 | 0.22328 | 0.3 | 0.26 | 0.2441 | 0.28 | 0.28 | 0.27
| Mean |  | 0.25842 | 0.30294 | 0.40667 | 0.29167 | 0.29 | 0.29833 | 0.31667 | 0.26 | 0.25429 | 0.305 | 0.25833 | 0.24237 | 0.28667 | 0.27833 | 0.26667
| Median |  | 0.25526 | 0.30545 | 0.4 | 0.29 | 0.29 | 0.3 | 0.315 | 0.26 | 0.25993 | 0.305 | 0.26 | 0.2429 | 0.29 | 0.28 | 0.27
| Std Dev. |  | 0.012486 | 0.0045911 | 0.012111 | 0.0075277 | 0 | 0.0040825 | 0.008165 | 0.0089443 | 0.028418 | 0.0054772 | 0.0040825 | 0.0016403 | 0.005164 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 4.832% | 1.516% | 2.978% | 2.581% | 0% | 1.368% | 2.578% | 3.44% | 11.18% | 1.796% | 1.58% | 0.6768% | 1.801% | 1.467% | 1.936%
| PDM3 |  | -7.462% | 8.478% | 45.62% | 4.441% | 3.845% | 6.829% | 13.39% | -6.898% | -8.943% | 9.216% | -7.495% | -13.21% | 2.651% | -0.333% | -4.511%
| Z-Score (Absolute) |  | 0.95 | 0.55 | 4.05 | 0.17 | 0.11 | 0.39 | 1.01 | 0.9 | 1.09 | 0.62 | 0.96 | 1.49 | 0 | 0.28 | 0.67
| Analytical results for Cd in OREAS 267 (Indicative Value 0.026 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.02 | < 0.2 | 0.05 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.17008 | < 0.05 | 0.03 | < 0.3 | < 0.1 | 0.03 |  | 1
| 1 | 2 | 0.016 | < 0.2 | 0.04 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.15937 | < 0.05 | 0.04 | < 0.3 | < 0.1 | 0.03 |  | 1
| 1 | 3 | 0.02 | < 0.2 | 0.05 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.14105 | < 0.05 | 0.03 | < 0.3 | < 0.1 | 0.03 |  | 16
| 1 | 4 | 0.013 | < 0.2 | 0.02 | 0.02 | 0.03 | 0.04 | 0.02 | 0.02 | 0.13743 | < 0.05 | 0.03 | < 0.3 | < 0.1 | 0.04 |  | 0.025796
| 1 | 5 | 0.019 | < 0.2 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.13444 | < 0.05 | 0.03 | < 0.3 | < 0.1 | 0.03 |  | 11
| 1 | 6 | 0.025 | < 0.2 | 0.03 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.11027 | < 0.05 | 0.03 | < 0.3 | < 0.1 | 0.03
| Mean |  | 0.018833 |  | 0.038333 | 0.02 | 0.025 | 0.026667 | 0.02 | 0.02 | 0.14211 |  | 0.031667 |  |  | 0.031667
| Median |  | 0.0195 |  | 0.04 | 0.02 | 0.025 | 0.025 | 0.02 | 0.02 | 0.13924 |  | 0.03 |  |  | 0.03
| Std Dev. |  | 0.0040702 |  | 0.01169 | 0 | 0.0054772 | 0.008165 | 0 | 0 | 0.020861 |  | 0.0040825 |  |  | 0.0040825
| Rel.Std.Dev. |  | 21.61% |  | 30.5% | 0% | 21.91% | 30.62% | 0% | 0% | 14.68% |  | 12.89% |  |  | 12.89%
| PDM3 |  | -26.99% |  | 48.6% | -22.47% | -3.087% | 3.374% | -22.47% | -22.47% | 450.9% |  | 22.76% |  |  | 22.76%
| Z-Score (Absolute) |  | 0.76 | 5.21 | 0.67 | 0.67 | 0.31 | 0.18 | 0.67 | 0.67 | 8.31 | 0.31 | 0.18 | 8.89 | 1.53 | 0.18
| Analytical results for Ce in OREAS 267 (Certified Value 31.9 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 30.97 | 30.6963 | 37.1 | 31.8 | 29.9 | 30.1 | 33.8 | 30.9 | 33.1 | 27.98 | 34.1 | 32 | 27.8 |  |  | 1
| 1 | 2 | 32.727 | 30.7619 | 39.3 | 32.6 | 30.9 | 30.5 | 33.7 | 29.9 | 33.9 | 28.62 | 34.6 | 32 | 28.8 |  |  | 26
| 1 | 3 | 32.475 | 30.3929 | 37.9 | 31.5 | 29.6 | 35.8 | 33.3 | 30.7 | 34.9 | 27.7 | 34.5 | 32 | 28.8 |  |  | 16
| 1 | 4 | 32.275 | 30.9047 | 38.5 | 32.8 | 28.8 | 32.7 | 32.7 | 31.8 | 33.6 | 26.96 | 33.8 | 32 | 28.5 |  |  | 31.8835
| 1 | 5 | 32.093 | 30.3445 | 37.7 | 32.9 | 30.4 | 31.1 | 34.6 | 30.8 | 35 | 28.13 | 32.1 | 32 | 28.7 |  |  | 79
| 1 | 6 | 31.793 | 31.4288 | 36.5 | 31.5 | 29.8 | 30.6 | 33.8 | 32.9 | 33 | 27.56 | 33.2 | 32 | 28.3
| Mean |  | 32.0555 | 30.7548 | 37.8333 | 32.1833 | 29.9 | 31.8 | 33.65 | 31.1667 | 33.9167 | 27.825 | 33.7167 | 32 | 28.4833
| Median |  | 32.184 | 30.7291 | 37.8 | 32.2 | 29.85 | 30.85 | 33.75 | 30.85 | 33.75 | 27.84 | 33.95 | 32 | 28.6
| Std Dev. |  | 0.62027 | 0.39484 | 0.99331 | 0.65549 | 0.71554 | 2.1596 | 0.62849 | 1.0424 | 0.86583 | 0.5627 | 0.9411 | 0 | 0.38687
| Rel.Std.Dev. |  | 1.935% | 1.284% | 2.625% | 2.037% | 2.393% | 6.791% | 1.868% | 3.345% | 2.553% | 2.022% | 2.791% | 0% | 1.358%
| PDM3 |  | 0.5395% | -3.54% | 18.66% | 0.9404% | -6.221% | -0.2619% | 5.541% | -2.248% | 6.377% | -12.73% | 5.75% | 0.3654% | -10.66%
| Z-Score (Absolute) |  | 0.02 | 0.51 | 2.38 | 0.07 | 0.86 | 0.08 | 0.67 | 0.34 | 0.78 | 1.71 | 0.7 | 0 | 1.44
| Analytical results for Co in OREAS 267 (Certified Value 7.02 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 6.135 | 7.1142 | 8.8 | 7 | 7.1 | 6.9 | 7 | 7.4 | 6.4727 | 7.7 | 7.3 | 7.3 | 6 | 5.8 |  | 1
| 1 | 2 | 6.433 | 6.9592 | 9.6 | 7.3 | 7.3 | 7 | 7.5 | 7.5 | 6.552 | 7.9 | 6.9 | 7.6 | 6 | 6 |  | 27
| 1 | 3 | 6.332 | 7.0228 | 9 | 6.8 | 6.9 | 8 | 6.7 | 7.2 | 5.4939 | 7.8 | 7 | 7.8 | 6 | 5.9 |  | 16
| 1 | 4 | 6.42 | 6.9929 | 9.1 | 7.1 | 7.3 | 7.3 | 7.1 | 7.6 | 6.5679 | 7.8 | 6.9 | 7.8 | 6 | 5.7 |  | 7.0171
| 1 | 5 | 6.556 | 7.1003 | 9.1 | 7.1 | 7.2 | 7 | 7.8 | 7.5 | 6.641 | 7.8 | 7.3 | 7.9 | 6 | 5.9 |  | 80
| 1 | 6 | 6.402 | 7.0473 | 9 | 7.1 | 7.1 | 7 | 6.9 | 7.8 | 5.2112 | 7.7 | 7.3 | 7.9 | 6 | 6.1
| Mean |  | 6.3797 | 7.0395 | 9.1 | 7.0667 | 7.15 | 7.2 | 7.1667 | 7.5 | 6.1564 | 7.7833 | 7.1167 | 7.7167 | 6 | 5.9
| Median |  | 6.411 | 7.035 | 9.05 | 7.1 | 7.15 | 7 | 7.05 | 7.5 | 6.5123 | 7.8 | 7.15 | 7.8 | 6 | 5.9
| Std Dev. |  | 0.14014 | 0.060396 | 0.26833 | 0.1633 | 0.15166 | 0.41473 | 0.40825 | 0.2 | 0.63137 | 0.075277 | 0.20412 | 0.23166 | 0 | 0.14142
| Rel.Std.Dev. |  | 2.197% | 0.858% | 2.949% | 2.311% | 2.121% | 5.76% | 5.696% | 2.667% | 10.26% | 0.9672% | 2.868% | 3.002% | 0% | 2.397%
| PDM3 |  | -9.083% | 0.3193% | 29.68% | 0.7071% | 1.895% | 2.607% | 2.132% | 6.882% | -12.26% | 10.92% | 1.42% | 9.97% | -14.49% | -15.92%
| Z-Score (Absolute) |  | 1.48 | 0.21 | 3.76 | 0.16 | 0 | 0.1 | 0.03 | 0.67 | 1.91 | 1.22 | 0.06 | 1.09 | NA | 2.41
| Analytical results for Cr in OREAS 267 (Certified Value 46.8 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 42.383 | 47.4075 | 52 | 47 | 46 | 47 | 49 | 49 | 47.4578 | 49 | 44 | 44 | 45.5 |  |  | 1
| 1 | 2 | 44.688 | 47.0479 | 56 | 47 | 46 | 47 | 49 | 49 | 48.1357 | 49 | 44 | 44 | 46.5 |  |  | 28
| 1 | 3 | 45.376 | 46.7449 | 57 | 45 | 46 | 47 | 48 | 48 | 38.0281 | 49 | 43 | 43 | 46.5 |  |  | 16
| 1 | 4 | 44.019 | 47.2996 | 57 | 46 | 46 | 47 | 49 | 51 | 48.1001 | 50 | 44 | 44 | 45.2 |  |  | 46.7753
| 1 | 5 | 45.498 | 46.9813 | 55 | 47 | 46 | 48 | 51 | 50 | 48.2596 | 50 | 45 | 44 | 45.4 |  |  | 81
| 1 | 6 | 44.726 | 47.6191 | 52 | 47 | 46 | 48 | 49 | 51 | 37.8812 | 49 | 44 | 44 | 46
| Mean |  | 44.4483 | 47.1834 | 54.8333 | 46.5 | 46 | 47.3333 | 49.1667 | 49.6667 | 44.6437 | 49.3333 | 44 | 43.8333 | 45.85
| Median |  | 44.707 | 47.1738 | 55.5 | 47 | 46 | 47 | 49 | 49.5 | 47.7789 | 49 | 44 | 44 | 45.75
| Std Dev. |  | 1.1443 | 0.31782 | 2.3166 | 0.83666 | 0 | 0.5164 | 0.98319 | 1.2111 | 5.1891 | 0.5164 | 0.63246 | 0.40825 | 0.56833
| Rel.Std.Dev. |  | 2.574% | 0.6736% | 4.225% | 1.799% | 0% | 1.091% | 2% | 2.438% | 11.62% | 1.047% | 1.437% | 0.9314% | 1.24%
| PDM3 |  | -4.975% | 0.8725% | 17.23% | -0.5885% | -1.657% | 1.193% | 5.113% | 6.181% | -4.557% | 5.469% | -5.933% | -6.29% | -1.978%
| Z-Score (Absolute) |  | 0.67 | 0.22 | 2.74 | 0 | 0.16 | 0.27 | 0.88 | 1.04 | 0.61 | 0.93 | 0.82 | 0.88 | 0.21
| Analytical results for Cs in OREAS 267 (Certified Value 1.18 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 1.187 | 1.2515 | 1.19 | 1.11 | 1.18 | 1.14 | 1.24 | 1.2 | 1.55 | 1.11 | 11.7 | 1.2 | 1.17 |  |  | 1
| 1 | 2 | 1.265 | 1.3022 | 1.3 | 1.14 | 1.22 | 1.12 | 1.28 | 1.19 | 1.51 | 1.1 | 11.02 | 1.2 | 1.2 |  |  | 29
| 1 | 3 | 1.116 | 1.1817 | 1.23 | 1.1 | 1.18 | 1.3 | 1.19 | 1.2 | 1.55 | 1.09 | 11.94 | 1.2 | 1.2 |  |  | 16
| 1 | 4 | 1.039 | 1.2686 | 1.23 | 1.13 | 1.2 | 1.22 | 1.21 | 1.22 | 1.55 | 1.09 | 11.88 | 1.2 | 1.2 |  |  | 1.1847
| 1 | 5 | 1.109 | 1.282 | 1.26 | 1.15 | 1.2 | 1.16 | 1.22 | 1.2 | 1.56 | 1.12 | 11.88 | 1.2 | 1.19 |  |  | 82
| 1 | 6 | 1.192 | 1.258 | 1.12 | 1.11 | 1.18 | 1.14 | 1.23 | 1.24 | 1.56 | 1.08 | 11.86 | 1.2 | 1.2
| Mean |  | 1.1513 | 1.2573 | 1.2217 | 1.1233 | 1.1933 | 1.18 | 1.2283 | 1.2083 | 1.5467 | 1.0983 | 11.7133 | 1.2 | 1.1933
| Median |  | 1.1515 | 1.2633 | 1.23 | 1.12 | 1.19 | 1.15 | 1.225 | 1.2 | 1.55 | 1.095 | 11.87 | 1.2 | 1.2
| Std Dev. |  | 0.079379 | 0.041235 | 0.061779 | 0.019664 | 0.01633 | 0.068118 | 0.030605 | 0.018348 | 0.018619 | 0.01472 | 0.34909 | 0 | 0.012111
| Rel.Std.Dev. |  | 6.895% | 3.28% | 5.057% | 1.75% | 1.368% | 5.773% | 2.492% | 1.518% | 1.204% | 1.34% | 2.98% | 0% | 1.015%
| PDM3 |  | -2.814% | 6.133% | 3.123% | -5.177% | 0.7317% | -0.3938% | 3.686% | 1.998% | 30.56% | -7.287% | 888.7% | 1.294% | 0.7317%
| Z-Score (Absolute) |  | 1.08 | 1.44 | 0.59 | 1.75 | 0.08 | 0.4 | 0.75 | 0.28 | 8.33 | 2.34 | NA | 0.08 | 0.08
| Analytical results for Cu in OREAS 267 (Certified Value 19.7 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 19.214 | 19.5413 | 21.5 | 18.9 | 19.2 | 19.8 | 19.5 | 20.5 | 22.8959 | 20 | 19.6 | 20 | 20.1 | 19.9 | 17.6 | 1
| 1 | 2 | 19.485 | 19.4014 | 23.1 | 19.7 | 19.4 | 19.8 | 20.3 | 20.4 | 20.0579 | 21 | 19.3 | 21 | 20.5 | 18.7 | 18 | 30
| 1 | 3 | 19.612 | 19.4603 | 22.7 | 18.6 | 18.8 | 22.2 | 19.4 | 20.4 | 17.5089 | 21 | 18.9 | 20 | 19.9 | 18.5 | 18.2 | 16
| 1 | 4 | 20.269 | 19.5577 | 22.6 | 19.2 | 19.3 | 22.5 | 19.8 | 20.7 | 21.1477 | 21 | 18.8 | 20 | 23.6 | 18.6 | 17.8 | 19.7317
| 1 | 5 | 20.126 | 19.368 | 23 | 19.4 | 19.4 | 19.8 | 20.1 | 20.6 | 23.1364 | 21 | 19.5 | 21 | 23.9 | 18.6 | 17.5 | 83
| 1 | 6 | 18.693 | 19.8256 | 21.7 | 19.2 | 19.1 | 19.7 | 19.8 | 21.3 | 18.0015 | 20 | 19.2 | 21 | 21.7 | 18.6 | 18
| Mean |  | 19.5665 | 19.5257 | 22.4333 | 19.1667 | 19.2 | 20.6333 | 19.8167 | 20.65 | 20.4581 | 20.6667 | 19.2167 | 20.5 | 21.6167 | 18.8167 | 17.85
| Median |  | 19.5485 | 19.5008 | 22.65 | 19.2 | 19.25 | 19.8 | 19.8 | 20.55 | 20.6028 | 21 | 19.25 | 20.5 | 21.1 | 18.6 | 17.9
| Std Dev. |  | 0.58353 | 0.16478 | 0.67429 | 0.38297 | 0.22804 | 1.3337 | 0.34303 | 0.33912 | 2.3868 | 0.5164 | 0.31885 | 0.54772 | 1.7691 | 0.53448 | 0.26646
| Rel.Std.Dev. |  | 2.982% | 0.8439% | 3.006% | 1.998% | 1.188% | 6.464% | 1.731% | 1.642% | 11.67% | 2.499% | 1.659% | 2.672% | 8.184% | 2.84% | 1.493%
| PDM3 |  | -0.8372% | -1.044% | 13.69% | -2.864% | -2.695% | 4.57% | 0.4307% | 4.654% | 3.681% | 4.738% | -2.61% | 3.894% | 9.553% | -4.637% | -9.536%
| Z-Score (Absolute) |  | 0.25 | 0.29 | 2.58 | 0.64 | 0.61 | 0.81 | 0 | 0.82 | 0.63 | 0.84 | 0.59 | 0.67 | 1.78 | 0.99 | 1.94
| Analytical results for Dy in OREAS 267 (Indicative Value 1.61 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 1.5592 | 1.6 | [13 empty] | 1
| 1 | 2 | 1.5539 | 1.7 | [13 empty] | 6
| 1 | 3 | 1.5236 | 1.7 | [13 empty] | 16
| 1 | 4 | 1.5183 | 1.7 | [13 empty] | 1.6076
| 1 | 5 | 1.5417 | 1.7 | [13 empty] | 12
| 1 | 6 | 1.5943 | 1.6
| Mean |  | 1.5485 | 1.6667
| Median |  | 1.5478 | 1.7
| Std Dev. |  | 0.027635 | 0.05164
| Rel.Std.Dev. |  | 1.785% | 3.098%
| PDM3 |  | -3.676% | 3.676%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Er in OREAS 267 (Indicative Value 0.66 ppm) | [16 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 0.63549 | 0.6 | [13 empty] | 1
| 1 | 2 | 0.66161 | 0.7 | [13 empty] | 7
| 1 | 3 | 0.65717 | 0.6 | [13 empty] | 16
| 1 | 4 | 0.62486 | 0.7 | [13 empty] | 0.65586
| 1 | 5 | 0.65317 | 0.7 | [13 empty] | 13
| 1 | 6 | 0.63807 | 0.7
| Mean |  | 0.64506 | 0.66667
| Median |  | 0.64562 | 0.7
| Std Dev. |  | 0.014389 | 0.05164
| Rel.Std.Dev. |  | 2.231% | 7.746%
| PDM3 |  | -1.647% | 1.647%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 267 (Indicative Value 0.57 ppm) | [16 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 0.54351 | 0.6 | [13 empty] | 1
| 1 | 2 | 0.54491 | 0.6 | [13 empty] | 8
| 1 | 3 | 0.53435 | 0.6 | [13 empty] | 16
| 1 | 4 | 0.54876 | 0.6 | [13 empty] | 0.5728
| 1 | 5 | 0.5376 | 0.6 | [13 empty] | 14
| 1 | 6 | 0.56452 | 0.6
| Mean |  | 0.54561 | 0.6
| Median |  | 0.54421 | 0.6
| Std Dev. |  | 0.010615 | 0
| Rel.Std.Dev. |  | 1.945% | 0%
| PDM3 |  | -4.748% | 4.748%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe in OREAS 267 (Certified Value 2.02 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 1.9232 | 2.0456 | 2.34 | 2.02 | 1.95 | 1.98 | 2.07 | 2.2 | 1.9531 | 2.32 | 1.95 | 1.7374 | 2.15 | 2 | 1.89 | 1
| 1 | 2 | 2.009 | 2.0445 | 2.44 | 2.01 | 1.96 | 2.01 | 2.08 | 2.19 | 1.9732 | 2.3 | 1.98 | 1.7332 | 2.18 | 2.02 | 1.91 | #N/A
| 1 | 3 | 1.9388 | 2.0402 | 2.46 | 1.96 | 1.98 | 2.08 | 2.05 | 2.15 | 1.5308 | 2.3 | 1.92 | 1.7131 | 2.21 | 2.01 | 1.9 | 16
| 1 | 4 | 2.0391 | 2.0379 | 2.36 | 2 | 1.95 | 2.07 | 2.07 | 2.25 | 1.9803 | 2.28 | 1.96 | 1.7342 | 2.11 | 1.99 | 1.88 | 2.0245
| 1 | 5 | 2.0457 | 2.0491 | 2.33 | 1.99 | 1.96 | 2.03 | 2.12 | 2.22 | 1.9846 | 2.33 | 1.94 | 1.7326 | 2.14 | 2.03 | 1.84 | 84
| 1 | 6 | 1.9334 | 2.0539 | 2.36 | 2.01 | 1.97 | 2.04 | 2.09 | 2.27 | 1.5396 | 2.29 | 1.98 | 1.7458 | 2.17 | 2.03 | 1.85
| Mean |  | 1.9816 | 2.0452 | 2.3817 | 1.9983 | 1.9617 | 2.035 | 2.08 | 2.2133 | 1.8269 | 2.3033 | 1.955 | 1.7327 | 2.16 | 2.0133 | 1.8783
| Median |  | 1.9739 | 2.0451 | 2.36 | 2.005 | 1.96 | 2.035 | 2.075 | 2.21 | 1.9631 | 2.3 | 1.955 | 1.7337 | 2.16 | 2.015 | 1.885
| Std Dev. |  | 0.056102 | 0.0058282 | 0.054559 | 0.02137 | 0.01169 | 0.037283 | 0.023664 | 0.043205 | 0.22623 | 0.018619 | 0.023452 | 0.010774 | 0.034641 | 0.01633 | 0.027869
| Rel.Std.Dev. |  | 2.831% | 0.285% | 2.291% | 1.069% | 0.5959% | 1.832% | 1.138% | 1.952% | 12.38% | 0.8083% | 1.2% | 0.6218% | 1.604% | 0.8111% | 1.484%
| PDM3 |  | -2.123% | 1.021% | 17.64% | -1.295% | -3.106% | 0.5163% | 2.739% | 9.325% | -9.761% | 13.77% | -3.435% | -14.41% | 6.691% | -0.5539% | -7.222%
| Z-Score (Absolute) |  | 0.32 | 0.32 | 3.73 | 0.15 | 0.52 | 0.22 | 0.67 | 2.02 | 1.89 | 2.93 | 0.59 | 2.84 | 1.48 | 0 | 1.37
| Analytical results for Ga in OREAS 267 (Certified Value 3.71 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 4.6019 | 4.51 | 3.47 | 3.48 | 3.44 | 3.72 | 3.57 | 3.1599 | 4.3 | 3.5 | 12.8 | 4 | 3.24 |  |  | 1
| 1 | 2 | 4.4422 | 4.91 | 3.59 | 3.56 | 3.46 | 3.79 | 3.65 | 4.3612 | 4.4 | 3.5 | 13.5 | 4 | 3.25 |  |  | 34
| 1 | 3 | 4.111 | 4.64 | 3.4 | 3.43 | 3.8 | 3.61 | 3.64 | 3.2045 | 4.3 | 3.4 | 13.4 | 4 | 3.22 |  |  | 16
| 1 | 4 | 3.8974 | 4.94 | 3.48 | 3.54 | 3.6 | 3.56 | 3.73 | 4.6304 | 4.4 | 3.5 | 13.8 | 4 | 3.13 |  |  | 3.7107
| 1 | 5 | 4.0842 | 4.97 | 3.53 | 3.56 | 3.53 | 3.71 | 3.69 | 4.9583 | 4.4 | 3.6 | 13.8 | 4 | 3.24 |  |  | 85
| 1 | 6 | 4.3078 | 4.48 | 3.51 | 3.54 | 3.59 | 3.67 | 3.82 | 2.8813 | 4.3 | 3.3 | 13.5 | 4 | 3.35
| Mean |  | 4.2407 | 4.7417 | 3.4967 | 3.5183 | 3.57 | 3.6767 | 3.6833 | 3.8659 | 4.35 | 3.4667 | 13.4667 | 4 | 3.2383
| Median |  | 4.2094 | 4.775 | 3.495 | 3.54 | 3.56 | 3.69 | 3.67 | 3.7828 | 4.35 | 3.5 | 13.5 | 4 | 3.24
| Std Dev. |  | 0.25873 | 0.22463 | 0.06377 | 0.052313 | 0.13023 | 0.082381 | 0.085713 | 0.88638 | 0.054772 | 0.10328 | 0.36697 | 0 | 0.070261
| Rel.Std.Dev. |  | 6.101% | 4.737% | 1.824% | 1.487% | 3.648% | 2.241% | 2.327% | 22.93% | 1.259% | 2.979% | 2.725% | 0% | 2.17%
| PDM3 |  | 14.29% | 27.78% | -5.767% | -5.183% | -3.791% | -0.9163% | -0.7366% | 4.184% | 17.23% | -6.576% | 262.9% | 7.797% | -12.73%
| Z-Score (Absolute) |  | 1.89 | 3.59 | 0.62 | 0.55 | 0.37 | 0.01 | 0.01 | 0.63 | 2.26 | 0.72 | 33.06 | NA | 1.49
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 267 (Indicative Value 2.43 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 2.203 | 2.5 | [13 empty] | 1
| 1 | 2 | 2.2451 | 2.8 | [13 empty] | 9
| 1 | 3 | 2.2472 | 2.5 | [13 empty] | 16
| 1 | 4 | 2.1984 | 2.7 | [13 empty] | 2.4256
| 1 | 5 | 2.2342 | 2.7 | [13 empty] | 15
| 1 | 6 | 2.2789 | 2.5
| Mean |  | 2.2344 | 2.6167
| Median |  | 2.2396 | 2.6
| Std Dev. |  | 0.030142 | 0.13292
| Rel.Std.Dev. |  | 1.349% | 5.08%
| PDM3 |  | -7.879% | 7.879%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ge in OREAS 267 (Indicative Value 0.073 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.06 | 0.05 | 0.07 | 1.2349 | 0.1 | < 0.5 | 0.9 | 0.1 |  |  |  | 1
| 1 | 2 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 1.6946 | 0.1 | < 0.5 | 0.9 | 0.11 |  |  |  | 2
| 1 | 3 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.06 | 0.05 | 0.07 | 1.2423 | 0.1 | < 0.5 | 0.9 | 0.09 |  |  |  | 16
| 1 | 4 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 1.8453 | 0.1 | < 0.5 | 0.9 | 0.08 |  |  |  | 0.0725
| 1 | 5 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.07 | 0.06 | 0.07 | 1.8977 | 0.1 | < 0.5 | 0.9 | 0.1 |  |  |  | 10
| 1 | 6 | < 0.2 | < 0.1 | < 0.05 | 0.05 | 0.07 | 0.05 | 0.08 | 1.1318 | 0.1 | < 0.5 | 0.9 | 0.09
| Mean |  |  |  |  | 0.05 | 0.063333 | 0.055 | 0.071667 | 1.5078 | 0.1 |  | 0.9 | 0.095
| Median |  |  |  |  | 0.05 | 0.06 | 0.055 | 0.07 | 1.4684 | 0.1 |  | 0.9 | 0.095
| Std Dev. |  |  |  |  | 7.6012e-18 | 0.005164 | 0.0054772 | 0.0040825 | 0.34268 | 1.5202e-17 |  | 0 | 0.010488
| Rel.Std.Dev. |  |  |  |  | 1.52e-14% | 8.154% | 9.959% | 5.696% | 22.73% | 1.52e-14% |  | 0% | 11.04%
| PDM3 |  |  |  |  | -31.03% | -12.64% | -24.14% | -1.149% | 1980% | 37.93% |  | 1141% | 31.03%
| Z-Score (Absolute) |  | 0.36 | 0.73 | 1.28 | 0.73 | 0.44 | 0.62 | 0.26 | 31.15 | 0.36 | 3.64 | 17.86 | 0.26
| Analytical results for Hf in OREAS 267 (Certified Value 0.48 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.5g |  |  |  | 2
| 1 | 1 | 0.42027 | < 0.1 | 0.48 | 0.45 | 0.42 | 0.44 | 0.45 | 1.2388 | 0.51 | 0.56 | 0.5 | 0.54 |  |  |  | 1
| 1 | 2 | 0.4319 | < 0.1 | 0.52 | 0.48 | 0.41 | 0.45 | 0.43 | 0.97341 | 0.53 | 0.57 | 0.5 | 0.53 |  |  |  | 20
| 1 | 3 | 0.41466 | < 0.1 | 0.52 | 0.46 | 0.52 | 0.42 | 0.43 | 1.1472 | 0.52 | 0.57 | 0.5 | 0.54 |  |  |  | 16
| 1 | 4 | 0.41629 | < 0.1 | 0.52 | 0.45 | 0.48 | 0.44 | 0.45 | 0.94403 | 0.53 | 0.58 | 0.5 | 0.53 |  |  |  | 0.48259
| 1 | 5 | 0.40919 | < 0.1 | 0.52 | 0.47 | 0.44 | 0.44 | 0.46 | 0.92466 | 0.52 | 0.58 | 0.5 | 0.54 |  |  |  | 86
| 1 | 6 | 0.41742 | < 0.1 | 0.51 | 0.45 | 0.42 | 0.43 | 0.44 | 0.81557 | 0.53 | 0.55 | 0.5 | 0.52
| Mean |  | 0.41829 |  | 0.51167 | 0.46 | 0.44833 | 0.43667 | 0.44333 | 1.0073 | 0.52333 | 0.56833 | 0.5 | 0.53333
| Median |  | 0.41685 |  | 0.52 | 0.455 | 0.43 | 0.44 | 0.445 | 0.95872 | 0.525 | 0.57 | 0.5 | 0.535
| Std Dev. |  | 0.0076139 |  | 0.016021 | 0.012649 | 0.043089 | 0.010328 | 0.012111 | 0.15616 | 0.008165 | 0.01169 | 0 | 0.008165
| Rel.Std.Dev. |  | 1.82% |  | 3.131% | 2.75% | 9.611% | 2.365% | 2.732% | 15.5% | 1.56% | 2.057% | 0% | 1.531%
| PDM3 |  | -13.32% |  | 6.026% | -4.681% | -7.098% | -9.516% | -8.134% | 108.7% | 8.443% | 17.77% | 3.608% | 10.52%
| Z-Score (Absolute) |  | 0.54 | 5.35 | 0.67 | 0 | 0.15 | 0.3 | 0.22 | 7.14 | 0.83 | 1.41 | NA | 0.96
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 267 (Certified Value 0.043 ppm) | [16 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.21 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.043 | < 1 | 0.05 | 0.03 | 0.04 | 0.05 | 0.05 | 0.04 | 0.017724 | 0.05 | 0.05 | 0.173 | 0.07 |  |  | 1
| 1 | 2 | 0.046 | < 1 | 0.05 | 0.03 | 0.05 | 0.04 | 0.05 | 0.04 | < 0.005 | 0.05 | 0.05 | 0.171 | 0.08 |  |  | 1
| 1 | 3 | 0.051 | < 1 | 0.04 | 0.04 | 0.04 | 0.06 | 0.04 | 0.04 | 0.0047731 | 0.04 | 0.05 | 0.157 | 0.07 |  |  | 16
| 1 | 4 | 0.042 | < 1 | 0.06 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.012569 | 0.04 | 0.06 | 0.154 | 0.08 |  |  | 0.042833
| 1 | 5 | 0.041 | < 1 | 0.04 | 0.03 | 0.03 | 0.05 | 0.04 | 0.04 | < 0.005 | 0.05 | 0.05 | 0.159 | 0.07 |  |  | 87
| 1 | 6 | 0.03 | < 1 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.0052747 | 0.05 | 0.05 | 0.159 | 0.08
| Mean |  | 0.042167 |  | 0.046667 | 0.031667 | 0.04 | 0.043333 | 0.043333 | 0.04 | 0.010085 | 0.046667 | 0.051667 | 0.16217 | 0.075
| Median |  | 0.0425 |  | 0.045 | 0.03 | 0.04 | 0.045 | 0.04 | 0.04 | 0.0089217 | 0.05 | 0.05 | 0.159 | 0.075
| Std Dev. |  | 0.006969 |  | 0.008165 | 0.0040825 | 0.0063246 | 0.012111 | 0.005164 | 0 | 0.0062148 | 0.005164 | 0.0040825 | 0.0078592 | 0.0054772
| Rel.Std.Dev. |  | 16.53% |  | 17.5% | 12.89% | 15.81% | 27.95% | 11.92% | 0% | 61.62% | 11.07% | 7.902% | 4.846% | 7.303%
| PDM3 |  | -1.556% |  | 8.949% | -26.07% | -6.615% | 1.167% | 1.167% | -6.615% | -76.46% | 8.949% | 20.62% | 278.6% | 75.1%
| Z-Score (Absolute) |  | 0.24 | 92.38 | 0.67 | 2.36 | 0.67 | 0 | 0 | 0.67 | 7.24 | 0.67 | 1.69 | 24.04 | 6.41
| Analytical results for Ho in OREAS 267 (Indicative Value 0.28 ppm) | [16 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 0.25699 | 0.3 | [13 empty] | 1
| 1 | 2 | 0.25978 | 0.3 | [13 empty] | 5
| 1 | 3 | 0.25968 | 0.3 | [13 empty] | 16
| 1 | 4 | 0.2578 | 0.3 | [13 empty] | 0.28012
| 1 | 5 | 0.26378 | 0.3 | [13 empty] | 11
| 1 | 6 | 0.2634 | 0.3
| Mean |  | 0.26024 | 0.3
| Median |  | 0.25973 | 0.3
| Std Dev. |  | 0.0028125 | 0
| Rel.Std.Dev. |  | 1.081% | 0%
| PDM3 |  | -7.097% | 7.097%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 267 (Certified Value 0.02 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | < 0.1 | 0.02 | 0.02 | 0.019 | 0.019 | 0.02 | 0.02 | 0.020825 | < 0.05 | < 0.02 | < 0.5 | 0.02 | 0.017 |  |  | 1
| 1 | 2 | < 0.1 | 0.03 | 0.019 | 0.019 | 0.022 | 0.021 | 0.021 | 0.024478 | < 0.05 | < 0.02 | < 0.5 | 0.02 | 0.017 |  |  | 22
| 1 | 3 | < 0.1 | 0.03 | 0.022 | 0.019 | 0.023 | 0.019 | 0.02 | 0.016488 | < 0.05 | < 0.02 | < 0.5 | 0.02 | 0.018 |  |  | 16
| 1 | 4 | < 0.1 | 0.03 | 0.021 | 0.018 | 0.056 | 0.021 | 0.02 | 0.021064 | < 0.05 | 0.02 | < 0.5 | 0.02 | 0.016 |  |  | 0.020233
| 1 | 5 | < 0.1 | 0.03 | 0.021 | 0.02 | 0.017 | 0.023 | 0.02 | 0.018913 | < 0.05 | < 0.02 | < 0.5 | 0.02 | 0.019 |  |  | 88
| 1 | 6 | < 0.1 | 0.03 | 0.02 | 0.019 | 0.023 | 0.022 | 0.023 | 0.017615 | < 0.05 | < 0.02 | < 0.5 | 0.02 | 0.016
| Mean |  |  | 0.028333 | 0.0205 | 0.019 | 0.026667 | 0.021 | 0.020667 | 0.019897 |  | 0.02 |  | 0.02 | 0.017167
| Median |  |  | 0.03 | 0.0205 | 0.019 | 0.0225 | 0.021 | 0.02 | 0.019869 |  | 0.02 |  | 0.02 | 0.017
| Std Dev. |  |  | 0.0040825 | 0.0010488 | 0.00063246 | 0.014569 | 0.0014142 | 0.0012111 | 0.0028636 |  |  |  | 0 | 0.001169
| Rel.Std.Dev. |  |  | 14.41% | 5.116% | 3.329% | 54.64% | 6.734% | 5.86% | 14.39% |  |  |  | 0% | 6.81%
| PDM3 |  |  | 40.04% | 1.32% | -6.094% | 31.8% | 3.791% | 2.144% | -1.66% |  | -1.151% |  | -1.151% | -15.15%
| Z-Score (Absolute) |  | 5.65 | 1.48 | 0.03 | 0.32 | 1.16 | 0.06 | 0 | 0.15 | 0.83 | 1.73 | 44.18 | 0.13 | 0.67
| Analytical results for K in OREAS 267 (Certified Value 0.158 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.13276 | 0.15696 | 0.17 | 0.15 | 0.15 | 0.15 | 0.16 | 0.16 | 0.20642 | 0.188 | 0.14 | 0.19 | 0.15 | 0.14 |  | 1
| 1 | 2 | 0.13357 | 0.15636 | 0.2 | 0.15 | 0.14 | 0.15 | 0.16 | 0.16 | 0.18949 | 0.188 | 0.15 | 0.19 | 0.15 | 0.14 |  | #N/A
| 1 | 3 | 0.13341 | 0.15577 | 0.19 | 0.14 | 0.15 | 0.15 | 0.16 | 0.16 | 0.16127 | 0.189 | 0.14 | 0.2 | 0.15 | 0.14 |  | 16
| 1 | 4 | 0.13544 | 0.15386 | 0.19 | 0.15 | 0.14 | 0.15 | 0.16 | 0.17 | 0.19813 | 0.188 | 0.15 | 0.21 | 0.15 | 0.13 |  | 0.15821
| 1 | 5 | 0.13543 | 0.15626 | 0.2 | 0.15 | 0.14 | 0.15 | 0.17 | 0.17 | 0.21211 | 0.188 | 0.14 | 0.22 | 0.15 | 0.14 |  | 89
| 1 | 6 | 0.12811 | 0.15886 | 0.18 | 0.15 | 0.14 | 0.15 | 0.16 | 0.17 | 0.16753 | 0.188 | 0.15 | 0.2 | 0.15 | 0.14
| Mean |  | 0.13312 | 0.15635 | 0.18833 | 0.14833 | 0.14333 | 0.15 | 0.16167 | 0.165 | 0.18916 | 0.18817 | 0.145 | 0.20167 | 0.15 | 0.13833
| Median |  | 0.13349 | 0.15631 | 0.19 | 0.15 | 0.14 | 0.15 | 0.16 | 0.165 | 0.19381 | 0.188 | 0.145 | 0.2 | 0.15 | 0.14
| Std Dev. |  | 0.0026909 | 0.0016262 | 0.01169 | 0.0040825 | 0.005164 | 0 | 0.0040825 | 0.0054772 | 0.02074 | 0.00040825 | 0.0054772 | 0.01169 | 0 | 0.0040825
| Rel.Std.Dev. |  | 2.021% | 1.04% | 6.207% | 2.752% | 3.603% | 0% | 2.525% | 3.32% | 10.96% | 0.217% | 3.777% | 5.797% | 0% | 2.951%
| PDM3 |  | -15.86% | -1.181% | 19.04% | -6.246% | -9.406% | -5.192% | 2.182% | 4.289% | 19.56% | 18.93% | -8.352% | 27.46% | -5.192% | -12.57%
| Z-Score (Absolute) |  | 1.25 | 0.2 | 2.19 | 0.3 | 0.61 | 0.2 | 0.53 | 0.74 | 2.24 | 2.18 | 0.51 | 3.02 | 0.2 | 0.92
| Analytical results for La in OREAS 267 (Certified Value 15.2 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 2
| 1 | 1 | 14.7484 | 18.7 | 15.8 | 15.2 | 15.2 | 15.3 | 15.4 | 17.1 | 13.6 | 32 | 15 | 12.6 |  |  |  | 1
| 1 | 2 | 14.9303 | 20.9 | 16 | 15.5 | 15.3 | 16.6 | 15.6 | 17.5 | 13.8 | 33.5 | 15 | 12.8 |  |  |  | 23
| 1 | 3 | 14.8521 | 19.4 | 15.6 | 15 | 17.5 | 14.9 | 15.3 | 17.8 | 13.4 | 33.1 | 14 | 12.8 |  |  |  | 16
| 1 | 4 | 14.6796 | 20.2 | 15.8 | 15.3 | 15.9 | 15.4 | 16.1 | 17.3 | 13.7 | 32.2 | 15 | 12.7 |  |  |  | 15.1627
| 1 | 5 | 14.713 | 20.1 | 16 | 15 | 15.8 | 16.6 | 15.6 | 17.4 | 14 | 32 | 15 | 12.6 |  |  |  | 90
| 1 | 6 | 14.9362 | 18.8 | 16.2 | 14.6 | 15.3 | 15.4 | 16.2 | 17.2 | 13.4 | 32.9 | 15 | 12.8
| Mean |  | 14.8099 | 19.6833 | 15.9 | 15.1 | 15.8333 | 15.7 | 15.7 | 17.3833 | 13.65 | 32.6167 | 14.8333 | 12.7167
| Median |  | 14.8003 | 19.75 | 15.9 | 15.1 | 15.55 | 15.4 | 15.6 | 17.35 | 13.65 | 32.55 | 15 | 12.75
| Std Dev. |  | 0.11168 | 0.86583 | 0.20976 | 0.30984 | 0.86641 | 0.72111 | 0.36878 | 0.24833 | 0.23452 | 0.63692 | 0.40825 | 0.098319
| Rel.Std.Dev. |  | 0.7541% | 4.399% | 1.319% | 2.052% | 5.472% | 4.593% | 2.349% | 1.429% | 1.718% | 1.953% | 2.752% | 0.7732%
| PDM3 |  | -2.326% | 29.81% | 4.863% | -0.4133% | 4.423% | 3.544% | 3.544% | 14.65% | -9.976% | 115.1% | -2.172% | -16.13%
| Z-Score (Absolute) |  | 0.68 | 3.06 | 0.15 | 0.46 | 0.1 | 0 | 0 | 1.29 | 1.57 | 12.99 | 0.67 | 2.29
| Analytical results for Li in OREAS 267 (Certified Value 5.56 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 5.247 | 5.4417 | 6.2 | 5.3 | 5.1 | 5 | 5.7 | 5.7 | 5.4964 | 6.4 | 5 | 6.5 | 4.4 | 3.9 |  | 1
| 1 | 2 | 5.335 | 5.6127 | 6.3 | 5.2 | 5.1 | 5.1 | 5.7 | 5.8 | 4.2379 | 6.5 | 5 | 6.4 | 4.5 | 3.9 |  | #N/A
| 1 | 3 | 5.623 | 5.5667 | 6.8 | 5 | 5.2 | 5.2 | 5.6 | 5.6 | 4.4996 | 6.6 | 5 | 6.8 | 4.8 | 3.9 |  | 16
| 1 | 4 | 5.572 | 5.6385 | 6.2 | 5.2 | 5.1 | 5.3 | 5.6 | 6 | 5.223 | 6.5 | 5 | 6.3 | 4.6 | 3.9 |  | 5.5561
| 1 | 5 | 5.511 | 5.4967 | 6 | 5.1 | 5.2 | 5.2 | 5.9 | 5.9 | 6.1627 | 6.5 | 5 | 6.4 | 4.8 | 3.9 |  | 91
| 1 | 6 | 5.243 | 5.533 | 6.3 | 5.2 | 5.2 | 5.2 | 5.8 | 6 | 4.717 | 6.4 | 5 | 6.3 | 4.2 | 4
| Mean |  | 5.4218 | 5.5482 | 6.3 | 5.1667 | 5.15 | 5.1667 | 5.7167 | 5.8333 | 5.0561 | 6.4833 | 5 | 6.45 | 4.55 | 3.9167
| Median |  | 5.423 | 5.5499 | 6.25 | 5.2 | 5.15 | 5.2 | 5.7 | 5.85 | 4.97 | 6.5 | 5 | 6.4 | 4.55 | 3.9
| Std Dev. |  | 0.16796 | 0.073348 | 0.26833 | 0.10328 | 0.054772 | 0.10328 | 0.1169 | 0.1633 | 0.7128 | 0.075277 | 0 | 0.18708 | 0.23452 | 0.040825
| Rel.Std.Dev. |  | 3.098% | 1.322% | 4.259% | 1.999% | 1.064% | 1.999% | 2.045% | 2.799% | 14.1% | 1.161% | 0% | 2.901% | 5.154% | 1.042%
| PDM3 |  | -2.416% | -0.1409% | 13.39% | -7.009% | -7.309% | -7.009% | 2.89% | 4.99% | -8.999% | 16.69% | -10.01% | 16.09% | -18.11% | -29.51%
| Z-Score (Absolute) |  | 0 | 0.23 | 1.62 | 0.47 | 0.5 | 0.47 | 0.54 | 0.76 | 0.67 | 1.96 | NA | 1.9 | 1.61 | 2.78
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 267 (Indicative Value 0.062 ppm) | [16 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.06 | 1.07 | 1.19 | 1.22 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.2g | 0.5g | 0.25g | 01g | [11 empty] | 3
| 1 | 1 | 0.064482 | < 0.1 | 0.06 | < 1 | [11 empty] | 1
| 1 | 2 | 0.060908 | < 0.1 | 0.06 | < 1 | [11 empty] | 6
| 1 | 3 | 0.056507 | < 0.1 | 0.06 | < 1 | [11 empty] | 16
| 1 | 4 | 0.061506 | < 0.1 | 0.07 | < 1 | [11 empty] | 0.062021
| 1 | 5 | 0.059966 | < 0.1 | 0.07 | < 1 | [11 empty] | 12
| 1 | 6 | 0.060887 | < 0.1 | 0.06 | < 1
| Mean |  | 0.060709 |  | 0.063333
| Median |  | 0.060897 |  | 0.06
| Std Dev. |  | 0.0025748 |  | 0.005164
| Rel.Std.Dev. |  | 4.241% |  | 8.154%
| PDM3 |  | -2.115% |  | 2.115%
| Z-Score (Absolute) |  | 0.13 | 1.22 | 0.13 | 44.3
| Analytical results for Mg in OREAS 267 (Certified Value 0.371 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.38485 | 0.37942 | 0.42 | 0.37 | 0.36 | 0.35 | 0.37 | 0.38 | 0.4145 | 0.41 | 0.35 | 0.37 | 0.39 | 0.34 |  | 1
| 1 | 2 | 0.38301 | 0.37749 | 0.53 | 0.37 | 0.36 | 0.36 | 0.37 | 0.38 | 0.38061 | 0.41 | 0.35 | 0.38 | 0.39 | 0.34 |  | #N/A
| 1 | 3 | 0.38189 | 0.37818 | 0.41 | 0.35 | 0.36 | 0.35 | 0.36 | 0.37 | 0.3247 | 0.41 | 0.34 | 0.38 | 0.38 | 0.34 |  | 16
| 1 | 4 | 0.38503 | 0.37333 | 0.5 | 0.36 | 0.36 | 0.36 | 0.37 | 0.4 | 0.39279 | 0.4 | 0.35 | 0.37 | 0.38 | 0.33 |  | 0.37136
| 1 | 5 | 0.38067 | 0.38265 | 0.49 | 0.36 | 0.36 | 0.36 | 0.39 | 0.39 | 0.43085 | 0.41 | 0.36 | 0.35 | 0.39 | 0.34 |  | 92
| 1 | 6 | 0.37569 | 0.37641 | 0.36 | 0.37 | 0.36 | 0.36 | 0.37 | 0.4 | 0.33378 | 0.4 | 0.35 | 0.37 | 0.39 | 0.33
| Mean |  | 0.38186 | 0.37791 | 0.45167 | 0.36333 | 0.36 | 0.35667 | 0.37167 | 0.38667 | 0.37954 | 0.40667 | 0.35 | 0.37 | 0.38667 | 0.33667
| Median |  | 0.38245 | 0.37784 | 0.455 | 0.365 | 0.36 | 0.36 | 0.37 | 0.385 | 0.3867 | 0.41 | 0.35 | 0.37 | 0.39 | 0.34
| Std Dev. |  | 0.0034593 | 0.003107 | 0.064936 | 0.008165 | 6.0809e-17 | 0.005164 | 0.0098319 | 0.012111 | 0.042739 | 0.005164 | 0.0063246 | 0.010954 | 0.005164 | 0.005164
| Rel.Std.Dev. |  | 0.9059% | 0.8221% | 14.38% | 2.247% | 1.689e-14% | 1.448% | 2.645% | 3.132% | 11.26% | 1.27% | 1.807% | 2.961% | 1.336% | 1.534%
| PDM3 |  | 2.827% | 1.765% | 21.63% | -2.161% | -3.058% | -3.956% | 0.08342% | 4.123% | 2.203% | 9.508% | -5.751% | -0.3654% | 4.123% | -9.341%
| Z-Score (Absolute) |  | 0.4 | 0.18 | 4.36 | 0.65 | 0.84 | 1.03 | 0.18 | 0.67 | 0.27 | 1.81 | 1.41 | 0.27 | 0.67 | 2.16
| Analytical results for Mn in OREAS 267 (Certified Value 0.015 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.013166 | 0.014839 | 0.0178 | 0.0149 | 0.0143 | 0.0144 | 0.0157 | 0.0153 | 0.012535 | 0.0156 | 0.0158 | 0.01507 | 0.0136 | 0.0136 |  | 1
| 1 | 2 | 0.013252 | 0.015021 | 0.019 | 0.0149 | 0.0143 | 0.0146 | 0.0155 | 0.0154 | 0.012641 | 0.0156 | 0.0161 | 0.01526 | 0.0138 | 0.0136 |  | 25
| 1 | 3 | 0.013371 | 0.014838 | 0.0195 | 0.0146 | 0.0146 | 0.0146 | 0.0154 | 0.0148 | 0.0098584 | 0.0156 | 0.0157 | 0.01553 | 0.0132 | 0.0134 |  | 16
| 1 | 4 | 0.013395 | 0.014966 | 0.0183 | 0.0148 | 0.0141 | 0.0147 | 0.0155 | 0.0158 | 0.012771 | 0.0155 | 0.016 | 0.01544 | 0.0136 | 0.0131 |  | 0.014759
| 1 | 5 | 0.013563 | 0.015077 | 0.0179 | 0.0147 | 0.0142 | 0.0149 | 0.0163 | 0.0156 | 0.012716 | 0.0156 | 0.0163 | 0.01511 | 0.0134 | 0.0133 |  | 93
| 1 | 6 | 0.013062 | 0.015095 | 0.0184 | 0.015 | 0.0141 | 0.0149 | 0.0158 | 0.0159 | 0.0098125 | 0.0155 | 0.0164 | 0.01556 | 0.0137 | 0.0134
| Mean |  | 0.013302 | 0.014973 | 0.018483 | 0.014817 | 0.014267 | 0.014683 | 0.0157 | 0.015467 | 0.011722 | 0.015567 | 0.01605 | 0.015328 | 0.01355 | 0.0134
| Median |  | 0.013312 | 0.014994 | 0.01835 | 0.01485 | 0.01425 | 0.01465 | 0.0156 | 0.0155 | 0.012588 | 0.0156 | 0.01605 | 0.01535 | 0.0136 | 0.0134
| Std Dev. |  | 0.00017893 | 0.00011338 | 0.00065549 | 0.0001472 | 0.00018619 | 0.00019408 | 0.00032863 | 0.00039833 | 0.0014638 | 5.164e-05 | 0.00027386 | 0.00021255 | 0.00021679 | 0.00018974
| Rel.Std.Dev. |  | 1.345% | 0.7573% | 3.546% | 0.9934% | 1.305% | 1.322% | 2.093% | 2.575% | 12.49% | 0.3317% | 1.706% | 1.387% | 1.6% | 1.416%
| PDM3 |  | -9.872% | 1.452% | 25.24% | 0.3938% | -3.333% | -0.5097% | 6.379% | 4.798% | -20.57% | 5.476% | 8.751% | 3.861% | -8.189% | -9.205%
| Z-Score (Absolute) |  | 1.45 | 0.07 | 3.28 | 0.07 | 0.57 | 0.19 | 0.74 | 0.52 | 2.9 | 0.61 | 1.05 | 0.4 | 1.23 | 1.36
| Analytical results for Mo in OREAS 267 (Certified Value 5.89 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 6.175 | 5.9946 | 5.55 | 5.95 | 5.76 | 6.27 | 5.82 | 5.88 | 18.5006 | 5.8 | 5.8 | 6.1 | 4.2 | 5.4 | 5.79 | 1
| 1 | 2 | 6.293 | 5.922 | 5.63 | 5.99 | 5.77 | 6.26 | 6.13 | 5.86 | 15.6931 | 5.8 | 5.89 | 6.1 | 5.3 | 5.5 | 5.86 | 26
| 1 | 3 | 6.122 | 5.8372 | 5.82 | 5.84 | 5.74 | 7.02 | 5.79 | 5.86 | 14.4661 | 5.8 | 5.78 | 5.9 | 4.7 | 5.4 | 5.83 | 16
| 1 | 4 | 6.227 | 5.9583 | 5.33 | 5.79 | 5.86 | 6.41 | 5.9 | 5.95 | 16.1487 | 5.8 | 5.87 | 6.1 | 4.9 | 5.4 | 5.81 | 5.8875
| 1 | 5 | 6.144 | 5.896 | 5.46 | 6.04 | 5.86 | 6.35 | 6.06 | 5.84 | 18.0239 | 5.7 | 5.85 | 6.1 | 4.3 | 5.5 | 5.86 | 94
| 1 | 6 | 6.247 | 6.0975 | 5.57 | 5.91 | 5.81 | 6.26 | 5.89 | 6.1 | 14.6257 | 5.7 | 5.71 | 6.1 | 4.3 | 5.5 | 5.94
| Mean |  | 6.2013 | 5.9509 | 5.56 | 5.92 | 5.8 | 6.4283 | 5.9317 | 5.915 | 16.243 | 5.7667 | 5.8167 | 6.0667 | 4.6167 | 5.45 | 5.8483
| Median |  | 6.201 | 5.9402 | 5.56 | 5.93 | 5.79 | 6.31 | 5.895 | 5.87 | 15.9209 | 5.8 | 5.825 | 6.1 | 4.5 | 5.45 | 5.845
| Std Dev. |  | 0.065454 | 0.089701 | 0.16468 | 0.093381 | 0.051769 | 0.29607 | 0.13497 | 0.098336 | 1.6943 | 0.05164 | 0.066833 | 0.08165 | 0.43089 | 0.054772 | 0.052694
| Rel.Std.Dev. |  | 1.055% | 1.507% | 2.962% | 1.577% | 0.8926% | 4.606% | 2.275% | 1.662% | 10.43% | 0.8955% | 1.149% | 1.346% | 9.333% | 1.005% | 0.901%
| PDM3 |  | 5.331% | 1.078% | -5.562% | 0.5523% | -1.486% | 9.186% | 0.7505% | 0.4674% | 175.9% | -2.052% | -1.203% | 3.043% | -21.59% | -7.431% | -0.665%
| Z-Score (Absolute) |  | 1.3 | 0.16 | 1.61 | 0.02 | 0.52 | 2.33 | 0.08 | 0 | 46.95 | 0.67 | 0.45 | 0.69 | 5.9 | 2.11 | 0.3
| Analytical results for Na in OREAS 267 (Certified Value 0.083 wt.%) | [16 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.084854 | 0.083945 | 0.107 | 0.09 | 0.08 | 0.07 | 0.09 | 0.09 | 0.027818 | 0.09 | 0.07 | 0.07 | 0.09 | 0.06 |  | 1
| 1 | 2 | 0.090043 | 0.08345 | 0.115 | 0.09 | 0.08 | 0.07 | 0.09 | 0.09 | 0.025704 | 0.08 | 0.07 | 0.07 | 0.1 | 0.06 |  | #N/A
| 1 | 3 | 0.082615 | 0.082821 | 0.111 | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | 0.021726 | 0.08 | 0.07 | 0.07 | 0.09 | 0.06 |  | 16
| 1 | 4 | 0.086041 | 0.083124 | 0.113 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.026651 | 0.08 | 0.07 | 0.07 | 0.09 | 0.06 |  | 0.082544
| 1 | 5 | 0.082769 | 0.084175 | 0.103 | 0.09 | 0.08 | 0.07 | 0.08 | 0.09 | 0.028807 | 0.09 | 0.07 | 0.07 | 0.09 | 0.06 |  | 95
| 1 | 6 | 0.095264 | 0.08437 | 0.106 | 0.09 | 0.08 | 0.08 | 0.09 | 0.09 | 0.022588 | 0.08 | 0.07 | 0.07 | 0.09 | 0.06
| Mean |  | 0.086931 | 0.083648 | 0.10917 | 0.09 | 0.08 | 0.075 | 0.085 | 0.088333 | 0.025549 | 0.083333 | 0.07 | 0.07 | 0.091667 | 0.06
| Median |  | 0.085447 | 0.083698 | 0.109 | 0.09 | 0.08 | 0.075 | 0.085 | 0.09 | 0.026177 | 0.08 | 0.07 | 0.07 | 0.09 | 0.06
| Std Dev. |  | 0.0049026 | 0.000614 | 0.0045789 | 1.5202e-17 | 0 | 0.0054772 | 0.0054772 | 0.0040825 | 0.002842 | 0.005164 | 0 | 0 | 0.0040825 | 0
| Rel.Std.Dev. |  | 5.64% | 0.734% | 4.194% | 1.689e-14% | 0% | 7.303% | 6.444% | 4.622% | 11.12% | 6.197% | 0% | 0% | 4.454% | 0%
| PDM3 |  | 5.314% | 1.336% | 32.25% | 9.032% | -3.083% | -9.14% | 2.975% | 7.013% | -69.05% | 0.9556% | -15.2% | -15.2% | 11.05% | -27.31%
| Z-Score (Absolute) |  | 0.32 | 0.01 | 2.36 | 0.6 | 0.32 | 0.78 | 0.14 | 0.44 | 5.32 | 0.01 | 1.24 | 1.24 | 0.75 | 2.16
| Analytical results for Nb in OREAS 267 (Certified Value 0.35 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 0.42191 | 0.1 | 0.31 | 0.25 | 0.3 | 0.39 | 0.31 | 0.43124 | 0.4 | 0.45 | 15.5 | 0.3 | 0.62 |  |  | 1
| 1 | 2 | 0.40259 | 0.2 | 0.33 | 0.28 | 0.31 | 0.4 | 0.28 | 0.27793 | 0.4 | 0.46 | 14.9 | 0.3 | 0.63 |  |  | 27
| 1 | 3 | 0.39007 | 0.1 | 0.35 | 0.26 | 0.41 | 0.36 | 0.29 | 0.49633 | 0.4 | 0.45 | 15.1 | 0.3 | 0.63 |  |  | 16
| 1 | 4 | 0.42698 | 0.2 | 0.32 | 0.26 | 0.38 | 0.37 | 0.28 | 0.29058 | 0.4 | 0.43 | 15.3 | 0.3 | 0.61 |  |  | 0.34868
| 1 | 5 | 0.39711 | 0.2 | 0.34 | 0.26 | 0.33 | 0.33 | 0.29 | 0.28179 | 0.4 | 0.45 | 14.5 | 0.3 | 0.62 |  |  | 96
| 1 | 6 | 0.4111 | 0.2 | 0.3 | 0.26 | 0.3 | 0.37 | 0.29 | 0.33916 | 0.4 | 0.42 | 14.3 | 0.3 | 0.65
| Mean |  | 0.40829 | 0.16667 | 0.325 | 0.26167 | 0.33833 | 0.37 | 0.29 | 0.35284 | 0.4 | 0.44333 | 14.9333 | 0.3 | 0.62667
| Median |  | 0.40685 | 0.2 | 0.325 | 0.26 | 0.32 | 0.37 | 0.29 | 0.31487 | 0.4 | 0.45 | 15 | 0.3 | 0.625
| Std Dev. |  | 0.014367 | 0.05164 | 0.018708 | 0.0098319 | 0.046224 | 0.024495 | 0.010954 | 0.091059 | 6.0809e-17 | 0.015055 | 0.46332 | 0 | 0.013663
| Rel.Std.Dev. |  | 3.519% | 30.98% | 5.756% | 3.757% | 13.66% | 6.62% | 3.777% | 25.81% | 1.52e-14% | 3.396% | 3.103% | 0% | 2.18%
| PDM3 |  | 17.1% | -52.2% | -6.792% | -24.96% | -2.968% | 6.114% | -16.83% | 1.191% | 14.72% | 27.15% | 4183% | -13.96% | 79.72%
| Z-Score (Absolute) |  | 0.67 | NA | 0.34 | 1.11 | 0.18 | 0.21 | 0.76 | 0 | NA | 1.1 | NA | NA | 3.33
Results from laboratories 7, 18 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 267 (Indicative Value 15.5 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 1
|  |  | 0.2g | 0.5g | [13 empty] | 1
| 1 | 1 | 15.0547 | 14.9 | [13 empty] | 1
| 1 | 2 | 15.1087 | 17.4 | [13 empty] | 7
| 1 | 3 | 14.8353 | 16.1 | [13 empty] | 16
| 1 | 4 | 14.7924 | 16.5 | [13 empty] | 15.5104
| 1 | 5 | 14.7379 | 16.3 | [13 empty] | 13
| 1 | 6 | 14.9963 | 15.4
| Mean |  | 14.9209 | 16.1
| Median |  | 14.9158 | 16.2
| Std Dev. |  | 0.15244 | 0.87407
| Rel.Std.Dev. |  | 1.022% | 5.429%
| PDM3 |  | -3.801% | 3.801%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 267 (Certified Value 39.6 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 36.089 | 37.5626 | 42 | 39.4 | 40.1 | 39.2 | 37.3 | 42.2 | 42.1666 | 40 | 41.5 | 21.1 | 31.6 | 36.7 |  | 1
| 1 | 2 | 36.646 | 37.2254 | 45.9 | 40.3 | 41.2 | 39.4 | 40.2 | 41.1 | 38.689 | 40 | 39.5 | 22.9 | 31.7 | 37.2 |  | 29
| 1 | 3 | 37.673 | 37.3449 | 45.5 | 38.1 | 39.6 | 42.3 | 37.4 | 41.1 | 33.0634 | 41 | 39.4 | 21.7 | 31.4 | 36.3 |  | 16
| 1 | 4 | 37.174 | 37.9981 | 45 | 39.4 | 39.9 | 38.8 | 39 | 42.1 | 39.7184 | 40 | 39 | 21.1 | 31.7 | 36.6 |  | 39.6436
| 1 | 5 | 37.951 | 37.7791 | 43.7 | 39.5 | 41.8 | 39.9 | 43.7 | 42.8 | 43.2893 | 41 | 41.2 | 21.2 | 31.7 | 36.1 |  | 97
| 1 | 6 | 36.632 | 37.6018 | 43.5 | 39.6 | 40.9 | 39.2 | 37.8 | 43.8 | 33.9363 | 41 | 40.8 | 21.6 | 31.7 | 35.8
| Mean |  | 37.0275 | 37.5853 | 44.2667 | 39.3833 | 40.5833 | 39.8 | 39.2333 | 42.1833 | 38.4772 | 40.5 | 40.2333 | 21.6 | 31.6333 | 36.45
| Median |  | 36.91 | 37.5822 | 44.35 | 39.45 | 40.5 | 39.3 | 38.4 | 42.15 | 39.2037 | 40.5 | 40.15 | 21.4 | 31.7 | 36.45
| Std Dev. |  | 0.70338 | 0.28145 | 1.4679 | 0.71391 | 0.85186 | 1.2759 | 2.4533 | 1.0342 | 4.202 | 0.54772 | 1.0596 | 0.68702 | 0.12111 | 0.49295
| Rel.Std.Dev. |  | 1.9% | 0.7488% | 3.316% | 1.813% | 2.099% | 3.206% | 6.253% | 2.452% | 10.92% | 1.352% | 2.634% | 3.181% | 0.3828% | 1.352%
| PDM3 |  | -6.599% | -5.192% | 11.66% | -0.6565% | 2.37% | 0.3945% | -1.035% | 6.406% | -2.942% | 2.16% | 1.488% | -45.51% | -20.21% | -8.056%
| Z-Score (Absolute) |  | 1.03 | 0.78 | 2.23 | 0.03 | 0.57 | 0.22 | 0.03 | 1.29 | 0.37 | 0.54 | 0.42 | 7.97 | 3.45 | 1.29
| Analytical results for P in OREAS 267 (Certified Value 0.032 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.031778 | 0.031558 | 0.034 | 0.032 | 0.03 | 0.032 | 0.033 | 0.033 | 0.0303 | 0.0321 | 0.032 | 0.028 | 0.0298 |  |  | 1
| 1 | 2 | 0.030853 | 0.031196 | 0.035 | 0.032 | 0.031 | 0.032 | 0.033 | 0.033 | 0.0308 | 0.0314 | 0.031 | 0.027 | 0.0306 |  |  | #N/A
| 1 | 3 | 0.031622 | 0.031568 | 0.035 | 0.03 | 0.031 | 0.032 | 0.032 | 0.032 | 0.0308 | 0.0319 | 0.031 | 0.027 | 0.0309 |  |  | 16
| 1 | 4 | 0.0329 | 0.031676 | 0.034 | 0.031 | 0.031 | 0.032 | 0.033 | 0.034 | 0.0302 | 0.0317 | 0.033 | 0.027 | 0.0299 |  |  | 0.031842
| 1 | 5 | 0.031433 | 0.032015 | 0.034 | 0.031 | 0.031 | 0.032 | 0.033 | 0.033 | 0.0311 | 0.0318 | 0.032 | 0.027 | 0.0299 |  |  | 98
| 1 | 6 | 0.030802 | 0.031106 | 0.034 | 0.032 | 0.03 | 0.032 | 0.033 | 0.034 | 0.0306 | 0.0323 | 0.033 | 0.027 | 0.03
| Mean |  | 0.031565 | 0.03152 | 0.034333 | 0.031333 | 0.030667 | 0.032 | 0.032833 | 0.033167 | 0.030633 | 0.031867 | 0.032 | 0.027167 | 0.030183
| Median |  | 0.031528 | 0.031563 | 0.034 | 0.0315 | 0.031 | 0.032 | 0.033 | 0.033 | 0.0307 | 0.03185 | 0.032 | 0.027 | 0.02995
| Std Dev. |  | 0.00076644 | 0.00033169 | 0.0005164 | 0.0008165 | 0.0005164 | 0 | 0.00040825 | 0.00075277 | 0.00033862 | 0.00031411 | 0.00089443 | 0.00040825 | 0.0004535
| Rel.Std.Dev. |  | 2.428% | 1.052% | 1.504% | 2.606% | 1.684% | 0% | 1.243% | 2.27% | 1.105% | 0.9857% | 2.795% | 1.503% | 1.503%
| PDM3 |  | -0.8702% | -1.011% | 7.825% | -1.597% | -3.69% | 0.4969% | 3.114% | 4.161% | -3.795% | 0.07818% | 0.4969% | -14.68% | -5.208%
| Z-Score (Absolute) |  | 0 | 0.03 | 2.08 | 0.17 | 0.67 | 0.33 | 0.95 | 1.2 | 0.7 | 0.23 | 0.33 | 3.3 | 1.04
| Analytical results for Pb in OREAS 267 (Certified Value 8.6 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 8.062 | 9.1491 | 10.2 | 8.2 | 7.7 | 8.5 | 9.2 | 9 | 5.6353 | 9.3 | 7.3 | 9 | 16.8 | 8.1 | 7.5 | 1
| 1 | 2 | 7.616 | 8.6036 | 11.1 | 8.2 | 7.8 | 8.6 | 9.3 | 9 | 5.0306 | 9.5 | 7.5 | 9 | 13.8 | 8.1 | 7.8 | 30
| 1 | 3 | 8.777 | 8.8161 | 10.8 | 8 | 7.7 | 9.2 | 9.1 | 8.8 | 3.7524 | 9.5 | 7.2 | 8 | 13.1 | 8.4 | 7.6 | 16
| 1 | 4 | 8.431 | 9.381 | 10.8 | 8.3 | 7.6 | 9.6 | 9.5 | 9 | 5.6613 | 9.4 | 7.4 | 9 | 14.2 | 8.1 | 7.7 | 8.6041
| 1 | 5 | 7.593 | 9.1345 | 10.3 | 8.2 | 7.8 | 8.8 | 9 | 9.2 | 5.102 | 9.4 | 7.7 | 9 | 14.5 | 8.1 | 7.6 | 99
| 1 | 6 | 7.707 | 9.024 | 10.4 | 8.1 | 7.6 | 8.6 | 9.3 | 9.2 | 3.8107 | 9.2 | 7.2 | 9 | 13.3 | 8.2 | 7.7
| Mean |  | 8.031 | 9.018 | 10.6 | 8.1667 | 7.7 | 8.8833 | 9.2333 | 9.0333 | 4.832 | 9.3833 | 7.3833 | 8.8333 | 14.2833 | 8.1667 | 7.65
| Median |  | 7.8845 | 9.0792 | 10.6 | 8.2 | 7.7 | 8.7 | 9.25 | 9 | 5.0663 | 9.4 | 7.35 | 9 | 14 | 8.1 | 7.65
| Std Dev. |  | 0.48714 | 0.27387 | 0.35214 | 0.10328 | 0.089443 | 0.43089 | 0.17512 | 0.15055 | 0.85486 | 0.1169 | 0.19408 | 0.40825 | 1.3408 | 0.12111 | 0.10488
| Rel.Std.Dev. |  | 6.066% | 3.037% | 3.322% | 1.265% | 1.162% | 4.851% | 1.897% | 1.667% | 17.69% | 1.246% | 2.629% | 4.622% | 9.387% | 1.483% | 1.371%
| PDM3 |  | -6.661% | 4.811% | 23.2% | -5.084% | -10.51% | 3.246% | 7.313% | 4.989% | -43.84% | 9.057% | -14.19% | 2.664% | 66.01% | -5.084% | -11.09%
| Z-Score (Absolute) |  | 0.43 | 0.43 | 1.83 | 0.32 | 0.73 | 0.32 | 0.62 | 0.45 | 3.25 | 0.75 | 1 | NA | 5.06 | 0.32 | 0.77
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 267 (Indicative Value 4.2 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 4.0074 | 4.3 | [13 empty] | 1
| 1 | 2 | 3.9298 | 4.6 | [13 empty] | 8
| 1 | 3 | 3.9019 | 4.5 | [13 empty] | 16
| 1 | 4 | 3.9998 | 4.5 | [13 empty] | 4.2042
| 1 | 5 | 3.9705 | 4.5 | [13 empty] | 14
| 1 | 6 | 4.0413 | 4.2
| Mean |  | 3.9751 | 4.4333
| Median |  | 3.9852 | 4.5
| Std Dev. |  | 0.05189 | 0.15055
| Rel.Std.Dev. |  | 1.305% | 3.396%
| PDM3 |  | -5.45% | 5.45%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 267 (Certified Value 11.2 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 9.959 | 11.8387 | 12 | 11.5 | 10.2 | 11 | 12.1 | 11.7 | 14.9 | 10.5 | 68.5 | 11.3 | 9.6 |  |  | 1
| 1 | 2 | 10.751 | 11.747 | 13.1 | 11.8 | 10.9 | 11.2 | 12.6 | 11.6 | 15.2 | 10.4 | 71.4 | 11.6 | 9.6 |  |  | 32
| 1 | 3 | 10.03 | 12.0746 | 12.2 | 11.1 | 10.5 | 13 | 11.8 | 12 | 15.1 | 10.3 | 71 | 11.4 | 9.7 |  |  | 16
| 1 | 4 | 10.465 | 12.2933 | 12.6 | 11.7 | 10.6 | 11.9 | 11.5 | 12 | 15 | 10.3 | 71.5 | 11.4 | 9.4 |  |  | 11.2477
| 1 | 5 | 10.885 | 11.8761 | 12.4 | 11.8 | 10.7 | 11.4 | 12.6 | 12.3 | 15.2 | 10.8 | 69.8 | 11.5 | 9.6 |  |  | 100
| 1 | 6 | 9.959 | 12.2297 | 11.9 | 11.2 | 10.3 | 11.2 | 11.9 | 12.6 | 15.3 | 10.1 | 68.3 | 11.7 | 9.8
| Mean |  | 10.3415 | 12.0099 | 12.3667 | 11.5167 | 10.5333 | 11.6167 | 12.0833 | 12.0333 | 15.1167 | 10.4 | 70.0833 | 11.4833 | 9.6167
| Median |  | 10.2475 | 11.9753 | 12.3 | 11.6 | 10.55 | 11.3 | 12 | 12 | 15.15 | 10.35 | 70.4 | 11.45 | 9.6
| Std Dev. |  | 0.41665 | 0.22322 | 0.44121 | 0.30605 | 0.2582 | 0.74409 | 0.4446 | 0.37238 | 0.1472 | 0.23664 | 1.4386 | 0.1472 | 0.13292
| Rel.Std.Dev. |  | 4.029% | 1.859% | 3.568% | 2.657% | 2.451% | 6.405% | 3.679% | 3.095% | 0.9737% | 2.275% | 2.053% | 1.282% | 1.382%
| PDM3 |  | -8.057% | 6.776% | 9.948% | 2.391% | -6.351% | 3.28% | 7.429% | 6.985% | 34.4% | -7.537% | 523.1% | 2.095% | -14.5%
| Z-Score (Absolute) |  | 1.15 | 0.35 | 0.67 | 0.09 | 0.97 | 0 | 0.42 | 0.37 | 3.15 | 1.09 | 52.57 | 0.12 | 1.8
| Analytical results for Re in OREAS 267 (Certified Value < 0.001 ppm) | [16 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.002 |  |  |  | 1
| 1 | 2 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | < 0.001 |  |  |  | 33
| 1 | 3 | < 0.01 | < 0.001 | < 0.001 | 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.001 |  |  |  | 16
| 1 | 4 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.002 |  |  |  | < 0.001
| 1 | 5 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.002 |  |  |  | 101
| 1 | 6 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.001
| Mean |  |  |  | 0.001 | 0.001 | 0.001 |  | 0.001 |  |  |  |  | 0.0016
| Median |  |  |  | 0.001 | 0.001 | 0.001 |  | 0.001 |  |  |  |  | 0.002
| Std Dev. |  |  |  | 0 |  | 0 |  | 0 |  |  |  |  | 0.00054772
| Rel.Std.Dev. |  |  |  | 0% |  | 0% |  | 0% |  |  |  |  | 34.23%
PDM3
| Z-Score (Absolute) |  | 6.61 | 0.67 | 0.4 | 0.54 | 0.13 | 0.67 | 0.27 | 38.98 | 0.13 | 403.1 | 2.56 | 0.81
| Analytical results for S in OREAS 267 (Certified Value 0.022 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.033686 | 0.021629 | 0.023 | 0.03 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.0239 | 0.01 | < 0.05 | 0.02 |  | 1
| 1 | 2 | 0.03356 | 0.021616 | 0.022 | 0.03 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.0229 | 0.01 | < 0.05 | 0.02 |  | 14
| 1 | 3 | 0.037447 | 0.021559 | 0.022 | 0.03 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.0226 | 0.01 | < 0.05 | 0.02 |  | 16
| 1 | 4 | 0.042723 | 0.021732 | 0.021 | 0.03 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.0246 | 0.01 | < 0.05 | 0.02 |  | 0.022195
| 1 | 5 | 0.036177 | 0.021498 | 0.021 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0233 | 0.01 | < 0.05 | 0.02 |  | 102
| 1 | 6 | 0.037472 | 0.021558 | 0.022 | 0.03 | 0.03 | 0.02 | 0.04 | 0.02 | 0.02 | 0.03 | 0.0238 | 0.01 | < 0.05 | 0.02
| Mean |  | 0.036844 | 0.021599 | 0.021833 | 0.03 | 0.021667 | 0.021667 | 0.038333 | 0.02 | 0.02 | 0.021667 | 0.023517 | 0.01 |  | 0.02
| Median |  | 0.036812 | 0.021587 | 0.022 | 0.03 | 0.02 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02355 | 0.01 |  | 0.02
| Std Dev. |  | 0.0033622 | 8.0492e-05 | 0.00075277 | 0 | 0.0040825 | 0.0040825 | 0.0040825 | 0 | 0 | 0.0040825 | 0.00073052 | 0 |  | 0
| Rel.Std.Dev. |  | 9.125% | 0.3727% | 3.448% | 0% | 18.84% | 18.84% | 10.65% | 0% | 0% | 18.84% | 3.106% | 0% |  | 0%
| PDM3 |  | 66% | -2.687% | -1.629% | 35.17% | -2.38% | -2.38% | 72.71% | -9.889% | -9.889% | -2.38% | 5.956% | -54.94% |  | -9.889%
| Z-Score (Absolute) |  | 6.14 | 0.03 | 0.07 | 3.37 | 0 | 0 | 6.74 | 0.67 | 0.67 | 0 | 0.75 | 4.72 | 1.35 | 0.67
| Analytical results for Sb in OREAS 267 (Indicative Value 1.66 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 1.67 | < 2.5 | 1 | 1.16 | 1.28 | 1.36 | 1.22 | 2.46 | 5.2997 | 3.11 | 1.61 | 1.5 | 5.5 | 4.21 | 13.1 | 1
| 1 | 2 | 1.799 | < 2.5 | 0.78 | 1.15 | 1.38 | 1.37 | 1.3 | 3 | 8.7846 | 2.94 | 1.54 | 1.28 | 6.9 | 2.63 | 10.3 | 9
| 1 | 3 | 1.774 | < 2.5 | 0.92 | 1.13 | 1.36 | 1.5 | 1.23 | 1.71 | 5.0484 | 2.42 | 1.57 | 1.14 | 8.4 | 2.77 | 13.5 | 16
| 1 | 4 | 2.178 | < 2.5 | 0.66 | 1.14 | 1.32 | 1.46 | 1.25 | 1.78 | 8.2053 | 2.84 | 1.58 | 1.58 | 6.1 | 3.05 | 13.1 | 1.658
| 1 | 5 | 1.715 | < 2.5 | 0.66 | 1.15 | 1.62 | 1.4 | 1.27 | 2.13 | 6.7256 | 2.58 | 1.56 | 1.26 | 4.7 | 2.28 | 10.7 | 15
| 1 | 6 | 1.649 | < 2.5 | 1.06 | 1.11 | 2.23 | 1.38 | 1.21 | 1.8 | 6.7431 | 2.84 | 1.6 | 1.53 | 5.7 | 2.2 | 17.1
| Mean |  | 1.7975 |  | 0.84667 | 1.14 | 1.5317 | 1.4117 | 1.2467 | 2.1467 | 6.8011 | 2.7883 | 1.5767 | 1.3817 | 6.2167 | 2.8567 | 12.9667
| Median |  | 1.7445 |  | 0.85 | 1.145 | 1.37 | 1.39 | 1.24 | 1.965 | 6.7344 | 2.84 | 1.575 | 1.39 | 5.9 | 2.7 | 13.1
| Std Dev. |  | 0.19517 |  | 0.17236 | 0.017889 | 0.36224 | 0.056006 | 0.033862 | 0.50437 | 1.4994 | 0.24935 | 0.02582 | 0.17826 | 1.2906 | 0.73369 | 2.4353
| Rel.Std.Dev. |  | 10.86% |  | 20.36% | 1.569% | 23.65% | 3.967% | 2.716% | 23.5% | 22.05% | 8.943% | 1.638% | 12.9% | 20.76% | 25.68% | 18.78%
| PDM3 |  | 8.415% |  | -48.93% | -31.24% | -7.618% | -14.86% | -24.81% | 29.48% | 310.2% | 68.18% | -4.904% | -16.67% | 275% | 72.3% | 682.1%
| Z-Score (Absolute) |  | 0.34 | 0.5 | 1.13 | 0.67 | 0.07 | 0.25 | 0.51 | 0.88 | 8.07 | 1.87 | 0 | 0.3 | 7.17 | 1.98 | 17.59
| Analytical results for Sc in OREAS 267 (Certified Value 2.23 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 1.829 | 2.2318 | 2.5 | 2.1 | 2.1 | 2 | 2.3 | 2.3 | 2 | 2 | 2.2 | 2.3 | 1.9 |  |  | 1
| 1 | 2 | 2.128 | 2.2092 | 2.9 | 2.2 | 2.1 | 2.1 | 2.4 | 2.3 | 2 | 2.3 | 2.2 | 2.3 | 1.8 |  |  | 35
| 1 | 3 | 1.93 | 2.2823 | 3 | 2.1 | 2.1 | 2.6 | 2.3 | 2.3 | 2 | 2.1 | 2.2 | 2.3 | 1.8 |  |  | 16
| 1 | 4 | 1.822 | 2.238 | 2.8 | 2.2 | 2.1 | 2.3 | 2.3 | 2.4 | 2 | 2.2 | 2.8 | 2.3 | 1.7 |  |  | 2.2298
| 1 | 5 | 1.979 | 2.3875 | 2.8 | 2.2 | 2.1 | 2.2 | 2.5 | 2.3 | 2 | 2.2 | 2.8 | 2.3 | 1.8 |  |  | 103
| 1 | 6 | 2.029 | 2.2231 | 2.6 | 2.2 | 2.1 | 2.1 | 2.3 | 2.4 | 2 | 1.9 | 2.8 | 2.3 | 1.7
| Mean |  | 1.9528 | 2.262 | 2.7667 | 2.1667 | 2.1 | 2.2167 | 2.35 | 2.3333 | 2 | 2.1167 | 2.5 | 2.3 | 1.7833
| Median |  | 1.9545 | 2.2349 | 2.8 | 2.2 | 2.1 | 2.15 | 2.3 | 2.3 | 2 | 2.15 | 2.5 | 2.3 | 1.8
| Std Dev. |  | 0.11844 | 0.066251 | 0.18619 | 0.05164 | 0 | 0.2137 | 0.083666 | 0.05164 | 0 | 0.1472 | 0.32863 | 4.8648e-16 | 0.075277
| Rel.Std.Dev. |  | 6.065% | 2.929% | 6.73% | 2.383% | 0% | 9.64% | 3.56% | 2.213% | 0% | 6.954% | 13.15% | 2.115e-14% | 4.221%
| PDM3 |  | -12.42% | 1.443% | 24.08% | -2.832% | -5.822% | -0.5896% | 5.39% | 4.642% | -10.31% | -5.074% | 12.12% | 3.148% | -20.02%
| Z-Score (Absolute) |  | 1.66 | 0.13 | 3.05 | 0.42 | 0.81 | 0.13 | 0.64 | 0.54 | NA | 0.71 | 1.51 | 0.35 | 2.64
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 267 (Indicative Value 0.27 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 0.457 | < 5 | 1 | 0.3 | 0.2 | 0.3 | < 0.2 | 0.4 | < 1 | < 1 | 0.41 | < 0.5 | 1.1 | < 0.2 |  | 1
| 1 | 2 | 0.452 | < 5 | 0.4 | 0.2 | 0.3 | 0.2 | 0.2 | 0.3 | < 1 | < 1 | 0.43 | < 0.5 | 0.9 | < 0.2 |  | 2
| 1 | 3 | 0.512 | < 5 | 1 | 0.3 | 0.3 | 0.3 | 0.2 | 0.2 | < 1 | < 1 | 0.44 | < 0.5 | 1 | 0.2 |  | 16
| 1 | 4 | 0.445 | < 5 | 0.3 | 0.2 | 0.3 | 0.2 | < 0.2 | 0.2 | < 1 | < 1 | 0.42 | < 0.5 | 1.1 | < 0.2 |  | 0.26976
| 1 | 5 | 0.49 | < 5 | 0.2 | 0.2 | 0.3 | 0.2 | 0.3 | 0.2 | < 1 | < 1 | 0.44 | < 0.5 | 1.2 | 0.2 |  | 10
| 1 | 6 | 0.582 | < 5 | 0.9 | 0.3 | 0.2 | 0.4 | 0.2 | 0.2 | < 1 | < 1 | 0.44 | < 0.5 | 1.1 | 0.2
| Mean |  | 0.48967 |  | 0.63333 | 0.25 | 0.26667 | 0.26667 | 0.225 | 0.25 |  |  | 0.43 |  | 1.0667 | 0.2
| Median |  | 0.4735 |  | 0.65 | 0.25 | 0.3 | 0.25 | 0.2 | 0.2 |  |  | 0.435 |  | 1.1 | 0.2
| Std Dev. |  | 0.051972 |  | 0.37238 | 0.054772 | 0.05164 | 0.08165 | 0.05 | 0.083666 |  |  | 0.012649 |  | 0.10328 | 3.3993e-17
| Rel.Std.Dev. |  | 10.61% |  | 58.8% | 21.91% | 19.36% | 30.62% | 22.22% | 33.47% |  |  | 2.942% |  | 9.682% | 1.7e-14%
| PDM3 |  | 81.52% |  | 134.8% | -7.326% | -1.147% | -1.147% | -16.59% | -7.326% |  |  | 59.4% |  | 295.4% | -25.86%
| Z-Score (Absolute) |  | 0.65 | 9.9 | 1.31 | 0.45 | 0.38 | 0.38 | 0.76 | 0.45 | 0.7 | 0.7 | 0.38 | 0.45 | 3.31 | 0.91
| Analytical results for Sm in OREAS 267 (Indicative Value 3 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | 2.8241 | 3.4 | [13 empty] | 1
| 1 | 2 | 2.7899 | 3.3 | [13 empty] | 36
| 1 | 3 | 2.6422 | 3.3 | [13 empty] | 16
| 1 | 4 | 2.7549 | 3.1 | [13 empty] | 3.0023
| 1 | 5 | 2.7764 | 3.4 | [13 empty] | 11
| 1 | 6 | 2.8395 | 2.9
| Mean |  | 2.7712 | 3.2333
| Median |  | 2.7832 | 3.3
| Std Dev. |  | 0.070366 | 0.19664
| Rel.Std.Dev. |  | 2.539% | 6.082%
| PDM3 |  | -7.697% | 7.697%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 267 (Certified Value 1.18 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 1.173 | < 2.5 | 1.38 | 1 | 1.1 | 1.1 | 1.2 | 1.2 | 0.98929 | 1.4 | 1.1 | 6.6 | 1.2 | 1 |  | 1
| 1 | 2 | 1.129 | < 2.5 | 1.54 | 1.1 | 1.1 | 1.2 | 1.3 | 1.1 | 1.0833 | 1.4 | 1.1 | 7.5 | 1.2 | 1 |  | 20
| 1 | 3 | 1.208 | < 2.5 | 1.42 | 1 | 1.1 | 1.3 | 1.2 | 1.1 | 0.79648 | 1.4 | 1.1 | 8.1 | 1.2 | 1 |  | 16
| 1 | 4 | 1.034 | < 2.5 | 1.46 | 1 | 1.1 | 1.4 | 1.2 | 1.1 | 0.95076 | 1.4 | 1.2 | 7.1 | 1.2 | 1 |  | 1.1762
| 1 | 5 | 0.946 | < 2.5 | 1.43 | 1 | 1.1 | 1.2 | 1.3 | 1.1 | 0.91078 | 1.4 | 1.2 | 8 | 1.2 | 1 |  | 104
| 1 | 6 | 1.09 | < 2.5 | 1.38 | 1.1 | 1.1 | 1.2 | 1.3 | 1.2 | 0.80848 | 1.4 | 1.1 | 8.7 | 1.2 | 1
| Mean |  | 1.0967 |  | 1.435 | 1.0333 | 1.1 | 1.2333 | 1.25 | 1.1333 | 0.92319 | 1.4 | 1.1333 | 7.6667 | 1.2 | 1
| Median |  | 1.1095 |  | 1.425 | 1 | 1.1 | 1.2 | 1.25 | 1.1 | 0.93077 | 1.4 | 1.1 | 7.75 | 1.2 | 1
| Std Dev. |  | 0.095853 |  | 0.059917 | 0.05164 | 2.4324e-16 | 0.10328 | 0.054772 | 0.05164 | 0.10966 | 2.4324e-16 | 0.05164 | 0.75542 | 0 | 0
| Rel.Std.Dev. |  | 8.74% |  | 4.175% | 4.997% | 2.211e-14% | 8.374% | 4.382% | 4.556% | 11.88% | 1.737e-14% | 4.556% | 9.853% | 0% | 0%
| PDM3 |  | -6.762% |  | 22% | -12.15% | -6.478% | 4.858% | 6.275% | -3.644% | -21.51% | 19.03% | -3.644% | 551.8% | 2.024% | -14.98%
| Z-Score (Absolute) |  | 0.7 | 0.34 | 1.58 | 1.12 | 0.67 | 0.22 | 0.34 | 0.45 | 1.87 | 1.35 | 0.45 | 43.61 | 0 | NA
Results from laboratory 1.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sr in OREAS 267 (Certified Value 30.6 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 1
| 1 | 1 | 24.769 | 31.6789 | 46.2 | 31 | 28.9 | 30.8 | 33.6 | 32.9 | 26.0414 | 33 | 23.9 | 32.5 | 30 | 24.7 |  | 1
| 1 | 2 | 25.796 | 32.3013 | 50.8 | 32.5 | 29.4 | 31.5 | 34.6 | 33.1 | 26.4457 | 33.2 | 25 | 33.3 | 30 | 24 |  | 21
| 1 | 3 | 24.691 | 31.5404 | 49.9 | 30.4 | 29.1 | 34.8 | 32.7 | 32.4 | 20.2009 | 32.7 | 24.5 | 33.6 | 30 | 24 |  | 16
| 1 | 4 | 24.883 | 31.5855 | 49.3 | 32.1 | 28.8 | 32.6 | 32.8 | 33.6 | 26.4594 | 32.7 | 24.8 | 33.1 | 30 | 23.6 |  | 30.6498
| 1 | 5 | 25.855 | 31.8564 | 47.6 | 31.6 | 28.8 | 31.8 | 35.3 | 33.6 | 25.9238 | 32.7 | 25.2 | 33.4 | 30 | 24.2 |  | 105
| 1 | 6 | 26.18 | 32.0077 | 47.8 | 31.8 | 29.2 | 31.5 | 33.6 | 34 | 20.4008 | 32.7 | 23.8 | 33 | 31 | 24.8
| Mean |  | 25.3623 | 31.8284 | 48.6 | 31.5667 | 29.0333 | 32.1667 | 33.7667 | 33.2667 | 24.2453 | 32.8333 | 24.5333 | 33.15 | 30.1667 | 24.2167
| Median |  | 25.3395 | 31.7676 | 48.55 | 31.7 | 29 | 31.65 | 33.6 | 33.35 | 25.9826 | 32.7 | 24.65 | 33.2 | 30 | 24.1
| Std Dev. |  | 0.65297 | 0.29014 | 1.6982 | 0.7607 | 0.24221 | 1.4152 | 1.0172 | 0.5785 | 3.0635 | 0.21602 | 0.5785 | 0.38341 | 0.40825 | 0.45789
| Rel.Std.Dev. |  | 2.575% | 0.9116% | 3.494% | 2.41% | 0.8343% | 4.399% | 3.012% | 1.739% | 12.64% | 0.6579% | 2.358% | 1.157% | 1.353% | 1.891%
| PDM3 |  | -17.25% | 3.845% | 58.57% | 2.992% | -5.274% | 4.949% | 10.17% | 8.538% | -20.9% | 7.124% | -19.96% | 8.157% | -1.576% | -20.99%
| Z-Score (Absolute) |  | 2.35 | 0.05 | 6.27 | 0.05 | 0.99 | 0.17 | 0.77 | 0.58 | 2.76 | 0.42 | 2.66 | 0.54 | 0.57 | 2.77
| Analytical results for Ta in OREAS 267 (Certified Value < 0.01 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.012612 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01 |  |  | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.0092436 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01 |  |  | 22
| 1 | 3 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.010279 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01 |  |  | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01356 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01 |  |  | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.016494 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01 |  |  | 106
| 1 | 6 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.0086916 | < 0.05 | < 0.05 | 0.9 | < 0.1 | < 0.01
| Mean | [8 empty] | 0.011813 |  |  | 0.9
| Median | [8 empty] | 0.011445 |  |  | 0.9
| Std Dev. | [8 empty] | 0.0029765 |  |  | 0
| Rel.Std.Dev. | [8 empty] | 25.2% |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 267 (Indicative Value 0.32 ppm) | [16 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.06 | 1.07 | 1.19 | 1.22 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.2g | 0.5g | 0.25g | 01g | [11 empty] | 2
| 1 | 1 | 0.29271 | 0.3 | 0.32 | 0.6 | [11 empty] | 1
| 1 | 2 | 0.28863 | 0.4 | 0.31 | 0.6 | [11 empty] | 6
| 1 | 3 | 0.27635 | 0.3 | 0.31 | 0.6 | [11 empty] | 16
| 1 | 4 | 0.30752 | 0.4 | 0.32 | 0.6 | [11 empty] | 0.32045
| 1 | 5 | 0.29607 | 0.4 | 0.32 | 0.7 | [11 empty] | 12
| 1 | 6 | 0.30679 | 0.3 | 0.32 | 0.7
| Mean |  | 0.29468 | 0.35 | 0.31667 | 0.63333
| Median |  | 0.29439 | 0.35 | 0.32 | 0.6
| Std Dev. |  | 0.011747 | 0.054772 | 0.005164 | 0.05164
| Rel.Std.Dev. |  | 3.986% | 15.65% | 1.631% | 8.154%
| PDM3 |  | -8.042% | 9.222% | -1.18% | 97.64%
| Z-Score (Absolute) |  | 0.94 | 0.41 | 0.41 | 7.31
| Analytical results for Te in OREAS 267 (Certified Value 0.053 ppm) | [16 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 |  | < 0.2 | 0.03 | 0.03 | 0.06 | 0.04 | 0.05 | 0.05 | 0.0667 | 0.06 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.05 | 1
| 1 | 2 | 0.09 | < 0.2 | < 0.02 | 0.04 | 0.05 | 0.07 | 0.04 | 0.05 | 0.063766 | 0.06 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.04 | 24
| 1 | 3 |  | < 0.2 | < 0.02 | 0.05 | 0.05 | 0.07 | 0.05 | 0.06 | 0.060201 | < 0.05 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.07 | 16
| 1 | 4 |  | < 0.2 | < 0.02 | 0.04 | 0.04 | 0.07 | 0.06 | 0.05 | 0.077186 | 0.06 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.07 | 0.053499
| 1 | 5 |  | < 0.2 | < 0.02 | 0.04 | 0.06 | 0.08 | 0.04 | 0.06 | 0.060141 | < 0.05 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.06 | 107
| 1 | 6 |  | < 0.2 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.064983 | < 0.05 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.04
| Mean |  | 0.09 |  | 0.045 | 0.041667 | 0.051667 | 0.063333 | 0.048333 | 0.053333 | 0.065496 | 0.06 |  |  |  |  | 0.055
| Median |  | 0.09 |  | 0.045 | 0.04 | 0.05 | 0.07 | 0.05 | 0.05 | 0.064374 | 0.06 |  |  |  |  | 0.055
| Std Dev. |  |  |  | 0.021213 | 0.0075277 | 0.0075277 | 0.015055 | 0.0075277 | 0.005164 | 0.0062941 | 0 |  |  |  |  | 0.013784
| Rel.Std.Dev. |  |  |  | 47.14% | 18.07% | 14.57% | 23.77% | 15.57% | 9.682% | 9.61% | 0% |  |  |  |  | 25.06%
| PDM3 |  | 68.23% |  | -15.89% | -22.12% | -3.425% | 18.38% | -9.656% | -0.3097% | 22.42% | 12.15% |  |  |  |  | 2.806%
| Z-Score (Absolute) |  | 2.22 | 2.79 | 1.73 | 0.58 | 0 | 0.67 | 0.19 | 0.1 | 0.8 | 0.53 | 1.54 | 2.12 | 11.46 | 0.1 | 0.19
| Analytical results for Th in OREAS 267 (Certified Value 5.96 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 6.417 | 5.8707 | 6.3 | 5.8 | 5.9 | 5.5 | 6.3 | 6 | 6.03 | 5.4 | 8.9 | 5.9 | 5.6 |  |  | 1
| 1 | 2 | 6.404 | 5.9204 | 6.7 | 5.6 | 5.9 | 5.5 | 6 | 6 | 6.1 | 5.5 | 9.2 | 6.1 | 5.7 |  |  | 25
| 1 | 3 | 6.309 | 5.9104 | 6.7 | 5.4 | 5.9 | 6.7 | 6.1 | 6 | 6.15 | 5.3 | 9.2 | 6.1 | 5.7 |  |  | 16
| 1 | 4 | 6.372 | 5.898 | 6.6 | 5.6 | 6.2 | 5.9 | 6.5 | 5.9 | 6.14 | 5.6 | 8.6 | 6 | 5.8 |  |  | 5.9634
| 1 | 5 | 6.316 | 5.8941 | 6.4 | 5.6 | 6 | 5.7 | 6.4 | 6 | 6.1 | 5.7 | 8.6 | 6.1 | 5.7 |  |  | 108
| 1 | 6 | 6.293 | 5.8975 | 6.4 | 5.5 | 6.1 | 5.5 | 6.2 | 6.2 | 6.02 | 5.4 | 8.3 | 6.1 | 5.7
| Mean |  | 6.3518 | 5.8985 | 6.5167 | 5.5833 | 6 | 5.8 | 6.25 | 6.0167 | 6.09 | 5.4833 | 8.8 | 6.05 | 5.7
| Median |  | 6.344 | 5.8978 | 6.5 | 5.6 | 5.95 | 5.6 | 6.25 | 6 | 6.1 | 5.45 | 8.75 | 6.1 | 5.7
| Std Dev. |  | 0.05283 | 0.01683 | 0.17224 | 0.13292 | 0.12649 | 0.46904 | 0.18708 | 0.098319 | 0.054406 | 0.1472 | 0.36332 | 0.083666 | 0.063246
| Rel.Std.Dev. |  | 0.8317% | 0.2853% | 2.643% | 2.381% | 2.108% | 8.087% | 2.993% | 1.634% | 0.8934% | 2.684% | 4.129% | 1.383% | 1.11%
| PDM3 |  | 6.514% | -1.087% | 9.278% | -6.373% | 0.6144% | -2.739% | 4.807% | 0.8939% | 2.124% | -8.05% | 47.57% | 1.453% | -4.416%
| Z-Score (Absolute) |  | 0.97 | 0.34 | 1.44 | 1.25 | 0.05 | 0.63 | 0.67 | 0 | 0.21 | 1.54 | 8.04 | 0.1 | 0.92
| Analytical results for Ti in OREAS 267 (Certified Value 0.06 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.043487 | 0.061391 | 0.08 | 0.061 | 0.057 | 0.058 | 0.066 | 0.064 | 0.060263 | 0.0757 | 0.05 | 0.1 | 0.062 | 0.05 |  | 1
| 1 | 2 | 0.045225 | 0.060926 | 0.09 | 0.061 | 0.056 | 0.059 | 0.064 | 0.063 | 0.054548 | 0.0755 | 0.05 | 0.11 | 0.063 | 0.05 |  | 26
| 1 | 3 | 0.043252 | 0.060111 | 0.08 | 0.06 | 0.058 | 0.06 | 0.063 | 0.061 | 0.043482 | 0.0735 | 0.05 | 0.11 | 0.064 | 0.05 |  | 16
| 1 | 4 | 0.044601 | 0.060388 | 0.08 | 0.061 | 0.056 | 0.061 | 0.063 | 0.065 | 0.057354 | 0.0727 | 0.05 | 0.11 | 0.062 | 0.04 |  | 0.059906
| 1 | 5 | 0.045013 | 0.061418 | 0.08 | 0.061 | 0.056 | 0.061 | 0.067 | 0.065 | 0.058565 | 0.0753 | 0.05 | 0.11 | 0.063 | 0.04 |  | 109
| 1 | 6 | 0.044856 | 0.060996 | 0.08 | 0.06 | 0.056 | 0.06 | 0.067 | 0.065 | 0.048351 | 0.0727 | 0.05 | 0.1 | 0.064 | 0.04
| Mean |  | 0.044406 | 0.060872 | 0.081667 | 0.060667 | 0.0565 | 0.059833 | 0.065 | 0.063833 | 0.053761 | 0.074233 | 0.05 | 0.10667 | 0.063 | 0.045
| Median |  | 0.044729 | 0.060961 | 0.08 | 0.061 | 0.056 | 0.06 | 0.065 | 0.0645 | 0.055951 | 0.0744 | 0.05 | 0.11 | 0.063 | 0.045
| Std Dev. |  | 0.00083142 | 0.00052916 | 0.0040825 | 0.0005164 | 0.00083666 | 0.001169 | 0.0018974 | 0.0016021 | 0.0065392 | 0.0014236 | 7.6012e-18 | 0.005164 | 0.00089443 | 0.0054772
| Rel.Std.Dev. |  | 1.872% | 0.8693% | 4.999% | 0.8512% | 1.481% | 1.954% | 2.919% | 2.51% | 12.16% | 1.918% | 1.52e-14% | 4.841% | 1.42% | 12.17%
| PDM3 |  | -25.87% | 1.613% | 36.33% | 1.271% | -5.685% | -0.1205% | 8.504% | 6.557% | -10.26% | 23.92% | -16.54% | 78.06% | 5.166% | -24.88%
| Z-Score (Absolute) |  | 1.96 | 0.01 | 2.5 | 0.01 | 0.51 | 0.11 | 0.51 | 0.37 | 0.84 | 1.61 | 1.29 | 5.49 | 0.27 | 1.89
| Analytical results for Tl in OREAS 267 (Certified Value 0.09 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.064 | 0.099871 | 0.1 | 0.08 | 0.09 | 0.09 | 0.1 | 0.09 | 0.077047 | 0.1 | 0.08 | 0.091 | < 1 | 0.11 | 0.07 | 1
| 1 | 2 | 0.061 | 0.097567 | 0.11 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.085875 | 0.11 | 0.09 | 0.092 | < 1 | 0.09 | 0.06 | 27
| 1 | 3 | 0.064 | 0.094716 | 0.11 | 0.08 | 0.09 | 0.09 | 0.1 | 0.09 | 0.063681 | 0.11 | 0.08 | 0.086 | < 1 | 0.1 | 0.09 | 16
| 1 | 4 | 0.074 | 0.097988 | 0.11 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.088271 | 0.11 | 0.09 | 0.089 | < 1 | 0.1 | 0.07 | 0.090064
| 1 | 5 | 0.098 | 0.09092 | 0.11 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.089183 | 0.11 | 0.09 | 0.096 | < 1 | 0.09 | 0.07 | 110
| 1 | 6 | 0.071 | 0.095188 | 0.1 | 0.08 | 0.09 | 0.09 | 0.1 | 0.09 | 0.061058 | 0.11 | 0.08 | 0.098 | < 1 | 0.1 | 0.06
| Mean |  | 0.072 | 0.096042 | 0.10667 | 0.085 | 0.09 | 0.091667 | 0.098333 | 0.09 | 0.077519 | 0.10833 | 0.085 | 0.092 |  | 0.098333 | 0.07
| Median |  | 0.0675 | 0.096377 | 0.11 | 0.085 | 0.09 | 0.09 | 0.1 | 0.09 | 0.081461 | 0.11 | 0.085 | 0.0915 |  | 0.1 | 0.07
| Std Dev. |  | 0.013638 | 0.0031462 | 0.005164 | 0.0054772 | 1.5202e-17 | 0.0040825 | 0.0040825 | 1.5202e-17 | 0.012523 | 0.0040825 | 0.0054772 | 0.0044272 |  | 0.0075277 | 0.010954
| Rel.Std.Dev. |  | 18.94% | 3.276% | 4.841% | 6.444% | 1.689e-14% | 4.454% | 4.152% | 1.689e-14% | 16.16% | 3.768% | 6.444% | 4.812% |  | 7.655% | 15.65%
| PDM3 |  | -20.06% | 6.637% | 18.43% | -5.623% | -0.07091% | 1.78% | 9.182% | -0.07091% | -13.93% | 20.29% | -5.623% | 2.15% |  | 9.182% | -22.28%
| Z-Score (Absolute) |  | 1.9 | 0.53 | 1.6 | 0.59 | 0.08 | 0.08 | 0.76 | 0.08 | 1.35 | 1.77 | 0.59 | 0.12 | NA | 0.76 | 2.11
| Analytical results for Tm in OREAS 267 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 1.06 | 1.07 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.2g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.1 | < 0.1 | [13 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | [13 empty] | 1
| 1 | 3 | < 0.1 | < 0.1 | [13 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | [13 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | [13 empty] | 13
| 1 | 6 | < 0.1 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 267 (Certified Value 0.9 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.906 | 0.9223 | 0.9 | 0.84 | 0.86 | 0.89 | 0.95 | 0.93 | 0.9 | 0.86 | 1.13 | 2 | 0.9 | 0.9 |  | 1
| 1 | 2 | 0.917 | 0.87295 | 0.9 | 0.87 | 0.85 | 0.89 | 0.93 | 0.94 | 0.91 | 0.91 | 1.13 | 2.1 | 0.9 | 0.91 |  | 29
| 1 | 3 | 0.872 | 0.88533 | 0.9 | 0.82 | 0.88 | 1.08 | 0.94 | 0.92 | 0.93 | 0.87 | 1.08 | 2.1 | 0.9 | 0.91 |  | 16
| 1 | 4 | 0.919 | 0.95534 | 0.9 | 0.87 | 0.85 | 0.96 | 0.94 | 0.92 | 0.91 | 0.88 | 1.12 | 2.6 | 0.9 | 0.98 |  | 0.89894
| 1 | 5 | 0.905 | 0.84556 | 0.9 | 0.85 | 0.86 | 0.89 | 0.91 | 0.96 | 0.91 | 0.92 | 1.19 | 2.6 | 0.9 | 0.9 |  | 111
| 1 | 6 | 0.912 | 0.89808 | 0.9 | 0.84 | 0.85 | 0.88 | 0.95 | 0.96 | 0.93 | 0.87 | 1.18 | 2.5 | 0.9 | 0.9
| Mean |  | 0.90517 | 0.89659 | 0.9 | 0.84833 | 0.85833 | 0.93167 | 0.93667 | 0.93833 | 0.915 | 0.885 | 1.1383 | 2.3167 | 0.9 | 0.91667
| Median |  | 0.909 | 0.89171 | 0.9 | 0.845 | 0.855 | 0.89 | 0.94 | 0.935 | 0.91 | 0.875 | 1.13 | 2.3 | 0.9 | 0.905
| Std Dev. |  | 0.017198 | 0.038482 | 0 | 0.019408 | 0.01169 | 0.078337 | 0.015055 | 0.018348 | 0.012247 | 0.02429 | 0.040702 | 0.27869 | 0 | 0.031411
| Rel.Std.Dev. |  | 1.9% | 4.292% | 0% | 2.288% | 1.362% | 8.408% | 1.607% | 1.955% | 1.339% | 2.745% | 3.576% | 12.03% | 0% | 3.427%
| PDM3 |  | 0.6924% | -0.2612% | 0.1176% | -5.63% | -4.517% | 3.64% | 4.196% | 4.382% | 1.786% | -1.551% | 26.63% | 157.7% | 0.1176% | 1.972%
| Z-Score (Absolute) |  | 0.31 | 0.57 | NA | 2.07 | 1.76 | 0.52 | 0.67 | 0.73 | 0 | 0.93 | 6.95 | NA | NA | 0.05
Results from laboratories 1.07, 1.22 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 267 (Certified Value 20.8 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 19.276 | 20.4834 | 24 | 21 | 20 | 20 | 22 | 21 | 21 | 21 | 28.3 | 20 | 17.5 |  |  | 1
| 1 | 2 | 20.479 | 20.525 | 24 | 21 | 20 | 21 | 22 | 21 | 22 | 22 | 27.4 | 20 | 17.5 |  |  | 30
| 1 | 3 | 19.616 | 20.3283 | 25 | 20 | 21 | 21 | 21 | 21 | 22 | 21 | 29.2 | 20 | 17.5 |  |  | 16
| 1 | 4 | 20.371 | 20.3707 | 24 | 20 | 20 | 21 | 21 | 22 | 21 | 22 | 29.7 | 20 | 16.9 |  |  | 20.8081
| 1 | 5 | 20.465 | 20.4841 | 23 | 20 | 21 | 21 | 22 | 22 | 22 | 21 | 28.2 | 20 | 17.5 |  |  | 112
| 1 | 6 | 19.534 | 20.5509 | 24 | 21 | 20 | 21 | 22 | 22 | 21 | 21 | 27.5 | 20 | 18
| Mean |  | 19.9568 | 20.4571 | 24 | 20.5 | 20.3333 | 20.8333 | 21.6667 | 21.5 | 21.5 | 21.3333 | 28.3833 | 20 | 17.4833
| Median |  | 19.9935 | 20.4838 | 24 | 20.5 | 20 | 21 | 22 | 21.5 | 21.5 | 21 | 28.25 | 20 | 17.5
| Std Dev. |  | 0.54054 | 0.088167 | 0.63246 | 0.54772 | 0.5164 | 0.40825 | 0.5164 | 0.54772 | 0.54772 | 0.5164 | 0.91524 | 0 | 0.34881
| Rel.Std.Dev. |  | 2.709% | 0.431% | 2.635% | 2.672% | 2.54% | 1.96% | 2.383% | 2.548% | 2.548% | 2.421% | 3.225% | 0% | 1.995%
| PDM3 |  | -4.091% | -1.687% | 15.34% | -1.48% | -2.281% | 0.1215% | 4.126% | 3.325% | 3.325% | 2.524% | 36.41% | -3.883% | -15.98%
| Z-Score (Absolute) |  | 0.89 | 0.38 | 3.2 | 0.34 | 0.51 | 0 | 0.84 | 0.67 | 0.67 | 0.51 | 7.64 | 0.84 | 3.39
| Analytical results for W in OREAS 267 (Certified Value 1.81 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 1.456 | 1.9772 | 1 | 1.73 | 1.58 | 2.08 | 1.9 | 1.7 | 3.0953 | 1.5 | 1.8 | 7.5 | 1.8 | 2.38 |  | 1
| 1 | 2 | 1.517 | 1.9006 | 0.7 | 1.78 | 1.67 | 2.02 | 1.87 | 1.67 | 2.3104 | 1.4 | 2 | 7.4 | 1.6 | 2.54 |  | 31
| 1 | 3 | 1.525 | 1.9339 | 0.8 | 1.82 | 1.6 | 2.2 | 1.97 | 1.69 | 2.5934 | 1.4 | 2 | 7.6 | 1.7 | 2.52 |  | 16
| 1 | 4 | 1.846 | 1.8559 | 0.6 | 1.68 | 1.68 | 2.11 | 1.95 | 1.56 | 2.139 | 1.5 | 2 | 8 | 1.8 | 2.5 |  | 1.8138
| 1 | 5 | 1.528 | 1.8745 | 0.6 | 1.84 | 1.64 | 1.93 | 1.79 | 1.7 | 2.1405 | 1.4 | 2.1 | 7.4 | 1.7 | 2.51 |  | 113
| 1 | 6 | 1.521 | 1.9252 | 0.9 | 1.74 | 1.58 | 1.92 | 1.97 | 1.7 | 2.4692 | 1.4 | 2.1 | 7.8 | 1.6 | 2.48
| Mean |  | 1.5655 | 1.9112 | 0.76667 | 1.765 | 1.625 | 2.0433 | 1.9083 | 1.67 | 2.458 | 1.4333 | 2 | 7.6167 | 1.7 | 2.4883
| Median |  | 1.523 | 1.9129 | 0.75 | 1.76 | 1.62 | 2.05 | 1.925 | 1.695 | 2.3898 | 1.4 | 2 | 7.55 | 1.7 | 2.505
| Std Dev. |  | 0.14004 | 0.043799 | 0.1633 | 0.059917 | 0.044609 | 0.10857 | 0.070545 | 0.055136 | 0.36026 | 0.05164 | 0.10954 | 0.24014 | 0.089443 | 0.056716
| Rel.Std.Dev. |  | 8.945% | 2.292% | 21.3% | 3.395% | 2.745% | 5.313% | 3.697% | 3.302% | 14.66% | 3.603% | 5.477% | 3.153% | 5.261% | 2.279%
| PDM3 |  | -13.69% | 5.369% | -57.73% | -2.692% | -10.41% | 12.65% | 5.21% | -7.93% | 35.51% | -20.98% | 10.26% | 319.9% | -6.276% | 37.19%
| Z-Score (Absolute) |  | 0.67 | 0.49 | 3.37 | 0 | 0.47 | 0.94 | 0.48 | 0.32 | 2.34 | 1.12 | 0.79 | NA | 0.22 | 2.44
| Analytical results for Y in OREAS 267 (Certified Value 6.29 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 5.744 | 6.2911 | 6.34 | 6.61 | 6.42 | 6.19 | 6.36 | 6.66 | 6.66 | 6.01 | 17.5 | 5.5 | 5.55 |  |  | 1
| 1 | 2 | 5.897 | 6.3052 | 7.37 | 6.79 | 6.53 | 6.24 | 6.89 | 6.61 | 6.78 | 6 | 17.4 | 5.5 | 5.55 |  |  | 32
| 1 | 3 | 6.228 | 6.3106 | 6.85 | 6.44 | 6.35 | 7.39 | 6.26 | 6.55 | 6.79 | 5.88 | 17.3 | 5.6 | 5.48 |  |  | 16
| 1 | 4 | 5.852 | 6.2802 | 6.89 | 6.6 | 6.53 | 6.72 | 6.53 | 6.58 | 6.75 | 5.91 | 17.9 | 5.5 | 5.39 |  |  | 6.2889
| 1 | 5 | 6.247 | 6.2717 | 6.93 | 6.68 | 6.53 | 6.39 | 6.93 | 6.7 | 6.86 | 5.99 | 17.6 | 5.6 | 5.49 |  |  | 114
| 1 | 6 | 5.833 | 6.3604 | 6.47 | 6.48 | 6.54 | 6.35 | 6.48 | 7.02 | 6.63 | 5.76 | 17.6 | 5.6 | 5.61
| Mean |  | 5.9668 | 6.3032 | 6.8083 | 6.6 | 6.4833 | 6.5467 | 6.575 | 6.6867 | 6.745 | 5.925 | 17.55 | 5.55 | 5.5117
| Median |  | 5.8745 | 6.2982 | 6.87 | 6.605 | 6.53 | 6.37 | 6.505 | 6.635 | 6.765 | 5.95 | 17.55 | 5.55 | 5.52
| Std Dev. |  | 0.21556 | 0.031622 | 0.36674 | 0.12853 | 0.079415 | 0.45293 | 0.27632 | 0.17201 | 0.085965 | 0.096488 | 0.20736 | 0.054772 | 0.076004
| Rel.Std.Dev. |  | 3.613% | 0.5017% | 5.387% | 1.947% | 1.225% | 6.919% | 4.203% | 2.572% | 1.275% | 1.628% | 1.182% | 0.9869% | 1.379%
| PDM3 |  | -5.121% | 0.2284% | 8.26% | 4.947% | 3.092% | 4.099% | 4.55% | 6.325% | 7.253% | -5.786% | 179.1% | -11.75% | -12.36%
| Z-Score (Absolute) |  | 1.67 | 0.65 | 0.9 | 0.26 | 0.1 | 0.1 | 0.18 | 0.52 | 0.7 | 1.8 | NA | 2.95 | 3.06
| Analytical results for Yb in OREAS 267 (Indicative Value 0.48 ppm) | [16 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.19 | 1.22 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.25g | 01g | [10 empty] | 2
| 1 | 1 | 0.438 | 0.47123 | 0.5 | 0.5 | < 0.5 | [10 empty] | 1
| 1 | 2 | 0.477 | 0.48862 | 0.5 | 0.5 | < 0.5 | [10 empty] | 8
| 1 | 3 | 0.465 | 0.48954 | 0.5 | 0.5 | < 0.5 | [10 empty] | 16
| 1 | 4 | 0.453 | 0.46665 | 0.5 | 0.5 | < 0.5 | [10 empty] | 0.48429
| 1 | 5 | 0.458 | 0.47985 | 0.5 | 0.5 | < 0.5 | [10 empty] | 14
| 1 | 6 | 0.46 | 0.47603 | 0.5 | 0.5 | < 0.5
| Mean |  | 0.4585 | 0.47865 | 0.5 | 0.5
| Median |  | 0.459 | 0.47794 | 0.5 | 0.5
| Std Dev. |  | 0.012942 | 0.0092253 | 0 | 0
| Rel.Std.Dev. |  | 2.823% | 1.927% | 0% | 0%
| PDM3 |  | -5.325% | -1.163% | 3.244% | 3.244%
| Z-Score (Absolute) |  | 0.64 | 0 | 0.67 | 0.67 | 7.22
| Analytical results for Zn in OREAS 267 (Certified Value 23.5 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 21.582 | 24.2047 | 30.7 | 25 | 23 | 24 | 21 | 25 | 22.7891 | 29 | 23 | 22 | 25.9 | 23 | 22.3 | 1
| 1 | 2 | 21.688 | 24.0401 | 32.3 | 25 | 24 | 24 | 21 | 25 | 23.2516 | 29 | 23 | 21 | 25.3 | 21 | 23.2 | 34
| 1 | 3 | 21.209 | 23.9678 | 31.9 | 24 | 24 | 22 | 21 | 24 | 18.0469 | 29 | 23 | 21 | 27.8 | 30 | 22.7 | 16
| 1 | 4 | 20.911 | 24.2419 | 31.8 | 25 | 23 | 31 | 22 | 26 | 23.3688 | 29 | 23 | 20 | 24.5 | 22 | 21.7 | 23.5071
| 1 | 5 | 21.596 | 24.1439 | 31.1 | 25 | 24 | 24 | 29 | 25 | 22.9353 | 30 | 24 | 21 | 24.8 | 22 | 22.6 | 115
| 1 | 6 | 21.94 | 25.1106 | 29.7 | 25 | 24 | 25 | 22 | 26 | 18.0656 | 29 | 23 | 21 | 24 | 22 | 21.5
| Mean |  | 21.4877 | 24.2848 | 31.25 | 24.8333 | 23.6667 | 25 | 22.6667 | 25.1667 | 21.4096 | 29.1667 | 23.1667 | 21 | 25.3833 | 23.3333 | 22.3333
| Median |  | 21.589 | 24.1743 | 31.45 | 25 | 24 | 24 | 21.5 | 25 | 22.8622 | 29 | 23 | 21 | 25.05 | 22 | 22.45
| Std Dev. |  | 0.36758 | 0.41722 | 0.95446 | 0.40825 | 0.5164 | 3.0984 | 3.1411 | 0.75277 | 2.6058 | 0.40825 | 0.40825 | 0.63246 | 1.3527 | 3.3267 | 0.64083
| Rel.Std.Dev. |  | 1.711% | 1.718% | 3.054% | 1.644% | 2.182% | 12.39% | 13.86% | 2.991% | 12.17% | 1.4% | 1.762% | 3.012% | 5.329% | 14.26% | 2.869%
| PDM3 |  | -8.591% | 3.309% | 32.94% | 5.642% | 0.679% | 6.351% | -3.575% | 7.06% | -8.923% | 24.08% | -1.448% | -10.67% | 7.982% | -0.739% | -4.993%
| Z-Score (Absolute) |  | 1.1 | 0.31 | 3.84 | 0.59 | 0 | 0.67 | 0.51 | 0.76 | 1.14 | 2.78 | 0.25 | 1.35 | 0.87 | 0.17 | 0.67
| Analytical results for Zr in OREAS 267 (Certified Value 21.3 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 21.1101 | 2.4 | 21.9 | 20.8 | 19.4 | 20.9 | 21.8 | 26.1837 | 21.7 | 21.1 | 139.5 | 20 | 20.8 |  |  | 1
| 1 | 2 | 21.4104 | 1.2 | 23.3 | 21.7 | 19.4 | 21.5 | 21.3 | 21.4995 | 22.6 | 22.7 | 124.8 | 20 | 20.7 |  |  | 35
| 1 | 3 | 20.985 | 2 | 22.9 | 21.4 | 24 | 20.5 | 21.1 | 23.3948 | 22.1 | 22.4 | 124.2 | 20.2 | 20.5 |  |  | 16
| 1 | 4 | 21.3652 | 0.9 | 22.8 | 22.2 | 21.7 | 20.8 | 22 | 19.8478 | 22.1 | 22.1 | 136.7 | 20 | 20.3 |  |  | 21.2834
| 1 | 5 | 21.2011 | 1.3 | 23.5 | 21.3 | 19.6 | 22.1 | 21.7 | 19.6894 | 22.3 | 22.5 | 127.1 | 19.6 | 20.6 |  |  | 116
| 1 | 6 | 21.483 | 2 | 22.5 | 21.2 | 19.4 | 21.3 | 21.8 | 16.9459 | 21.7 | 21.8 | 132.3 | 20 | 21
| Mean |  | 21.2592 | 1.6333 | 22.8167 | 21.4333 | 20.5833 | 21.1833 | 21.6167 | 21.2602 | 22.0833 | 22.1 | 130.7667 | 19.9667 | 20.65
| Median |  | 21.2832 | 1.65 | 22.85 | 21.35 | 19.5 | 21.1 | 21.75 | 20.6736 | 22.1 | 22.25 | 129.7 | 20 | 20.65
| Std Dev. |  | 0.19235 | 0.58195 | 0.57417 | 0.4761 | 1.902 | 0.57417 | 0.34303 | 3.2227 | 0.34881 | 0.5831 | 6.4186 | 0.19664 | 0.2429
| Rel.Std.Dev. |  | 0.9048% | 35.63% | 2.516% | 2.221% | 9.241% | 2.71% | 1.587% | 15.16% | 1.58% | 2.638% | 4.908% | 0.9848% | 1.176%
| PDM3 |  | -0.1138% | -92.33% | 7.204% | 0.7046% | -3.289% | -0.47% | 1.566% | -0.109% | 3.759% | 3.837% | 514.4% | -6.187% | -2.976%
| Z-Score (Absolute) |  | 0 | 20.58 | 1.63 | 0.18 | 0.71 | 0.08 | 0.37 | 0 | 0.86 | 0.88 | NA | 1.36 | 0.64
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 267 (Indicative Value 10.1 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.09 | 11
| Mean |  | 10.095 | 16
| Median |  | 10.095 | 10.095
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.07005%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 267 (Indicative Value 1.59 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.59 | 1
| 1 | 2 | 1.58 | 12
| Mean |  | 1.585 | 16
| Median |  | 1.585 | 1.585
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.4461%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 267 (Indicative Value 4.25 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.25 | 1
| 1 | 2 | 4.24 | 13
| Mean |  | 4.245 | 16
| Median |  | 4.245 | 4.245
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.1666%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 267 (Indicative Value 1.49 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.49 | 1
| 1 | 2 | 1.49 | 14
| Mean |  | 1.49 | 16
| Median |  | 1.49 | 1.49
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 267 (Indicative Value 1.6 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.61 | 1
| 1 | 2 | 1.59 | 11
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.8839%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 267 (Indicative Value 0.03 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.03 | 1
| 1 | 2 | 0.03 | 12
| Mean |  | 0.03 | 16
| Median |  | 0.03 | 0.03
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 267 (Indicative Value 0.109 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.107 | 13
| Mean |  | 0.1085 | 16
| Median |  | 0.1085 | 0.1085
| Std Dev. |  | 0.0021213 | 19
| Rel.Std.Dev. |  | 1.955%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 267 (Indicative Value 0.021 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0208 | 1
| 1 | 2 | 0.0208 | 14
| Mean |  | 0.0208 | 16
| Median |  | 0.0208 | 0.020823
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.1124%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 267 (Indicative Value 77.06 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 77.05 | 1
| 1 | 2 | 77.07 | 11
| Mean |  | 77.06 | 16
| Median |  | 77.06 | 77.06
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.01835%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 267 (Indicative Value 0.7 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.7 | 12
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 267 (Indicative Value 2.15 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.11 | 1
| 1 | 2 | 2.18 | 16
| Mean |  | 2.145 | 16
| Median |  | 2.145 | 2.145
| Std Dev. |  | 0.049497 | 22
| Rel.Std.Dev. |  | 2.308%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 267 (Indicative Value 0.06 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.05 | 18
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0.014142 | 24
| Rel.Std.Dev. |  | 23.57%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 267 (Indicative Value 0.015 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | 0.02 | 14
| Mean |  | 0.015 | 16
| Median |  | 0.015 | 0.015
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 267 (Indicative Value 7.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.1 | 1
| 1 | 2 | 7.3 | 20
| Mean |  | 7.2 | 16
| Median |  | 7.2 | 7.2
| Std Dev. |  | 0.14142 | 26
| Rel.Std.Dev. |  | 1.964%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 267 (Indicative Value 158 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 159 | 1
| 1 | 2 | 156 | 21
| Mean |  | 157.5 | 16
| Median |  | 157.5 | 157.5
| Std Dev. |  | 2.1213 | 27
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 267 (Indicative Value 388 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 390 | 1
| 1 | 2 | 386 | 22
| Mean |  | 388 | 16
| Median |  | 388 | 388
| Std Dev. |  | 2.8284 | 28
| Rel.Std.Dev. |  | 0.729%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 267 (Indicative Value 1.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 1.8 | 23
| Mean |  | 1.9 | 16
| Median |  | 1.9 | 1.9
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 267 (Indicative Value 1.63 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.64 | 1
| 1 | 2 | 1.62 | 24
| Mean |  | 1.63 | 16
| Median |  | 1.63 | 1.63
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 0.8676%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 267 (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 1
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 31
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 267 (Indicative Value 60 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 59.9 | 1
| 1 | 2 | 60.2 | 26
| Mean |  | 60.05 | 16
| Median |  | 60.05 | 60.05
| Std Dev. |  | 0.21213 | 32
| Rel.Std.Dev. |  | 0.3533%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 267 (Indicative Value 11.9 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.8 | 1
| 1 | 2 | 11.9 | 27
| Mean |  | 11.85 | 16
| Median |  | 11.85 | 11.85
| Std Dev. |  | 0.070711 | 33
| Rel.Std.Dev. |  | 0.5967%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 267 (Indicative Value 112 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 112 | 1
| 1 | 2 | 112 | 28
| Mean |  | 112 | 16
| Median |  | 112 | 112
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 267 (Indicative Value 6.31 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.3 | 1
| 1 | 2 | 6.31 | 29
| Mean |  | 6.305 | 16
| Median |  | 6.305 | 6.305
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.1122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 267 (Indicative Value 31 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 30 | 30
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 1.4142 | 36
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 267 (Indicative Value 3.79 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.83 | 1
| 1 | 2 | 3.74 | 6
| Mean |  | 3.785 | 16
| Median |  | 3.785 | 3.785
| Std Dev. |  | 0.06364 | 37
| Rel.Std.Dev. |  | 1.681%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 267 (Indicative Value 2.04 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.01 | 1
| 1 | 2 | 2.06 | 7
| Mean |  | 2.035 | 16
| Median |  | 2.035 | 2.035
| Std Dev. |  | 0.035355 | 38
| Rel.Std.Dev. |  | 1.737%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 267 (Indicative Value 1.05 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.04 | 1
| 1 | 2 | 1.05 | 8
| Mean |  | 1.045 | 16
| Median |  | 1.045 | 1.045
| Std Dev. |  | 0.0070711 | 39
| Rel.Std.Dev. |  | 0.6767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 267 (Indicative Value 17 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17 | 1
| 1 | 2 | 16.9 | 34
| Mean |  | 16.95 | 16
| Median |  | 16.95 | 16.95
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.4172%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 267 (Indicative Value 4.35 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.36 | 1
| 1 | 2 | 4.33 | 9
| Mean |  | 4.345 | 16
| Median |  | 4.345 | 4.345
| Std Dev. |  | 0.021213 | 41
| Rel.Std.Dev. |  | 0.4882%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 267 (Indicative Value 1.35 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.35 | 2
| Mean |  | 1.35 | 16
| Median |  | 1.35 | 1.35
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 267 (Indicative Value 4.81 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.84 | 1
| 1 | 2 | 4.78 | 20
| Mean |  | 4.81 | 16
| Median |  | 4.81 | 4.81
| Std Dev. |  | 0.042426 | 26
| Rel.Std.Dev. |  | 0.882%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 267 (Indicative Value 0.75 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.73 | 5
| Mean |  | 0.745 | 16
| Median |  | 0.745 | 0.745
| Std Dev. |  | 0.021213 | 27
| Rel.Std.Dev. |  | 2.847%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 267 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 22
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 267 (Indicative Value 31.3 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.3 | 1
| 1 | 2 | 31.2 | 23
| Mean |  | 31.25 | 16
| Median |  | 31.25 | 31.25
| Std Dev. |  | 0.070711 | 29
| Rel.Std.Dev. |  | 0.2263%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 267 (Indicative Value 0.28 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.27 | 6
| Mean |  | 0.275 | 16
| Median |  | 0.275 | 0.275
| Std Dev. |  | 0.0070711 | 30
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 267 (Indicative Value 0.029 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0287 | 1
| 1 | 2 | 0.0289 | 25
| Mean |  | 0.0288 | 16
| Median |  | 0.0288 | 0.0288
| Std Dev. |  | 0.00014142 | 31
| Rel.Std.Dev. |  | 0.491%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 267 (Indicative Value 6.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.8 | 1
| 1 | 2 | 6.4 | 26
| Mean |  | 6.6 | 16
| Median |  | 6.6 | 6.6
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 267 (Indicative Value 16.3 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.4 | 1
| 1 | 2 | 16.2 | 27
| Mean |  | 16.3 | 16
| Median |  | 16.3 | 16.3
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 0.8676%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 267 (Indicative Value 27.4 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.6 | 1
| 1 | 2 | 27.1 | 7
| Mean |  | 27.35 | 16
| Median |  | 27.35 | 27.35
| Std Dev. |  | 0.35355 | 34
| Rel.Std.Dev. |  | 1.293%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 267 (Indicative Value 57 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 58 | 1
| 1 | 2 | 56 | 29
| Mean |  | 57 | 16
| Median |  | 57 | 57
| Std Dev. |  | 1.4142 | 35
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 267 (Indicative Value 18 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 18 | 30
| Mean |  | 18 | 16
| Median |  | 18 | 18
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 267 (Indicative Value 7.29 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.3 | 1
| 1 | 2 | 7.28 | 8
| Mean |  | 7.29 | 16
| Median |  | 7.29 | 7.29
| Std Dev. |  | 0.014142 | 37
| Rel.Std.Dev. |  | 0.194%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 267 (Indicative Value 77 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 78 | 1
| 1 | 2 | 76.6 | 32
| Mean |  | 77.3 | 16
| Median |  | 77.3 | 77.3
| Std Dev. |  | 0.98995 | 38
| Rel.Std.Dev. |  | 1.281%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 267 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 33
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 39
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 267 (Indicative Value 7.9 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.3 | 1
| 1 | 2 | 7.5 | 9
| Mean |  | 7.9 | 16
| Median |  | 7.9 | 7.9
| Std Dev. |  | 0.56569 | 40
| Rel.Std.Dev. |  | 7.161%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 267 (Indicative Value 10.3 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.2 | 1
| 1 | 2 | 10.3 | 35
| Mean |  | 10.25 | 16
| Median |  | 10.25 | 10.25
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 267 (Indicative Value 5.24 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.22 | 1
| 1 | 2 | 5.26 | 36
| Mean |  | 5.24 | 16
| Median |  | 5.24 | 5.24
| Std Dev. |  | 0.028284 | 42
| Rel.Std.Dev. |  | 0.5398%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 267 (Indicative Value 8.7 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.6 | 1
| 1 | 2 | 8.8 | 20
| Mean |  | 8.7 | 16
| Median |  | 8.7 | 8.7
| Std Dev. |  | 0.14142 | 26
| Rel.Std.Dev. |  | 1.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 267 (Indicative Value 151 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 151 | 1
| 1 | 2 | 151 | 21
| Mean |  | 151 | 16
| Median |  | 151 | 151
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 267 (Indicative Value 1.69 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.71 | 1
| 1 | 2 | 1.66 | 22
| Mean |  | 1.685 | 16
| Median |  | 1.685 | 1.685
| Std Dev. |  | 0.035355 | 28
| Rel.Std.Dev. |  | 2.098%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 267 (Indicative Value 0.66 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.66 | 1
| 1 | 2 | 0.66 | 6
| Mean |  | 0.66 | 16
| Median |  | 0.66 | 0.66
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 267 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 267 (Indicative Value 9.55 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.48 | 1
| 1 | 2 | 9.62 | 25
| Mean |  | 9.55 | 16
| Median |  | 9.55 | 9.55
| Std Dev. |  | 0.098995 | 31
| Rel.Std.Dev. |  | 1.037%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 267 (Indicative Value 0.415 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.414 | 1
| 1 | 2 | 0.416 | 26
| Mean |  | 0.415 | 16
| Median |  | 0.415 | 0.415
| Std Dev. |  | 0.0014142 | 32
| Rel.Std.Dev. |  | 0.3408%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 267 (Indicative Value 0.3 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.2 | 27
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 267 (Indicative Value 0.3 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.29 | 1
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 4.714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 267 (Indicative Value 2.07 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.07 | 1
| 1 | 2 | 2.07 | 29
| Mean |  | 2.07 | 16
| Median |  | 2.07 | 2.07
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 267 (Indicative Value 72 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.2 | 1
| 1 | 2 | 71.7 | 30
| Mean |  | 71.95 | 16
| Median |  | 71.95 | 71.95
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 0.4914%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 267 (Indicative Value 8.38 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.4 | 1
| 1 | 2 | 8.35 | 31
| Mean |  | 8.375 | 16
| Median |  | 8.375 | 8.375
| Std Dev. |  | 0.035355 | 37
| Rel.Std.Dev. |  | 0.4222%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 267 (Indicative Value 19.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.3 | 1
| 1 | 2 | 19.2 | 32
| Mean |  | 19.25 | 16
| Median |  | 19.25 | 19.25
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.3673%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 267 (Indicative Value 2.03 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.04 | 1
| 1 | 2 | 2.01 | 8
| Mean |  | 2.025 | 16
| Median |  | 2.025 | 2.025
| Std Dev. |  | 0.021213 | 39
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 267 (Indicative Value 50 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 34
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 267 (Indicative Value 173 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 173 | 1
| 1 | 2 | 172 | 35
| Mean |  | 172.5 | 16
| Median |  | 172.5 | 172.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 0.4099%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.